32239 | 1220648-2023-05535 | 18378673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18378673&pc=PYX | 11/21/2023 | 11/05/2024 |  | 12/21/2023 |  |  | 11/21/2023 | 12/21/2023 |  |  | 08/31/2025 |  | 12/21/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024395790 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 11/21/2023 |  |  |  |  | The user facility reported a (b)(6) female was implanted with an impella rp device for mechanical circulatory support. the complainant reported that the impella rp flows abruptly diminished with loss of pa waveform. the patient was taken for positioning and the right ventricle was assessed under echo and x-ray. a clot was suspected by the team. the rp was weaned and explanted. no additional information was provided. | [5 empty] | The impella device has not yet been received from the customer. however, device evaluation is anticipated. upon investigation closure, a supplemental mdr will be filed. | The investigation of low pump flow has been completed. the impella device was not returned for evaluation. the root cause of the low pump flow is most likely biomaterial based on the log analysis. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05535. h3 device was not returned to manufacturer for evaluation.
32240 | 1220648-2023-05536 | 18378682 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18378682&pc=OZD | 11/21/2023 | 02/22/2024 |  | 12/21/2023 |  |  |  | 12/21/2023 |  |  | 09/30/2025 |  | 12/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024398214 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/21/2023 |  |  |  |  | The user facility reported a 68-year-old male with cardiomyopathy was to be implanted with an impella 5.5 for mechanical circulatory support. the team tried multiple attempts, multiple wires and insertion attempts/techniques but the delivery failed. the team reviewed the anatomy and concluded the calcified vasculature of the native anatomy had prevented placement. the team aborted 5.5 insertion and placed a cp instead via left axillary graft without further issues. | [5 empty] | The investigation into the reported delivery issue been completed since the original report was filed. the cause of the delivery issue was patient condition related with pump unable to advance past the ascending aorta due to the presence of plaque. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32241 | 1220648-2023-05537 | 18378695 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18378695&pc=OZD | 11/22/2023 | 11/29/2024 |  | 12/21/2023 |  |  |  | 12/21/2023 |  |  | 09/30/2025 |  | 12/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024401546 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/22/2023 |  |  |  |  | The user facility reported a 49-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the impella 5.5 had purge leaking at connection site of the purge tubing to the one way stop valve of the cassette. the leak was coming from an apparent crack in the connection. purge pressure and flow were stable and the nurse was directed to change the tubing which resolved the issue. | [5 empty] | The investigation for the purge cassette leak has been completed. no product was returned for investigation. clinical details were insufficient to determine the root cause. the cause of the purge leak cassette was not determined since product was not returned. corrections to the initial mfr report: d4 model number, added d6a/d6b, f6/f8 should have been left blank. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
32242 | 1220648-2023-05534 | 18378852 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18378852&pc=OZD | 12/20/2023 | 10/24/2024 |  | 12/21/2023 |  |  |  | 12/21/2023 |  |  |  |  | 12/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | XXXXXX |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional |  |  |  |  | 0 DA |  | Unknown |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Life Threatening | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/20/2023 |  |  |  |  | The abiomed team in japan located a publication entitled ... "intracerebral hematoma in patients with impella ventricular assist device placement for cardiogenic shock: report of three cases" ... noted a cerebral bleed and death on cp support. the event is reportable as the causal relationship can not be confirmed nor denied. the patient had an evacuation and transfusion prior to death from mof. the case is being followed up with team in japan for case and pump particulars. | [5 empty] | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrections to the initial mfr report: d4-model number f6/f8 should have been left blank. h1- type of reportable event changed to death. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32243 | 1220648-2023-05521 | 18379873 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18379873&pc=OZD | 12/17/2023 | 01/17/2024 |  | 12/21/2023 |  |  | 12/17/2023 | 12/21/2023 |  |  | 03/31/2025 |  | 12/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED INC. |  |  | US |  | 2024343800 |  | 0048-0045 |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/17/2023 |  |  |  |  | The user facility reported a 67-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient on impella cp experienced delivery issues. the patient was on impella support for 0.13 hours before the procedure was aborted. it was reported the sheath could not completely be inserted due to freshly placed iliac stent. the physician made the decision to abort and place intra-aortic balloon pump instead of impella cp. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported delivery issue has been completed since the original report was filed. no product was returned. the root cause of the delivery issue was most likely patient condition related.
32244 | 1220648-2023-05538 | 18379880 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18379880&pc=OZD | 12/16/2023 | 09/20/2024 |  | 12/21/2023 |  |  | 12/16/2023 | 12/21/2023 |  |  | 10/31/2024 |  | 12/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024291753 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/16/2023 |  |  |  |  | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 61-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient was being transferred to icu when the impella migrated into the left ventricle. the patient became bradycardic and asystole occurred. cpr was initiated immediately, and rosc (return of spontaneous circulation) was achieved. the impella position was verified under echo guidance and was 2.7cm below valve. the patient is stable. |  |  |  |  | The investigation for the reported bradycardia has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the bradycardia and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-05538 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32245 | 1220648-2023-05511 | 18368643 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18368643&pc=OZD | 12/01/2023 | 12/31/2024 |  | 12/20/2023 |  |  | 12/01/2023 | 12/20/2023 |  |  | 09/30/2025 |  | 12/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 PUMP |  |  | OZD |  |  | IMPELLA 5.5 PUMP |  |  | ABIOMED, INC. |  |  | US |  | 2024394423 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/01/2023 |  |  |  |  | The user facility reported a 47-year-old male was implanted with an impella 5.5 for mechanical circulatory support. approximately 2 weeks after implant the patient developed an acute hematoma and access site bleeding. it was noted the patient was sleeping on that side. patient returned to the operating room for emergent vascular repair which required the removal of the impella 5.5 pump. while in the operating room the patient received 4 units of packed red blood cells (prbc), 2 fresh frozen plasma (ffp), and platelets. the vascular damage was damaged at the anastomosis. the patient did not hemodynamically tolerate the removal of the impella 5.5 pump and had to have a new impella 5.5 pump placed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. assess access site for bleeding and hematoma: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿ potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ vascular dissection/perforation: cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ | The investigation for the access site bleed and the vascular damage has been completed. pump was not returned for investigation. as per clinical information provided, the blood had accumulated in the patients drain at the cutdown site which was resolved with suturing and protamine. protamine is given to patient with high act value to neutralize heparin' s effects. therefore, the root cause of the access site bleeding issue was most likely anticoagulation management due to high patient act values. clinical information mentions that the patient was sleeping on the side of the access site. a hematoma and damage to graft anastomosis was noted on the same side. therefore, the root cause of the vascular damage - peripheral vessel issue was most likely use related to patient movement. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. added h6 component code 925 corrections to the initial mfr report: added d6a/d6b. f6/f8 should have been left blank. g1 mdr reporting contact email updated.
32246 | 1220648-2023-05512 | 18368826 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18368826&pc=OZD | 11/27/2023 | 02/12/2024 |  | 12/20/2023 |  |  |  | 12/20/2023 |  |  | 09/30/2025 |  | 12/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 EU |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2024407003 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 69-year-old male noted as high-risk percutaneous cardiac intervention was implanted with an impella 5.5 device for mechanical circulatory support. during patient support the impella 5.5 was switched to another impella 5.5 due to a broken purge filter. the patient was stable during the procedure. | [5 empty] | The investigation into the purge leak ¿ pump has been completed. the device was not returned for investigation. based on the clinical information provided, a purge leak was discovered, and the cause was determined to be a broken purge filter. the pump was replaced with another 5.5 to resolve the issue. no logs were returned for data analysis. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
32247 | 1220648-2023-05514 | 18369133 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18369133&pc=OZD | 11/28/2023 | 01/17/2024 |  | 12/20/2023 |  |  | 11/28/2023 | 12/20/2023 |  |  | 10/31/2025 |  | 12/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024409581 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had an inability to deliver the impella 5.5 via the 8x30 vascutek gel weave graft. the anatomic limitations of the 59-year-old male patient made the team remove and replace the pump. there was no reported patient harm or injury. a new impella 5.5 was able to be placed via direct access and support initiated. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation into the delivery issues- anatomy has been completed. the device was not returned for benchtop evaluation and investigation. the root cause of the delivery issue was determined to be patient condition because of the patient anatomy.
32248 | 1220648-2023-05515 | 18369374 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18369374&pc=OZD | 12/04/2023 | 09/17/2024 |  | 12/20/2023 |  |  |  | 12/20/2023 |  |  | 09/30/2025 |  | 12/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024398872 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/04/2023 |  |  |  |  | The user facility reported a 58-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. large hematoma noted at insertion site. femstop applied to right groin site. (b)(4) unit packed red blood cells were administered. femostop removed the following morning with resolution of hematoma. the pump remained on for another 5 days of support. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- h6 component code 925 in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32249 | 1220648-2023-05516 | 18369507 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18369507&pc=OZD | 12/01/2023 | 03/14/2024 |  | 12/20/2023 |  |  |  | 12/20/2023 |  |  | 09/30/2025 |  | 12/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024404438 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/01/2023 |  |  |  |  | The user facility reported a 60-year-old female was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella cp was unsuccessfully placed. upon removing the cp, a right subclavian dissection occurred. the right subclavian artery was repaired. the patient is stable. | [5 empty] | The investigation for the reported peripheral vessel damage has been completed. the impella device was not returned for evaluation. the root cause of the vessel damage and its relation to the impella were not determined since no product was returned and limited clinical information was provided. d4-updated model number. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
32250 | 1220648-2023-05480 | 18354371 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18354371&pc=OZD | 12/09/2023 | 05/29/2024 |  | 12/19/2023 |  |  | 12/09/2023 | 12/19/2023 |  |  | 08/31/2025 |  | 12/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024395189 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  |  | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | There is not enough information to associate use of device with outcome. | The complainant reported that 65-year-old male patient in critical condition on impella cp support, had permanent suction alarms. pump was noted to be in good position but running with low flows. the physician suspected thrombus and recommended pump exchange. however, pump exchange was not done as the patient was through to be too unstable to tolerate being off pump. |  |  |  |  | The investigation into the low pump flow issue has been completed since the original report was filed. the pump was returned for investigation. no physical abnormalities were observed. the pump ran as expected during the test. data logs show low pump flow throughout the case with several suction alarms. the cause of the low pump flow issue was not determined since sufficient clinical information was not provided, and the returned pump ran as expected without reported physical abnormalities. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
32251 | 1220648-2023-05479 | 18354491 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18354491&pc=OZD | 11/25/2023 | 04/09/2024 |  | 12/19/2023 |  |  | 11/25/2023 | 12/19/2023 |  |  | 09/30/2025 |  | 12/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024394517 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/25/2023 |  |  |  |  | The user facility reported an impella 5.5 was implanted in a male of unknown age for mechanical circulatory support due to congestive heart failure and cardiogenic shock. there was high purge pressure on the pump. there were no kinks observed in the purge line. no patient harm was reported. no additional information was provided. | [5 empty] | The impella device has not yet been received from the customer. however, a device evaluation is anticipated. upon investigation closure, a supplemental mdr will be filed. the instructions for use states the following about high purge pressure: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure. which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation into the high purge pressure issue has been completed since the original report was filed. the pump was returned, tested, and evaluated. a significant amount of biomaterial was observed on the outflow cage and purge gap. the cause of the high purge pressure issue was biomaterial based on return product investigation and data log review.
32252 | 1220648-2023-05481 | 18354590 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18354590&pc=OZD | 11/22/2023 | 11/22/2024 |  | 12/19/2023 |  |  | 11/22/2023 | 12/19/2023 |  |  | 08/31/2025 |  | 12/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024384886 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Activation, Positioning or Separation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 11/22/2023 |  |  |  |  | The user facility reported an attempt to implant an impella 5.5 in a 69yo female for mechanical circulatory support due to acute myocardial infarction and cardiogenic shock. impella 5.5 was opened, prepped, and inserted, but due to stenosis at the distal subclavian artery, the pump was unable to pass. the pump was removed and balloon angioplasty was attempted. a second attempt to place the impella was tried to no avail. impella cp was then placed via axillary artery for support. no additional information was provided. | [5 empty] | The investigation into delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related. instructions for use as follows: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings and cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings and cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: axillary insertion of the impella 5.5 catheter ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings and cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05481. d6a and d6b revised from the initial submission in accordance with updated procedures. e4 should have been left blank f6/f8-should have been left blank. g1 reporting contact facility name missing g1 reporting contact fax number missing. h2 device evaluation was selected in error h3 device evaluated by manufacturer section incorrect h6 code 4629 and 2906 were reported incorrectly. new codes were added to health effect - impact code and medical device problem code h10 instructions for use missing. | The impella device has not yet been received from the customer. device evaluation is anticipated. upon investigation closure, a supplemental mdr will be filed.
32253 | 1220648-2023-05482 | 18355669 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18355669&pc=OZD | 11/28/2023 | 02/07/2024 |  | 12/19/2023 |  |  | 11/28/2023 | 12/19/2023 |  |  | 08/31/2025 |  | 12/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024394522 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/28/2023 |  |  |  |  | The user facility reported a 53-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient was on impella cp support and there was high purge pressure post implant. staff tried troubleshooting and changed the purge cassette to no avail. staff added tissue plasminogen activator in the purge solution which seemed to resolve the issue. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported high purge pressure has been completed since the original report was filed. there is a spike in motor current and the purge pressure begins to rise after 5 hours. after 5 hours the ¿purge system blocked and the high purge pressure¿ alarm occurred. purge flows were low. tpa was added to the purge and then the purge flows stabilized. the root cause of the high purge pressure is most likely due to biomaterial.
32254 | 1220648-2023-05483 | 18355685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18355685&pc=OZD | 11/18/2023 | 09/29/2025 |  | 12/19/2023 |  |  | 11/18/2023 | 12/19/2023 |  |  | 06/30/2025 |  | 12/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024360239 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hematoma | Hemolysis | Ischemia | [9 empty] | Required Intervention | Hospitalization | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 11/28/2023 |  |  |  |  | The user facility reported a 76-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient¿s distal pulse was poor in the impella site. fem/fem bypass was done, and leg warmer used. support continued with the implanted pump following the intervention. | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry indicating this patient to have endured hemolysis with a "definite" relationship to the impella device. as a result, hospitalization was prolonged. the patient's family ultimately made the decision to withdraw care and the patient expired. | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured a hematoma of right femoral impella access site with a "definite" relationship to the impella device used: ¿concern for visual hematoma of the right femoral artery accompanied with a drop in hemoglobin (hgb) while on heparin. ct reveals stranding in the area of concern, but negative for obvious hematoma. the patient received 2 units of packed red blood cells (prbc), stabilizing hemoglobin. groin impella access site was stable. after prbc transfusion the patient's hgb improved to 10.2 g/dl. heparin was stopped and resumed once the patient's hgb was stabilized. loqi also reported that the patient endured intracranial hemorrhage- subarachnoid/intraparenchymal with a "possible" relationship to the impella device used: ¿ct-head without contrast (21 nov 2023) reveals acute subarachnoid/intraparenchymal hemorrhage. acute intraparenchymal hemorrhage in right basal ganglia ruptures into the right sylvian fissure; subarachnoid hemorrhage of suprasellar, prepontine, interpeduncular and peri-mesencephalic cisterns. anticoagulation/antiplatelet medications held. per loqi the patient also endured ischemic stroke with a "possible" relationship to the impella device used: ¿patient developed left sided hemiparesis. ct reveals right mca infarct and occlusion of proximal right m2 branch with reconstitution in m3 branch. embolectomy was unsuccessful. follow-up ct reveals large subarachnoid hemorrhage. care measures were transitioned to comfort care. time of death was called (b)(6) 2023 at 0640.¿ |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/pos, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | New information was received and recorded in b(5). upon completion of our investigation, a supplemental mdr will be filed. | When making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition. any concomitant medication. vascular size or variations thereof. detailed description of the symptoms experienced by the patient. physical examination findings, including pulse assessment and visual examination of the affected limb. diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. laboratory test results, including blood tests for markers of inflammation or clotting disorders. any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. pump was not returned for investigation. data logs were investigated but no suction alarm, motor current spike, positioning issues, etc. were observed. data logs were inconclusive. therefore, the root cause of the hemolysis issue was not determined since product was not returned and insufficient clinical information provided. this report is being filed as part of a retrospective review of historical records. | B5: additional narrative added for new information received for stroke and hematoma. h6: coding for health effect, health impact, investigation findings and investigation conclusions added for new information received and investigated. the investigation for the reported issues has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. b2: date of death was corrected from (b)(6) 2023. h1: corrected to event type death.
32255 | 1220648-2023-05484 | 18355704 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18355704&pc=OZD | 11/28/2023 | 01/16/2024 |  | 12/19/2023 |  |  | 11/28/2023 | 12/19/2023 |  |  | 08/31/2025 |  | 12/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC |  |  | US |  | 2024391556 |  | 0048-0371-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/28/2023 |  |  |  |  | The user facility in japan reported a male patient of an unknown age presented to the hospital with chest pain and elevated st and was to be implanted with an impella cp device for mechanical circulatory support. the impella cp implant insertion was attempted in the left leg, but the pig tail did not enter the left ventricle (lv) easily. an angiography was performed and there was a type a dissection. the clinical representative confirmed the dissection was not impella cp related. the impella cp implant was aborted. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. clinical details states that the patient had an aortic dissection, that was found during the insertion of impella when used the angiography, which revealed that the patient had type a dissociation and the aortic dissection was not related to impella. hence, the root cause was determined to be patient condition. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32256 | 1220648-2023-05485 | 18356043 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18356043&pc=OZD | 12/17/2023 | 02/22/2024 |  | 12/19/2023 |  |  | 12/17/2023 | 12/19/2023 |  |  | 10/31/2025 |  | 12/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024411991 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/17/2023 |  |  |  |  | The user facility reported a 72-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient had right femoral artery insertion of impella cp, and no right pedal pulse could be found after the procedure. md injected dye down, re-access port of reperfusion sheath and no flow seen under fluoroscopy. md elected to remove due to improved hemodynamics. the patient was on impella support for two hours. the patient is stable. | [5 empty] | The investigation into the limb ischemia reversible issue has been completed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32257 | 1220648-2023-05486 | 18356051 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18356051&pc=OZD | 11/29/2023 | 02/01/2024 |  | 12/19/2023 |  |  | 11/29/2023 | 12/19/2023 |  |  | 07/31/2025 |  | 12/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024391369 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/29/2023 |  |  |  |  | The user facility reported an 87-year-old male in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support after a percutaneous coronary intervention (pci). following the pci procedure, the physician advanced impella cp into the left ventricle and when advancing repositioning sheath, the impella folded in the left ventricle and did not straighten out. the impella cp was removed and a new one was placed. the new impella insertion was successful and there was no patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. clinical details states that while advancing the repositioning sheath, the impella folded in the left ventricle, did not straighten out, and had to be removed. the root cause of the delivery issue was determined to be use related as the pump migrated too deep into the ventricle when the repo unit was advanced. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32258 | 1220648-2023-05487 | 18357428 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18357428&pc=OZD | 12/07/2023 | 04/17/2024 |  | 12/19/2023 |  |  | 12/07/2023 | 12/19/2023 |  |  | 06/30/2025 |  | 12/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024379687 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/07/2023 |  |  |  |  | The user facility reported a 84-year-old female was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient was experiencing a cold leg. doppler study ordered. it was further reported that the patient¿s right femoral artery was oozing. the dressing was changed. the impella was explanted after reperfusion cannula placed with excessive bleeding. three units of prbc administered for vascular injury. the patient is stable. | [5 empty] | The investigation for the reported limb ischemia and access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia and access site bleed could not be determined. d4-updated model number. h6 impact code added 4627, 4641. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32259 | 1220648-2023-05488 | 18357447 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18357447&pc=OZD | 11/25/2023 | 11/05/2024 |  | 12/19/2023 |  |  | 11/25/2023 | 12/19/2023 |  |  | 02/28/2025 |  | 12/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024337424 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had impella cp placed for a 40-year-old male stemi patient with a ventricular fibrillation arrest. the impella cp was placed but, the team had an access site bleed at the site. the team treated the bleed with a femostop placement. additionally the team saw signs of hemolysis. the team made choice to explant after just over 15hours. the discussion was made to place instead a 5.5 via the axillary but, instead just utilized medical management. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation of access site bleeding and hemolysis has been completed. the impella device was not returned for evaluation. the root cause of the hemolysis issue was improper positioning of the device. the root cause of the access site bleeding issue was high patient activated clotting time (act) values. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05488. d6a and d6b revised from the initial submission in accordance with updated procedures. g1 reporting contact fax number missing. h6 component code revised from the initial submission in accordance with updated procedures.
32260 | 1220648-2023-05489 | 18358299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18358299&pc=OZD | 12/06/2023 | 12/19/2023 |  | 12/19/2023 |  |  | 12/06/2023 | 12/19/2023 |  |  | 05/31/2025 |  | 12/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024358090 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that female patient of unknown age on impella cp support had purge filter leakage. the purge pressure and purge flow were noted to be sufficient for support, and the pump continued to run. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
32261 | 1220648-2023-05490 | 18358633 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18358633&pc=OZD | 11/29/2023 | 09/21/2024 |  | 12/19/2023 |  |  | 11/29/2023 | 12/19/2023 |  |  | 10/31/2025 |  | 12/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 |  |  | ABIOMED, INC. |  |  | US |  | 2024408272 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup |  |  | 11/29/2023 |  |  |  |  | The user facility reported an impella 5.5 in a 61yo male for mechanical circulatory support during coronary artery bypass graft. after placing the pump and support was initiated, the team observed neuro changes one day after implant. a cerebral vascular accident (cva) was suspected and it was embolic in nature. the patient underwent intervention of thrombectomy and craniotomy. the pump remained on for support with condition being monitored. no additional information was provided. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device has been returned for evaluation, but has not yet been evaluated. investigation is in process. a supplemental report will be submitted upon completion of the investigation. the instructions for use identifies cva as a potential event. it states "¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6. | The investigation into the high purge pressure has been completed since the original report was filed. the product was returned to the lab for investigation. there was biomaterial found wrapped around the impeller blades. in the data logs, there was a motor current spike, suction alarms, and decreased flow at the same time a stroke was reported. the root cause of the stroke/cva is most likely due to biomaterial ingested. the relationship to the impella is unknown.
32262 | 1220648-2023-05491 | 18358853 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18358853&pc=OZD | 11/21/2023 | 12/19/2023 |  | 12/19/2023 |  |  |  | 12/19/2023 |  |  | 04/30/2025 |  | 12/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024366215 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/21/2023 |  |  |  |  | The user facility reported a 64-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported there was a motor current increase during weaning of the impella. purge flows dropped below 2 and pressure was above 1000. the pump was explanted for the low flows. | [5 empty] | Impella cp pump was returned for investigation. visually inspected and no biomaterial or other defects were observed. pump was connected to console and pump ran at auto with flow of 3.7 l/min. the purge flow and pressure were normal between 15-17 ml/hr and 370 to 450 mmhg range without reproducing high purge pressure. no other abnormalities were found. data logs were reviewed, and high purge pressure alarmed at the end of case with motor current spike seen along with gradual rise in purge pressure and a slight saturation in flow waveform. the cause of the high purge pressure issue was most likely biomaterial ingestion with motor current spike followed by gradual rise in purge pressure and slight saturation in the flow at end of case per log review.
32263 | 1220648-2023-05492 | 18359911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18359911&pc=OZD | 12/08/2023 | 09/23/2024 |  | 12/19/2023 |  |  | 12/08/2023 | 12/19/2023 |  |  |  |  | 12/19/2023 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1735461 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant in the eu reported automated impella controller (aic) stopped working during the procedure. there was no reported harm. | [5 empty] | The investigation for automated impella controller (aic) - a/c power issues has been completed. after reviewing data logs and functional testing of the console, the cause of the aic issue was a defective pbm board. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05492: d2 common device name. g1 reporting contact fax number missing. | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
32264 | 1220648-2023-05493 | 18360019 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18360019&pc=OZD | 11/30/2023 | 01/30/2024 |  | 12/19/2023 |  |  |  | 12/19/2023 |  |  |  |  | 12/19/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1679798 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Mechanical Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/30/2023 |  |  |  |  | The user facility reported a 47-year-old male with cardiomyopathy was implanted with an impella for mechanical circulatory support. the automated impella controller's (aic) blue piece, where the pressure disk slides into, was broken. the pressure disk kept inflating and alarming low pressure and expanding diaphragm pushing blue piece further out appearing broken. this ruptured diaphragm of purge disk. staff replaced the purge cassette, but when the purge disk slid out, the blue piece was visually broken and would not snap into place. the aic was replaced successfully. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. imc logs from the reported occurrence date revealed that the console did issue purge pressure low alarms. there was also an instance where a purge disc not detected alarm was issued. after further analysis of the imc logs, there was evidence that the console was able to detect the purge disc in some cases but would receive alarms for purge pressure low. the console was returned and tested after arriving. the purge pressure low alarm issue was able to be reproduced, and the purge retainer was confirmed to be broken. the portion of the purge retainer that was supposed to be fastened to the purge assembly was no longer intact. the cause of the issue was a broken purge retainer.
32265 | 1220648-2023-05494 | 18360174 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18360174&pc=OZD | 09/03/2023 | 12/19/2023 |  | 12/19/2023 |  |  | 11/21/2023 | 12/19/2023 |  |  | 08/14/2025 |  | 12/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5, - EU |  |  | ABIOMED, INC. |  |  | US |  | 2024374613 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/21/2023 |  |  |  |  | The user facility reported a patient (unknown age, race and gender) with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, the impella 5.5 had high purge pressure alarms. tissue plasminogen activator was added to the purge and that resolved the high purge pressure. | [5 empty] | The investigation into the high purge pressure has been completed. the device was not returned for evaluation. the root cause of the high purge pressure is most likely due to biomaterial as tissue plasminogen activator was added to the purge and that resolved the high purge pressure. biomaterial is considered a patient condition.
32266 | 1220648-2023-05497 | 18360293 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18360293&pc=PYX | 11/21/2023 | 09/21/2024 |  | 12/19/2023 |  |  | 11/21/2023 | 12/19/2023 |  |  | 09/30/2024 |  | 12/19/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023270986 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  |  | Death | Yes | Yes | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/21/2023 |  |  |  |  | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 66-year-old male (race unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. during support, the impella rp pump stopped and the team was unable to restart. the pump was removed. the patient was on bipella support and continued with support from an impella cp. |  |  |  |  | The investigation into the pump stop issue has been completed since the initial report was submitted. the product was not returned for investigation. the cause of the pump stop issue was most likely biomaterial, based on data log review. | B2: date of patient death is unknown. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-05497 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
32267 | 1220648-2023-05495 | 18360316 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18360316&pc=OZD | 12/05/2023 | 02/06/2024 |  | 12/19/2023 |  |  |  | 12/19/2023 |  |  | 05/31/2025 |  | 12/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024363160 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 37 YR |  | Male |  |  |  | Malfunction | Yes | Yes | Infusion or Flow Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported 37-year-old male on impella cp support had constant suction alarms and questionable thrombus. therefore, the pump was explanted and a new impella was placed. there was no reported patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported low pump flow has been completed since the original report was filed. the pump was returned for investigation due to a significant amount of biomaterial observed on the impeller. no other physical abnormalities were detected. during testing, zero pump flow was recorded while running on p9. no logs were provided. as per clinical, pump was explant done due to consistent suction alarm and doubtful thrombus. the cause of the low pump flow issue was biomaterial wrapped around the impeller.
32268 | 1220648-2023-05496 | 18360326 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18360326&pc=OZD | 12/05/2023 | 02/02/2024 |  | 12/19/2023 |  |  |  | 12/19/2023 |  |  |  |  | 12/19/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1081667 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Activation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/05/2023 |  |  |  |  | No patient involvement: the us complainant had a patient being prepped for a hrpci (high risk percutaneous coronary intervention). the aic (automated impella controller) was found to have a purge disc that was not recognized. the aic was replaced and support proceeded without any harm or injury from delay. | [5 empty] | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs revealed several purge disc not detected alarms. the console was evaluated and it was confirmed that the purge disc detect flag had broken off on the console. the root cause of this issue was a broken purge retainer. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32269 | 1220648-2023-05499 | 18360341 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18360341&pc=OZD | 11/21/2023 | 11/26/2024 |  | 12/19/2023 |  |  |  | 12/19/2023 |  |  | 01/31/2025 |  | 12/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024329787 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/21/2023 |  |  |  |  | The user facility reported a 77-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient urine output was dark brown and plasma free was greater than 300. the patient was taken back to the operation room and impella cp was removed and intra-aortic balloon pump was placed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the hemolysis has been completed. pump was not returned for investigation. data logs were investigated and there was no motor current spikes. there was slight downward trend observed in placement signal. clinical mentioned that the patient had a cpr and "impella in aorta" alarms were observed right after that and downward placement signal trend was observed following that. therefore, the root cause of the hemolysis issue was most likely improper positioning of the device due to improper technique (cpr given to the patient and ramped up to p5 without echo). added h6 component code 925 corrections to the initial mfr report: added d6a/d6b f6/f8 should have been left blank.
32270 | 1220648-2023-05498 | 18360382 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18360382&pc=OZD | 11/24/2023 | 10/07/2024 |  | 12/19/2023 |  |  |  | 12/19/2023 |  |  | 07/31/2023 |  | 12/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022114315 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Mitral Perforation | Thrombosis/Thrombus | Mitral Valve Insufficiency/ Regurgitation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/24/2023 |  |  |  |  | The complainant reported that a 78-year-old male patient developed an apical thrombus and worsening mitral regurgitation following impella cp explant. anticoagulants were initially utilized to treat the thrombus; however, a thrombectomy was eventually required to remove the clot. upon attempting to treat the mitral regurgitation, it was determined that the chordae tendineae had been ruptured during pump removal. a mitral valvuloplasty was performed to treat the condition. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation of heart valve damage and thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05498: d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact fax number missing. h6 component code revised from the initial submission in accordance with updated procedures.
32271 | 1220648-2023-05500 | 18360390 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18360390&pc=OZD | 11/21/2023 | 02/29/2024 |  | 12/19/2023 |  |  |  | 12/19/2023 |  |  | 07/31/2025 |  | 12/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024391301 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Disability | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/21/2023 |  |  |  |  | The user facility reported a 38-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp and ecmo (extracorporeal membrane oxygenation) support ongoing for a patient admitted in a vfib (ventricular fibrillation) arrest. the patient developed limb ischemia and bilateral reperfusion sheaths. the team additionally stated they were going to explant and replace with a 5.5 pump via the axillary site. the cp and ecmo remain on at this time. | [5 empty] | The investigation into the limb limb ischemia ¿ irreversible has been completed. the device and data were not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32272 | 1220648-2023-05501 | 18360568 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18360568&pc=OZD | 11/24/2023 | 09/20/2024 |  | 12/19/2023 |  |  | 11/24/2023 | 12/19/2023 |  |  | 09/30/2025 |  | 12/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024407276 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/24/2023 |  |  |  |  | The user facility reported a 68-year-old female patient was implanted with an impella device for mechanical circulatory support. the complainant reported that after implantation of the impella, a small hematoma was noted, and manual pressure was held until hemostasis was achieved. intraprocedural the impella site hematoma was again noted and manual pressure was held. it was determined that the impella would be pulled due to continued bleeding and the patient tolerating weaning. the impella was successfully explanted. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- h6 component code 925 d4-corrected model number added- d6a/d6b f6/f8 should have been left blank in the initial report.
32273 | 1220648-2023-05502 | 18360627 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18360627&pc=OZD | 11/25/2023 | 05/07/2024 |  | 12/19/2023 |  |  | 11/25/2023 | 12/19/2023 |  |  | 08/31/2025 |  | 12/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024385570 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/25/2023 |  |  |  |  | The user facility reported a 35-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the pump was delivered but a hematoma developed and the team treated the axillary site by placement of a jackson pratt drain and removed clot burden. the pump remains on for support while being monitored. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned by the user facility for investigation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. it was reported possible increased hematoma in right chest with decreased right upper extremity (rue) function. patient was taken to or for hematoma evacuation by doctor. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32274 | 1220648-2023-05442 | 18343845 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18343845&pc=OZD | 11/21/2023 | 02/22/2024 |  | 12/18/2023 |  |  | 11/21/2023 | 12/18/2023 |  |  | 09/30/2025 |  | 12/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024408656 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that upon finishing the case of a 85-year-old female patient, an angiogram was performed and no flow or pulse was found to the left upper extremity (lue). the impella cp was therefore removed. | [5 empty] | The investigation into the reported limb ischemia been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32275 | 1220648-2023-05443 | 18343851 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18343851&pc=OZD | 11/20/2023 | 04/17/2024 |  | 12/18/2023 |  |  | 11/20/2023 | 12/18/2023 |  |  | 06/30/2025 |  | 12/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024377906 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had impella cp placed for support of a patient admitted with a stroke and other underlying cardiac and systemic comorbidities. when the patient was taken to the catheterization lab and coronary artery disease was diagnosed, the cp was placed. the cp was removed and replaced due to the low flows and suction that was not resolved. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ | The investigation for the low pump flow has been completed. the impella device was not returned for evaluation. data logs were reviewed which showed low pump flow with several suction alarms during the case run. the cause of the issue was not determined since product was not returned and sufficient clinical information was not provided.
32276 | 1220648-2023-05444 | 18343859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18343859&pc=OZD | 11/19/2023 | 08/18/2025 |  | 12/18/2023 |  |  | 11/19/2023 | 12/18/2023 |  |  | 09/30/2025 |  | 12/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024404371 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had impella cp support that was ongoing for a 70-year-old male patient admitted in cgs. the team noted an access site bleed on day of insertion. the bleed was treated by a stitch by vascular at the patient's bedside. the team additionally noted there was signs of hemolysis. the team tried pump repositioning and administered blood products and fluids. the pump was explanted the following day. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. d6a and d6b revised from the initial submission in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number. h6 component code and investigation conclusions revised from the initial submission in accordance with updated procedures. h10 additional manufacture narrative revised from the initial submission in accordance with updated procedures.
32277 | 1220648-2023-05391 | 18344410 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18344410&pc=OZD | 11/05/2023 | 06/19/2024 |  | 12/18/2023 |  |  |  | 12/18/2023 |  |  | 08/31/2025 |  | 12/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024394526 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hematoma | Sepsis | Multiple Organ Dysfunction Syndrome | [9 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/13/2023 |  |  |  |  | Long-term outcome and quality indicator impella (loqi) registry notification received alleging multi-organ failure with a possible relationship to the impella cp device. the report narratives are as follows. multi-organ failure: "subject's cause of death is multi-organ failure." sepsis: "wbc trending upward. patient hypotensive but no fever. ua and covid-19 test negative. vanco and zosyn started." hematoma: "small hematoma @ impella insertion site w/oozing out of r chest cp drain. hgb went from 9.9. to 8.7. received 1 unit prbc. ct surgery consulted. stated no active bleed but ctm. q1 hr flushes and pressure drain around jp site. chest tube less than 2l output within 24 hrs." cardiac arrest: "subject went to vtach and unresponsive. performed cpr but was not able to resuscitate (asystolic to v fib). ctb @ 0716." | [5 empty] | The investigation into the reported issues has been completed. the device was not returned for investigation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. as the necessary information was not provided, the root cause of the vt/vf (ventricular tachycardia/ventricular fibrillation) was not determined. because evidence related to the infection has not been provided, the root cause could not be determined. instructions for use for the related event are as follows: impella cp with smartassist. section: table 3.2 impella catheter components. ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b.7 revised as this information was inadvertently omitted from the previously submitted report. d.4 revised model number from the initial submission in accordance with updated procedures. d.6a and d.6b revised from the initial submission in accordance with updated procedures. g.1 reporting contact facility name was inadvertently left off the previously submitted report and was added. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures. g.3 date was reported incorrectly on the previously submitted report. the correct date is 13-nov-2023. h.5 revised as previously entered incorrectly. h.6 added codes 4111 and 4115 as they were inadvertently left off the previously submitted manufacture device report. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: hematoma: ¿assess access site for bleeding and hematoma.¿ sepsis: ¿infusion filter the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿
32278 | 1220648-2023-05445 | 18345334 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18345334&pc=OZD | 11/21/2023 | 11/05/2024 |  | 12/18/2023 |  |  | 11/21/2023 | 12/18/2023 |  |  | 04/30/2025 |  | 12/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024330690 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/21/2023 |  |  |  |  | The user facility reported a 49-year-old male patient was implanted with an impella device for mechanical circulatory support. the us complainant had a 5.5 placed at community hospital and was then transferred to tertiary center for continued care and monitoring. the patient exhibited signs of hemolysis during the support. the team treated by rbc (red blood cell) infusion. the pump support level was also adjusted. the patient remains on for support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the hemolysis has been completed. pump was not returned for investigation. data logs were reviewed and no motor current spikes, suction alarms or positioning issue symptoms were observed. clinical mentions labs were hemolyzed and red urine was observed. p-level was dropped but the result of that is unknown. therefore, the root cause of the hemolysis issue was not determined since product was not returned, limited clinical information was provided and data log review was inconclusive. added- h6 component code 925 corrections to the initial mfr report: d4-model number, added d6a/d6b, f6/f8 should have been left blank.
32279 | 1220648-2023-05446 | 18350064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18350064&pc=OZD | 11/18/2023 | 03/22/2024 |  | 12/18/2023 |  |  | 11/18/2023 | 12/18/2023 |  |  | 07/31/2025 |  | 12/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024375623 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the 62-year-old male patient pulled their impella cord during impella 5.5 support, which caused the red impella plug to detach from the catheter and wires. the pump subsequently stopped and the decision was made to remove the device. there was no patient harm as a result of the noted issue. | [5 empty] | The investigation into the pump stop issue has been completed since the initial report was submitted. the device was not returned for investigation. according to clinical details, the patient pulled the impella cord, causing the red handle to detach from the catheter. the cause of the pump stop issue was an open electrical line at red handle, based on given clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
32280 | 1220648-2023-05448 | 18350572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18350572&pc=OZD | 11/19/2023 | 01/29/2024 |  | 12/18/2023 |  |  | 11/19/2023 | 12/18/2023 |  |  | 09/30/2025 |  | 12/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024404417 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/19/2023 |  |  |  |  | The user facility reported a 59-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp support ongoing when the pump had signs of access site bleeding. the bleed prompted the team to stop the systemic heparin. the cp remained on despite this for another day. | [5 empty] | The investigation into the access site bleeding - major has been completed. the pump was not returned for investigation. the root cause of the access site bleeding - major was not determined since product was not returned and there is a lack of clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32281 | 1220648-2023-05447 | 18350574 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18350574&pc=OZD | 11/18/2023 | 04/30/2024 |  | 12/18/2023 |  |  | 11/18/2023 | 12/18/2023 |  |  | 09/30/2025 |  | 12/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024403253 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported male patient of unknown age on impella cp support had an access site hematoma. additional information pertaining to the hematoma was not provided. support with the implanted pump was maintained despite the condition. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. b in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
32282 | 1220648-2023-05449 | 18350812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18350812&pc=OZD | 11/20/2023 | 04/05/2024 |  | 12/18/2023 |  |  | 11/20/2023 | 12/18/2023 |  |  | 09/30/2025 |  | 12/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024397145 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/20/2023 |  |  |  |  | The user facility reported a 48-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported there was a high purge alarm on the impella. the staff checked for kinks and none were found. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation into the high purge pressure issue has been completed since the initial report was submitted. the device was not returned for investigation. the cause of the high purge pressure issue was most likely biomaterial, based on data log review.
32283 | 1220648-2023-05450 | 18351451 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18351451&pc=OZD | 11/19/2023 | 04/09/2024 |  | 12/18/2023 |  |  | 11/19/2023 | 12/18/2023 |  |  |  |  | 12/18/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1133147 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant had impella 5.5 support ongoing when the automated impella controller (aic) was found to be exhibiting a yellow alarm for placement signal and lv signal. support remained ongoing despite the aic alarming. there was no reported patient harm. | [5 empty] | The investigation into the aic - controller error (yellow alarm) issue has been completed since the initial report was submitted. the console was returned, visually inspected, and tested. the reported failure was confirmed in the data logs; however, the failure was unable to be reproduced during testing. the cause of the issue was not determined since the issue could not be reproduced. | The automated impella controller (aic) was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32284 | 1220648-2023-05451 | 18351605 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18351605&pc=PYX | 11/19/2023 | 04/15/2024 |  | 12/18/2023 |  |  | 11/19/2023 | 12/18/2023 |  |  | 08/31/2025 |  | 12/18/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024395807 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Positioning Problem | Migration | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella rp for support of a patient was was found down and receiving cpr in the field. the patient was admitted and had percutaneous coronary intervention (pci). the patient was then transferred to tertiary center for the rp placement. the team had issues with placement, but eventually pump delivered. upon attempting to remove the sheath, the pump dislodged from the appropriate position. the mispositioned pump was unable to be re-delivered and was therefore explanted. the explanted pump was not replaced. | [5 empty] | The investigation for the reported positioning issues the issues with the placement signal are related to the position of the device as when the pump was started the cvp was negative with placement signal issues. when the pump was repositioned, the placement signal and metrics became stable. the root cause of the positioning issues is most likely related to use as the pump was not in the proper position upon starting the device causing the optical issues and then accidentally pulled back. added-h6 medical device product problem 3009. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
32285 | 1220648-2023-05452 | 18351800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18351800&pc=OZD | 11/19/2023 | 10/30/2024 |  | 12/18/2023 |  |  | 11/19/2023 | 12/18/2023 |  |  | 05/31/2025 |  | 12/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024356628 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the 51-year-old male patient suffered a vascular injury during impella cp support. five percloses were initially deployed, but were deemed inadequate for the condition. vascular surgery was subsequently called for balloon placement and vessel repair. the pump was explanted as part of the repair and an impella 5.5 was placed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05452. d6a and d6b revised from the initial submission in accordance with updated procedures. g1 reporting contact fax number missing. h6 component code revised from the initial submission in accordance with updated procedures.
32286 | 1220648-2023-05453 | 18351983 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18351983&pc=OZD | 11/20/2023 | 03/01/2024 |  | 12/18/2023 |  |  | 11/20/2023 | 12/18/2023 |  |  | 09/30/2025 |  | 12/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024401054 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 pump could not be successfully placed during attempted delivery. it was noted that due to the patient's anatomy, the pump kept deflecting to the descending aorta and was unable to track into the ascending aorta. as a result, the decision was made to abort the implant. there was no patient harm. | [5 empty] | The investigation into the delivery issues- anatomy issue has been completed. the device was not returned for investigation. the cause of the delivery issue was most likely patient condition related. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
32287 | 1220648-2023-05454 | 18352053 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18352053&pc=OZD | 11/20/2023 | 07/17/2024 |  | 12/18/2023 |  |  | 11/20/2023 | 12/18/2023 |  |  | 05/31/2025 |  | 12/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024351417 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  | White | Malfunction | No | Yes | Positioning Failure | Migration | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during transport of 66-year-old male patient on impella cp support, the cp was noted to no longer be in the correct position and was unable to be repositioned without guidewire. therefore, the impella was removed and replaced with a new one. there was no reported patient harm. | [5 empty] | The investigation for the reported positioning issue has been completed. neither the device nor data logs were returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the positioning issues was most likely due to patient movement during transport. added- h6 impact code 4628, h6 med dev prob code 1158 added- d6a/d6b g1-mfr site fax number corrected. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32288 | 1220648-2023-05455 | 18352061 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18352061&pc=OZD | 11/21/2023 | 11/08/2024 |  | 12/18/2023 |  |  | 11/21/2023 | 12/18/2023 |  |  | 09/30/2025 |  | 12/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024407085 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/21/2023 |  |  |  |  | The user facility reported a 69-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed hemolysis during impella cp support. the impella was noted to be shallow and intertwined in the mitral valve at the time of the noted condition. support levels for the pump were subsequently dropped to p2 temporarily to allow for some support while limiting hemolysis. the pump was explanted the following day. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the hemolysis has been completed. pump was not returned for investigation. data logs were reviewed and no positioning issues, suction alarms or motor current spikes were observed. the clinical information provided cannot be confirmed in the data logs. therefore, the root cause of the hemolysis issue was not determined since product was not returned and data logs were inconclusive. added- h6 component code 925 corrections to the initial mfr report: added d6a/d6b. f6/f8 should have been left blank.
32289 | 1220648-2023-05456 | 18352132 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18352132&pc=OZD | 11/21/2023 | 11/11/2024 |  | 12/18/2023 |  |  | 11/21/2023 | 12/18/2023 |  |  | 02/28/2025 |  | 12/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024338466 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Cardiac Perforation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the 70-year-old female patient suffered a left ventricle (lv) perforation when the pigtail perforated the apex of the lv during anastomosis of the distal vessels. the perforation was successfully repaired and support continued as planned. the following day, the physician noted a lack of distal doppler pulse in the right foot. the pump was ultimately explanted later that day. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Pump was not returned for investigation. clinical information provided was not sufficient to investigate. therefore, the root cause of the ventricle perforation and limb ischemia issue was not determined since product was not returned and insufficient clinical information provided. instructions for use: impella cp with smartassist system with smartassist for use during cardiogenic shock section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings and cautions: warnings. section: pre-support evaluation. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist system for use during cardiogenic shock with automated impella controller. section: warnings and cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05456. d6a and d6b revised from the initial submission in accordance with updated procedures. g1 reporting contact fax number missing. h6 component code revised from the initial submission in accordance with updated procedures. h10 instructions for use missing.
32290 | 1220648-2023-05457 | 18352172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18352172&pc=OZD | 11/16/2023 | 03/04/2024 |  | 12/18/2023 |  |  | 11/21/2023 | 12/18/2023 |  |  | 08/31/2025 |  | 12/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024393300 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pulmonary Embolism | Stroke/CVA | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/22/2023 |  |  |  |  | The user facility reported a 57-year-old male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the patient suffered a pulmonary embolism and stroke during impella support. it was noted that there were multiple foci of hypoattenuation involving the right parietotemporal lobe, right medial frontal gyrus, and left frontal lobe. support was maintained with the implanted pump. | [5 empty] | The investigation into the reported stroke/cva has been completed since the original report was filed. as the necessary clinical information or the device was not provided, the root cause of the stroke/cva could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32291 | 1220648-2023-05458 | 18352264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18352264&pc=OZD | 09/27/2023 | 09/19/2024 |  | 12/18/2023 |  |  | 11/22/2023 | 12/18/2023 |  |  | 07/31/2025 |  | 12/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024368128 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/22/2023 |  |  |  |  | The user facility reported a 54-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support developed sepsis and passed away. it is unknown if the impella caused/contributed to the infection. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ | The investigation for sepsis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05458. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8- should have been left blank. g1 reporting contact fax number missing. h6 component code revised from the initial submission in accordance with updated procedures.
32292 | 1220648-2023-05459 | 18352371 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18352371&pc=OZD | 11/22/2023 | 11/25/2025 |  | 12/18/2023 |  |  | 11/22/2023 | 12/18/2023 |  |  |  |  | 12/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1579730 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 | [18 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/22/2023 |  |  |  |  | The user facility reported a 63-year-old female patient was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support when the aic (automated impella controller) had a "controller error". the aic was replaced successfully. | [5 empty] | The impella device was received from the customer and an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. | Data analysis: on the complaint date, during support, the console displayed ¿controller error (loss of communication with keypad or between the kbd and i2c for kbd hw revision 1) ¿ alarm,¿ event #24. the issue was never resolved for the rest of the case (refer to imc_logs_(b)(6).pdf). this confirms the reported failure mode. device analysis: the console was tested upon arrival at fs, and no issue was found with the ribbon cable between the 4-in-1 and the glass keyboard. upon performing the aic 200 power cycle, the issue could not be reproduced. root cause: the root cause for the controller error could not be determined due to the inability to reproduce. d2 common device name updated. d4 model number updated. d10 concomitant medical products and therapy dates updated. f6 and f8 should have been left blank on the initial report. g1 reporting contact email and reporting contact fax number updated.
32293 | 1220648-2023-05460 | 18352383 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18352383&pc=OZD | 11/22/2023 | 10/28/2025 |  | 12/18/2023 |  |  | 11/22/2023 | 12/18/2023 |  |  |  |  | 12/18/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1431034 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Analyze Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/22/2023 |  |  |  |  | The complainant reported that the automated impella controller (aic) failed to recognize and installed purge disc during a routine purge bag exchange. troubleshooting was attempted, but the aic still failed to register the purge disc. the aic was subsequently swapped and support with the implanted pump was maintained. there was no patient harm. | The user facility reported a 61-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the automated impella controller (aic) failed to recognize an installed purge disc during a routine purge bag exchange. troubleshooting was attempted, but the aic still failed to register the purge disc. the aic was subsequently swapped and support with the implanted pump was maintained. there was no patient harm. |  |  |  |  | Investigation was completed. data analysis: the aic log from the complaint date shows that the console (ic8709) received multiple instances of "purge disc not detected - alarm issued" (event set #402), which confirms the reported failure. the log also shows that the user removed the purge cassette and inserting a different one, but the problem persisted. device analysis: the console was tested after arriving in fs. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be dislodged. the purge disc retainer was observed to be not broken, but a small crack is present. no dextrose residue was found. the root cause for the purge disc not detected ¿ alarm was due to the dislodged disk detection flag. before returning this console back to the customer, fs was instructed to perform the following, install new purge retainer and purge flag (and foil) [3100-0256]. validate sensor accuracy via section 12 of the pm procedure (0042-7309). perform a full functional check on the console and optical system (0042-7316). | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
32294 | 1220648-2023-05461 | 18352437 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18352437&pc=OZD | 11/22/2023 | 11/14/2024 |  | 12/18/2023 |  |  | 11/22/2023 | 12/18/2023 |  |  | 08/31/2025 |  | 12/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024384873 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Migration | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/22/2023 |  |  |  |  | The user facility reported a 63-year-old male with cardiomyopathy was implanted with an impella for mechanical circulatory support. the surgeon had difficulty positioning the impella 5.5 post-implant because the pump repeatedly "flipped" on echo toward mitral valve. no suction, reduced flows, or sustained arrhythmias noted due to positioning. during examination, the surgeon noted that the junction between the white plastic of the locking mechanism and the blue hub allowed for rotation of the white plastic locking mechanism. once identified, the surgeon placed a heavy tie around the blue rubber hub to maintain placement. issue was resolved. support continued with the implanted pump following the issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of positioning issues has been completed. the root cause of the positioning issue was not determined since product was not returned and insufficient clinical information provided. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05461. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank.
32295 | 1220648-2023-05462 | 18352441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18352441&pc=OZD | 11/22/2023 | 01/30/2024 |  | 12/18/2023 |  |  | 11/22/2023 | 12/18/2023 |  |  |  |  | 12/18/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1368535 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/22/2023 |  |  |  |  | The complainant reported that an automated impella controller (aic) failed to recognize an installed purge disc. despite troubleshooting attempts, the failure persisted, and the decision was made to swap out the console. there was no patient harm. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. after reviewing the imc logs from the reported occurrence date, the issues with purge disc not being able to be recognized was confirmed. the console was tested after arriving. the issue with the console not being able to detect the purge disc was reproduced. upon examining the automated impella controller (aic) for any visible defects, it was evident that the purge flag was broken. the cause of the issue was a broken purge flag.
32296 | 1220648-2023-05463 | 18352448 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18352448&pc=OZD | 11/22/2023 | 01/18/2024 |  | 12/18/2023 |  |  | 11/22/2023 | 12/18/2023 |  |  | 08/31/2025 |  | 12/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024398777 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Advance | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/22/2023 |  |  |  |  | The user facility reported an 80-year-old male patient was implanted with an impella device for mechanical circulatory support. the us complainant was attempting to deliver a cp to support a patient admitted for a hrpci (high risk percutaneous coronary intervention). the cp failed to deliver and kinked with the attempted placement. the pump was explanted. no harm or injury noted. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported delivery issue has been completed since the original report was filed. no product was returned. the root cause of the delivery issue was most likely patient condition related.
32297 | 1220648-2023-05464 | 18352485 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18352485&pc=OZD | 11/22/2023 | 02/12/2024 |  | 12/18/2023 |  |  | 11/22/2023 | 12/18/2023 |  |  | 08/31/2025 |  | 12/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024395721 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Migration | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/22/2023 |  |  |  |  | The user facility reported an 86-year-old male noted for high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the impella cp pulled back into the aorta and the physician was unable to advance impella into proper position. percutaneous coronary intervention was performed without impella support. the impella was removed. the patient is stable. | [5 empty] | The investigation into the positioning issues has been completed. the device was not returned for investigation. based on the clinical information provided, while performing pci on the mlad the pump was accidentally pulled back into the aorta. an attempt was made to advance the impella back into proper position without success. no other issues or product damage was noted afterwards. log analysis was not applicable for this investigation. the most likely cause of the positioning issue was use related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32298 | 1220648-2023-05465 | 18352526 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18352526&pc=OZD | 11/22/2023 | 01/17/2024 |  | 12/18/2023 |  |  | 11/22/2023 | 12/18/2023 |  |  | 08/31/2025 |  | 12/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024389778 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A patient of unknown age was scheduled to receive hemodynamic support from an impella cp heart pump due to acute myocardial infarction/cardiogenic shock. attempts to insert the impella cp were unsuccessful due to the patient's anatomy. the user was unable to guide the pump through the aortic valve. the patient suffered no harm as a result of the incident and further treatment was carried out without support from an impella pump. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation into the reported delivery issue -anatomy has been completed since the original report was filed. the root cause of delivery issue is determined to be patient condition because of the patient anatomy.
32299 | 1220648-2023-05466 | 18352551 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18352551&pc=OZD | 11/22/2023 | 12/18/2023 |  | 12/18/2023 |  |  | 11/27/2023 | 12/18/2023 |  |  |  |  | 12/18/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, JP |  |  | ABIOMED, INC. |  |  | US |  | 1422863 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 11/27/2023 |  |  |  |  | The user facility reported a 60-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the automated impella controller (aic) did not recognize the purge set. the aic was replaced. there was no patient harm. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
32300 | 1220648-2023-05467 | 18352554 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18352554&pc=OZD | 11/22/2023 | 03/04/2024 |  | 12/18/2023 |  |  | 11/22/2023 | 12/18/2023 |  |  | 09/30/2025 |  | 12/18/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 - EU |  |  | ABIOMED, INC. |  |  | US |  | 2024401188 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/22/2023 |  |  |  |  | The user facility reported a 75-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. during support, there was a pump stop and retrograde flow. restart was not possible. the physician pulled the impella into the aorta and replaced the impella cp. there was no patient harm. | [5 empty] | The investigation into the pump stop issue has been completed. the device was not returned for investigation. the root cause of the pump stop was not determined since product and data logs were not returned and there is a lack of clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
32301 | 1220648-2023-05468 | 18352557 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18352557&pc=OZD | 11/24/2023 | 10/25/2024 |  | 12/18/2023 |  |  | 11/24/2023 | 12/18/2023 |  |  | 08/31/2025 |  | 12/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024394486 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/24/2023 |  |  |  |  | The user facility reported a 70-year-old male of unknown ethnicity and race in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the cp was placed however the team observed signs of hemolysis. the team attempted to troubleshoot with p level adjustment and pump repositioning. the team additionally removed the heparin but, to date did not yet infuse any blood products for the hemolysis. the pump remains on for support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the hemolysis has been completed. no product was returned for analysis. the data logs were reviewed and there were no suction alarms, positioning issues or motor current spikes. additionally, there was no trend in the placement signal. the root cause of the hemolysis cannot be determined due to lack of product returned and insufficient clinical details. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added h6 component code 925 corrections to the initial mfr report: added d6a/d6b f6/f8 should have been left blank.
32302 | 1220648-2023-05469 | 18352558 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18352558&pc=OZD | 12/11/2023 | 01/17/2024 |  | 12/18/2023 |  |  | 12/11/2023 | 12/18/2023 |  |  | 10/20/2025 |  | 12/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024405081 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/11/2023 |  |  |  |  | The user facility reported a 76-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the sheath kinked so motor housing could not pass through the kink. troubleshooting was unsuccessful and the case was aborted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was due to patient condition.
32303 | 1220648-2023-05470 | 18352560 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18352560&pc=OZD | 12/05/2023 | 01/17/2024 |  | 12/18/2023 |  |  | 12/05/2023 | 12/18/2023 |  |  |  |  | 12/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | XXXXXX |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/05/2023 |  |  |  |  | The us complainant had a cp pump (of unknown/undocumented s/n) that failed to deliver. the team was unable to deliver across and through the aorta. the team instead made the decision to implant an iabp (intra-aortic balloon pump) . cp was aborted due to anatomic limitations. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely due to patient condition.
32304 | 1220648-2023-05471 | 18352562 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18352562&pc=OZD | 11/26/2023 | 11/01/2024 |  | 12/18/2023 |  |  | 11/26/2023 | 12/18/2023 |  |  | 10/31/2025 |  | 12/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024406264 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Material Frayed | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/26/2023 |  |  |  |  | A 48 year old male with cardiomyopathy was implanted with impella 5.5 for mechanical circulatory support. during support, the distal contamination sleeve was compromised and was wrapped with gauze and tegaderm. it was labeled do not remove. patient maintained stability throughout implant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of product damage (sterile sleeve damage) has been completed. no product returned. the root cause of the product damage sterile sleeve damage was not determined as no product/images returned for investigation. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05471. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank.
32305 | 1220648-2023-05472 | 18352566 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18352566&pc=OZD | 12/07/2023 | 05/29/2024 |  | 12/18/2023 |  |  | 12/07/2023 | 12/18/2023 |  |  | 03/31/2025 |  | 12/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024344806 |  | 0048-0045 |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/07/2023 |  |  |  |  | The patient's outcome was reviewed by abiomed's medical safety officer (mso) it was determined that there was not enough information to associate use of device with outcome. | The user facility reported a 70-year-old female was implanted with an impella device for mechanical circulatory support. the us complainant had a cp placed for support of a patient admitted in cgs/ami (cardiogenic shock / acute myocardial infarction) . the pump was placed femorally and the peel-away was removed. the patient developed a hematoma at that site which prompted intervention. manual hold and 2 units of rbcs were given. the pump remained on for support until md spoke to family and the patient expired. |  |  |  |  | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use reported in the initial report is from ifu for impella cp with smart assist system ,section: potential adverse events (united states), which now includes statement "cardiac or vascular injury (including ventricular perforation).¿ added-b5 mso review d4-updated model number added- d6a/d6b. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32306 | 1220648-2023-05473 | 18352569 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18352569&pc=OZD | 11/27/2023 | 04/05/2024 |  | 12/18/2023 |  |  | 11/27/2023 | 12/18/2023 |  |  | 09/30/2025 |  | 12/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024400666 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/27/2023 |  |  |  |  | The user facility reported a 56-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the pump was cross-clamped. this clamping in surgical mode caused/contributed to the pump stopping. the clamp was adjusted and pump restarted. post coronary artery bypass graft when the patient woke up he revealed that he was not able to move his left side. the head scan was positive for stroke. the pump was explanted without any issue, interventional radiology performed thrombectomy. patient is stable post intervention. | [5 empty] | The investigation into the pump stop and the stroke/cva issues has been completed since the original report was filed. the pump was returned, tested, and evaluated. it was found biomaterial wrapped around the impeller. the cause of the pump stop issue was biomaterial. as the necessary clinical information was not provided, the root cause of the stroke/cva could not be determined. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
32307 | 1220648-2023-05474 | 18352570 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18352570&pc=OZD | 11/21/2023 | 02/05/2024 |  | 12/18/2023 |  |  | 11/21/2023 | 12/18/2023 |  |  | 09/30/2025 |  | 12/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024405268 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | Migration | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/21/2023 |  |  |  |  | The user facility reported a 62-year-old male with cardiomyopathy was implanted with an impella for mechanical circulatory support. the 5.5 had issues with the delivery. the team used many troubleshooting measures to get the pump delivered to the left ventricle. the pump was supporting for 3 days when the pump was explanted with concerns for flow saturation and abnormal readings on the automated impella controller (aic). patient stable after explant awaiting the new heart. no patient harm reported due to the issue. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the delivery issue and saturated flow has been completed. the impella device and data logs were returned for evaluation. visual inspection found biomaterial within the cannula and interacting with the impeller. a cannula kink was also observed near the outflow cage. the logs confirmed that the false pump in aorta alarms triggered after a motor current spike. saturated flows followed afterwards as well. the motor current spike is evident of biomaterial ingestion. the cause of the saturated flow is biomaterial. false impella in aorta alarms can be triggered when a clot is ingested. according to the clinical, the patient had small vessel size. the product was returned and visually inspected. no abnormalities related to delivery issue were found. the cause of the delivery issue is patient condition.
32308 | 1220648-2023-05475 | 18352588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18352588&pc=OZD | 11/21/2023 | 02/10/2025 |  | 12/18/2023 |  |  | 11/21/2023 | 12/18/2023 |  |  |  |  | 12/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1194346 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/21/2023 |  |  |  |  | The user facility reported a 63-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the automated impella controller (aic) soft touch button wouldn't respond during pump prime. the aic was replaced successfully. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - kbd/rotary knob issues has been completed. the logs were reviewed and confirmed that there were no issues during pump priming, as well as there were no issues with keyboard communication. no alarms were observed. the console was tested, and the keyboard was confirmed to be working correctly. the reported complaint could not be reproduced. the following have been reported incorrectly or missing from the manufacturer device report 1220648-2023-05475. d2 common device name revised in accordance with updated procedures. d4 model number. f6/f8-should have been left blank. g1 reporting contact email revised in accordance with updated procedures. g1 reporting contact fax number missing.
32309 | 1220648-2023-05476 | 18352592 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18352592&pc=OZD | 12/08/2023 | 02/22/2024 |  | 12/18/2023 |  |  | 12/08/2023 | 12/18/2023 |  |  | 09/30/2025 |  | 12/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024410769 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Consumer | Health Professional |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Disability | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/08/2023 |  |  |  |  | The user facility reported a 60-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient on ecmo (extracorporeal membrane oxygenation) and impella cp support is experiencing limb ischemia. the patient was taken to the operation room for an emergency fasciotomy and antegrade sheath was placed with no improvement. perfusion to the left lower extremity (lle) with arterial cannula and right upper extremity (rue) with brachial arterial line was also severely compromised. the patient is in critical condition. | [5 empty] | The investigation into the limb ischemia irreversible issue has been completed. the device was not returned for investigation. the root cause of the limb ischemia was not determined as the necessary clinical information was not provided and the device was not returned. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32310 | 1220648-2023-05477 | 18352599 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18352599&pc=OZD | 12/12/2023 | 11/01/2024 |  | 12/18/2023 |  |  | 12/12/2023 | 12/18/2023 |  |  | 09/30/2025 |  | 12/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024406514 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Consumer | Health Professional |  |  |  | 59 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/12/2023 |  |  |  |  | The user facility reported a 59-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the connection of the yellow purge tubing and the cassette was leaking. there was no resolution from changing the purge tubing. the clinician will place a piece of gauze around the tubing connection and change it hourly. it was further reported that the sterile sleeve distal attachment was broken, and the sleeve was hanging loose. a sterile tegaderm was applied to hold the connection of the sleeve in place. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure. reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation for the product damage and the purge leak has been completed. no product returned for investigation. clinical reports the connection of the yellow purge tubing and the cassette was leaking and the sterile sleeve was detached. the root cause of the product damage sterile sleeve damage was not determined because product/images were not returned for review. the root cause of the purge leak- pump was not determined because product was not returned for investigation to verify the leak or location. corrections to the initial mfr report: f6/f8 should have been left blank.
32311 | 1220648-2023-05478 | 18352605 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18352605&pc=OZD | 11/21/2023 | 02/29/2024 |  | 12/18/2023 |  |  | 11/21/2023 | 12/18/2023 |  |  | 08/31/2025 |  | 12/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024394460 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had impella cp support ongoing for a patient with a stemi. the team placed a cp for hemodynamic support via the femoral artery. the site developed a hematoma. the hematoma required red blood cell (rbc) infusion and vascular surgeon to step in and suture. the pump remained on despite the bleed for 2 days when the patient expired. there was no allegation that the cp caused or contributed to the death. patient was maxed on pressors and discussing care escalation beyond the cp. | [5 empty] | The investigation into the reported access site bleeding -major has been completed since the original report was filed. product was not returned for investigation. as per the clinical details, bleeding occurred during impella access as a hematoma and site ooze developed while suturing the repositioned sheath in place. the cause of the access site bleeding issue was most likely use related, based on given clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32312 | 1220648-2023-05398 | 18334772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18334772&pc=OZD | 11/15/2023 | 10/03/2024 |  | 12/15/2023 |  |  | 11/15/2023 | 12/15/2023 |  |  | 08/31/2025 |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC |  |  | US |  | 2024389551 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported patient of unknown age on impella cp for hemodynamic support was bleeding from the access site. the repositioning sheath was reported to not be correct. the physician fixed it and the bleeding stopped. support was maintained with the implanted pump. | The patient expired while on support. per medical safety officer review use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was most likely use related; the repositioning sheet was not correct/completely introduced as per the clinical description provided. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05398: b5 patient outcome and mso statement added. d3 manufacturer name, address, city and state, country, and fax number. d6a and d6b revised from the initial submission in accordance with updated procedures. h6 component code revised from the initial submission in accordance with updated procedures.
32313 | 1220648-2023-05359 | 18334836 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18334836&pc=OZD | 11/15/2023 | 11/19/2025 |  | 12/15/2023 |  |  | 11/15/2023 | 12/15/2023 |  |  |  |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1169580 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | Unexpected Shutdown | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/15/2023 |  |  |  |  | The complainant reported an aic (automated impella controller) shut off, and there was a blank screen saying there was a controller error. after, the controller started, but then there was a controller error that stated to use a back up controller and the aic was replaced. | [5 empty] | The impella device was received from the customer and an evaluation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. | The investigation has been completed. data analysis: imc logs confirmed the reported failure mode given there was an instance of a controller error alarm (opm comm crc error ¿ event set # 1045) triggered during the clinical procedure. real time (rt) logs confirmed that signals received from optical bench ( snr, contrast, gain) are represented as -16384 values due to the crc error. although it was reported that ic3399 shutoff, there were no instances of an unexpected reboot or shutdown observed in the logs. device analysis: the console was operated to specification when a known good pump and purge line was run with it. conclusion: the controller error was due to an optical bench fw communication protocol deficiency, resulting in misalignment of data communication packets received by epc. console sw operated as designed. subsequent preventative maintenance performed on the console are sufficient to ensure that the console works to specification. d4: corrected the catalog number. g1: corrected the manufacturer email address. h3: added information regarding the investigation status. h6: added codes per the results of the investigation. h10: added the results of the investigation.
32314 | 1220648-2023-05399 | 18334972 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18334972&pc=OZD | 11/16/2023 | 03/06/2024 |  | 12/15/2023 |  |  | 11/16/2023 | 12/15/2023 |  |  | 08/31/2025 |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024388061 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 67-year-old male patient was attempted to be implanted with an impella cp smartassist heart pump. however, the procedure had to be aborted due to the aortic loop anatomy. there was no reported patient harm. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was returned for evaluation. visual inspection of the pump found the cannula kinked near the outflow cage. no other abnormalities were observed. additionally, clinical details provided contributed to determining the root cause. therefore, the the cause of the delivery issue was patient condition related as clinical details states the insertion was unsuccessful due to aortic loop anatomy. d3-updated mfr name, city, state
32315 | 1220648-2023-05400 | 18334995 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18334995&pc=OZD | 11/17/2023 | 10/30/2024 |  | 12/15/2023 |  |  | 11/17/2023 | 12/15/2023 |  |  | 09/30/2025 |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2024405894 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/17/2023 |  |  |  |  | The user facility reported a 66-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed hemolysis during impella cp support. the p-level was adjusted in an attempt to troubleshoot the issue, but a definitive resolution was not confirmed. support with the implanted pump remains ongoing. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the hemolysis has been completed. pump was not returned for investigation. data logs were investigated. there were no suction alarms seen, motor current spikes or a trend in the placement signal. it was stated in the clinical that the p-level was reduced for support but no blood products were given to the patient. the logs show that the p-level was reduced because the flows were not optimal. however, the clinical information and data logs are inconclusive. therefore, the root cause of the hemolysis issue was not determined since product was not returned and data logs were inconclusive. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- h6 component code 925 corrections to the initial mfr report: d4-model number added- d6a/d6b f6/f8 should have been left blank.
32316 | 1220648-2023-05401 | 18335261 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18335261&pc=OZD | 11/17/2023 | 11/17/2025 |  | 12/15/2023 |  |  | 11/17/2023 | 12/15/2023 |  |  |  |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1232458 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/17/2023 |  |  |  |  | The complainant reported that upon powering on their aic (automated impella controller) to check for updates, the console displayed a system self check failure. there was no patient involvement at the time of the noted issue. | [5 empty] | The impella device was received from the customer and an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed.
32317 | 1220648-2023-05402 | 18335488 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18335488&pc=OZD | 11/17/2023 | 12/09/2024 |  | 12/15/2023 |  |  | 11/17/2023 | 12/15/2023 |  |  | 10/31/2025 |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024406433 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Cardiac Perforation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had impella 5.5 pump delivered to escalate care from an impella cp. the delivery was complicated and multiple wires and catheters had to be utilized to access the left ventricle. the site developed a hematoma and was treated by a jp drain placement. the patient additionally noted an effusion after the 5.5 was delivered. the team also performed a pericardiocentesis. support with the implanted pump was maintained. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the vascular damage and the ventricle perforation has been completed. no product was returned for investigation. data logs show impella position unknown alarms at the start of the case and some ventricular waveforms were noted at the beginning of support. the root cause of the ventricle perforation was most likely due to improper positioning of the device. the root cause of the vascular damage - great vessel was most likely use related due to the numerous wire and catheter exchanges. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).
32318 | 1220648-2023-05403 | 18335545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18335545&pc=OZD | 11/16/2023 | 10/30/2024 |  | 12/15/2023 |  |  | 11/16/2023 | 12/15/2023 |  |  | 09/30/2025 |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024399661 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/16/2023 |  |  |  |  | The user facility reported a 67-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. when the impella was explanted, the decision was made to take the patient to the have vascular repair due to noted right femoral artery ischemia. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation for the limb ischemia and vascular damage has been completed. clinical details states that the patient was noted with right femoral artery (rfa) ischemia that required vascular repair as the patient had tortuous vessels. the root cause of the vascular damage - peripheral vessels is most likely due to the patients condition. the root cause of the limb ischemia could not be determined without additional clinical details. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- h6 component code 925 corrections to the initial mfr report: added d6a/d6b f6/f8 should have been left blank.
32319 | 1220648-2023-05404 | 18335575 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18335575&pc=OZD | 11/17/2023 | 01/29/2024 |  | 12/15/2023 |  |  | 11/17/2023 | 12/15/2023 |  |  |  |  | 12/15/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1084304 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Unexpected Shutdown | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/17/2023 |  |  |  |  | The complainant reported that the automated impella controller (aic ) shutdown unexpectedly during patient support. battery still had full charge. aic was removed from service. no patient harm reported due to this issue. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the aic - shutdown or restart issue has been completed since the original report was filed. the console was returned for investigation, but the failure mode seen on 11/16 was not reproduced. the pbm was examined for any contamination that could have impacted communication, but none was found. additionally multiple attempts to run a pump and modulate the p-level did not reproduce the failure mode nor point to what caused the initial unexpected restart. the cause of the issue was not determined since the failure mode was not reproduced.
32320 | 1220648-2023-05405 | 18335650 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18335650&pc=OZD | 11/17/2023 | 02/28/2024 |  | 12/15/2023 |  |  | 11/17/2023 | 12/15/2023 |  |  | 04/30/2025 |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024352929 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/17/2023 |  |  |  |  | The user facility reported a 65-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient had no distal flow in the lower left extremity following impella cp implant. due to the noted condition, the decision was made to place an external fem-fem bypass. support with the implanted pump continued following the noted intervention. | [5 empty] | The investigation into the limb ischemia ¿ reversible has been completed. the device and data were not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/pos, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32321 | 1220648-2023-05406 | 18335662 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18335662&pc=OZD | 11/17/2023 | 02/02/2024 |  | 12/15/2023 |  |  | 11/17/2023 | 12/15/2023 |  |  | 08/31/2025 |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024393138 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/17/2023 |  |  |  |  | The user facility reported a 71-year-old female patient was implanted with an impella device for mechanical circulatory support. the us complainant had a cp support ongoing for a patient with multiple cardiac and systemic comorbidities. the pump was placed via femoral artery access and this caused the loss of pulses. the limb ischemia caused the team to explant and escalate to the axillary 5.5 pump. the access site needed thrombus treatment and a cutdown. the 5.5 was placed and patient survived. | [5 empty] | The investigation into the reported, limb ischemia has been completed since the original report was filed. no product was returned. as per clinical patient the cause of the limb ischemia issue was patient condition related with clot in the femoral artery observed per imaging. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32322 | 1220648-2023-05407 | 18335717 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18335717&pc=OZD | 11/17/2023 | 11/08/2024 |  | 12/15/2023 |  |  | 11/17/2023 | 12/15/2023 |  |  | 09/30/2025 |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024397106 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 67-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient had reports of hemolysis while on impella support. patient had continuous renal replacement therapy and the patient was supported by impella and ecmo. clinical representative confirmed that the hemolysis was likely caused by all supporting devices. the pump was explanted the following day after the patient showed no signs of conditional improvement. | It was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the hemolysis has been completed. pump was not returned for investigation. data logs were returned and there was no positioning alarms observed. there were no suction alarms or motor current spikes as well. the data logs are inconclusive of any evidence. without the pump, the hemolysis cannot be investigated. therefore, the root cause of the hemolysis issue was not determined since product was not returned and insufficient clinical information provided. added- h6 component code 925 corrections to the initial mfr report: b2-selected death and added date of death. b5-mso review. added d6a/d6b. h1-type of reportable event changed to death. h6 impact code 1802.
32323 | 1220648-2023-05408 | 18335770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18335770&pc=OZD | 11/17/2023 | 04/03/2024 |  | 12/15/2023 |  |  | 11/17/2023 | 12/15/2023 |  |  | 09/30/2025 |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024405887 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 pump insertion had to be aborted due to the patient's anatomy. a second impella 5.5 pump was later attempted and inserted successfully. there was no patient harm resulting from the noted issue. | [5 empty] | The investigation into the delivery issues- anatomy issue has been completed since the initial report was submitted. the product was not returned for investigation. the root cause of the delivery issue was not determined since insufficient clinical information is provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
32324 | 1220648-2023-05349 | 18335772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18335772&pc=OZD | 11/14/2023 | 01/12/2024 |  | 12/15/2023 |  |  | 11/14/2023 | 12/15/2023 |  |  |  |  | 12/15/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1207023 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/14/2023 |  |  |  |  | The complainant reported that when attempting to insert a purge disc into their (automated impella controller) aic, the grey holder was found sheared off from the console and loose in the purge area. the aic was replaced as a result of the noted damage. there was no patient harm. | [5 empty] | The impella device was received from the customer and an evaluation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. | The investigation for the console purge disc/flag issues has been completed. the console was returned for evaluation. the data logs were not relevant to the reported event. the console was visually inspected and confirmed that the purge retainer was broken. according to clinical details, the clinical staff reported that the ¿holder¿ was sheared off the console they meant the purge retainer. therefore, the root cause of the reported issue was due to a broken retainer.
32325 | 1220648-2023-05347 | 18335773 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18335773&pc=OZD | 11/13/2023 | 12/04/2025 |  | 12/15/2023 |  |  | 11/13/2023 | 12/15/2023 |  |  |  |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1245097 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/13/2023 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) had a red alarm, controller failure. the unit was removed from service to be returned for evaluation. there was no patient involvement. | [5 empty] | Data analysis: the case associated with the reported complaint was performed on october 12, 2023. due to a pmd communication issue, the console rebooted automatically. after the reboot, the console displayed the ¿system self check failed¿ screen (imc_logs_ic4399.pdf) due to the persistent pmd communication issue. this was reported as a ¿total system failure.¿ device analysis: the console was tested upon arrival at fs, and the reported failure mode was reproduced. fwcheck via telnet showed the sau_motor_1 hardware revision as ¿n.a." after replacing the original impellatronic board with a known good one, the console booted up as intended. root cause: the root cause of the ¿system self-check failed¿ screen was determined to be an impellatronic malfunction. there was no evidence in the log to confirm the ¿aic - controller failure (red alarm).¿ the ¿system self-check failed¿ screen was misreported as ¿aic - controller failure (red alarm).¿ d2 common device name updated. d4 model number updated. f6 and f8 should have been left blank in the initial report. g1 reporting contact email updated. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
32326 | 1220648-2023-05346 | 18335779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18335779&pc=OZD | 11/13/2023 | 06/20/2024 |  | 12/15/2023 |  |  | 11/13/2023 | 12/15/2023 |  |  | 09/30/2025 |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024399456 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Myocardial Infarction | Ventricular Fibrillation | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 64-year-old male noted as high risk for coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient went into cardiogenic shock during impella cp support. it was noted that the patient's condition developed after an anterior myocardial infarction and multiple defibrillations. there were no noted malfunction with the pump at the time of the reported incident. despite this, the staff believed that the pump could have contributed to the noted condition. the patient was able to recover with native heart function. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the vf/vt/af/at (v-fib) and the access site bleeding - major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined. d6a, d6b
32327 | 1220648-2023-05348 | 18335781 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18335781&pc=OZD | 11/14/2023 | 09/20/2024 |  | 12/15/2023 |  |  | 11/14/2023 | 12/15/2023 |  |  | 09/30/2025 |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024402461 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/14/2023 |  |  |  |  | The user facility reported a 45-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed some oozing at their access site following impella implant. manual pressure was initially applied to the site; however, the site worsened after the patient was moved. after an hour of manual pressure, the hematoma that developed was expressed and the site was injected with lido/epi. when the bleeding still continued, the decision was ultimately made to remove the pump. a new pump was then placed for planned support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for access site bleeding has been completed. the impella device was not returned for evaluation. the cause of the issue was high patient activated clotting time (act) values than recommended during the use of impella. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05348. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact fax number missing. h6 component code revised from the initial submission in accordance with updated procedures.
32328 | 1220648-2023-05344 | 18335782 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18335782&pc=OZD | 11/13/2023 | 08/01/2024 |  | 12/15/2023 |  |  | 11/13/2023 | 12/15/2023 |  |  | 03/31/2025 |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024327069 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Material Twisted/Bent | [5 empty] | Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the 60-year-old female patient suffered a left ventricle perforation during impella 5.5 support. it was stated that the physician was unable to cross clamp during a scheduled coronary artery bypass graft (cabg) and pushed the impella forward to allow for proper cross clamping. as a result, the pump perforated the left ventricle. the pump was pulled back and the perforation was surgically closed without issue. following the intervention, repositioning of the pump was attempted, but the catheter kinked and could no longer advance. as such, the pump was removed and replaced with another impella 5.5 pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation of positioning issues and ventricle perforation has been completed. the impella device was not returned for evaluation. the root cause of the ventricle perforation issue was related to the patient's condition with a diseased left ventricle, based on the given clinical details. the root cause of the positioning issue was most likely due to use based on the clinical details provided. instructions for use as follows: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings and cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance d6a/d6b added the following have been reported incorrectly or missing from manufacturer device report: in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. g1 reporting contact fax number missing.
32329 | 1220648-2023-05345 | 18335783 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18335783&pc=OZD | 11/13/2023 | 12/03/2024 |  | 12/15/2023 |  |  | 11/13/2023 | 12/15/2023 |  |  | 06/30/2025 |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024357279 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/13/2023 |  |  |  |  | The user facility reported a 58-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced persistent hemolysis during impella cp support. repositioning was performed following the initial diagnosis, but the condition persisted into the following day. plans were subsequently made to wean the patient off the pump. the device was explanted two days later. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05345. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. h6 component code revised from the initial submission in accordance with updated procedures.
32330 | 1220648-2023-05409 | 18336001 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18336001&pc=OZD | 11/18/2023 | 02/29/2024 |  | 12/15/2023 |  |  | 11/18/2023 | 12/15/2023 |  |  | 05/31/2025 |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024369050 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  |  |  | Injury | Yes | Yes | Fluid/Blood Leak | Material Split, Cut or Torn | [5 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/18/2023 |  |  |  |  | The user facility reported a 53-year-old male (race unknown) undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. blood was leaking through the repositioning sheath of the impella cp. the impella was replaced with a new impella cp. | [5 empty] | The investigation into the reported access site bleeding -major has been completed since the original report was filed. pump was returned for investigation. no physical damage were observed on reposition sheath. as per clinical notes, blood was leaking through reposition sheath and pump was replaced with new impella cp. the cause of the access site bleeding issue was not determined since sufficient clinical information was not provided. correction : section b1 and b2 are updated as they were inadvertently marked incorrect in the initial report. section h6: the clinical code has been updated as it was inadvertently coded incorrect in the initial report. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
32331 | 1220648-2023-05410 | 18336034 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18336034&pc=OZD | 11/22/2023 | 12/15/2023 |  | 12/15/2023 |  |  | 11/22/2023 | 12/15/2023 |  |  |  |  | 12/15/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1347548 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 11/22/2023 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) had a purge disc failure. the team effectively troubleshot by replacement of the console. no patient involvement. | [5 empty] | The investigation into the automated impella controller (aic) purge disc/flag issues has been completed. the console was tested after arriving in field service. the purge disc retainer was noted to be raised. the failure mode was reproduced which makes sense because if the purge retainer is lifted off, the purge disc will not make good contact with the purge flag so the console will not be able to recognize the purge disc. the cause of the issue was a broken purge retainer.
32332 | 1220648-2023-05411 | 18336038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18336038&pc=OZD | 11/18/2023 | 10/03/2024 |  | 12/15/2023 |  |  | 11/18/2023 | 12/15/2023 |  |  | 09/30/2025 |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024403001 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/18/2023 |  |  |  |  | The user facility reported a 71-year-old male patient was implanted with an impella device for mechanical circulatory support. the complainant reported that due to no dopplerable pulses and the need to maintain support, the impella cp was removed overnight, and the patient was escalated to the impella 5.5. | It was further reported that the patient expired sometime after the procedure. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added-h6 component code 925 corrections to the initial mfr report: b5 mso review. added d6a/d6b. f6/f8 should have been left blank.
32333 | 1220648-2023-05412 | 18336119 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18336119&pc=OZD | 11/17/2023 | 12/04/2025 |  | 12/15/2023 |  |  | 11/17/2023 | 12/15/2023 |  |  |  |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1266273 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Use of Device Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/17/2023 |  |  |  |  | The medical center reported the patient was on impella support and automated impella controller (aic) had a broken purge clip. the aic was replaced and the console was sent back for investigation. | [5 empty] | Console returned to manufacturer and is pending results of evaluation at this time. upon investigation closure, a supplemental mdr will be filed. | Data analysis: imc log is extracted and reviewed, but it did not provided any additional information as the issue was clearly evident during device inspection and consistent with complaint description. device analysis: the console was tested after arriving in fs. the purge disc retainer was confirmed to be broken, and the purge flag was missed. the purge motor shaft was not immobilized. the cause of the issue was a defective component: broken purge retainer d2 common device name updated. d4 model number updated. f6 and f8 should have been left blank in the initial report. g1 reporting contact email and reporting contact fax number updated.
32334 | 1220648-2023-05413 | 18336403 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18336403&pc=OZD | 11/20/2023 | 12/09/2024 |  | 12/15/2023 |  |  | 11/20/2023 | 12/15/2023 |  |  | 08/31/2025 |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024392968 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Remove | Detachment of Device or Device Component | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/20/2023 |  |  |  |  | The user facility reported a 64-year-old male post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella for mechanical circulatory support. upon removal of the impella 5.5, sentinel device was placed to protect from cerebral embolization due to clot on impella seen on echo. md began to pull impella out of left ventricle and into the graft and resistance was initially met on first attempt. md then rotated the catheter and pulled again, and catheter came all the way back to the blue suture hub, but catheter broke off at the motor housing at this point. under fluoro the impella was seen up by the clavicles at the proximal end of the graft. the impella was secured with a clamp, control of the graft was obtained for hemostasis. the blue hub was removed and impella was pulled out of graft successfully. the graft was then stapled, and the incision closed. the patient tolerated the removal well and hemodynamically is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the vascular damage, product damage, and thrombus has been completed. pump was not returned for investigation. no images were returned as well. based off the clinical details, the root cause of the product damage - detached inflow/outflow - peripheral vessels is most likely due to use as resistance was met in the graft his the removal issues. the root cause of the vascular damage - peripheral vessels is most likely due to use. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).
32335 | 1220648-2023-05414 | 18336786 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18336786&pc=OZD | 11/03/2023 | 02/10/2025 |  | 12/15/2023 |  |  | 11/17/2023 | 12/15/2023 |  |  | 07/31/2025 |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024379790 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Encephalopathy | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Sepsis | Cardiac Tamponade | Cardiogenic Shock | Thrombocytopenia | Unspecified Hepatic or Biliary Problem |  |  |  | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | Followup | 11/17/2023 |  |  |  |  | New information alleges hemolysis and hepatic failure. for hemolysis, the reported narrative indicates ldh and pfhg labs were trending upwards, resulting in removal of the device on (b)(6) 2023. regarding hepatic failure, the narrative states, "trend for alt/ast trend upward. lactulose started for ischemic hepatopathy. ctm and watch lab trends." each of of the above adverse events were alleged with a "possible" relationship to the impella device. | A notification was received via abiomed's long-term outcome and quality indicator impella registry on (b)(6) 2023. the report alleged renal failure, cardiac tamponade, and ventricular arrhythmia had each occurred with a possible relationship to the impella 5.5 device used on this patient. renal failure was treated via dialysis and ventricular arrhythmia was treated via two (2) cardioversions. while cardiac tamponade was alleged, and blood products were delivered, it's unclear if the physician "tamponaded" the patient's procedural chest bleed, or an actual cardiac tamponade occurred. we are actively following up for more information. | New information alleges sepsis and acute encephalopathy with a possible relationship to the impelladevice. for sepsis, the reported narrative indicates the patient's white blood cell count had trended upwards after the procedure and antibiotics were required. regarding acute encephalopathy, the narrative states, "after impella implant, patient developed acute encephalopathy. eeg - w/o seizure (b)(6) 2023 and ct scan negative. patient's cause of death was multi-organ failure." | New information alleges cardiogenic shock and hemolysis with a possible relationship to the impella device. for cardiogenic shock, the reported narrative indicates the patient developed said condition after impella insertion. epi was initiated and patient's lactate trends were monitored. regarding hemolysis, the patient's ldh and pfhg were trending upwards, leading to device removal on (b)(6) 2023. | New information alleges bleeding type 3b cardiac tamponade and thrombocytopenia. for bleeding, the reported narrative indicates a small area of bleeding developed at the takeoff of the surgical graft to "innominate and l common carotid arteries, found upon cardiac tamponade surgery. this was repaired with several 3/0 pledgeted prolene sutures. 1 unit of prbc was documented given also." regarding thrombocytopenia, the patient's labs showed, "trending downward platelets. hit panel was negative and asa was held." | According to information received from abiomed's long-term outcome and quality indicator impella registry the patient expired on an unknown date due to multi-organ failure. per medical safety officer review there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | Added new information to b(5) via loqi registry. the investigation is still in progress. a supplemental will be filed upon completion. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | Added new information to b(5) via loqi registry. the investigation is still in progress. a supplemental will be filed upon completion. | Added new information to b(5) via loqi registry. the investigation is still in progress. a supplemental will be filed upon completion. | Since the original report, our investigation is still in progress. upon closure, a supplemental mdr will be filed. new information received has been added to b(5) and additional codes to section h(6). | The investigation of hemolysis, ventricle perforation, vf/vt/af/at, access site bleeding, thrombocytopenia, and infection/sepsis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. access site bleeding failure deleted since it is a chain event of ventricle perforation. instructions for use: sepsis impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ thrombocytopenia impella 5.5 for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ ventricle perforation impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella 5.5 with smartassist for use during cardiogenic shock with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ hemolysis impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05414. b2 outcome; death. b5 updated with mso statement and patient outcome. d4 model number d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact email revised in accordance with updated procedures g1 reporting contact facility name missing g1 reporting contact fax number missing. h1 type of reportable event h6 health effect - impact code death added h6 component code revised from the initial submission in accordance with updated procedures. h10 instructions for use missing.
32336 | 1220648-2023-05415 | 18336805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18336805&pc=OZD | 11/17/2023 | 03/25/2024 |  | 12/15/2023 |  |  | 11/17/2023 | 12/15/2023 |  |  | 09/30/2023 |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024404540 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/17/2023 |  |  |  |  | The user facility reported a 56-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. bleeding at groin site with a hematoma noted. impella was explanted. patient is stable post explant. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32337 | 1220648-2023-05416 | 18337027 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18337027&pc=OZD | 11/19/2023 | 02/08/2024 |  | 12/15/2023 |  |  | 11/19/2023 | 12/15/2023 |  |  | 09/30/2025 |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2024400381 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/19/2023 |  |  |  |  | The user facility reported a 64-year-old male patient was implanted with an impella device for mechanical circulatory support. us complainant reported the impella cp had a pump stop and the impella was removed. the patient was escalated to 5.5 support. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop issue has been completed since the original report was filed. it was reported that impella cp pump stopped suddenly with no active alarms leading up to pump stop after 18 days of use. failure was observed on day 18 of use, which is beyond the 8-day design life per design trace matrix (b)(4). the pump was returned for evaluation, and a large purge gap was observed. data logs confirm no alarms leading up to a sudden high motor current pump stop. the root cause of the pump stop was bearing wear due to use beyond indicated design life.
32338 | 1220648-2023-05417 | 18337163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18337163&pc=OZD | 11/20/2023 | 09/21/2024 |  | 12/15/2023 |  |  | 11/20/2023 | 12/15/2023 |  |  | 09/30/2025 |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024400493 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/20/2023 |  |  |  |  | The user facility reported a 43-year-old male patient was implanted with an impella device for mechanical circulatory support. the complainant reported that an implanted pump experienced positioning issues during patient support. it was noted that the device was mispositioned in the papillary and left ventricle cavitation. upon attempting to reposition the pump, the device came across the valve and could not be re-advanced. the pump was subsequently left in the descending aorta at p0 until it could be explanted. there was no patient harm. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. |  |  |  |  | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-05417 was provided. note: corrected information provided in h6 is an addition to previous coding.
32339 | 1220648-2023-05418 | 18337572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18337572&pc=OZD | 11/20/2023 | 01/29/2025 |  | 12/15/2023 |  |  | 11/20/2023 | 12/15/2023 |  |  | 06/30/2025 |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024376392 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/20/2023 |  |  |  |  | The user facility reported a 75-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed hemolysis due to improper positioning during impella cp support. attempts were made to reposition the pump, but the device would not stay out of the mitral valve. as a result of the persistent issue, the decision was made to escalate to impella 5.5 support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the hemolysis and the positioning issue has been completed. pump was not returned for investigation. data logs were investigated. data logs showed multiple position in aorta alarms. clinical also mentioned that pump's inlet was found in the mitral valve. repositioning was not successful and after pump explant, hemolysis started resolving as per clinical. therefore, the root cause of the hemolysis issue was improper positioning of the device du to improper technique. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrections to the initial mfr report: added d6a/d6b. f6/f8 should have been left blank. g1 mdr reporting contact email.
32340 | 1220648-2023-05419 | 18338037 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18338037&pc=OZD | 12/14/2023 | 02/22/2024 |  | 12/15/2023 |  |  | 12/14/2023 | 12/15/2023 |  |  | 07/31/2025 |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024390720 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Disability | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/14/2023 |  |  |  |  | The team had implanted an impella cp via the femoral access to support a patient admitted with cgs/ami (cardiogenic shock / acute myocardial infarction) and underlying cardiac and systemic comorbidities, like lower extremity edema and cad (coronary artery disease). the pump supported the patient for 12 days when then explanted. at the time of explant the limb became ischemic and vascular was consulted to intervene. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ irreversible issue has been completed. the device was not returned for investigation. the root cause of the limb ischemia was not determined as the necessary clinical information was not provided and the device was not returned.
32341 | 1220648-2023-05420 | 18338877 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18338877&pc=OZD | 11/24/2023 | 02/22/2024 |  | 12/15/2023 |  |  | 11/24/2023 | 12/15/2023 |  |  | 09/30/2025 |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024403689 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/24/2023 |  |  |  |  | The user facility reported a 56-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support post mitral valve repair. the patient lost pulses in the limb at explant. the vascular surgical team performed surgical cutdown to restore blood flow. the patient was escalated to an impella 5.5. | [5 empty] | The investigation into the limb ischemia irreversible issue has been completed. the device was not returned for investigation. the root cause of the limb ischemia was not determined as the necessary clinical information was not provided and the device was not returned. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32342 | 1220648-2023-05421 | 18340348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18340348&pc=OZD | 11/18/2023 | 11/05/2024 |  | 12/15/2023 |  |  | 11/18/2023 | 12/15/2023 |  |  | 03/31/2025 |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024329686 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/18/2023 |  |  |  |  | The us complainant reported that the 61-year-old male patient was on impella cp support. the patient reportedly had hemolysis and the impella was repositioned. the impella was pulled back to achieve 3.5cm for left ventricle measurement. the hemolysis started to clear. the hemolysis was to be also treated with continuous veno-venous hemofiltration (cvvh) and explant after weaning. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the hemolysis has been completed. pump was not returned for investigation. data logs were investigated and no motor current spike or position issues observed. suction was present. clinical mentioned that pump was seated at 5.5 cm in the ventricle but no alarms were observed. no other information provided. data logs were inconclusive. therefore, the root cause of the hemolysis issue was not determined since product was not returned and data logs were inconclusive. the instructions for use for the impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions including catheter position, pre-existing medical conditions, and small left ventricular volumes may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ added- h6 component code 925 corrections to the initial mfr report: added d6a/d6b f6/f8 should have been left blank.
32343 | 1220648-2023-05422 | 18340588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18340588&pc=OZD | 11/15/2023 | 07/01/2024 |  | 12/15/2023 |  |  | 11/15/2023 | 12/15/2023 |  |  |  |  | 12/15/2023 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1554466 |  | 0042-0010 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that there was not enough information to associate use of the device with the outcome. | The complainant reported the automated impella controller (aic) could not recognize the purge cartridge. another console and another cassette was successfully used. |  |  |  |  | The investigation for the reported console (aic) has been completed. data logs were reviewed and the issue with the aic not being able to detect the purge disc was confirmed. there was a single instance of purge disc not detected alarm found in the data logs on the reported occurrence date. issue was quickly resolved after inserting the purge disc. suspect that the user ran the case start wizard and waited too long to insert the purge disc which prompted the purge disc not detected alarm. the alarm cleared after the user inserted the purge disc but because a red alarm was issued, it is suspected that the user reported this occurrence as a precaution. the console was returned for evaluation and the issue detecting purge disc was unable to be reproduced. purge disc flag and retainer did not appear to have any issues. no visible issues found with the purge disc flag and retainer. no problem was found with the device. | The automated impella controller (aic) controller was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32344 | 1220648-2023-05423 | 18340590 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18340590&pc=OZD | 11/15/2023 | 01/17/2024 |  | 12/15/2023 |  |  | 11/15/2023 | 12/15/2023 |  |  | 07/31/2025 |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024389445 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Unraveled Material | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/15/2023 |  |  |  |  | The user facility reported a 78-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp being placed independent of an abiomed rep in the ccl. the team was notified that there was a pump-wire interaction. the .018 abiomed wire "unraveled". the team also placed a second wire, the steelcore, which also unraveled. the team disposed of all product and explanted the pump. the team additionally was told the patient was sent to the icu and expired. the cod was not shared, but no allegation made of impella cause/contributing to the patient outcome. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the product damage (guidewire damage - peripheral vessels) has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the root cause of the product damage (guidewire damage - peripheral vessels) was most likely use related.
32345 | 1220648-2023-05424 | 18340596 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18340596&pc=OZD | 11/15/2023 | 04/01/2024 |  | 12/15/2023 |  |  | 11/15/2023 | 12/15/2023 |  |  |  |  | 12/15/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1662880 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/15/2023 |  |  |  |  | The user facility reported that during a case preparation, the automated impella controller, (aic) had a purge system open alarm. a replacement loaner was sent to the facility. no patient is involved. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the aic - purge issue ¿ other has been completed since the initial report was submitted. the console was returned, visually inspected, and tested. the reported failure mode was unable to reproduce. the software team was able to confirm that the reported issue was due to a software defect. the root cause of the persistent purge system open alarm was a software defect.
32346 | 1220648-2023-05425 | 18340600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18340600&pc=OZD | 11/15/2023 | 02/02/2024 |  | 12/15/2023 |  |  | 11/15/2023 | 12/15/2023 |  |  | 08/23/2025 |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024393001 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 11/15/2023 |  |  |  |  | The user facility reported a 83-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported there were multiple attempts by the doctor to get the impella to pass through the subclavian artery but was not successful. the doctor ultimately aborted the 5.5 placement and will leave the patient on ecmo (extracorporeal membrane oxygenation) for the time being. no known patient harm or injury. | [5 empty] | The investigation into the delivery issues- anatomy has been completed. the device was not returned for investigation. as per clinical, the pump was unable to pass through vasculature with the patient having transcatheter aortic valve replacement (tavr). the cause of the delivery issue was patient condition related due to the pump unable to pass the tavr valve per clinical. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32347 | 1220648-2023-05426 | 18340601 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18340601&pc=OZD | 11/15/2023 | 09/24/2024 |  | 12/15/2023 |  |  | 11/15/2023 | 12/15/2023 |  |  | 05/31/2025 |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024358530 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/15/2023 |  |  |  |  | The user facility reported a 68-year-old male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the impella access site had significant bloody drainage, dressing was removed, the catheter was advanced and pulled back to prolapse the hemostatic valve of the peel away sheath. the site was redressed and secured. additionally, the patient had acute left arm pain. there was a concern for thrombus so, brachial cutdown with vessel loops for control was performed. patient was stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed and thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed and thrombus could not be determined. added- h6 component code 925 corrections for the initial mfr report: f6/f8 should have been left blank.
32348 | 1220648-2023-05427 | 18340611 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18340611&pc=OZD | 11/14/2023 | 01/29/2024 |  | 12/15/2023 |  |  | 11/16/2023 | 12/15/2023 |  |  | 02/28/2025 |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 - JP |  |  | ABIOMED, INC. |  |  | US |  | 2024311704 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Non Hispanic | Asian | Injury | No | Yes | Pumping Stopped | Use of Device Problem | [5 empty] | Appropriate Term / Code Not Available | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/16/2023 |  |  |  |  | The user facility reported a 75-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the pump had been on for long-term use beyond indication when the pump stopped after 239 days. the pump was not replaced but explanted as the patient was made a dnr. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop has been completed. the impella pump nor data logs could be evaluated as they were not returned for evaluation. the root cause of the pump stop was unable to be determined as limited clinical details were provided and no product was returned for review. however, pump stop occurred after 239 days of run time, which is well beyond the design life of 60 days.
32349 | 1220648-2023-05428 | 18340616 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18340616&pc=OZD | 11/16/2023 | 11/11/2025 |  | 12/15/2023 |  |  | 11/16/2023 | 12/15/2023 |  |  |  |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1759996 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 11/16/2023 |  |  |  |  | Us complainant reported that in the process of training, the demo catheters were used to attempt to connect the aic, but it failed. the second aic was used and it was successful with all demo catheters used. no patient is involved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. a.1 added information was as it was inadvertently omitted from the initial manufacturer device report number. a.2 age should have been left blank in accordance with updated procedures on the initial manufacturer device report number as no patient was involved. a.3 na should have been selected on the initial manufacturer device report number as no patient was involved. d.2 common device name were revised in accordance with updated procedures since the initial manufacturer device report number was submitted. d.4 serial number and unique identifier (udi) # were revisedas they were reported incorrectly on the initial manufacturer device report number. e.1 phone number was reported incorrectly on the previously submitted manufacturer device report number. e.2 revised selection in accordance with updated procedures since the initial manufacturer device report number was submitted. f.6 and f.8 dates were reported on manufacturer device report number and should not have been. g.1 revised mdr reporting contact email since the initial submission of manufacturer device report number in accordance with updated procedures. added reporting contact fax number in accordance with updated procedures since the initial manufacturer device report number was submitted. h.4 revised date as it was reported incorrectly on the initial manufacturer device report number. h.6 investigation findings and investigation conclusions were revised since the initial manufacturer device report number was submitted in accordance with updated procedures. h.10 additional manufacture narrative was revised since the initial manufacturer device report number was submitted in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of the reported failure mode has been completed. the console and data logs were received for investigation. data analysis: the logs repeatedly had 1036 errors when demo pumps were connected in the field, resulting in placement signal not reliable alarms. the logs confirmed the near zero snr as optical signal went to -3276.8 and lamp was 100. this was indicative of repeated air gaps. device analysis: the complaint was confirmed. the optical pump connector was found to have a broken piece of debris in the pump connector preventing good optical connection despite fine electrical connection. the debris was removed and the console returned functionality. as this was an out of box failure, (b)(4) was opened for this complaint. root cause: the optical signal issue was caused by a piece of debris in the optical pump connector causing an air gap preventing good optical connection. d9. device returned to manufacturer date added d2 common device name revised on manufacturer device report 1220648-2023-05428 in accordance with updated procedures. e2. was erroneously selected on the initial manufacturer device report number 1220648-2023-05428. f6/f8. should have been left blank on the initial manufacturer device report number 1220648-2023-05428 in accordance with updated procedures. g1. reporting contact email revised on manufacturer device report 1220648-2023-05428 in accordance with updated procedures. g1. reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2023-05428.
32350 | 1220648-2023-05429 | 18340619 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18340619&pc=OZD | 11/16/2023 | 09/04/2024 |  | 12/15/2023 |  |  | 11/16/2023 | 12/15/2023 |  |  |  |  | 12/15/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1271193 |  | 40042-0000 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/16/2023 |  |  |  |  | The user facility reported a 45-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported as the impella was implanted, the aic(automated impella controller) went black and was alarming "system self check failed change console immediately". there was an aic replacement. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The engineer found that this automated impella controller aic was incorrectly sent for investigation. there was strong relation between both the complaint description and logs from reported controller serial number (including the alleged pump). the serial number of the console sent for investigation was never in use with the reported pump serial number. the console that was sent was investigated and there were no issues found. due to the wrong console being provided for investigation a regulatory report is no longer necessary.
32351 | 1220648-2023-05430 | 18340623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18340623&pc=OZD | 11/16/2023 | 01/23/2024 |  | 12/15/2023 |  |  | 11/16/2023 | 12/15/2023 |  |  |  |  | 12/15/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1081052 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during a routine automated impella controller (aic) cassette change, while supporting a patient, the purge disc holder broke off. the aic was replaced successfully. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the console was returned for investigation and was tested upon return. the cause of the issue was a defective component. | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
32352 | 1220648-2023-05431 | 18340625 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18340625&pc=OZD | 11/17/2023 | 01/16/2024 |  | 12/15/2023 |  |  | 11/17/2023 | 12/15/2023 |  |  | 06/30/2025 |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024372082 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Female |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A patient of unknown age was scheduled to receive hemodynamic support from an impella cp smartassist heart pump during high-risk percutaneous coronary intervention (hrpci). due to the patient's anatomy, the pump could not be inserted via the femoral approach as there was stenosis and arteriosclerosis and the femoral artery was kinked. it was decided to perform the coronary intervention without support from an impella pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation into the reported delivery issue -anatomy has been completed since the original report was filed. the root cause of the delivery issue- anatomy was determined to be patient condition because of the kinking of right femoral artery.
32353 | 1220648-2023-05432 | 18340626 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18340626&pc=OZD | 11/17/2023 | 03/04/2024 |  | 12/15/2023 |  |  | 11/17/2023 | 12/15/2023 |  |  | 03/31/2025 |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024327580 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/17/2023 |  |  |  |  | The user facility reported a (b)(6) male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 had an easy delivery and no excessive use of torque note but when the level was ramped to p4 a high purge pressure alarm occurred. troubleshooting was initiated with no resolution. a purge system blocked alarm followed. the team elected to exchange the impella. | [5 empty] | The investigation into the reported high purge pressure has been completed since the original report was filed. impella 5.5 was returned. pump was visually inspected and biomaterial ingestion observed in the purge gap and outflow cage. no visible catheter kink or purge line kink observed. the purge fluid was connected to the pump and high purge pressure reproduced after initiating purge. the catheter was cut close to the motor housing, near to red plug kink protector and the purge fluid did not pass through. the red impella plug was opened to inspect the purge line link and no bond issue or kinked purge line observed. the purge line was cut across the monkey cap side and also at the large purge lumen near stainless steel connector but still the purge fluid did not pass and still purge system blocked alarm reproduced. later after a while it started to purge through the cut section of purge lumen near stainless steel connector. no blood ingress in the purge line was observed. all cut section of the purge lumen were initiated purging with manual injection with and also using systemic purge connection. the purge fluid was unable to completely flow through the purge lumen due to dextrose buildup. the purge line closer to motor housing started to purge halfway and did not flow further through the purge gap due to biomaterial blockage. no other abnormalities were found. data logs were reviewed, and gradual purge pressure rise trend observed within 7 to 8 minutes with purge flow reduced to less than 2 ml/hr. the cause of the high purge pressure issue was biomaterial ingested over the purge gap per field investigation and log review. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
32354 | 1220648-2023-05433 | 18340628 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18340628&pc=OZD | 11/17/2023 | 01/29/2024 |  | 12/15/2023 |  |  | 11/17/2023 | 12/15/2023 |  |  |  |  | 12/15/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1656412 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/17/2023 |  |  |  |  | The user facility reported that during a case preparation, the automated impella controller, (aic) had a broken purge clip.. a replacement loaner was sent to the facility. no patient is involved. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the console was tested after arriving in fs. upon examining aic for any visible defects, it was evident that the purge retainer was broken. the root cause of the aic - purge disc/flag issues was a broken purge retainer.
32355 | 1220648-2023-05434 | 18340629 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18340629&pc=OZD | 11/17/2023 | 12/10/2024 |  | 12/15/2023 |  |  | 11/17/2023 | 12/15/2023 |  |  | 08/31/2025 |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024399162 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the 52-yearl-old male patient had hemolysis before the impella support, however the hemolysis continued with impella support. initially, the hemolysis was improving however, after impella support the patient's hemolysis continued and worsened. the clinical rep. suggested to provide blood volume in order to improve the hemolysis, however the doctor did not want to due to high central venous pressure. additionally, the patient had a hematoma that was soft and not growing. no intervention was performed for the hematoma. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into hemolysis has been completed. pump was not returned for investigation. data logs were reviewed and no positioning issue was observed. there was no motor current spike observed but suction alarms were present with no change in p-level. no placement signal trend observed. as per clinical, patient was given blood products to try to resolve the hemolysis. therefore, as per the provided clinical information and data log review, the root cause of the hemolysis issue was most likely patient condition related. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05434 d6a and d6b revised from the initial submission in accordance with updated procedures. g1 reporting contact email g1 reporting contact fax number missing. h6 component code revised from the initial submission in accordance with updated procedures.
32356 | 1220648-2023-05436 | 18340630 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18340630&pc=OZD | 11/17/2023 | 09/25/2025 |  | 12/15/2023 |  |  | 11/17/2023 | 12/15/2023 |  |  | 05/31/2025 |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024358718 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [11 empty] | Death | Other |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/17/2023 |  |  |  |  | The user facility reported a 65-year-old male with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. there was difficulty delivering. the team noted mild kinking. the pump was delivered but upon use, the optical signal was lost. the 5.5 remained on for support for just a day when then explanted and patient was implanted with an impella cp. | Patient was hemodynamically unstable and continued to deteriorate and the procedural outcome was that the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-05436 was provided in sections b1, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. pump was not returned for investigation. data logs were returned and placement signal not reliable (psnr) alarm was observed. the root cause of the placement signal issue was most likely a damaged optical fiber line due to excessive force. clinical has no information on patient¿s vessel condition but no percutaneous coronary intervention was performed prior to insertion. due to experienced resistance during pump implant, a kink was observed under echo. therefore, the root cause of the delivery issue was most likely use related. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 adverse event updated. b2 updated to death. h1 updated to death. corrections that were made from the initial manufacturer device report number 1220648-2023-05436. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated. f6 and f8 should not have had entries in the initial report. g1 reporting contact email updated.
32357 | 1220648-2023-05438 | 18340636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18340636&pc=OZD | 11/18/2023 | 09/20/2024 |  | 12/15/2023 |  |  | 11/18/2023 | 12/15/2023 |  |  | 07/31/2025 |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2024389383 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Hispanic |  | Death | Yes | Yes | Restricted Flow rate | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/18/2023 |  |  |  |  | Additional information noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 69-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the impella was placed in the left femoral artery with 1.0l of flow. the patient went into vf (ventricular fibrillation) arrest. there was 1 shock of pulseless electrical activity and cpr (cardiopulmonary resuscitation) continued. return of spontaneous circulation was achieved. the staff attempted to reposition the device. there was a kink that was notable on fluro. and they were unable to achieve flows greater than 1.0l despite several attempts at repositioning. the impella was replaced. |  |  |  |  | The investigation into the low pump flow issue has been completed since the original report was filed. pump was returned and investigated. cannula kink was observed which happened during delivery due to patients short ascending aorta. the root cause of the low flow issue was a kinked cannula due to patient anatomy. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-05438 was provided in sections b2, b5, d6a, d6b, h1, and h6. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32358 | 1220648-2023-05435 | 18340637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18340637&pc=OZD | 11/17/2023 | 04/16/2024 |  | 12/15/2023 |  |  | 11/17/2023 | 12/15/2023 |  |  | 08/31/2025 |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - EU |  |  | ABIOMED, INC. |  |  | US |  | 2024395218 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/17/2023 |  |  |  |  | The user facility reported a 67-year-old male (race unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was bleeding heavily from the impella puncture site with haematoma formation. hemoglobin dropped from 14 to 12. ptt at 180, leg cold, pulses palpable. coagulation recommendations discussed with md. heparin was paused and it was recommended to apply purgfluid with nabic. abiomed best practice given to bandage. patient to have ct scan. the md noted if the bleeding does not improve, the impella may have to be explanted. the patient remained on support for seven additional days and was successfully weaned. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major and the limb ischemia ¿ reversible issues has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. the root cause of the access site bleeding issue was not determined since product was not returned and there is limited clinical information provided.
32359 | 1220648-2023-05437 | 18340638 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18340638&pc=OZD | 11/17/2023 | 05/07/2024 |  | 12/15/2023 |  |  | 11/17/2023 | 12/15/2023 |  |  | 01/31/2025 |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024333933 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/17/2023 |  |  |  |  | The user facility reported an 81-year-old female enrolled in the protect iv study noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the perclose failed after implant. md held manual pressure for 45 minutes and used a femstop as a precautionary measure after manual pressure. the patient received a blood transfusion. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32360 | 1220648-2023-05439 | 18340721 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18340721&pc=OZD | 11/21/2023 | 11/25/2024 |  | 12/15/2023 |  |  | 11/21/2023 | 12/15/2023 |  |  | 10/31/2025 |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024408775 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/21/2023 |  |  |  |  | The user facility reported a 64-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the patient was having signs of hemolysis with plasma free hemoglobin of 220. it was advised for the pump to be repositioned, but no action was taken at that time. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into hemolysis has been completed. the impella device was not returned for evaluation. the root cause of the hemolysis issue was improper positioning of the device due to improper technique. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05439. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. h6 component code revised from the initial submission in accordance with updated procedures.
32361 | 1220648-2023-05440 | 18341405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18341405&pc=OZD | 12/06/2023 | 05/03/2024 |  | 12/15/2023 |  |  | 12/06/2023 | 12/15/2023 |  |  | 12/31/2024 |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024300612 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/06/2023 |  |  |  |  | The us complainant in 2/2023 had an impella 5.5 support for 17 days. the patient had the support for ischemic cardiomyopathy until the pump was explanted, and patient went for an lvad (the left ventricular assist device); per abiomed field rep. later in (b)(6) 2023 the medical center published a paper reviewing multiple 5.5 pump supports. the author noted the patient had an "on post-operative day 4, he developed an axillary artery thrombosis, requiring a surgical thrombectomy via right brachial cut-down and the removal of the impella." | [5 empty] | The investigation into the thrombus issue has been completed since the original report was submitted. the device was not returned by the user facility for investigation. as the necessary information was not provided, the root cause of the thrombus was not determined. d6a, d6b. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32362 | 1220648-2023-05441 | 18341406 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18341406&pc=OZD | 12/11/2023 | 03/11/2025 |  | 12/15/2023 |  |  | 12/11/2023 | 12/15/2023 |  |  | 09/30/2025 |  | 12/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024408900 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | Unspecified Heart Problem | [10 empty] | Disability | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/11/2023 |  |  |  |  | The user facility reported a 66-year-old female was implanted with an impella cp device for mechanical circulatory support. the us complainant had a cp placed for support of a patient admitted/transferred to the tertiary center from the community hospital. the cp was explanted after 6 hours due to limb ischemia and heart block. the patient was also treated with 2 units of rbc and crrt (continuous renal replacement therapy) due to hemolysis. the team escalated the care to an axillary placed 5.5 pump. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | Product not returned. data logs show suction alarms near the end of the case. no mc spikes outside of p level changes and no other abnormalities. the cause of the hemolysis was not determined due to product not returned, lack of clinical details, and inconclusive data logs. as all the necessary information was not provided, the root cause of the limb ischemia was not determined." the cause of the vt/vf was not determined since product was not returned and lack of clinical details.
32363 | 1220648-2023-05392 | 18325710 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18325710&pc=OZD | 12/02/2022 | 02/26/2024 |  | 12/14/2023 |  |  | 12/02/2022 | 12/14/2023 |  |  | 05/31/2024 |  | 12/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023229974 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient in the eu had their impella cp implant procedure aborted due to patient anatomy. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation into the delivery issues has been completed. the device was not returned for investigation. per the provided clinical information, the root cause of the delivery issue was patient condition related to patient¿s small/ diseased vessels (aorta).
32364 | 1220648-2023-05393 | 18326174 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18326174&pc=OZD | 12/14/2022 | 02/28/2024 |  | 12/14/2023 |  |  | 12/14/2022 | 12/14/2023 |  |  | 10/31/2024 |  | 12/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023275286 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  |  | Injury | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 63-year-old male noted as high risk coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient had an air in purge system alarm while being moved from the cath lab to icu. the issue did not resolve even after venting multiple times. upon arrival to the icu, the p-level was still on auto mc at 1107/906 ma and an emergency shutdown was imminent. the impella had to be restarted twice, after which the purge flow and pressure improved and the mc values normalized. the impella was scheduled for removal the next day. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the reported high purge pressure has been completed since the original report was filed. pump and data logs were not available for review. as per the clinical review, air in purge alarm was reported during pump use. the cause of the priming issue was not determined since product was not returned and sufficient clinical information was not provided.
32365 | 1220648-2023-05394 | 18326683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18326683&pc=OZD | 10/13/2023 | 12/14/2023 |  | 12/14/2023 |  |  | 10/13/2023 | 12/14/2023 |  |  |  |  | 12/14/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1574621 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  |  |  | Malfunction | No | Yes | Material Fragmentation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a 78-year-old male patient noted as high risk coronary intervention was implanted with an impella cp device for mechanical circulatory support. after pump explant a part of the cover of connection between pump and automated impella controller (aic) was fallen off. there was no reported patient harm. | [5 empty] | The investigation into the reported 'aic - damage" has been completed. product was returned for investigation. the device was inspected and found the optical connector cover was broken. after replacing the broken part, the console passed functional check. during data analysis, the automated impella controller (aic) and rt logs neither confirmed nor denied the complaint given its nature. the root cause of the failure was a broken optical connector cover. this report is being filed as part of a retrospective review of historical records.
32366 | 1220648-2023-05395 | 18327441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18327441&pc=OZD | 08/03/2023 | 08/15/2025 |  | 12/14/2023 |  |  | 11/15/2023 | 12/14/2023 |  |  | 03/31/2020 |  | 12/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1338989 |  | 005062-RO |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Male |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Valve Insufficiency/ Regurgitation | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/15/2023 |  |  |  |  | The complainant reported that a 29-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. after almost 3 years post use of device, the patient is suffering from atrial fibrillation and thyrotoxicosis-induced heart failure and developed progressive aortic insufficiency due to prolonged use of an implanted impella 5.0 pump. it was noted that the patient had undergone multiple interventions as a result of the pre-existing condition including extracorporeal membrane oxygenation (ecmo) and multiple valve repairs/replacements, support was maintained with the implanted pump throughout the course of needed support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | D.6a and d.6b added since the initial submission of manufacturer device report number 1220648-2023-05395 in accordance with updated procedures. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2023-05395 and should not have been. g.1 revised mdr reporting contact email since the initial submission of manufacturer device report number 1220648-2023-05395 in accordance with updated procedures. added reporting contact fax number in accordance with updated procedures since the initial manufacturer device report number 1220648-2023-05395 was submitted. h.6 code 3221 was reported incorrectly on the initial manufacturer device report number 1220648-2023-05395 in accordance with updated procedures. component code was added since the initial submission of manufacturer device report number 1220648-2023-05395 in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report number 1220648-2023-05395 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures.
32367 | 1220648-2023-05396 | 18331061 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18331061&pc=OZD | 11/15/2023 | 11/07/2024 |  | 12/14/2023 |  |  | 11/15/2023 | 12/14/2023 |  |  | 07/31/2024 |  | 12/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023255687 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/15/2023 |  |  |  |  | The user facility reported a 50-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced hemolysis the hemolysis was never able to be resolved or troubleshot effectively. the labs had lactate dehydrogenase (ldh) rise. hemolysis issues persisted until pump explanted. patient is stable post explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05396. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact fax number missing. h6 component code revised from the initial submission in accordance with updated procedures. h10 instructions for use.
32368 | 1220648-2023-05397 | 18331069 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18331069&pc=OZD | 11/15/2023 | 11/18/2025 |  | 12/14/2023 |  |  | 11/15/2023 | 12/14/2023 |  |  |  |  | 12/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1383042 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 11/15/2023 |  |  |  |  | The user facility reported a 50-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp support ongoing when the team made observation of the aic (automated impella controller) having a loose purge disc/flag issue. the console remained on for the support til weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | Corrections to the initial mfr report: g1. mdr reporting contact email. d2. device name updated. d4. model number updated. f6/ f8. should have been left blank. h3. has been updated to device eval completed. h6. type of investigation code 10 added.
32369 | 1220648-2023-05369 | 18318275 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18318275&pc=OZD | 11/13/2023 | 08/26/2024 |  | 12/13/2023 |  |  | 11/13/2023 | 12/13/2023 |  |  | 08/31/2025 |  | 12/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC |  |  | US |  | 2024389345 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Power Up | Failure to Pump | Material Deformation |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a 70-year-old female patient was implanted with an impella cp for mechanical circulatory support after papillary muscle rupture due to acute myocardial infarction. however, the pump reportedly did not start. the cp was therefore explanted and replaced with another cp. upon removal, the blades in the pump appeared to be damaged and some tissue matter was found. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop has been completed since the original report was filed. the device was returned to the lab for investigation. the root cause of the pump did not start was determined to be a damaged outflow cage. | The root cause of the pump did not start issue was most likely due to a damaged outflow cage but the exact cause of the deformity is unknown due to a lack of clinical details. d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05369: h.6 code 2199, 1476, 2976, and 4741 were reported incorrectly. new code was added medical device problem code. h6 type of investigation missing. h6 213 and 19 were reported incorrectly. new codes were added investigation findings and investigation conclusions.
32370 | 1220648-2023-05370 | 18318336 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18318336&pc=OZD | 11/13/2023 | 02/10/2025 |  | 12/13/2023 |  |  | 11/13/2023 | 12/13/2023 |  |  |  |  | 12/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1637336 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/13/2023 |  |  |  |  | The user facility reported an unspecified patient was implanted with an impella device for mechanical circulatory support. the complainant reported that an aic (automated impella controller) experienced a control device error. the aic was removed from service as a result of the noted issue and a replacement was requested for the customer site. there was no reported patient harm. | [5 empty] | The impella device was received from the customer and an evaluation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. | The investigation for the console (aic) controller error-yellow alarm has been completed. data logs were confirmed the reported failure mode given there was a controller error alarm (opm comm crc error ¿ event set # 1045) triggered during the clinical procedure. real time (rt) logs confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) are represented as -16384 values due to the crc error. console was received for evaluation. the console was found to have low snr during the investigation which unrelated to the complaint. the root cause for the controller error was due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. corrections to the initial mfr report: d2 updated common device name. d4-updated model number. f6/f8 should have been left blank. g1 mdr reporting contact email.
32371 | 1220648-2023-05371 | 18318669 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18318669&pc=OZD | 11/09/2023 | 03/04/2024 |  | 12/13/2023 |  |  | 11/13/2023 | 12/13/2023 |  |  | 07/31/2025 |  | 12/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP JP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024374352 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp support ongoing along with ecmo when the male patient suffered a cva/stroke and expired. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported stroke /cva has been completed since the original report was filed. as the necessary clinical information was not provided, the true root cause of the stroke/cva could not be determined.
32372 | 1220648-2023-05372 | 18319291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18319291&pc=OZD | 11/13/2023 | 04/16/2024 |  | 12/13/2023 |  |  | 11/13/2023 | 12/13/2023 |  |  | 02/28/2025 |  | 12/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024321033 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 65-year-old male was implanted with an impella cp device for mechanical circulatory support. the complainant had placed a cp along with va-ecmo for patient with atrial clot. the impella cp was placed and distal perfusion sheath was additionally placed. the repo-sheath site had bleeding issues that prompted the team to explant and replace with a larger bore perfusion, via a 16fr sheath. the impella was later explanted that same day. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ irreversible and access site bleeding - minor issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was most likely due to lack of vessel recoil since best practices were followed and bleeding occurred onto the repositioning sheath and hemostasis wasn't achieved until a 16fr sheath was placed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
32373 | 1220648-2023-05373 | 18319318 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18319318&pc=OZD | 11/13/2023 | 01/24/2025 |  | 12/13/2023 |  |  | 11/13/2023 | 12/13/2023 |  |  | 09/30/2025 |  | 12/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024403162 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/13/2023 |  |  |  |  | The user facility reported a 83-year-old female with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the support was ongoing and had signs of hemolysis and an access site bleed. labs were hemolyzing and the effluent from the continuous renal replacement therapy was tinged. pump was repositioned under echo to address the hemolysis, impella dropped to p5 after repositioning as patient was stable and to further reduce hemolysis. the access site ooze prompted the team to hold the anticoagulant heparin. bleeding resolved at both sites post intervention. | It was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the access site bleed and hemolysis has been completed. pump was not returned but data logs were investigated. data logs show no positioning issue, suction alarm, motor current spike or placement signal trend. blood products were not given to the patient as per the clinical information. with the limited clinical information and no product, it is not possible to investigate the complaint further. the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical information provided. therefore, the root cause of the hemolysis issue was not determined since product was not returned and insufficient clinical information provided. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrections to the initial mfr report: b2-selected death and added date of death. b5-patient outcome and mso review. f6/f8 should have bene left blank. g1 mdr reporting contact email. h1-changed to death. h6-impact code updated to death. h6 component code 925.
32374 | 1220648-2023-05374 | 18319453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18319453&pc=OZD | 11/16/2023 | 03/22/2024 |  | 12/13/2023 |  |  | 11/16/2023 | 12/13/2023 |  |  | 08/31/2025 |  | 12/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024397355 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the 70-year-old male patient had the impella placed in the setting of cardiogenic shock s/p eight shocks for ventricular fibrillation (vf) and prolonged out of hospital cpr. the automated impella controller (aic) alarmed for low purge flow, high purge pressure. the staff de-aired the tubing to ensure that fluid was able to pass freely when disconnected from the side arm, and it did. the pump stopped for the first time. it was able to be restarted, but it stopped again for the final time. the interventionalist was at bedside to remove it and close the groin with perclose. no clot was visualized on the pump, and the physician stated he did not see left ventricle (lv) thrombus on the echo prior to impella insertion. | [5 empty] | The investigation for the high purge pressure and pump stop has been completed. the pump was returned for evaluation. visual inspection revealed biomaterial in the purge gap and motor housing. data logs indicate a gradual rise in purge pressure within a 5-hour period. also, several pump stop alarms, indicating high motor current, were reported during the purge system block alarm. the cause of the pump stop issue was biomaterial. the root cause of the high purge pressure issue was biomaterial. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
32375 | 1220648-2023-05375 | 18319566 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18319566&pc=OZD | 11/16/2023 | 11/19/2025 |  | 12/13/2023 |  |  | 11/16/2023 | 12/13/2023 |  |  |  |  | 12/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1616687 |  | AUTOMATED IMPELLA CONTROLLER |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/16/2023 |  |  |  |  | The user facility reported a 76-year-old male patient was implanted with an impella device for mechanical circulatory support. the complainant reported that their aic (automated impella controller) displayed an "emergency shutdown imminent" message during high-risk percutaneous coronary intervention. although the issue was initially self-resolved, the failure recurred, and the decision was made to replace the aic with a backup console. there was no patient harm as a result of the failure. | [5 empty] | The impella device was received from the customer and an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. | The investigation has been completed. data analysis: console imc logs confirm that emergency shutdown imminent alarm was issued after shutdown was requested while pump was connected. the console was then shutdown. device analysis: the consoles internal circuitry was inspected with no evidence of anomaly. console operated to specification when run with a known good pump and purge line. conclusion: the root cause of the failure mode couldn't be determined due to the failure mode not being able to be reproduced. d2: corrected the brand name and common device. d4: corrected the catalog number. g1: corrected the manufacturer email address. h3: added information regarding the investigation status. h6: added codes per the results of the investigation. h10: added the results of the investigation.
32376 | 1220648-2023-05376 | 18320169 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18320169&pc=OZD | 11/16/2023 | 12/13/2023 |  | 12/13/2023 |  |  | 11/16/2023 | 12/13/2023 |  |  |  |  | 12/13/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1399871 |  | 0042-0010 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant in the eu reported the purge disc holder of the automated impella controller (aic) was broken. there was no reported patient harm. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
32377 | 1220648-2023-05377 | 18320234 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18320234&pc=OZD | 11/16/2023 | 10/14/2024 |  | 12/13/2023 |  |  | 11/16/2023 | 12/13/2023 |  |  | 03/31/2025 |  | 12/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024331329 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported patient of unknown age was attempted to be implanted with an impella cp for hemodynamic support. however, there was an ¿impella in ventricle¿ alarm displayed. the physician attempted to unsuccessfully reposition the impella since it was noted to be in the aorta. the device was ultimately explanted. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation for the positioning issue has been completed. pump was not returned for investigation. data logs were returned and "impella in aorta" and "impella in ventricle" alarms were observed. clinical information provided is limited and the pump was explanted due to no longer being able to reposition. therefore, the root cause of the positioning issue was use related to improper technique. added- h6 impact code 4628 corrections to the initial mfr report: added-d6a/d6b
32378 | 1220648-2023-05378 | 18320274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18320274&pc=OZD | 11/14/2023 | 09/12/2024 |  | 12/13/2023 |  |  | 11/14/2023 | 12/13/2023 |  |  | 09/30/2025 |  | 12/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024403195 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a 53-year-old male patient on ongoing impella cp support following a stemi. the patient had imaging done on the second day of support and apical clot was observed in the left ventricle (lv). the pump remained on for two more days despite the thrombosis. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- h6 component code 925 in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
32379 | 1220648-2023-05379 | 18320328 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18320328&pc=OZD | 11/14/2023 | 02/22/2024 |  | 12/13/2023 |  |  | 11/14/2023 | 12/13/2023 |  |  | 11/30/2024 |  | 12/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024294883 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  |  | Malfunction | No | Yes | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 75-year-old male was implanted with an impella cp device for mechanical circulatory support. the complainant had an impella cp support ongoing when there was an observation made of an access site bleed that needed a femostop to control the bleed. additionally the team an alarm occurring for positioning issues. the pump in aorta alarm did prompt the team to perform echo to evaluate the position. the pump was in the aorta. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation into the positioning issues and the access site bleeding ¿ minor issues has been completed. the device was not returned for investigation. as per the clinical details, a femstop was placed due to bleeding at the access site. the cause of the positioning issue was most likely use related. the cause of the access site bleeding issue was not determined since product was not returned and sufficient clinical information was not provided.
32380 | 1220648-2023-05380 | 18320342 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18320342&pc=OZD | 11/14/2023 | 04/05/2024 |  | 12/13/2023 |  |  | 11/14/2023 | 12/13/2023 |  |  | 07/31/2025 |  | 12/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024391689 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/14/2023 |  |  |  |  | The user facility reported a 64-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported that during the initial setup for the impella, no purge fluid was noted exiting the pump outlet and the purge pressure was above 1100 and purge fluid was below 1. no kinks or obstructions were found. purge cassette was replaced with no resolution. . a new impella 5.5 was placed successfully. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation into the high purge pressure issue has been completed since the initial report was submitted. the device was not returned for investigation. the cause of the high purge pressure issue was not determined since product was not returned and sufficient clinical information was not provided.
32381 | 1220648-2023-05381 | 18320413 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18320413&pc=OZD | 11/16/2023 | 01/19/2024 |  | 12/13/2023 |  |  | 11/16/2023 | 12/13/2023 |  |  |  |  | 12/13/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1370139 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Unexpected Shutdown | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/16/2023 |  |  |  |  | The complainant reported that the automated impella controller (aic) made a noise that it usually does not make after it booted up. as the connection was being checked, the controller unexpectedly shut down. the aic was replaced. no patient is involved. | [5 empty] | The investigation into the reported aic shut down or restart issue has been completed since the original report was filed. data showed device was powered on. subsystems were initialized, but then the console reboots. there was an unexpected shutdown alarm after the reboot and the console was requested to shut down by the user. analysis of the device showed the pump was run overnight and the console was power cycled 100 times but neither test reproduced the clinical failure. due to some pbm duplicate commands in the logs, the pbm was examined under a microscope for potential corrosion but there was none visible. the cause of the issue was not determined since the clinical failure was not reproduced. due to the beeps heard in the clinical setting, the 4-in-1 or the pbm powering the 4-in-1 could have malfunctioned, causing the reboot. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
32382 | 1220648-2023-05382 | 18320486 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18320486&pc=OZD | 11/18/2023 | 09/21/2024 |  | 12/13/2023 |  |  | 11/18/2023 | 12/13/2023 |  |  | 09/30/2025 |  | 12/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024398847 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Hispanic |  | Death | Yes | Yes | Material Twisted/Bent | [6 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/18/2023 |  |  |  |  | The user facility reported a 69-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient had their impella replaced with the impella cp with the existing sheath. the new pump was placed with improved flows up to 1.5l. however, the patient then coded again requiring cpr. pneumothorax diagnosed requiring chest tube placement. return of spontaneous circulation was achieved. position-checked device slightly kinked again. staff attempted to reposition and the patient coded again and the patient expired. no allegations were made towards the impella for cause of death. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b1, b2, b5, h1, and h6. | The investigation for the kinked cannula has been completed. the cannula was returned for evaluation. upon visual inspection, a kink was noted on the returned pump. no other physical abnormalities were observed. during testing, the pump ran as expected. the root cause of the product damage-cannula kink issue was use related.
32383 | 1220648-2023-05383 | 18320508 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18320508&pc=OZD | 11/16/2023 | 01/23/2024 |  | 12/13/2023 |  |  | 11/16/2023 | 12/13/2023 |  |  |  |  | 12/13/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1299518 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/16/2023 |  |  |  |  | The complainant reported a battery failure on the aic. the aic was removed from service. there was no patient or user harm. | [5 empty] | The impella device was received from the customer and an evaluation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - battery failure (red alarm) has been completed since the original report was filed. the console was returned for investigation and was tested upon return. data logs confirmed the reported failure. the batteries and charging system were tested and confirmed to be working correctly with the batttest utility. the unexpected controller shutdown alarm was determined to be caused by a previous console shutdown triggered by disconnecting the ac power cord while the battery transport switch was disabled. the cause of the issue was use related. the battery transport switch was disabled.
32384 | 1220648-2023-05384 | 18320539 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18320539&pc=OZD | 11/22/2023 | 03/01/2024 |  | 12/13/2023 |  |  | 11/22/2023 | 12/13/2023 |  |  | 05/31/2025 |  | 12/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024374376 |  | 0048-0044 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Infusion or Flow Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/22/2023 |  |  |  |  | The user facility reported a 55-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient was on impella cp support when the motor current became elevated and the left ventricle systolic number was elevated (higher than aortic pressure) and falsely elevated impella flow. the doctor turned the flow to p8. there was a risk of the pump stopping. additionally, the patient was overly anticoagulated and the patient experienced hemolysis and access site bleeding. the pump was removed due to the patient's complications. patient is stable post explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major, the hemolysis and the saturated flow has been completed. the device was not returned for investigation. per the clinical information provided, the patient was overly anticoagulated, leading to access site bleeding. the root cause of the access site bleeding issue was anticoagulation management related due to high patient act values as patient was overly anticoagulated. the root cause of the saturated flow and the hemolysis issues was not determined since product and data logs were not returned for investigation.
32385 | 1220648-2023-05385 | 18320552 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18320552&pc=OZD | 11/16/2023 | 04/03/2024 |  | 12/13/2023 |  |  | 11/16/2023 | 12/13/2023 |  |  | 09/30/2025 |  | 12/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024403732 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/16/2023 |  |  |  |  | The user facility reported a 67-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed persistent oozing from the access site following impella implant. manual pressure was being held to treat the condition, but the decision was ultimately made to explant the pump. patient is stable post explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the initial report was submitted. the product was not returned for investigation. as per clinical details, bleeding at access site was resolved after applied manual pressure. the patient had a stenosis vessel condition. the cause of the access site bleeding issue was most likely related to the patient's condition, specifically the presence of stenosis.
32386 | 1220648-2023-05386 | 18320609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18320609&pc=OZD | 11/16/2023 | 12/13/2023 |  | 12/13/2023 |  |  | 11/16/2023 | 12/13/2023 |  |  | 04/30/2024 |  | 12/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023216430 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Information from protect iv study: a patient of unknown age and gender received hemodynamic support from an impella cp smartassist heart pump during high-risk percutaneous coronary intervention (hrpci). access for the coronary intervention was via the femoral access for the impella heart pump (single access technique). coronary intervention was successfully performed and the impella cp was removed. access site infection was reported to occur 15 days after impella heart pump removal. the infection was not considered serious and was treated with a sterile wound dressing. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32387 | 1220648-2023-05387 | 18320649 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18320649&pc=OZD | 11/16/2023 | 09/12/2024 |  | 12/13/2023 |  |  | 11/16/2023 | 12/13/2023 |  |  | 07/31/2025 |  | 12/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024389546 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/16/2023 |  |  |  |  | The user facility reported a 56-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. md notified significant bleeding from impella access site. angle matching done and manual compression applied. the bleed/hematoma was successfully evacuated and manual hold and femstop applied and bleeding was resolved. patient is stable post intervention. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely use issue since angle matching resolved the issue. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- h6 component code 925 added- d6a/d6b. f6/f8 should have been left blank in the initial report.
32388 | 1220648-2023-05388 | 18320756 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18320756&pc=OZD | 11/16/2023 | 10/30/2024 |  | 12/13/2023 |  |  | 11/16/2023 | 12/13/2023 |  |  | 09/30/2025 |  | 12/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024403893 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  |  |  | Injury | Yes | Yes | Migration | [6 empty] | Atrial Fibrillation | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/16/2023 |  |  |  |  | The user facility reported a 80-year-old female of unknown race in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the pump migrated out of the left ventricle when patient went into atrial fibrilation/rapid ventricular response. the pump was removed at bedside. patient is stable post explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the positioning issues and the atrial fibrillation has been completed. pump was not returned for investigation. data logs were returned and "impella in aorta" and "impella in ventricle" alarms were observed confirming the clinical details. therefore, the root cause of the positioning issue was most likely use related to improper technique. the root cause of the atrial fibrillation could not be determined without additional clinical details. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrections to initial mfr report: added d6a/d6b f6/f8 should have been left blank. h6 clinical code 2130 changed to 1729.
32389 | 1220648-2023-05389 | 18320843 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18320843&pc=OZD | 11/16/2023 | 10/30/2024 |  | 12/13/2023 |  |  | 11/16/2023 | 12/13/2023 |  |  | 06/30/2025 |  | 12/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - APAC |  |  | ABIOMED, INC. |  |  | US |  | 2024386709 |  | 0048-0047 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | Distributor |  | 82 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/16/2023 |  |  |  |  | The user facility reported via distributor an 82-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient had sepsis and hemolysis leading to pump removal. the patient was treated with blood products and antibiotics. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related events are as follows: the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the hemolysis and the sepsis has been completed. as per the clinical information provided, hemolysis was treated with blood products but no concrete evidence provided if that resolved hemolysis. data logs were not provided. therefore, the root cause of the hemolysis issue was most likely patient condition related. the root cause of the sepsis could not be determined without additional clinical details. added- h6 component code 925 corrections to the initial mfr report: added d6a/d6b f6/f8 should have been left blank.
32390 | 1220648-2023-05390 | 18320863 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18320863&pc=OZD | 09/27/2023 | 06/25/2024 |  | 12/13/2023 |  |  | 11/13/2023 | 12/13/2023 |  |  | 08/31/2025 |  | 12/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024384520 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 31 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/13/2023 |  |  |  |  | A loqi notification was received reporting a 31 year old female patient presenting with cardiomyopathy for support via the impella cp device. during support, thrombocytopenia was endured. per report, "stable thrombocytopenia. multiple reasons for this given multiple mcs devices, zosyn, ppi, sepsis. hitt sent x2, neg x2." a possible relationship with the impella device was noted. | There is not enough information to associate use of the device with the outcome |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the thrombocytopenia issue has been completed since the original report was submitted. the device was not returned for investigation. the cause of the thrombocytopenia was not determined due to insufficient clinical details. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
32391 | 1220648-2023-05342 | 18312117 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18312117&pc=OZD | 11/13/2023 | 01/12/2024 |  | 12/12/2023 |  |  | 11/13/2023 | 12/12/2023 |  |  | 07/31/2025 |  | 12/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024391182 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/13/2023 |  |  |  |  | The user facility reported a 78-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the implanted impella cp optical pump experienced persistent high purge pressure throughout the course of patient support. it was noted that purge flows were as low as 1-2 ml/hr which triggered high purge pressure alarms. the pump was eventually explanted with no harm to the patient. | [5 empty] | The investigation for the reported high purge pressure has been completed. impella cp pump and purge cassette was returned for evaluation. visually inspected the pump and no catheter kink was observed, but biomaterial was ingested over the purge gap. purge fluid was connected to purge cassette, primed and pressurized purge. the purge pressure raised close to 900 mmhg with flow of 2.5 ml/hr after 10 minutes of purging. but slowly after 20-25 mins of purging the blockage was easing out and the purge fluid passed through the purge gap with purge pressure reducing back to 650 mmhg. no other abnormalities were found. data logs were reviewed, and purge pressure rise was seen after 1 hour with likely gradual biomaterial ingestion due to saturated flow waveforms. no motor current spike was seen. the cause of the high purge pressure issue was due to biomaterial ingestion over purge gap per field investigation and data log review. the instructions for use states "¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
32392 | 1220648-2023-05343 | 18312203 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18312203&pc=OZD | 11/13/2023 | 02/06/2024 |  | 12/12/2023 |  |  |  | 12/12/2023 |  |  | 09/30/2025 |  | 12/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024401745 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male | 56 KG | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/13/2023 |  |  |  |  | The user facility reported a 56-year-old male patient was implanted with an impella 5.5 device for mechanical circulatory support. the impella motor current increased followed by a pump stop. pump was successfully restarted on p2. a few minutes later the pump stopped again. the physician decided not to replace the pump. after the case, the pump was inspected. visible biomaterial was noticed in the cannula just distal to the outlet windows. no patient harm reported due to the issue. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the reported pump stop has been completed since the original report was filed. the motor current is not very pulsatile and there is low flow. there is an immediate hmc pump stop and then it is able to run and then stopped again and would not turn back on. this is most likely indicative of biomaterial. upon investigation of the returned product, there was hardened calcium caught in the impeller blades. it was removed and soaked in a cup of warm water to confirm it was calcium and not dextrose. the pump was then run in a bucket after and it ran for 10 minutes with no issues. the root cause of the pump stop was due to biomaterial.
32393 | 1220648-2023-05350 | 18312491 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18312491&pc=OZD | 11/14/2023 | 02/10/2025 |  | 12/12/2023 |  |  | 11/14/2023 | 12/12/2023 |  |  |  |  | 12/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1142509 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Priming Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/14/2023 |  |  |  |  | The user facility reported a 73-year-old male patient was implanted with an impella device for mechanical circulatory support. the complainant reported that when the impella was inserted through the sheath and was in the iliac, the purge pressure low alarm was triggered. the purge not detected alarm as well was triggered. the staff noticed purge fluid leaking from the purge disc when they opened the door. the md then removed the impella from the body. the staff opened a new cassette and it failed to prime. a new console was used with the same new cassette and the impella primed appropriately. the patient was stable. | [5 empty] | The aic was received from the customer and an evaluation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. | The investigation for the automated impella controller (aic) purge disc/flag issues has been completed. the console logs confirmed that the purge disc not detected alarm was issued but quickly resolved on the first insertion of the pump. the console was then able to proceed through purge priming as well as optical bench calibration. the pump was unplugged and re-plugged into the console. the console was shortly restarted after unplugging the pump a second time. it was on the third restart of the console that multiple instances of purge disc not detected alarms were issued as well as other purge issues. real time logs confirmed anomalous purge tick behavior when multiple purge issue alarms were issued. inspection of the console revealed a crack in the purge retainer but it is not related to the complaint. the console operated to specification when a know good pump and purge were run with it. the root cause for the purge issues is most likely due to an air gap introduced into the purge system during change from one purge cassette to another. the failure mode could not be reproduced in the lab. d.1 and d.2 revised brand name and common device name since initial manufacturer device report number 1220648-2023-05350 was submitted in accordance with updated procedures. d.4 revised model number, catalog number, serial number and unique identifier (udi) # as they were previously entered incorrectly on the initial manufacturer deivce report number 1220648-2023-05350. e.1 revised us phone number as it was previously entered incorrectly on the initial manufacturer deivce report number 1220648-2023-05350. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2023-05350 and should not have been. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2023-05350 was submitted in accordance with updated procedures. revised reporting contact fax number as it was inadvertently omitted from the initial manufacturer device report number 1220648-2023-05350. h.4 revised device manufacture date as it was entered incorrectly on the initial manufacturer deivce report number 1220648-2023-05350. h.6 codes 4627 and 2199 were reported entered incorrectly on the initial manufacturer deivce report number 1220648-2023-05350. a new code was added.
32394 | 1220648-2023-05351 | 18312820 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18312820&pc=OZD | 11/14/2023 | 01/18/2024 |  | 12/12/2023 |  |  | 11/14/2023 | 12/12/2023 |  |  | 05/31/2025 |  | 12/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2024354789 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that due to the 64-year-old female patient's anatomy, the physician was unable to advance the impella 5.5 even though the 21f dilator was able to progress. the vessel was noted to be only 6.2mm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the root cause of the delivery issue was most likely patient condition related. the failure mode will be monitored and trended.
32395 | 1220648-2023-05352 | 18312847 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18312847&pc=OZD | 11/14/2023 | 10/30/2024 |  | 12/12/2023 |  |  | 11/14/2023 | 12/12/2023 |  |  | 09/30/2025 |  | 12/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024398878 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient suffered a minor dissection of the left femoral/iliac artery during insertion of the introducer. the pump supported the patient as intended once inserted; however, upon explant, it was noted that the perclose sutures failed along with the 8fr angioseal. to treat the conditions and properly seal the access site, a balloon tamponade was utilized. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the vascular damage and the access site bleed has been completed. per the clinical information provided, the sheath was swapped for a longer sheath due to stenosis of patients access vasculature which would likely have caused the sheath to dissect the femoral artery. the bleeding at the access site was a result of the minor dissection of the "uncontrolled arteriotomy". due to it not being adequately closed/ controlled, the bleeding continued. therefore, the root cause of the vascular damage issue was most likely patient condition related to patients diseased vessels. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- h6 component code 925 corrections to the initial mfr report: added-d6a/d6b.
32396 | 1220648-2023-05353 | 18312936 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18312936&pc=OZD | 11/14/2023 | 12/12/2023 |  | 12/12/2023 |  |  | 11/15/2023 | 12/12/2023 |  |  |  |  | 12/12/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, JP |  |  | ABIOMED, INC. |  |  | US |  | 1336513 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 11/15/2023 |  |  |  |  | The user facility reported a patient of unknown age and gender in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the automated impella controller (aic) failed to recognize an installed purge cassette. the purge cassette was replaced but the purge alarm persisted. the aic was replaced and the issue was resolved. there was no patient harm. | [5 empty] | The investigation into the automated impella controller (aic) malfunction has been completed. data analysis: (b)(6) was able to recognize the purge disc until purge cassette sn (b)(6) is swapped out for sn (b)(6) on (b)(6). the purge cassette was removed and reinserted three more times with no change in ability to recognize purge disc. finally, purge cassette (b)(6) was swapped in but again the purge cassette was not recognized so the pump was unplugged and moved to another aic. device analysis: the console was tested after arriving in fs japan. the purge flag was confirmed to be missing which explains why, regardless of the purge cassette, the console was unable to recognize the purge disc. root cause: the cause of the issue was a missing purge flag.
32397 | 1220648-2023-05354 | 18313034 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18313034&pc=OZD | 11/15/2023 | 12/12/2023 |  | 12/12/2023 |  |  | 11/15/2023 | 12/12/2023 |  |  | 05/31/2025 |  | 12/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 EU |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2024353103 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  |  | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a 67-year-old male patient was on an impella cp for mechanical circulatory support. however, the purge filter was noted to be broken. a purge filter bypass was initially completed; however, the device was subsequently replaced as a result of the issue. there was no reported patient harm. | [5 empty] | The pump was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
32398 | 1220648-2023-05355 | 18313044 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18313044&pc=OZD | 11/15/2023 | 12/12/2023 |  | 12/12/2023 |  |  | 11/15/2023 | 12/12/2023 |  |  | 09/30/2025 |  | 12/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024399852 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Puncture/Hole | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella cp placed on 84-year-old male patient. the pump was prematurely explanted due to concerns of pump damage/hole from the single access technique. the md explanted prior to the intervention to the right coronary artery (rca) due to the "interference". there was no harm or injury noted and the patient survived without another cp being placed. | [5 empty] | The device was received for evaluation. the investigation into the reported "product damage (hole in catheter)" has been completed. the device was inspected and found there was a small puncture noted around the 30cm mark of the catheter. the root cause of the product damage (hole in catheter) is most likely due to use.
32399 | 1220648-2023-05356 | 18313085 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18313085&pc=OZD | 11/15/2023 | 03/04/2024 |  | 12/12/2023 |  |  |  | 12/12/2023 |  |  | 09/30/2025 |  | 12/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 JP |  |  | OZD |  |  | IMPELLA 5.5 JP |  |  | ABIOMED, INC. |  |  | US |  | 2024393880 |  | 1000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 pump met resistance during attempted delivery for 71-year-old male patient support. following insertion, it was stated that the pump was unable to traverse the patient's tortuous vasculature. the 5.5 implant was aborted and an impella cp pump was subsequently placed for patient support. there was no patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues- anatomy issue has been completed. the device was not returned for investigation. the cause of the delivery issue was most likely patient condition related.
32400 | 1220648-2023-05357 | 18313106 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18313106&pc=OZD | 11/15/2023 | 10/20/2025 |  | 12/12/2023 |  |  | 11/15/2023 | 12/12/2023 |  |  |  |  | 12/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1398838 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 40 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Analyze Signal | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/15/2023 |  |  |  |  | The user facility reported a 40-year-old female patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. when pump was connected to the aic, the optical sensor was not recognized. troubleshooting was unsuccessful so the automated impella controller (aic) was replaced. the new aic operated without any issues. no patient harm reported. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
32401 | 1220648-2023-05358 | 18313225 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18313225&pc=OZD | 11/15/2023 | 03/19/2024 |  | 12/12/2023 |  |  | 11/15/2023 | 12/12/2023 |  |  | 08/31/2025 |  | 12/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024401048 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that as a 59-year-old male patient was coming off the utilized bypass, the motor current on the implanted impella 5.5 pump spiked and the pump stopped. multiple attempts were made unsuccessfully to restart the device and the decision was ultimately made to replace the device with a new pump via direct access. there was no patient harm. | [5 empty] | The investigation into the pump stop issue has been completed since the original report was filed. the device was returned and tested after arriving in fs. upon physical inspection of the returned pump, catheter kink was found near the red handle, and biomaterial was observed on the impeller and inlet cage. the root cause of the pump stop issue was biomaterial observed on impeller. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
32402 | 1220648-2023-05360 | 18313484 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18313484&pc=OZD | 11/15/2023 | 09/12/2024 |  | 12/12/2023 |  |  |  | 12/12/2023 |  |  | 09/30/2025 |  | 12/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024401569 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/15/2023 |  |  |  |  | The user facility reported a 61-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. bleeding was noted at arteriotomy. patient restless on table and physician noted calcium at insertion site on ultrasound. attempted manual pressure and angle matching with minimal results. the physician decided to remove impella. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access site bleeding was most likely patient condition since patient had calcium at insertion site on ultrasound. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- d6a/d6b, h6 component code 925 f6/f8 should have been left blank in the initial report.
32403 | 1220648-2023-05361 | 18313654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18313654&pc=OZD | 11/15/2023 | 12/12/2023 |  | 12/12/2023 |  |  |  | 12/12/2023 |  |  | 09/30/2025 |  | 12/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024398743 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/15/2023 |  |  |  |  | The user facility reported a 42-year-old male patient was scheduled for impella cp implant for mechanical circulatory support. during impella implant via left femoral access, wire placement was lost and attempts to dilate the artery resulted in a common iliac/femoral dissection. the affected area was ballooned for stent placement. manual pressure was held while the impella sheath was removed. the pump was subsequently implanted via the right femoral artery for patient support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon clinical review of the provided information, the cause of the noted issue was determined to be use related due to the reattempt to dilate the artery after wire access was lost. instructions for use for the related event are as follows: ¿potential adverse events (united states) arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation.¿ ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿
32404 | 1220648-2023-05362 | 18314280 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18314280&pc=OZD | 11/13/2023 | 09/03/2024 |  | 12/12/2023 |  |  | 11/13/2023 | 12/12/2023 |  |  | 05/31/2025 |  | 12/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024344550 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 66-year-old male noted as high risk coronary intervention was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported the patient had bleeding and a hematoma noted at impella access site. three units of blood products were administered. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h6 component code added g1 reporting contact fax number missing from manufacturer device report 1220648-2023-05362.
32405 | 1220648-2023-05363 | 18314433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18314433&pc=OZD | 12/11/2023 | 02/20/2024 |  | 12/12/2023 |  |  | 12/11/2023 | 12/12/2023 |  |  | 09/30/2025 |  | 12/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2024407020 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | Valvular Insufficiency/ Regurgitation | [11 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/11/2023 |  |  |  |  | The user facility reported a 85-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. complainant in japan reported the patient was on impella cp support and had mild aortic regurgitation (ar) before impella insertion, but severe aortic regurgitation (ar) occurred after impella removal. there was no valve leaflet damage, so it is unclear whether the ar was caused by impella. aortic valve replacement (avr) was performed for the ar. originally, open-heart surgery for ventricular septal perforation and cabg (coronary artery bypass grafting¿) was performed. avr was additionally implemented. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation into the heart valve damage has been completed. the device was not returned for investigation. the root cause of the heart valve damage was unable to be determined and the relationship to impella cannot be determined since no product or data logs were returned and there is a lack of clinical details.
32406 | 1220648-2023-05364 | 18314493 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18314493&pc=PYX | 12/10/2023 | 03/08/2024 |  | 12/12/2023 |  |  | 12/10/2023 | 12/12/2023 |  |  | 08/31/2025 |  | 12/12/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | RP FLEX W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2024397347 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | Hemolysis | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/10/2023 |  |  |  |  | The user facility reported a 66-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a rp flex placed to support a patient who had decompensated post mitral procedure. the rp flex was placed but, had signs of hemolysis. the rp was removed and replaced with this 2nd pump that was delivered but, stopped after mc rising. the rp was explanted after 31 minutes. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ | The investigation into the pump stop issue has been completed since the original report was filed. the device was returned and tested after arriving in fs. data logs show that purge pressure starts to rise immediately upon starting pump and continues to rise until the pump was explanted. the pump was unable to be restarted. upon investigation of the pump, biomaterial was observed wrapped around the impeller and the purge gap. the root cause of the pump stop was determined to be biomaterial.
32407 | 1220648-2023-05365 | 18314502 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18314502&pc=OZD | 11/17/2023 | 03/27/2024 |  | 12/12/2023 |  |  | 11/17/2023 | 12/12/2023 |  |  | 09/30/2025 |  | 12/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024405968 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/17/2023 |  |  |  |  | The user facility reported a 59-year-old male post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. the impella 5.5 needed to be explanted due to a pump stop. the patient was in icu after impella insertion, extubated and progressing well. then suddenly the patient mps dropped to the 50s and drips were increased. the patient was tachypneic. stat echo had been ordered to check heart function and impella position. echo showed large pericardial effusion. the heart failure team was at the bedside and began stripping chest tubes with significant draining coming out. labs sent and massive transfusion ordered. bleeding from chest tubes was constantly filling up the pleurovacs. abg came back with ph of 7.0 and patient was emergently intubated. heparin stopped and transfusions began. packed red blood cells, platelets and fresh frozen plasma and well as albumen given. effusion continued to return if stripping of the chest tubes stopped. the patient was taken to the or and as soon as chest open a large gush of blood came from the lower aspect of the anastomosis of the graft onto the aorta. the impella level was dropped to p1 and surgeon emergently pulled impella back in order to clap and control large amount of bleeding. the impella subsequently alarmed and the motor stopped. the graft was repaired but unable to use the current impella as it stopped and in the graft. the impella was removed, and surgeon made the decision to support the patient with inotropes and pressors as the graft was not usable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related events are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation into the heart valve damage and the pump stop issues has been completed since the initial report was submitted. the product was not returned for investigation. it was reported that there was significant bleeding from the graft's lower aspect onto the aorta. the cause of the pump stop issue was most likely biomaterial, based on data log review. the cause of the access site bleeding issue was not determined since product was not returned and sufficient clinical information was not provided.
32408 | 1220648-2023-05366 | 18314525 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18314525&pc=OZD | 11/21/2023 | 09/22/2024 |  | 12/12/2023 |  |  | 11/21/2023 | 12/12/2023 |  |  | 09/30/2025 |  | 12/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024402857 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  |  | Death | No | Yes | Pumping Stopped | [6 empty] | Insufficient Information | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/21/2023 |  |  |  |  | The user facility reported a 66-year-old male of unknown race and ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella had a pump stop alarm. two attempts to restart were unsuccessful. on the third attempt the pump started. the team wanted to continue with the therapy even though there was a risk of the pump stopping again and will closely monitor. | Additional information noted that the procedure outcome was that patient expired its unknow what was the reason. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop issue has been completed since the initial report was submitted. the product was not returned for investigation. the cause of the pump stop issue was not determined since product was not returned and sufficient clinical information was not provided. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. as noted in the impella ifu: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrected information for the initial mfr 1220648-2023-05366 was provided in sections b2, b5, d6, h1, and h6.
32409 | 1220648-2023-05367 | 18314530 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18314530&pc=OZD | 11/11/2023 | 04/04/2024 |  | 12/12/2023 |  |  | 11/11/2023 | 12/12/2023 |  |  | 06/30/2025 |  | 12/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024360106 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/21/2023 |  |  |  |  | The user facility reported a 35-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the echo was obtained, and left ventricle thrombus was observed. decision was made to remove the impella. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the thrombus issue has been completed since the initial report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined.
32410 | 1220648-2023-05368 | 18314543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18314543&pc=PYX | 08/12/2023 | 09/22/2024 |  | 12/12/2023 |  |  | 08/12/2023 | 12/12/2023 |  |  | 06/30/2024 |  | 12/12/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023232117 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/12/2023 |  |  |  |  | The user facility reported a 62-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. the impella rp was prepared in a normal fashion for patient implant. the right femoral vein access was obtained and serial dilation to 23fr sheath. flow directed catheter and .027 was unable to be passed into the left pa due to patients anatomy. attempts to place rp were aborted. | The procedural outcome noted that the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The investigation into the reported delivery issues- anatomy has been completed . product was not return for investigation. the cause of the delivery issue was most likely patient condition related challenging vasculature, based on provided clinical details. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-05368 was provided in sections b1, b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
32411 | 1220648-2023-05329 | 18300243 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18300243&pc=OZD | 12/07/2023 | 02/29/2024 |  | 12/11/2023 |  |  | 12/07/2023 | 12/11/2023 |  |  | 05/31/2025 |  | 12/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024356993 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/07/2023 |  |  |  |  | The user facility reported a 76-year-old male was implanted with an impella device for mechanical circulatory support. the us complainant had placed a cp for support of the patient admitted with a stemi (st elevation myocardial infarction). the patient was supported for a cabg (coronary artery bypass grafting) procedure. the pump had been crossclamped when in the or for the cabg for the multivessel cad (coronary artery disease). the pump stopped repeatedly on day 10 of the support. the pump was able to be restarted but was explanted. at the time of explant, the team additionally treated the limb with fogarty embolectomy to the sfa (superficial femoral artery) and iliac. | [5 empty] | The investigation into the reported pump stop and thrombus has been completed since the original report was filed. the pump was returned for investigation. no physical abnormalities were observed on the returned pump, except for a small amount of biomaterial observed in the inlet cage. the pump successfully passed the electrical testing. during the test in a bucket of water, the pump ran as expected. data logs showed several motor current spikes with impella stop-119 alarms. the pump continued to run for 4 hours after the pump stop events without any reported issues. the cause of the pump stop issue was not determined without sufficient clinical information. the cause of the thrombus issue was most likely patient condition related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
32412 | 1220648-2023-05330 | 18300248 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18300248&pc=OZD | 11/13/2023 | 08/26/2024 |  | 12/11/2023 |  |  | 11/13/2023 | 12/11/2023 |  |  | 08/31/2025 |  | 12/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024388830 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Female |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient expired while on support. per medical safety officer review use of the device cannot conclusively be associated with the outcome. | The complainant reported a 70-year-old female patient was implanted with an impella cp for mechanical circulatory support. however, blood was noted to be dripping out of the red impella plug. there was no reported patient harm. |  |  |  |  | The investigation of purge leak ¿ pump has been completed. the root cause of the product damage (hole in catheter) was not determined as no product was returned for inspection. b5 updated with additional information d6a and d6b added | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
32413 | 1220648-2023-05332 | 18302506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18302506&pc=OZD | 11/13/2023 | 09/19/2024 |  | 12/11/2023 |  |  | 11/16/2023 | 12/11/2023 |  |  | 04/30/2025 |  | 12/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2024339289 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/16/2023 |  |  |  |  | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported an 83-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on impella support, the patient underwent a chest hemostasis surgery due to bleeding. the relationship between the bleeding and the impella is unknown. the patient passed away the night of the surgery. the cause of death was not related to the impella. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-05332 was provided in sections b2, b5, d6a, d6b, h1, and h6. | The investigation into patient's vascular damage-peripheral vessel has been completed. no product or data was received from the customer. per the clinical investigation, the root cause of the vascular damage - peripheral vessel is most likely due to the patients condition.
32414 | 1220648-2023-05331 | 18302516 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18302516&pc=OZD | 09/11/2023 | 01/11/2024 |  | 12/11/2023 |  |  | 11/16/2023 | 12/11/2023 |  |  | 11/30/2024 |  | 12/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024301380 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 55 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | Thrombocytopenia | [10 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/16/2023 |  |  |  |  | The principal investigator of the protect iv study alleges a definite relationship for the use of the impella cp and patient experiencing renal failure, kidney failure, and thrombocytopenia. the impella cp had supported the patient to allow fro the coronary angioplasty. the cp had supported for 1.7 hours when then weaned and explanted. | [5 empty] | The investigation for the reported renal failure, kidney failure, and thrombocytopenia has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the reported thrombocytopenia was most likely patient condition as it occurred two months after treatment with impella. the relationship to the impella is unknown. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32415 | 1220648-2023-05333 | 18302540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18302540&pc=OZD | 11/17/2023 | 02/29/2024 |  | 12/11/2023 |  |  | 11/17/2023 | 12/11/2023 |  |  | 08/31/2025 |  | 12/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024394440 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/17/2023 |  |  |  |  | The user facility reported a 64-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support the patient developed cold left lower extremity and was brought back to the lab for ipsilateral distal perfusion cannula (dpc) with restoration of flow. the patient is now making purposeful movements, good impella flows, and good limb perfusion. patient was on device support for additional 2 days and is stable. | [5 empty] | The investigation into the limb ischemia ¿ reversible has been completed. the device and data were not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/pos, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32416 | 1220648-2023-05334 | 18302614 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18302614&pc=OZD | 11/13/2023 | 03/11/2024 |  | 12/11/2023 |  |  | 11/13/2023 | 12/11/2023 |  |  |  |  | 12/11/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1721936 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported that during preparation in the operating room, the circulating nurse spilled d5w on the cassette and automated impella controller (aic) while spiking the bag. the d5w was cleaned off and the pump was prepped per best practice. about 5-10 minutes into pump priming, a critical ¿pump system failure¿ alarm appeared. due to availability of secondary aic, the decision was made to switch aics before opening new purge cassette. aic-to-aic transfer made without incident. | [5 empty] | The investigation for the reported console "purge system failure" alarm has been completed. the console was returned for evaluation. the logs were reviewed which revealed multiple purge system failure errors. the console was evaluated finding that the purge motor would not fully rotate and would not correctly respond to commands from the purge unit driver. the most likely cause of the ribbon cable being incorrectly seated was that the cable header locking mechanism was defective and therefore allowed the ribbon cable to become loose. | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
32417 | 1220648-2023-05335 | 18302626 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18302626&pc=OZD | 11/13/2023 | 11/11/2025 |  | 12/11/2023 |  |  | 11/13/2023 | 12/11/2023 |  |  |  |  | 12/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1088165 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a battery failure on an automated impella controller (aic). there was a a battery failure (red alarm). the aic was replaced. there was no reported patient harm. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
32418 | 1220648-2023-05336 | 18302860 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18302860&pc=OZD | 11/12/2023 | 01/17/2024 |  | 12/11/2023 |  |  | 11/12/2023 | 12/11/2023 |  |  | 09/30/2025 |  | 12/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024396372 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/12/2023 |  |  |  |  | The user facility reported a 59-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella 5.5 insertion through the axillary artery met a lot of anatomy resistance. after multiple attempts the case was aborted. after removal of pump the catheter shaft of pump was kinked and bent. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported delivery issue ¿ use has been completed since the original report was filed. the root cause of the delivery issue was most likely due to patient condition.
32419 | 1220648-2023-05337 | 18307770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18307770&pc=OZD | 09/20/2023 | 04/25/2024 |  | 12/11/2023 |  |  | 11/11/2023 | 12/11/2023 |  |  | 08/31/2025 |  | 12/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024384964 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 24 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 11/11/2023 |  |  |  |  | Two notifications were received via abiomed¿s long-term outcome and quality indicator impella registry indicating a patient to have endured both ischemia and wide complex arrhythmia with a possible relationship to the impella device. regarding ischemia, the patient reportedly had "worsening lactic acidosis, likely from ischemic lower extremity with clotting of reperfusion sheath." as a result, the impella was removed on (b)(6) 2023 with improving lactic acidosis. the patient underwent a successful below knee amputation with vascular surgery on (b)(6) 2023. regarding wide complex arrhythmia that occurred on the same day as the reported ischemia, an amiodarone bolus was given as treatment. after the patient did not have any further episodes of wide complex rhythm on telemetry, amiodarone delivery was stopped. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported limb ischemia and arrhythmia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia and arrhythmia could not be determined.
32420 | 1220648-2023-05338 | 18307825 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18307825&pc=OZD | 09/27/2023 | 09/24/2024 |  | 12/11/2023 |  |  | 11/11/2023 | 12/11/2023 |  |  | 06/30/2025 |  | 12/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024372290 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Perforation of Vessels | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/11/2023 |  |  |  |  | Abiomed¿s long-term outcome and quality indicator impella registry notification received indicating a possible relationship between ventricular tachycardia and access site bleeding with the impella device. regarding ventricular tachycardia, it was alleged that on (b)(6) 2023, the "patient went down to cath lab for decannulation of ecmo and impella. course complicated by left groin bleeding and decreased right heart function with cv collapse requiring cpr. patient went into vtach then requiring shocks." | Additional information: the patient experienced a left groin bleed and was taken to the operating room for right femoral artery repair. the vascular surgeon discovered a full thickness injury to the left superficial femoral artery (sfa) and the left common femoral vein (cfv), both were repaired. the patient recovered with sequelae. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿assess access site for bleeding and hematoma.¿ | The investigation of vf/vt/af/at and access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined d6a and d6b added h6 component code added the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05338: d4 model number. f6/f8-should have been left blank. g1 reporting contact fax number missing. g1 reporting contact facility name missing. g2 report source company representative inadvertently was not selected. g4 pma/510(k)number. h6 health effect - impact code 4647 was inadvertently not included. | The root cause of the vascular damage was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ failure mode for vascular damage - peripheral vessel inadvertently was not included in manufacturer device report 1220648-2023-05338. b5 and h6 updated to include omitted failure mode.
32421 | 1220648-2023-05339 | 18307897 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18307897&pc=OZD | 11/16/2023 | 04/09/2024 |  | 12/11/2023 |  |  | 11/16/2023 | 12/11/2023 |  |  | 09/30/2025 |  | 12/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024397191 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 45-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the complainant reported the impella had high purge pressure/system blocked alarms. there was no kinks that were observed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation into the high purge pressure issue has been completed since the initial report was submitted. the device was not returned for investigation. as per clinical details, the doctor reported a high purge pressure alarm with no visible kink. the cause of the high purge pressure issue was not determined since product was not returned and limited clinical information was provided.
32422 | 1220648-2023-05340 | 18307903 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18307903&pc=PYX | 11/29/2023 | 12/11/2023 |  | 12/11/2023 |  |  | 11/29/2023 | 12/11/2023 |  |  | 06/06/2023 |  | 12/11/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2024345483 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Male |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/29/2023 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. rp fell out of place three times, creating positional issues. | [5 empty] | Pump was not returned for investigation. the cause of the positioning issue was not determined since product was not returned and sufficient clinical information was not available.
32423 | 1220648-2023-05341 | 18307906 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18307906&pc=OZD | 12/07/2023 | 04/12/2024 |  | 12/11/2023 |  |  | 12/07/2023 | 12/11/2023 |  |  | 10/31/2025 |  | 12/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024409454 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/07/2023 |  |  |  |  | The user facility reported a 66-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support due to not tolerating coming off bypass well. the impella 5.5 was not able to deliver more than 0.5lpm of support at p-2. anything above p-2 resulted in a pump output of 0lpm. a clot was suspected and the pump was removed. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the device was returned for evaluation. the pump could not be tested as it was returned with the cannula cut in half while the other half remained in the sheath. data logs were reviewed and do not indicate any biomaterial interactions. however, clinical information states the pump was implanted again after difficult insertion. the pump began experiencing low flow accompanying positioning alarms, as it was most likely in the improper position. therefore, the root cause of the positioning issue was use related. added-d9 device return date d4-updated model number | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32424 | 1220648-2023-05294 | 18290520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18290520&pc=OZD | 12/05/2023 | 07/16/2024 |  | 12/08/2023 |  |  | 12/05/2023 | 12/08/2023 |  |  | 09/30/2025 |  | 12/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024403747 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/05/2023 |  |  |  |  | The user facility reported a 67-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp placed to support a patient found down/unresponsive in the field and sent to the hospital with ems and cpr. the patient had the cp placed and had run for a support of 69hours. the pump was removed and found to have adhered to clot burden. during the support there were signs of limb ischemia and hemolysis, but no treatment was given. | [5 empty] | The investigation for the reported limb ischemia and the thrombus has been completed. no product was returned for investigation. the root cause of the thrombus was most likely patient condition related. the root cause of the limb ischemia could not be determined without additional clinical details. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ d4-corrected model number. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32425 | 1220648-2023-05295 | 18290562 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18290562&pc=OZD | 12/06/2023 | 04/15/2025 |  | 12/08/2023 |  |  | 12/06/2023 | 12/08/2023 |  |  | 08/31/2025 |  | 12/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2024387493 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Life Threatening | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/06/2023 |  |  |  |  | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility in japan reported an implant of an impella device for mechanical circulatory support. complainant in japan reported the patient had an impella cp implant and when the impella was temporarily pushed in there was a ventricle perforation that was caused by the impella. the surgeon refused to operate on the patient and repair the ventricle, impella support was continued, and blood transfusions were stopped. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the pump was stated to have been pushed into the left ventricle during removal of peel away; this area was said to have already had heart damage and was worsened by the pump. the root cause of the ventricle perforation was most likely use related as it was pushed deep during peel away removal. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
32426 | 1220648-2023-05296 | 18290571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18290571&pc=OZD | 12/02/2023 | 12/03/2024 |  | 12/08/2023 |  |  | 12/07/2023 | 12/08/2023 |  |  | 08/31/2025 |  | 12/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024384928 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/07/2023 |  |  |  |  | The user facility reported a 64-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support for 99 % blocked left anterior descending artery. the patient began having hemolysis and staff decided to remove the impella. reportable due to patient's hemolysis and necessary intervention. | [5 empty] | The investigation into hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05296. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
32427 | 1220648-2023-05297 | 18290593 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18290593&pc=OZD | 12/06/2023 | 02/16/2024 |  | 12/08/2023 |  |  | 12/06/2023 | 12/08/2023 |  |  | 08/31/2025 |  | 12/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024395933 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Pumping Stopped | Improper Flow or Infusion |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/06/2023 |  |  |  |  | The user facility reported a 69-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp placed for a patient with multiple cardiac comorbidities, such as aaa (abdominal aortic aneurysm), ai (aortic insufficiency) and cad (coronary artery disease). the cp was placed for support when a 5.5 pump failed to deliver. the cp ran for 5 days when it stopped. the pump stop occurred after saturated flows. the patient survived the pump stop. | [5 empty] | The investigation into the pump stop and the saturated flow issues has been completed since the original report was filed. the product was returned for investigation. no physical abnormalities were observed on the returned pump, and no biomaterial was observed on the impeller or cannula kink. saturated flow was reproduced while running in a bucket of water without any purge-related difficulties. upon further teardown of the motor, biomaterial was observed in the motor housing. the cause of the pump stop issue was biomaterial observed inside the motor housing, which also led to saturated flow during the case. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
32428 | 1220648-2023-05298 | 18290614 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18290614&pc=OZD | 12/01/2023 | 07/22/2024 |  | 12/08/2023 |  |  | 12/01/2023 | 12/08/2023 |  |  | 09/10/2025 |  | 12/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024398976 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Handling Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/01/2023 |  |  |  |  | The user facility reported a 70-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had admitted a patient with extensive cardiac and systemic comorbidities and covid +. the team knew of the pad and had issues with access and sheath placement. the long 25cm sheath kinked and was replaced. a new sheath was placed, and the cp delivered for pci (percutaneous coronary intervention) of the multivessel coronary artery disease. | [5 empty] | The investigation for the introducer damage has been completed. the impella device was not returned for evaluation. per the given clinical information, patient had calcified vessels. the root cause of the introducer damage issue was patient condition related to small/ diseased vessels. d4-corrected model number added- d6a/d6b f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32429 | 1220648-2023-05299 | 18292808 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18292808&pc=OZD | 11/13/2023 | 12/17/2024 |  | 12/08/2023 |  |  | 11/13/2023 | 12/08/2023 |  |  | 05/31/2025 |  | 12/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024369475 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | Material Deformation |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/13/2023 |  |  |  |  | The user facility reported a 53-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the impella cp was inserted without complication, but impella would only achieve average flows of 1.4l. when checking position, the catheter appeared kinked about 1.5 cm below the motor housing. the physician tried several times to unkink the cp and was unsuccessful. the decision was made to explant and place a new device. the patient¿s left ventricle was small, and the aortic valve had moderate stenosis. upon visual inspection the catheter was intact. there was no harm to the patient | [5 empty] | The investigation into product damage (cannula kink) has been completed. the pump was not returned for investigation. the data log shows low pump flow during activation, with an active suction alarm. as per the clinical notes, a cannula kink was reported due to a small left ventricle (lv). the product damage (cannula kink) failure mode was changed to low pump flow as investigated failure mode. the low pump flow issue was most likely caused by the patient's condition, specifically a small ventricle, which led to a kink in the cannula, based on the given clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32430 | 1220648-2023-05300 | 18294515 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18294515&pc=OZD | 11/11/2023 | 01/15/2024 |  | 12/08/2023 |  |  | 11/11/2023 | 12/08/2023 |  |  |  |  | 12/08/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1397383 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/11/2023 |  |  |  |  | The user facility reported a 70-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. automated impella controller (aic) stuck in priming purge cassette screen and the team was unable to resolve. swapped to back up console and all resolved. no patient harm reported due to the issue. | [5 empty] | The investigation for the reported console purge issue has been completed. the console was returned for evaluation. upon visual inspection, there were signs of glucose ingress on the purge motor shaft, which could cause the purge motor shaft to be unable to rotate which would prevent the cassette from successfully deairing. however, the failure mode was unable to be reproduced with multiple purge cassettes, as they were able to run without issue. review of the data showed that the cassette did not finish deairing, so it was removed and reinserted. this did not allow the purge to complete deairing, and additional attempts did not change the results. additionally, there was a single piston blocked alarm. when the console was swapped for a known good console, the procedure was able to be completed with purge cassette. therefore, the cause of the purge issue was most likely a glucose ingress. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
32431 | 1220648-2023-05301 | 18294536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18294536&pc=OZD | 11/11/2023 | 03/05/2024 |  | 12/08/2023 |  |  | 11/11/2023 | 12/08/2023 |  |  | 09/30/2025 |  | 12/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024398196 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | Activation, Positioning or Separation Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/11/2023 |  |  |  |  | The user facility reported a 37-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the impella pump was unable to pass through the anastomoses. a new impella was placed successfully. no known patient harm or injury. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the cause of the first 5.5 pump delivery issue was most likely use related due to improper technique, as the patient had anastamoses per clinical details. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32432 | 1220648-2023-05302 | 18294548 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18294548&pc=OZD | 11/11/2023 | 01/11/2024 |  | 12/08/2023 |  |  | 11/11/2023 | 12/08/2023 |  |  | 07/31/2025 |  | 12/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024380186 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp placed for support of a ami/cgs (b)(6) male patient. the pump was supporting when on day 10 the pump stopped and despite all measures and troubleshooting the pump could not be re-started. the team explanted the pump and did not replace the pump as the patient was stable. | [5 empty] | The investigation into the reported pump stop has been completed since the original report was filed. impella cp pump was returned for evaluation. it was visually inspected and a large purge gap with bearing wear failure was found. no other abnormalities such as biomaterial, damaged outflow, were found. as per clinical, the impella cp had pump stop during 11th day of support. data log was reviewed, and high motor current pump stop was observed on 11th day of support. the cause of the pump stop was bearing wear with large purge gap observed per field investigation and data log review. the impella cp optical device is not indicated for support more than 8 days of design life period according to design trace matrix. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
32433 | 1220648-2023-05303 | 18294604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18294604&pc=OZD | 11/11/2023 | 04/16/2024 |  | 12/08/2023 |  |  | 11/11/2023 | 12/08/2023 |  |  | 05/31/2025 |  | 12/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024356463 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/11/2023 |  |  |  |  | The user facility reported a 75-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after the pump was placed in the left ventricle and started in auto mode, the impella cp was not able to flow and had high motor current. clot ingestion was assumed and immediately replaced with a new impella cp. there was no harm to the patient. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the reported low pump flow has been completed. the cp pump was visually inspected. no cannula kink observed. no broken blades found. biomaterial around impella and purge gap observed. the cause of the low pump flow issue was biomaterial ingestion. the instructions for use states ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿
32434 | 1220648-2023-05304 | 18294746 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18294746&pc=OZD | 11/11/2023 | 01/11/2024 |  | 12/08/2023 |  |  | 11/11/2023 | 12/08/2023 |  |  | 07/31/2025 |  | 12/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024389524 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/11/2023 |  |  |  |  | The user facility reported a 71-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the patient developed left ventricle thrombus. a decision was made to change to iabp and explant impella to avoid aspirating thrombus into the pump. impella removed without any complications and patient is stable post explant. | [5 empty] | The investigation into the thrombus has been completed. the device was not returned for benchtop evaluation and investigation. per the clinical information provided, the root cause of the thrombus was determined to be patient condition because thrombus was observed in the left ventricle. no issues with the function were reported and the impella was removed as a precaution. the relationship between the thrombus and the impella was unknown. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32435 | 1220648-2023-05306 | 18294825 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18294825&pc=OZD | 11/09/2023 | 10/07/2024 |  | 12/08/2023 |  |  | 11/09/2023 | 12/08/2023 |  |  | 06/30/2025 |  | 12/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024374897 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a (b)(6) male noted as high risk coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient was on impella cp support and there were placement signal alarms. tte performed multiple times, device repositioned, impella was in decent position per providers, but due to patient anatomy staff could not get in most optimal position. additionally, the patient was given albumin for suction alarms, and the pump was repositioned multiple times. | [5 empty] | The investigation for the reported positioning issue has been completed. no product was returned for investigation. data logs confirm placement signal low alarms. the 5s parameters were found to be stable and within normal ranges. impella position in ventricle alarms were observed at the same time as suction alarms. the root cause of the positioning issues was most likely patient condition related due to the patient's challenging anatomy. added- h6 impact code 4628 corrections to the initial mfr report: added- d6a/d6b. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
32436 | 1220648-2023-05305 | 18294826 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18294826&pc=OZD | 11/11/2023 | 07/02/2024 |  | 12/08/2023 |  |  | 11/11/2023 | 12/08/2023 |  |  | 09/30/2025 |  | 12/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024402858 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/11/2023 |  |  |  |  | Additional information: the patient's care was withdrawn and the patient expired. per abiomed medical safety office, abiomed does not have enough information to associate use of device with outcome. | The user facility reported a 59-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding from their access site during impella cp support. as a result of the condition, systemic heparin was put on hold. the condition was able to resolve and support with the implanted pump continued. |  |  |  |  | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding issue was anticoagulation management due to high patient activated clotting time values; heparin was stopped to achieve hemostasis as per the clinical description. b5 additional information added d6a, d6b added f6 and f8 should be blank h6 added component code corrected information for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32437 | 1220648-2023-05307 | 18295127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18295127&pc=OZD | 05/23/2023 | 03/10/2025 |  | 12/08/2023 |  |  | 11/10/2023 | 12/08/2023 |  |  | 01/31/2025 |  | 12/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024333640 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/10/2023 |  |  |  |  | Abiomed¿s long-term outcome and quality indicator impella registry notification received 2023-11-10 indicating a possible relationship between hemolysis and this impella device: according to the report, the 72 year old female patient presented with cardiomyopathy and was placed on impella support. during support, hemolysis was noted, as evident by rising ldh and pfhg levels. she had received 1 unit of prbc and switched to dobutamine, added epinephrine and norepinephrine. | [5 empty] | No product was returned. data logs revealed there were no abnormal pump performance metrics or alarms on the day of the reported hemolysis. the cause of the hemolysis was not determined since neither the product nor sufficient clinical information were provided. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32438 | 1220648-2023-05308 | 18295588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18295588&pc=OZD | 11/03/2023 | 08/20/2024 |  | 12/08/2023 |  |  | 11/03/2023 | 12/08/2023 |  |  | 04/30/2025 |  | 12/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024353444 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/17/2023 |  |  |  |  | The user facility reported a 62-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. when patient was being moved off the table the pump was pulled out of position. physician opted to perform a femoral cutdown and remove the pump. the malposition prompted the team to explant the pump .patient is stable post explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another" | The investigation for the positioning issue has been completed. according to the clinical details, the patient was taken to the or emergently and while moving the patient off of the cath table the pump was pulled back into the aorta. the decision was made to go to the or and place the pump back across the valve at the end of the procedure. however, after the procedure was performed the physician changed their mind and opted to perform a femoral cutdown and remove the pump. no product was returned for an investigation. data log analysis was not necessary for this complaint. the most likely cause of the positioning issue was patient movement related. the instructions for use for the impella cp with smartassist states the following: section: general patient considerations: when transferring a patient with the device in place: be careful not to pull on the impella catheter when transferring a patient from one bed to another added - f6/f8 should have been left blank in the initial report.
32439 | 1220648-2023-05309 | 18297270 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18297270&pc=OZD | 11/08/2023 | 02/28/2024 |  | 12/08/2023 |  |  | 11/08/2023 | 12/08/2023 |  |  | 07/31/2025 |  | 12/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024378453 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 63 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/08/2023 |  |  |  |  | The user facility reported a 63-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient arrived in the operating room from the intensive care unit emergently for impella 5.5 insertion. after the impella was placed, the patient had access site bleeding and heparin was paused. | [5 empty] | The investigation into the access site bleeding - major has been completed. the device was not returned for investigation. the root cause of the access site bleeding - major was not determined since product was not returned and there is a lack of clinical details. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32440 | 1220648-2023-05310 | 18297272 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18297272&pc=OZD | 11/08/2023 | 10/24/2025 |  | 12/08/2023 |  |  | 11/09/2023 | 12/08/2023 |  |  |  |  | 12/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1465594 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/09/2023 |  |  |  |  | The user facility reported a 52-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the automated impella controller (aic) displayed a "controller error" alarm and then the alarm disappeared for approximately one minute. during that time, the position waveform also stopped appearing. the aic was swapped out with another aic successfully. the aic will be returned for evaluation. there was no patient harm related to this event. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
32441 | 1220648-2023-05311 | 18297275 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18297275&pc=OZD | 11/09/2023 | 01/26/2024 |  | 12/08/2023 |  |  | 11/09/2023 | 12/08/2023 |  |  | 08/31/2024 |  | 12/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023257631 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Injury | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported patient of unknown age received hemodynamic support from an impella cp smartassist heart pump. following implant, a purge system alarm was reported. there was no y connector in the system and no leak noted. the automated impella controller (aic) was exchanged without success. after exchange, fluid was noted to be leaking from the purge filter. pump change was also recommended, but the recommendation was initially not followed. pump was ultimately explanted the day after. there was no reported patient harm. | [5 empty] | The investigation into the purge leak ¿ pump has been completed. the device was not returned for benchtop investigation. based on the clinical information provided, doctor reported fluid was leaking from purge filter and pump was explanted next day morning. the cause of the purge leak issue was most likely a leak from the sidearm, but the exact location of the leak was unknown since product was not returned and due to insufficient clinical information. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
32442 | 1220648-2023-05312 | 18297312 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18297312&pc=OZD | 11/09/2023 | 01/11/2024 |  | 12/08/2023 |  |  | 11/09/2023 | 12/08/2023 |  |  |  |  | 12/08/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1115981 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/09/2023 |  |  |  |  | No patient involvement: the us complainant had a console being used for training. the aic (automated impella controller) had purge disc failure with multiple pc during the training. | [5 empty] | The investigation for the reported console purge disc/flag issues has been completed. the console and data logs were returned for evaluation. upon visual inspection of the console, the purge flag was found broken. the data logs were reviewed and confirmed not being able to detect the purge disc. there were multiple instances of purge disc not detected alarms on the reported occurrence date. therefore, the root cause of the reported issue was due to a broken purge flag. updated: d4 serial number. | The impella device was received from the customer and an evaluation of the device is underway. upon investigation closure, a supplemental mdr will be filed.
32443 | 1220648-2023-05314 | 18297314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18297314&pc=OZD | 11/09/2023 | 10/08/2024 |  | 12/08/2023 |  |  | 11/09/2023 | 12/08/2023 |  |  | 09/30/2025 |  | 12/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024394581 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Improper Flow or Infusion | Material Deformation | Positioning Problem |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/09/2023 |  |  |  |  | The user facility reported a 61-year-old male undergoing coronary artery bypass grafting (cabg) was implanted with an impella 5.5 device for mechanical circulatory support. after the triple-vessel cabg was completed, the pump was withdrawn into the aorta as there was difficulty achieving an adequate seal with aortic cross-clamp. the pump was re-advanced across valve wirelessly with some difficulty. after adequate pump position was achieved, several kinks were noted in the catheter portion of the pump. also noted high purge pressures in the 900¿s. due to difficulty in positioning and with patient¿s dependence on the pump, team agreed to continue use with this pump with close monitoring of purge pressures going forward. the patient was stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the high purge pressure, positioning issue, and the product damage has been completed. as per the clinical information provided, the pump was tried to be wirelessly repositioned which lead to several catheter kinks. therefore, the root cause of the product damage issue was a catheter kink due to improper technique. clinical mentions that difficulty was experienced to achieve an adequate aortic cross clamp seal leading to pump pulled in aorta after a cabg procedure. data logs confirm "pump in aorta" alarm. therefore, the root cause of the positioning issue was use related to improper technique. data logs were reviewed and the chain of events observed in conjunction to clinical details were as follows - pump was pulled in aorta and was repositioned, followed by suction event and motor current spike, followed by a gradual rise in purge pressure (over the span of approx. 9 hours) leading to high purge pressure and purge system blocked alarms. therefore, the root cause of the high purge pressure issue was most likely biomaterial. the instructions for use for the impella 5.5 with smartassist states the following: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ added- h6 impact code 4628 correction to the initial mfr report: added- d6a/d6b and h6 med dev prob code 2954 changed to 1491.
32444 | 1220648-2023-05315 | 18297315 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18297315&pc=OZD | 11/09/2023 | 03/11/2025 |  | 12/08/2023 |  |  | 11/09/2023 | 12/08/2023 |  |  | 08/31/2025 |  | 12/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024395077 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/09/2023 |  |  |  |  | The user facility reported a 61-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was impella cp support and the patient sat up during percutaneous coronary intervention and access site bleeding occurred. a femostop was placed right when the patient arrived to icu (after pci / percutaneous coronary intervention), and four units of blood were given to the patient. vascular surgical repair was performed on the artery and the bleeding was resolved. the site looked good and patient had distal pulses. reportable due to vascular damage and necessary intervention. | [5 empty] | Introducer was not returned for investigation. important details related to the introducer placement and leading to the vascular damage is not available. therefore, the root cause of the vascular damage - peripheral vessel issue was not determined since product was not returned and insufficient clinical information was provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
32445 | 1220648-2023-05313 | 18297346 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18297346&pc=OZD | 11/09/2023 | 12/03/2025 |  | 12/08/2023 |  |  | 11/09/2023 | 12/08/2023 |  |  |  |  | 12/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1664272 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/09/2023 |  |  |  |  | The user facility's biomedical engineer reported the gray piece broke off of the automated impella controller (aic) when cassette was being removed. there was no patient involvement. the aic was sent back for evaluation and repair. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Data analysis: there was no evidence in the logs to confirm the reported failure. device analysis: the console was tested upon arrival at fs, and the purge disc retainer was confirmed to be broken. this was reported as ¿grey piece broke off aic (b)(4).¿ root cause: the root cause of the broken retainer was a defective retainer. d2 common device name updated. d4 model number updated. f6 and f8 should have been left blank. g1 reporting contact email updated.
32446 | 1220648-2023-05316 | 18297349 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18297349&pc=OZD | 11/10/2023 | 12/08/2023 |  | 12/08/2023 |  |  | 11/10/2023 | 12/08/2023 |  |  |  |  | 12/08/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1461597 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) did not complete priming. the purge cassette was installed during priming, however the priming did not start. the purge cassette was replaced twice and it was unsuccessful. the aic was replaced successfully. | [5 empty] | The investigation into the reported "aic - purge issue - other" has been completed. product was returned for investigation. the console was tested after arriving in fs. the purge flag was confirmed to be missing. data analysis indicates that ic7896 tried to start case wizard on november 10th but was unable to detect the purge disc repeatedly. this confirms the complaint because, if the purge disc could not be detected, then the case wizard could not go from step 2 to step 3 and therefore priming could not start. the root cause of the issue was determined to be a missing purge flag.
32447 | 1220648-2023-05318 | 18297356 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18297356&pc=OZD | 11/10/2023 | 10/04/2024 |  | 12/08/2023 |  |  | 11/10/2023 | 12/08/2023 |  |  | 09/30/2025 |  | 12/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024397079 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/10/2023 |  |  |  |  | The user facility reported a 73-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient kept moving their legs on impella support and the patient had access site groin bleeding. a femstop was placed. additionally, the patient had extreme hemolysis. 3 units of blood were given for the hemolysis and access site bleeding. additionally, the bleeding continued, and 4 more units were given. reportable due to patient harm and necessary intervention. | [5 empty] | The investigation for access site bleeding and hemolysis has been completed. the impella device was not returned for evaluation. the root cause of the hemolysis issue was not determined since product was not returned and insufficient clinical information was provided. the root cause of the access site bleeding issue was most likely patient movement. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05318. d4 model number. d6a and d6b revised from the initial submission in accordance with updated procedures. g1 reporting contact fax number missing. h6 component code revised from the initial submission in accordance with updated procedures. h6 code 1884 was reported incorrectly. new code was added to health effect - clinical code and health effect - impact code. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma. remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site. potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it. obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine. hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis. suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure. patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
32448 | 1220648-2023-05327 | 18297363 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18297363&pc=OZD | 11/12/2023 | 03/04/2024 |  | 12/08/2023 |  |  | 11/12/2023 | 12/08/2023 |  |  | 08/31/2025 |  | 12/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024386820 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/12/2023 |  |  |  |  | The user facility reported an 84-year-old male post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. on day three post surgery, the team restarted the anticoagulant heparin. once heparin was added, the patient developed a hematoma. the heparin was again stopped. the patient remained on impella support and was successfully weaned two days later. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed. the device was not returned for investigation. as per the clinical details, the cause of the access site bleeding issue was high patient act values since the bleeding was reported after administering heparin and was resolved after stopping heparin and applying a sandbag, based on given clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32449 | 1220648-2023-05328 | 18297365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18297365&pc=OZD | 11/12/2023 | 10/17/2024 |  | 12/08/2023 |  |  | 11/12/2023 | 12/08/2023 |  |  | 09/30/2025 |  | 12/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024400674 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [5 empty] | 7FR X 45CM TERUMO PINNACLE DESTINATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/12/2023 |  |  |  |  | The user facility reported a 72-year-old male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. there was a left femoral artery dissection and a peel away sheath was placed tamponading the area to address the vascular damage. staff stated that they would place a covered stent upon explant. patient is stable post this issue. | [5 empty] | The investigation for the reported vascular damage has been completed. the impella was not returned for evaluation. however, clinical details were sufficient to determine the root cause. the root cause of vascular damage was most likely due to use issue since non-abiomed recommended product (sheath) was used. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ added-d10-concomitant device, h6 component code 925 corrections to the initial mfr report: added- d6a/d6b f6/f8 should have been left blank. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
32450 | 1220648-2023-05322 | 18297376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18297376&pc=OZD | 11/11/2023 | 04/30/2024 |  | 12/08/2023 |  |  | 11/11/2023 | 12/08/2023 |  |  | 09/30/2025 |  | 12/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024403144 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 71-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that when the impella cp was explanted, the doctor noticed there was no bleed back from the canula or at the groin site. no distal pulses were found on the doppler so vascular was called in to complete the repair. it was also noted that thrombus was observed in the patient or on the device. | [5 empty] | The investigation for the reported limb ischemia and thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia and thrombus could not be determined as the device was not returned nor sufficient clinical details. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32451 | 1220648-2023-05323 | 18297382 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18297382&pc=OZD | 11/11/2023 | 04/01/2024 |  | 12/08/2023 |  |  | 11/11/2023 | 12/08/2023 |  |  | 09/30/2025 |  | 12/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024404522 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/11/2023 |  |  |  |  | The user facility reported a 73-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella cp was removed due to access site bleeding that occurred during transport. patient is stable post removal. | [5 empty] | The investigation into the access site bleeding ¿ major has been completed since the initial report was submitted. the product was not returned for investigation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32452 | 1220648-2023-05325 | 18297383 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18297383&pc=OZD | 11/12/2023 | 12/12/2024 |  | 12/08/2023 |  |  | 11/12/2023 | 12/08/2023 |  |  | 08/31/2025 |  | 12/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024396409 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/12/2023 |  |  |  |  | The user facility reported a 64-year-old male (race unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient started to have pink urine output in the foley. patient coded and was given tissue plasminogen activator. the clinical representative was not able to determine whether or not the pink urine was from the impella or a traumatic foley. | [5 empty] | The hemolysis has been completed. pump was returned in damaged condition (catheter was cut off and missing pump plug). visual inspection found no issues on the returned parts. no biomaterial or any other abnormalities were observed. the pump was unable to be run or further tested due to the cut catheter. data logs were investigated and no issues were observed. therefore, the root cause of the hemolysis issue was not determined since product and data log investigation was inconclusive. added h6 component code 925, h6 impact code 4644 corrections to the initial mfr report: added d6a/d6b f6/f8 should have been left blank. g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
32453 | 1220648-2023-05317 | 18297394 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18297394&pc=OZD | 11/10/2023 | 07/12/2024 |  | 12/08/2023 |  |  | 11/10/2023 | 12/08/2023 |  |  | 11/30/2023 |  | 12/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023240246 |  | 005042-RO |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 94 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/10/2023 |  |  |  |  | It was further reported that the patient expired and no other details were provided around the death. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there was not enough information to associate use of the device with the outcome. | The user facility reported a 94-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella 2.5 device for mechanical circulatory support. the patient had the 2.5 pump for the multi-vessel intervention. the pump support was for 30 minutes when the pump was explanted and the patient suffered from ventricular fibrillation, coded, and expired. the physician noted the death was not related to the impella 2.5. |  |  |  |  | The investigation into the vf/vt/af/at (chest pain and suffered ventricular fibrillation and cardiac arrest) issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary information was not provided, the root cause of the vt/vf was not determined. . | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32454 | 1220648-2023-05319 | 18297407 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18297407&pc=OZD | 11/10/2023 | 09/26/2025 |  | 12/08/2023 |  |  | 11/10/2023 | 12/08/2023 |  |  | 08/31/2025 |  | 12/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024382972 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | Material Deformation | [5 empty] | Unspecified Vascular Problem | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/10/2023 |  |  |  |  | The user facility reported a 57-year-old male post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. during support, there was damage/kinking at the red impella plug. the three-point fixation was not intact and nurse re-secured it. it was further reported placement signal not reliable alarm. the nurse disabled audio and will be monitoring the motor current flows. | Decision made to withdraw care, and the patient expired. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into product damage (hole in catheter) and the placement signal issues has been completed. the data stick was returned, but the device was not returned for investigation. based on the clinical information provided, the root cause of the optical signal issue was most likely damaged optical fiber line per log and clinical review. | B1 adverse event updated b2 updated to death b5 updated additional information for death h6 updated to include optical signal coding, death corrections that were made from the initial manufacturer device report number 1220648-2023-05319. b7 other relevant history, including preexisting medical conditions updated. d6a / d6b updated. d10 concomitant medical products and therapy dates updated. f6 and f8 should have been left blank in the initial report. g1 reporting contact email updated.
32455 | 1220648-2023-05321 | 18297413 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18297413&pc=OZD | 11/11/2023 | 09/21/2024 |  | 12/08/2023 |  |  | 11/11/2023 | 12/08/2023 |  |  |  |  | 12/08/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1635221 |  | 0042-0040 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Computer Operating System Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 11/11/2023 |  |  |  |  | The user facility reported a 69-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp support that was ongoing when the pump stopped on the 2nd day of support. the cp stopped when the controller had a failure. the failure prompted the team to remove the console and the patient expired. the console swap had occurred. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b1, b2, b5, h1, and h6. | The investigation into the reported aic -controller error has been completed since the original report was filed. the analysis of data showed the case associated with the reported complaint was started on 10th nov 2023 (console: ic9473, pump sn: (b)(6)), and on 11th nov 2023, when the pump was running, the console received the following logs before the ¿impella stopped. controller failure (pmd comm loss) ¿ alarm¿, event set #1047, which confirms the reported failure. the console was tested on an engineering lab bench. the reported failure symptoms were reproduced. upon manually unplugging the 20 v dc power cable at j001.2 while the pump was running, the same sequence of logs and the controller failure alarm (event set #1047) were reproduced as they occurred on the complaint date. no issue was found with the ribbon cable between impellatronic and the 4-in-1. found a hair crack in the purge retainer (unrelated to the reported failure). the root cause for the controller error was an intermittent 20 v dc power interruption, most likely due to either pbm malfunction or impellatronic malfunction. correction : section d1 /d2 : the brand name and the common device name was submitted incorrect inadvertently in the initial report which has been corrected to this report.
32456 | 1220648-2023-05320 | 18297420 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18297420&pc=OZD | 11/11/2023 | 09/20/2024 |  | 12/08/2023 |  |  | 11/11/2023 | 12/08/2023 |  |  | 07/31/2025 |  | 12/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024387401 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 69-year-old male was implanted with an impella cp device for mechanical circulatory support. impella cp support was ongoing when the pump stopped on the second day of support. the cp stopped when the controller had a failure. the failure prompted the team to remove the console and the patient expired. the console swap had occurred. | Additional information was provided that noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | This complaint was alleged against the pump and the investigation confirmed the pump stop was due to a confirmed automated impella controller (aic) related issue. the investigation can be found in manufacturer device report number 1220648-2023-05321. a regulatory report is no longer necessary for the impella cp pump. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-05320 was provided in sections b1, b2, b5, d6a, d6b, h1, and h6.
32457 | 1220648-2023-05324 | 18297422 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18297422&pc=OZD | 11/11/2023 | 12/11/2024 |  | 12/08/2023 |  |  | 11/11/2023 | 12/08/2023 |  |  | 09/30/2025 |  | 12/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024400767 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | [6 empty] | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/11/2023 |  |  |  |  | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient became restless and agitated overnight and pulled impella out of left ventricle. the impella was explanted at bedside with good hemostasis observed. the patient is stable. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. based on the clinical details the root cause of the positioning issue is most likely due to use as the patient pulled the impella out. corrections to the initial mfr report: d4 model number added d6a/d6b f6/f8 should have been left blank. g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32458 | 1220648-2023-05326 | 18297425 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18297425&pc=PYX | 11/12/2023 | 12/17/2024 |  | 12/08/2023 |  |  | 11/12/2023 | 12/08/2023 |  |  | 02/28/2025 |  | 12/08/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024333408 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/12/2023 |  |  |  |  | The user facility reported a 67-year-old male with cardiomyopathy was implanted with an impella rp for mechanical circulatory support. during support, purge pressure alarms occurred. purge fluid was changed to d5w with 25 units of heparin. pressure did drop to 500s but flow was still low. team checked purge tubing for kinks and none were noted. alteplase purge infusion was started per protocol. purge flow improved to 6 ml/hr after 30 minutes. the patient was weaned and impella rp was explanted two days later. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿¿ do not clean with or expose any part of the clear sidearm of the impella catheter (eg, infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir.¿ | The investigation for the purge leak has been completed. the pump was not returned for investigation. the total case run time was 17 days. the data log shows a motor current spike on the second day of the implant. two days before explant, high purge pressure was noted, with a gradual rise in purge pressure, which then returned to normal within 1 hour with a gradual decrease trend. the same event occurred two more times during the case. as per the clinical notes, high purge pressure returned to normal after heparin was added to the purge solution. the reported failure mode of purge leak was changed to high purge pressure after investigation. the cause of the high purge pressure issue was most likely biomaterial, based on data log review and provided clinical details. corrections to the initial mfr report: added d6a/d6b f6/f8 should have been left blank. g1 mdr reporting contact email.
32459 | 1220648-2023-05286 | 18284218 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18284218&pc=OZD | 11/12/2023 | 09/13/2024 |  | 12/07/2023 |  |  | 11/12/2023 | 12/07/2023 |  |  | 08/31/2025 |  | 12/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024398357 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/12/2023 |  |  |  |  | The user facility reported a 77-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. on the second day of support, the team withheld heparin due to an access site bleed. the pump remained on for six more days and was then weaned and removed. | [5 empty] | The investigation for access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05286. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. h6 component code added; revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32460 | 1220648-2023-05287 | 18285624 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18285624&pc=OZD | 12/01/2023 | 02/10/2025 |  | 12/07/2023 |  |  | 12/01/2023 | 12/07/2023 |  |  |  |  | 12/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1356952 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 12/01/2023 |  |  |  |  | The us complainant had an aic (automated impella controller) with a battery failure, red alarm, as reported by the hospital biomed team. there was no patient use and no harm or injury was possible. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - battery failure (red alarm) has been completed. according to the device and data analysis the root cause of this issue was an internal cell imbalance found in battery 1. the following have been reported incorrectly or missing from the manufacturer device report 1220648-2023-05287. d2 common device name revised in accordance with updated procedures. d4 model number. e2 health professional should have been no. f6/f8-should have been left blank. g1 reporting contact email revised in accordance with updated procedures. g1 reporting contact fax number missing. g4 pma/510(k)number. h3 device not returned to manufacturer for evaluation; should have been no.
32461 | 1220648-2023-05288 | 18287628 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18287628&pc=OZD | 11/07/2023 | 12/07/2023 |  | 12/07/2023 |  |  | 11/07/2023 | 12/07/2023 |  |  |  |  | 12/07/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1150343 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 11/07/2023 |  |  |  |  | The user facility reported that during a case preparation, the automated impella controller, (aic) had a battery failure . a replacement loaner was sent to the facility. no patient is involved. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
32462 | 1220648-2023-05289 | 18287639 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18287639&pc=OZD | 11/07/2023 | 02/10/2025 |  | 12/07/2023 |  |  | 11/08/2023 | 12/07/2023 |  |  |  |  | 12/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1376689 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Mechanical Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller (aic) had a rattling noise upon start up. the device will be inspected by field service team. there was no reported patient harm. | [5 empty] | The investigation for the automated impella controller (aic) damage has been completed. data logs were reviewed but were not necessarily for the complaint. the console was returned for evaluation and the failure was confirmed for unusual sound coming from the fan assembly of the console. the ball bearing of the fan was inspected and found to have unknown debris. the root cause of the unusual sound was debris of unknown origin within the ball bearing of the defective fan. the cause of the aic issue was contamination. a.3 revised as selection was inadvertently omitted from the initial manufacturer device report number 1220648-2023-05289. b.3 revised as date of event was previously entered incorrectly on the initial manufacturer device report number 1220648-2023-05289. d.1 and d.2 revised brand name and common device name since initial manufacturer device report number 1220648-2023-05289 was submitted in accordance with updated procedures. d.4 revised model number, catalog number and serial number were previously entered incorrectly on the initial manufacturer device report number 1220648-2023-05289. d.9 revised as date device was returned was previously entered incorrectly on the initial manufacturer device report number 1220648-2023-05289. e.1 revised address line 2 as it was inadvertently omitted from the initial manufacturer device report number 1220648-2023-05289. g.1 revised reporting contact email since he initial manufacturer device report number 1220648-2023-05289 was submitted in accordance with updated procedures. revised reporting contact fax number as it was inadvertently omitted from initial manufacturer device report number 1220648-2023-05289. | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
32463 | 1220648-2023-05290 | 18287725 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18287725&pc=OZD | 11/13/2023 | 09/12/2024 |  | 12/07/2023 |  |  | 11/13/2023 | 12/07/2023 |  |  | 08/31/2025 |  | 12/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024389404 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 63-year-old male in st elevated myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. the complainant had a 5.5 support ongoing for care escalation from the impella cp. the team placed the 5.5 via the axillary. the team observed a hematoma at day five of the support. the hematoma was treated by an evacuation on day eight and again on day twelve. the pump remained on for 6 more days, was weaned and explanted. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ added- h6 component code 925 added-d6a/d6b. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32464 | 1220648-2023-05291 | 18288711 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18288711&pc=OZD | 08/21/2023 | 02/29/2024 |  | 12/07/2023 |  |  | 11/07/2023 | 12/07/2023 |  |  | 04/30/2025 |  | 12/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024351120 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/07/2023 |  |  |  |  | Abiomed¿s long-term outcome and quality indicator impella registry (loqi) notification received indicating a 75 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient endured a thrombotic vascular complication prompting intervention, with a probable relationship to the impella device: "on august 21, he lost the pulse in his right dorsalis pedis (dp). the impella was gradually discontinued and removed due to suspected thrombus formation. an angiogram revealed a thrombus in the superficial femoral artery (sfa), which was subsequently treated with thrombectomy. while in the cicu, the patient was found to have an arterial occlusion on the right leg from the sfa to the popliteal artery. vascular surgery was consulted and stented the right sfa and peroneal arteries, leading to the restoration of pulses in the right dp and pt arteries." | [5 empty] | The investigation into the limb ischemia reversible has been completed. the device and data were not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32465 | 1220648-2023-05292 | 18288726 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18288726&pc=OZD | 11/08/2023 | 11/18/2025 |  | 12/07/2023 |  |  | 11/08/2023 | 12/07/2023 |  |  |  |  | 12/07/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1301114 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/08/2023 |  |  |  |  | The user facility reported a 60-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. us complainant reported that the purge cassette broke off of the aic. the cassette and the automated impella controller (aic) were replaced to resolve the issue. no patient harm reported due to the issue. | [5 empty] | The investigation has been completed. device analysis: (b)(4) was returned and visually inspected under a digital microscope. a broken purge disc retainer was identified. conclusion: the root cause of this issue was a broken purge disc retainer. g1: corrected the manufacturer email address h3: added information regarding the investigation status. h6: added codes per the results of the investigation h10: added the results of the investigation | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
32466 | 1220648-2023-05293 | 18288732 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18288732&pc=OZD | 11/08/2023 | 02/28/2024 |  | 12/07/2023 |  |  | 11/08/2023 | 12/07/2023 |  |  | 08/31/2025 |  | 12/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024398003 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/08/2023 |  |  |  |  | The user facility reported a 80-year-old male with multiple underlying comorbidities in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the team placed the pump at the femoral despite known peripheral artery disease. the team made choice to leave sheath and pump despite the limb ischemia and lost pulses as the condition was poor and end of life decisions were being made. the family made choice to withdraw care and the patient expired. the death is not attributed to the impella device. | [5 empty] | The investigation into the limb ischemia ¿ reversible has been completed. the device and data were not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32467 | 1220648-2023-05264 | 18273885 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18273885&pc=OZD | 11/06/2023 | 03/01/2024 |  | 12/06/2023 |  |  | 11/06/2023 | 12/06/2023 |  |  | 08/06/2025 |  | 12/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024398391 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Great Vessel Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/06/2023 |  |  |  |  | The user facility reported a 80-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support.patient went to the operating room for emergent coronary artery bypass graft with cp across aortic valve with peel away moved and repositioning sheath sutured in. perfusion gave antegrade cardioplegia and the cp was put on p2. doctor stated that the aorta dissected and needed repair. the doctor had the interventional cardiologist to come pull back impella cp across the aortic valve into the aorta so that they could start aorta repair. once interventional cardiologist pulled it back so that the motor touched the repositioning sheath, cardio vascular surgeon stated that it was still in their way and proceeded to cut the impella catheter shaft and remove it from the body after it verbalized that it was not suggested cutting the impella. once the impella was cut, the aorta was repaired, and doctor continued successfully with coronary artery bypass graft surgery. at the end of the case, axillary cp was was removed and closure was successful with no issues | [5 empty] | The investigation into the removal issues and the vascular damage - great vessel has been completed. the device was not returned for investigation. the root cause of the removal issue and the vascular damage - great vessel issue was not determined since product was not returned and the pump catheter was purposely cut to enable visibility during surgery. the relation to impella is unknown due to insufficient clinical information. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows:¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
32468 | 1220648-2023-05279 | 18275285 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18275285&pc=OZD | 04/19/2022 | 03/10/2025 |  | 12/06/2023 |  |  | 04/19/2022 | 12/06/2023 |  |  | 12/31/2023 |  | 12/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023178703 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/19/2022 |  |  |  |  | The user facility reported a 46-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was on impella cp support when there was a "position in aorta" alarm. left ventricle and aorta were super imposed over each other, and mitral valve was flat. initial attempts to reposition were unsuccessful, however the next morning the pump was repositioned and the pump was still deep at 5.4cm post repositioning. the pump was explanted later that day. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | Due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
32469 | 1220648-2023-05280 | 18275303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18275303&pc=OZD | 11/25/2023 | 10/03/2025 |  | 12/06/2023 |  |  | 11/25/2023 | 12/06/2023 |  |  | 09/30/2025 |  | 12/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024400510 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic |  | Death | Yes | Yes | Activation, Positioning or Separation Problem | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Ischemia | [12 empty] | Life Threatening | Death | Disability |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/25/2023 |  |  |  |  | The procedural outcome noted that the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 59-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had cp support ongoing for patient when there was harm/injury noted. the limb ischemia was treated by bypass placement/dpc placed from ecmo (extra-corporeal membrane oxygenation) limb. the limb flow improved but, the patient did observe the pump to have "in aorta" positioning alarms. the pump remained on for another day until the patient expired on support. no allegation made against pump for death. |  |  |  |  | B2: date of patient death is unknown. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-05280 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition, any concomitant medication, vascular size or variations thereof, detailed description of the symptoms experienced by the patient, physical examination findings, including pulse assessment and visual examination of the affected limb, diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography, laboratory test results, including blood tests for markers of inflammation or clotting disorders, any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension, patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. pump was not returned for investigation. data logs were investigated and multiple "impella in aorta" alarms were observed along with a lot of suction alarms. clinical mentions that echo was done and pump showed good position. this is a case of false pump in aorta alarm. the root cause of the placement signal issue was patient condition related to false aorta alarms. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. h6 updated to include optical signal coding. d4 model number updated. d10 concomitant medical products and therapy dates updated. d6a / d6b updated. f6 and f8 should have been left blank in the initial report. g1 reporting contact email and reporting contact fax number updated.
32470 | 1220648-2023-05281 | 18275351 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18275351&pc=OZD | 11/24/2023 | 11/08/2024 |  | 12/06/2023 |  |  | 11/24/2023 | 12/06/2023 |  |  | 09/30/2025 |  | 12/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024403648 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/24/2023 |  |  |  |  | The user facility reported a 56-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient had an impella 5.5 implant and the physician was unable to clearly visualize position of wire/pigtail inside of ventricle due to severe artifact on tee (transesophageal echocardiography). the physician believed the pump was in the correct position, however when the pump was turned on there were no flows and position in aorta alarm was seen on aic (automated impella controller). the pump was removed and it was determined that the impella had become folded over on top of the aortic valve. a new pump was successfully inserted. | It was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. clinical mentions that the patient has severe mvr. while implanting pump via tee the doctor faced a lot of artifacts which made the insertion process complex. due to the artifacts the tee was obstructed and the doctor felt as per his experience that the pump had ventricular access. the artifacts are a result of mostly patient condition and the patients diseased mitral valve. therefore, the root cause of the positioning issue was patient condition related to diseased valve. corrections to the initial mfr report: b1-selected adverse event. b2-selected death and added date of death. b5-mso review. d4-model number. added d6a/d6b. f6/f8 should have been left blank. h1-type of reportable event changed to death. h6 added impact code 1802.
32471 | 1220648-2023-05282 | 18275391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18275391&pc=OZD | 11/22/2023 | 01/17/2024 |  | 12/06/2023 |  |  | 11/22/2023 | 12/06/2023 |  |  | 10/31/2024 |  | 12/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023276020 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/22/2023 |  |  |  |  | The user facility in japan reported a 78-year-old male was implanted with an impella device for mechanical circulatory support. complainant in japan reported the patient was on impella cp support and upon removal of the impella, thrombus was observed. when the impella was removed, a large blood clot was found in the cannula. no treatment was required for the patient. | [5 empty] | The investigation into the thrombus issue has been completed. the device was not returned for benchtop evaluation and investigation. the root cause of the thrombus was most likely patient condition related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32472 | 1220648-2023-05283 | 18278638 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18278638&pc=OZD | 11/06/2023 | 04/05/2024 |  | 12/06/2023 |  |  | 11/07/2023 | 12/06/2023 |  |  | 02/28/2025 |  | 12/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024325571 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | ECMO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/07/2023 |  |  |  |  | The user facility reported a 86-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. upon turning on the pump it was noted that the wire had had not yet been removed which resulted in pump failure. . the wire was removed, and the physician tried to start the impella again, but it would not start. the patient was supported by ecmo following the failure of the pump but later passed away. as per the physician the device or the issue is not related to patients death. | [5 empty] | The investigation into the pump did not start issue has been completed since the initial report was submitted. the device was not returned for investigation. according to clinical findings, the pump started without the guidewire being removed, but it still did not start even after the guidewire was removed. the cause of the pump did not start issue was use related since doctor started pump without removing guidewire. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: to prevent device failure, do not start the impella catheter until the guidewire has been removed.
32473 | 1220648-2023-05284 | 18278652 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18278652&pc=OZD | 11/07/2023 | 01/17/2024 |  | 12/06/2023 |  |  | 11/07/2023 | 12/06/2023 |  |  | 07/31/2024 |  | 12/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2023252126 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/07/2023 |  |  |  |  | The user facility reported a 78-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient had an attempted impella cp implant. multiple attempts were made to pass impella into left ventricle, though resistance was met due to tortuosity and stenosis in the iliac artery. | [5 empty] | The investigation into the reported delivery issue has been completed since the original report was filed. no product was returned. the root cause of the delivery issue was most likely patient condition related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32474 | 1220648-2023-05285 | 18278930 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18278930&pc=OZD | 06/28/2023 | 07/30/2024 |  | 12/06/2023 |  |  | 11/06/2023 | 12/06/2023 |  |  | 04/30/2025 |  | 12/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024349703 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | User facility |  |  |  |  | 60 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Cardiac Perforation | [11 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 11/06/2023 |  |  |  |  | Abiomed¿s long-term outcome and quality indicator impella registry notification received regarding a 60 year old male patient presenting with acute myocardial infarction and cardiogenic shock. the report indicates, "as his lactate continued to elevate and cvp/pap began to equalize, bedside echo performed showed worsening tamponade physiology. patient subsequently taken to the catheterization lab for pericardiocentesis after pericardiocentesis/drain his pericardial effusion decreased and lvef improved." a possible relationship to the impella device was alleged. an additional allegation alleged that, "he also has some evidence of hemolysis from the impella cp. after discussion with the heart failure team, we felt that given his reasonable rv function, exchanging his mechanical support to axillary impella only would be a good strategy." | [5 empty] | The investigation for the ventricle perforation and hemolysis has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hemolysis and ventricle perforation could not be determined. the instructions for use for the impella cp with smartassist system states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ added-h6 component code 925 d4-corrected model number. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8 should have been left blank in the initial report. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions: including catheter position, pre-existing medical conditions, and small left ventricular volumes may also play a role in patient susceptibility to hemolysis.¿
32475 | 1220648-2023-05231 | 18267761 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18267761&pc=OZD | 11/06/2023 | 01/11/2024 |  | 12/05/2023 |  |  | 11/06/2023 | 12/05/2023 |  |  | 07/31/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024370785 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Hispanic | White | Injury | Yes | Yes | Failure to Advance | [6 empty] | Vascular Dissection | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that resistance was met during attempted delivery of an impella 5.5 pump due to a patient's tight anastomosis. it was noted that the device was unable to advance via right axillary approach. when resistance was met, the doctor noted that the artery had been dissected. the artery was surgically repaired; however clots were discovered following the intervention. a fogarty catheter was utilized to remove the clots and the implant was subsequently aborted. | [5 empty] | The investigation into the vascular damage - peripheral vessel, the thrombus, and the delivery issues- anatomy has been completed. the device was not returned for investigation. based on the clinical information provided, the cause of the vascular damages was patient condition related as per the doctor's comments, the impella would not cross due to tight anastomosis and possible small artery. the cause of thrombus was patient condition related as thrombus was present upon insertion of impella. the cause of the delivery issue was patient condition related because there were issues with the size of the artery per the clinical notes. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32476 | 1220648-2023-05232 | 18267770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18267770&pc=OZD | 11/07/2023 | 09/24/2024 |  | 12/05/2023 |  |  | 11/07/2023 | 12/05/2023 |  |  | 05/31/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024365076 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/07/2023 |  |  |  |  | The user facility reported a 74-year-old male patient was implanted with an impella cp device for mechanical circulatory support. the patient was taken back to cardiac cath lab for repositioning the impella cp. the impella would not flow over p6 without having suction. after attempted repositioning of device in ccl flows improved; primary care team not satisfied with position, as inflow was still along posterior wall while evaluating position on echo. further attempts were made to free the device from ventricular structures, but ultimately the device pulled back across the aortic valve. decision was made to explant device and implant new impella cp. no patient harm reported due to the issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for positioning issues has been completed. no product was returned for an investigation. data log analysis was not necessary for this complaint. the most likely cause of the positioning issues was use related. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05232:
32477 | 1220648-2023-05233 | 18267836 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18267836&pc=OZD | 11/07/2023 | 09/23/2024 |  | 12/05/2023 |  |  | 11/07/2023 | 12/05/2023 |  |  | 08/31/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024387186 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | Positioning Problem | [5 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/07/2023 |  |  |  |  | The user facility reported a 67-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. there was a hematoma at the access site. hematoma was roughly the size of a baseball and was observed around the axillary cutdown. sandbags were applied to the hematoma, and it resolved. patient sub therapeutic and systemic heparin stopped due to hematoma formation. tpa used in purge due to positive trending purge pressure and negatively trending purge flow. purge pressure and flow look to have improved. it was further reported that the physician attempted to reposition the pump due to positional suction but was unable to manipulate the device. blue button was pressed; however, catheter did not move. the team concluded that they believe clot has formed in the graft and was preventing the pump from moving. the physician stated, to leave the pump as is for now. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the positioning issues and access site bleed has been completed. there was no product returned for investigation. according to the clinical, a baseball-sized hematoma was discovered at the access site. sandbags were applied to the hematoma and it later resolved. there were no other clinical details provided for the bleeding. the systemic anticoagulation was held. the most likely cause of the major access site bleeding was use related. the data logs were reviewed and at the time of the reported complaint there was a gradual increase in purge pressure, a decrease in purge flow and a decrease in flow with suction alarms present. the tissue plasminogen activator (tpa) was have said to improve these flows, and they remained stable for the duration of the case. the root cause of the high purge pressure is most likely due to biomaterial. based upon clinical details, the most likely cause of the positioning issues was a cath anchor use issue. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ added- h6 impact code 4644, 4628. added- h6 med dev prob code 1491 corrections for the initial mfr report: added- d6a/d6b.
32478 | 1220648-2023-05234 | 18267854 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18267854&pc=OZD | 11/07/2023 | 03/22/2024 |  | 12/05/2023 |  |  | 11/07/2023 | 12/05/2023 |  |  | 07/31/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024385960 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/07/2023 |  |  |  |  | The user facility reported a 63-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient developed a hematoma following impella cp implant. manual pressure was initially applied at the access site to treat the condition; however, the bleeding persisted and the decision was ultimately made to explant the pump. the patient was taken to the operating room where the pump was removed and the femoral artery was repaired. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32479 | 1220648-2023-05235 | 18267881 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18267881&pc=OZD | 11/07/2023 | 09/13/2024 |  | 12/05/2023 |  |  | 11/07/2023 | 12/05/2023 |  |  | 05/31/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024365594 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/07/2023 |  |  |  |  | The user facility reported a 71-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced persistent bleeding/oozing from their access site during impella cp support. frequent dressing changes were completed and the patient was transfused 8 units of packed red blood cells and fresh frozen plasma to treat the condition. additional sutures were subsequently placed to resolve the noted oozing. support with the implanted pump continued following the intervention. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely use related, impella groin ooze subsided after doctor placed a suture at the site as per the clinical description provided. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- h6 component code 925 added- d6a/d6b. f6/f8 should have been left blank in the initial report.
32480 | 1220648-2023-05236 | 18267930 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18267930&pc=OZD | 11/07/2023 | 07/10/2024 |  | 12/05/2023 |  |  | 11/07/2023 | 12/05/2023 |  |  | 07/31/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024377547 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/07/2023 |  |  |  |  | The user facility reported a 70-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the patient developed a thrombus in their right subclavian during impella 5.5 support. heparin administration was temporarily on hold due to a pre-existing condition; however it was reintroduced to the patient roughly two weeks into support. the patient's condition was said to be improving and support with the pump was maintained. | [5 empty] | The investigation for the reported thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombus could not be determined. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock in the section: direct aortic insertion states ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ it also states in section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ h6 component code 925 added in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32481 | 1220648-2023-05237 | 18267961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18267961&pc=OZD | 11/07/2023 | 10/02/2024 |  | 12/05/2023 |  |  | 11/07/2023 | 12/05/2023 |  |  | 08/31/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024393172 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that resistance was met during attempted impella cp implant due to a patient's significant tortuous vessels. upon observation, it was noted that the peel away sheath had kinked in the vessel. the physician stated that they had used force while trying to advance the pump and believed that the impella may also be damaged. as a result, both the introducer and pump were removed and a second attempt was made with new devices. there was no patient harm as a result of the noted issue. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the introducer damage and the delivery issues has been completed. according to the clinical, a long peel away sheath was used to insert the pump. however, the patient had very tortuous vessels so when the device was inserted it was met with resistance. it was then discovered that the sheath had been kinked in the vessel. after the device was removed and replaced another attempt was made with the same results. on the third attempt, the pump was successfully inserted. no product was returned for an investigation. data log analysis was not necessary for this complaint. the most likely cause of the delivery issues was patient condition related. the most likely cause of the introducer damage was patient condition related. corrections to the initial mfr report: added- d6a/d6b
32482 | 1220648-2023-05238 | 18268078 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18268078&pc=OZD | 11/07/2023 | 09/13/2024 |  | 12/05/2023 |  |  | 11/07/2023 | 12/05/2023 |  |  | 06/30/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024379226 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp pump got caught on an existing ebu guide catheter during attempts to advance the pump around the aortic arch. as a result, the decision was made to abort the implant and an intra-aortic balloon pump (iabp) was subsequently placed for ongoing support. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The root cause of the delivery issue was patient condition related to interaction with other device (ebu guide catheter). the root cause of the delivery issue was patient condition related to interaction with other device (ebu guide catheter). the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05238: d6a and d6b revised from the initial submission in accordance with updated procedures g1 reporting contact fax number missing h.6 code 4642 and 4327 were reported incorrectly. new code was added to medical device problem code. instruction for use: impella cp with smart assist for use during cardiogenic shock and high-risk pci. section: potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings and cautions: warnings: section: pre-support evaluation. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.
32483 | 1220648-2023-05239 | 18268112 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18268112&pc=OZD | 11/07/2023 | 08/30/2024 |  | 12/05/2023 |  |  | 11/07/2023 | 12/05/2023 |  |  | 08/31/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024392944 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following implant of an impella cp optical pump, it was noted via echo that the pump was too deep. the pigtail was stuck in the papillary muscle and, upon attempted repositioning, the pump came back across the valve. after being unable to re-advance the device into position, the decision was made to replace the pump. there was no patient harm. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the positioning issue was use related. added- h6 impact code 4628 added- d6a/d6b. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32484 | 1220648-2023-05240 | 18268140 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18268140&pc=PYX | 11/08/2023 | 08/21/2024 |  | 12/05/2023 |  |  | 11/08/2023 | 12/05/2023 |  |  | 03/31/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC |  |  | US |  | 2024394988 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/08/2023 |  |  |  |  | The user facility reported a 73-year-old male patient was implanted with an impella device for mechanical circulatory support. the complainant reported that following impella rp flex implant, manual pressure was held at the site to remove the peel away sheath and the pump temporarily stopped. upon lightening the pressure at the sight, the alarm resolved and support with the implanted pump was able to continue. there was no patient harm resulting from the failure. | [5 empty] | The investigation of pump stop has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05240. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
32485 | 1220648-2023-05241 | 18268156 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18268156&pc=OZD | 11/10/2023 | 04/05/2024 |  | 12/05/2023 |  |  | 11/10/2023 | 12/05/2023 |  |  | 06/30/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024371816 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/10/2023 |  |  |  |  | The user facility reported a 71-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the patient had bleeding at the impella access site and the impella was removed. the patient was stable. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the initial report was submitted. the device was not returned for investigation. based on clinical description, the root cause of the access site bleeding issue was not determined since product was not returned and not much information is provided in the clinical description. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32486 | 1220648-2023-05242 | 18268165 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18268165&pc=OZD | 11/08/2023 | 04/01/2024 |  | 12/05/2023 |  |  | 11/08/2023 | 12/05/2023 |  |  | 09/30/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024398193 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/08/2023 |  |  |  |  | The user facility reported a 59-year-old male patient was implanted with an impella device for mechanical circulatory support. us complainant reported that the patient had a nosebleed and an impella access site bleed. heparin was withheld to manage the bleeding. support continued with the implanted pump. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. clinical details states that the bleeding was resolved by holding heparin. therefore, the root cause of the access site bleed was most likely due to anticoagulation management. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32487 | 1220648-2023-05243 | 18268248 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18268248&pc=OZD | 11/13/2023 | 02/13/2025 |  | 12/05/2023 |  |  | 11/13/2023 | 12/05/2023 |  |  | 07/31/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024370480 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/13/2023 |  |  |  |  | The user facility reported a 59-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the patient had been on support for two months and began to have intermittent suction with positioning. the flow had to be decreased to p4. echo showed pump had flipped and was directed towards the mitral valve, but they were unable to torque pump into correct position. patient was dependent on pump. the patient was successfully transplanted several days later. no patient harm reported due to the issue. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the positioning issue has been completed. pump was not returned for investigation. data logs were returned and "suction" and "impella position unknown" alarms were observed. as per clinical information, imaging was done and pump was found to be flipped. the root cause of the positioning issue was unable to be determined as limited details on the fixation of the pump was provided and no product was returned for analysis. corrections to the initial mfr report: f6/f8 should have been left blank.
32488 | 1220648-2023-05244 | 18268249 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18268249&pc=OZD | 11/08/2023 | 08/18/2025 |  | 12/05/2023 |  |  | 11/08/2023 | 12/05/2023 |  |  |  |  | 12/05/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1561776 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/08/2023 |  |  |  |  | The complainant reported that an automated impella controller( aic ) experienced an unspecified error. the device was removed from service as a result of the failure and a loaner was sent to the customer site for use as needed. no patient is involved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The impella device will not be returned from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. e2 health professional revised on manufacturer device report in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number. h6 investigation findings 3233 should not be on initial manufacturer device report in accordance with updated procedures. h6 investigation conclusions revised in accordance with updated procedures. h10 updated with additional information.
32489 | 1220648-2023-05245 | 18268304 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18268304&pc=OZD | 11/08/2023 | 09/03/2024 |  | 12/05/2023 |  |  | 11/08/2023 | 12/05/2023 |  |  | 12/31/2024 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024316010 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/08/2023 |  |  |  |  | The user facility reported a 72-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella site was reassessed and large hematoma expanding distal to insertion site and bleeding around insertion occurred. the patient with vagal response to the pressure so modified purse string was added by the physician. the patient was transported to another facility and upon arriving blood was ordered and given along with fluids. vascular was brought in to do a cutdown and evacuate the site. the impella was weaned and removed but impella sheath left as team felt patient was in hemorrhagic shock. the patient is stable post intervention. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was most likely due to patient movement since the patient had large hematoma after the transport. d6a and d6b added h6 component code added the following have been reported incorrectly or missing from manufacturer device report. f6/f8-should have been left blank g1 reporting contact fax number missing. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32490 | 1220648-2023-05246 | 18268311 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18268311&pc=OZD | 11/08/2023 | 02/28/2024 |  | 12/05/2023 |  |  | 11/08/2023 | 12/05/2023 |  |  | 09/30/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024402129 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 45-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported ischemia developing in the patient's impella inserted limb, during support. a distal perfusion catheter was placed, as treatment. | [5 empty] | The investigation into the limb ischemia ¿ reversible has been completed. the device and data were not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32491 | 1220648-2023-05247 | 18268347 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18268347&pc=OZD | 11/12/2023 | 04/04/2024 |  | 12/05/2023 |  |  | 11/13/2023 | 12/05/2023 |  |  | 08/31/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024397987 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 25 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/13/2023 |  |  |  |  | The user facility reported a 25-year-old female with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the patient had oozing and slight bleeding from swan site and impella femoral access site. the patient was given 1 unit of blood at night and 1 unit of platelets in the morning. patient is stable and was on pump support for additional 10 days. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the access site bleeding issue was not determined since product was not returned and not much information is provided in the clinical description. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32492 | 1220648-2023-05248 | 18268391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18268391&pc=OZD | 11/09/2023 | 12/05/2023 |  | 12/05/2023 |  |  | 11/09/2023 | 12/05/2023 |  |  |  |  | 12/05/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1554428 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Unable to Obtain Readings | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 11/09/2023 |  |  |  |  | The user facility reported a 68-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the automated impella controller (aic ) stopped showing purge flow data after the quantity of heparin being administered was adjusted. support was maintained with the device, and there was no patient injury resulting from the noted issue. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
32493 | 1220648-2023-05249 | 18268466 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18268466&pc=OZD | 11/09/2023 | 01/12/2024 |  | 12/05/2023 |  |  | 11/09/2023 | 12/05/2023 |  |  | 06/30/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024399125 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/09/2023 |  |  |  |  | The user facility reported a 67-year-old female with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the impella 5.5 pump was unable to advance beyond 5 cm and would not turn down the innominate artery during attempted delivery. after several unsuccessful attempts, the surgeon felt that the artery was too calcified and torturous and the decision was made to abort the implant. there was no patient harm. | [5 empty] | The investigation into the delivery issues- anatomy has been completed. the device was not returned for benchtop evaluation and investigation. per the clinical information provided, the root cause of the delivery issue was patient condition related to over calcified vessels. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32494 | 1220648-2023-05252 | 18268524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18268524&pc=OZD | 11/09/2023 | 02/28/2024 |  | 12/05/2023 |  |  | 11/09/2023 | 12/05/2023 |  |  | 06/30/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024375629 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/09/2023 |  |  |  |  | The user facility reported a 60-year-old female in acute myocardial infarction/cardiogenic shock was emergently implanted with an impella cp device for mechanical circulatory support. the patient developed ischemia during impella cp support. it was noted that an iliac/femoral angio showed no distal flow around the impella introducer. as a result, the decision was made to explant the pump. | [5 empty] | The investigation into the limb ischemia reversible has been completed. the device and data were not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32495 | 1220648-2023-05250 | 18268532 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18268532&pc=OZD | 11/09/2023 | 03/12/2024 |  | 12/05/2023 |  |  | 11/09/2023 | 12/05/2023 |  |  | 08/31/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024398014 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/09/2023 |  |  |  |  | The user facility reported an 87-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the flow rates were low. the team discovered the impella cp had kinked. repositioning was attempted, however the flow rates did not improve. as a result, the decision was made to replace the pump. there was no patient harm. | [5 empty] | The investigation into the low pump flow issue has been completed since the original report was filed. the device was returned and tested after arriving in fs. visual inspection found a cannula kink on the device. the root cause of low flow was a cannula kink. the cause of kinked cannula was most likely a patient condition. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
32496 | 1220648-2023-05251 | 18268534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18268534&pc=OZD | 11/14/2023 | 05/15/2024 |  | 12/05/2023 |  |  |  | 12/05/2023 |  |  | 06/30/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024361225 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/14/2023 |  |  |  |  | The user facility reported a 67-year-old female with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. there was leaking from purge side arm on impella 5.5. there were no visible cracks, and the purge flow and purge pressure were within range. the patient is stable and no patient harm was reported. | [5 empty] | The investigation into the purge leak ¿ pump issue has been completed since the original report was filed. the pump was returned for analysis. the failure was reproducible. the root cause of the purge leak - pump was due to a damaged sidearm reservoir. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
32497 | 1220648-2023-05253 | 18268563 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18268563&pc=OZD | 11/10/2023 | 09/26/2024 |  | 12/05/2023 |  |  | 11/10/2023 | 12/05/2023 |  |  | 05/31/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024356260 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Hispanic |  | Death | Yes | Yes | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/10/2023 |  |  |  |  | The user facility reported a 63-year-old male (race unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was in route to another hospital via ground ems when the impella experienced a copious amount of suction alarms. attempts made by ems personnel to break suction were unsuccessful. the patient was brought back to the original hospital to the emergency department (ed) and expired in the ed. | [5 empty] | The investigation into the low pump flow issue has been completed since the original report was submitted. the device was not returned by the user facility for investigation. the data log showed several suction alarms and indicated that the pump was unable to achieve adequate flow while running above p-level 2 at the end of the case. the placement signal returned to normal from a decreasing trend during the low flow event. no position-related issues or motor current spikes were reported during the case. according to clinical records, the doctor failed to resolve the suction alarms. the cause of the low pump flow issue was not determined since product was not returned and sufficient clinical information was not returned. a.5 ethnicity was inadvertently left off the previously submitted report and was added. b.1 product problem was inadvertently left off the previously submitted report and was added. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.1 revised as previously entered incorrectly. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted report incorrectly. no ifus are needed to be reported for this report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in section b2.
32498 | 1220648-2023-05254 | 18268566 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18268566&pc=PYX | 11/10/2023 | 01/16/2024 |  | 12/05/2023 |  |  | 11/10/2023 | 12/05/2023 |  |  | 02/28/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP FLEX W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2024337524 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/10/2023 |  |  |  |  | The user facility reported a 67-year-old female patient was implanted with an impella device for mechanical circulatory support. the us complainant was placing an rp flex when there was resistance in placement. the placement of the sheath did occur but shortly thereafter the md noted a large collection of blood in the pleural space. the blood was evacuated, and the team believed the pulmonary vein had perforated. the team stitched the damage, and a thoracotomy was performed. the team gave blood products and medical therapy in the icu and was stabilized. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation for the reported vascular damage was completed. the impella device was not returned for investigation. as per the clinical notes, resistance was encountered during the insertion of an impella device via the right internal jugular vein while advancing the 23 french sheath. the sheath was subsequently removed from the neck, and the site was closed with a silk stitch, repairing visible vessel damage during a thoracotomy. the cause of the vascular damage issue was most likely patient condition related since resistance was encountered during insertion due to vascular anatomy.
32499 | 1220648-2023-05255 | 18268582 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18268582&pc=OZD | 11/10/2023 | 04/09/2024 |  | 12/05/2023 |  |  | 11/10/2023 | 12/05/2023 |  |  |  |  | 12/05/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1701427 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/10/2023 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) displayed a battery failure (red) alarm during patient support. the aic was successfully replaced. there are no patient details available. it was reported there was no patient harm. | [5 empty] | The investigation into the aic - battery failure (red alarm) issue has been completed since the initial report was submitted. the console was returned, visually inspected, and tested. the reported failure was confirmed in the data logs; the failure was unable to be reproduced during testing. the cause of the issue was not determined since the issue could not be reproduced. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
32500 | 1220648-2023-05256 | 18268624 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18268624&pc=OZD | 11/10/2023 | 06/08/2024 |  | 12/05/2023 |  |  |  | 12/05/2023 |  |  | 12/31/2024 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024292774 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/10/2023 |  |  |  |  | The user facility reported a 67-year-old male with cardiomyopathy was to be implanted with an impella 5.5 for mechanical circulatory support. due to the patient¿s anatomy, the physician was unsuccessful in crossing the aortic valve. the case was aborted. the physician then decided to place the patient on va ecmo and was successful. | [5 empty] | The investigation into the delivery issues- anatomy has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue was determined to be patient condition as it was reported that the delivery was unsuccessful due to patient anatomy. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist : section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ a.5 revised ethnicity as previously entered incorrectly. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 codes 2199, 4627 and 4114 were reported incorrectly on the previously submitted report. a new code has been added to type of investigation codes. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. additional manufacturer narrative on the previously submitted report stated that the device was not returned but the device was discarded. type of investigation codes has been updated to reflect discarded. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32501 | 1220648-2023-05257 | 18268714 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18268714&pc=OZD | 11/10/2023 | 10/18/2024 |  | 12/05/2023 |  |  | 11/10/2023 | 12/05/2023 |  |  | 08/31/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024394489 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/10/2023 |  |  |  |  | The user facility reported a 76-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed hemolysis during impella cp support. the patient was subsequently escalated to an impella 5.5 device as a result of the noted condition. support continued as planned with the implanted 5.5 device. | The patient did not survive post procedure. per medical safety officer review use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation of hemolysis has been completed. pump was not returned for investigation. data logs were reviewed and suction was observed, followed by high purge pressure and motor current spike. these symptoms suggests that there was likely a biomaterial ingestion event. therefore, as per data log review, the root cause of the hemolysis issue was most likely biomaterial. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05257: b5 patient outcome and mso statement added. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. h6 component code revised from the initial submission in accordance with updated procedures.
32502 | 1220648-2023-05258 | 18268983 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18268983&pc=OZD | 10/30/2023 | 02/15/2024 |  | 12/05/2023 |  |  | 10/30/2023 | 12/05/2023 |  |  | 05/31/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024351583 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 65-year-old male patient was implanted with an impella cp for mechanical circulatory support. when the guidewire was pulled back, the impella cp was noted to have slipped put of the ventricle. therefore, a new impella was placed. there was no reported patient harm. | [5 empty] | The investigation into the positioning issues has been completed. the device was not returned for investigation. as per the clinical, impella was slipped out of the ventricle during guidewire removal. the cause of the positioning issue was most likely use related since pump was pulled back while removing guidewire. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
32503 | 1220648-2023-05259 | 18269063 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18269063&pc=OZD | 11/09/2023 | 10/07/2024 |  | 12/05/2023 |  |  | 11/09/2023 | 12/05/2023 |  |  | 05/31/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024358527 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 63-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported the impella cp was removed and vascular surgery was needed to be performed. the cutdown and repair of the femoral artery was performed with two compartment fasciotomies. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vascular damage could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrections to the initial mfr report: added- d6a/d6b.
32504 | 1220648-2023-05260 | 18269074 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18269074&pc=PYX | 11/09/2023 | 04/29/2024 |  | 12/05/2023 |  |  | 11/09/2023 | 12/05/2023 |  |  | 01/31/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024325150 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 59-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella rp flex device for mechanical circulatory support. the complainant reported a drop in impella pump flow, occurring after decannulating va ecmo. troubleshooting to restore proper pump flow was not successful, therefore, the pump was removed/replaced. a clot was noted in the outflow cage upon removal of the device. | [5 empty] | The investigation for the reported low pump flow has been completed. the impella pump was returned for evaluation. upon investigation of the pump, a biomaterial clot was observed around the impeller. data logs show a high motor current spike followed by low pump flows. the root cause of the low pump flow issue was determined to be biomaterial. added-d6a/d6b in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
32505 | 1220648-2023-05261 | 18269106 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18269106&pc=OZD | 11/10/2023 | 01/17/2024 |  | 12/05/2023 |  |  | 11/10/2023 | 12/05/2023 |  |  | 08/31/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024385557 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/10/2023 |  |  |  |  | The user facility reported a 47-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the patient developed a hematoma at their impella access site during support. heparin was stopped temporarily and manual pressure was held at the site until hemostasis was achieved. once the condition resolved, heparin was resumed. support with the implanted pump was maintained. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding - major has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the cause of the access site bleeding issue was not determined due to insufficient clinical information and since product was not returned.
32506 | 1220648-2023-05262 | 18270463 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18270463&pc=OZD | 11/06/2023 | 04/22/2024 |  | 12/05/2023 |  |  | 11/06/2023 | 12/05/2023 |  |  | 08/31/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024395780 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/06/2023 |  |  |  |  | The user facility reported a 71-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the motor current was flat with low flows. unable to increase p level. volume given and attempted to reposition and no change. the decision was made to replace the device. the patient is stable. | [5 empty] | The investigation into the low pump flow issue has been completed since the original report was filed. the pump was returned, tested, and evaluated. visual inspection found biomaterial inside the pump housing & wrapped around the impeller. the root cause of low flow was due to biomaterial ingestion. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32507 | 1220648-2023-05263 | 18270474 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18270474&pc=OZD | 11/06/2023 | 09/23/2024 |  | 12/05/2023 |  |  | 11/06/2023 | 12/05/2023 |  |  | 09/30/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024400645 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/06/2023 |  |  |  |  | The user facility reported a 81-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the cp placement was delayed. the pump delivery was complicated by the interaction of the pump and .018 wire. the .018 guidewire kinked and was replaced with no further complications. no patient harm reported due to the issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for product damage (guidewire damage - peripheral vessels) has been completed. the cause of the guidewire damage was not determined because no product was returned and not enough clinical information was given. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05263: d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact fax number missing.
32508 | 1220648-2023-05265 | 18270545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18270545&pc=OZD | 11/06/2023 | 09/03/2024 |  | 12/05/2023 |  |  | 11/06/2023 | 12/05/2023 |  |  | 08/31/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024381530 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Use of Device Problem | Material Split, Cut or Torn | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/06/2023 |  |  |  |  | The user facility reported a 69-year-old male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, the patient was agitated and there was an "impella in aorta" alarm. the physician was able to re-advance across the aortic valve with echo but the sterile sleeve was torn and was patched with tegaderm. additional information received stated that the sterile sleeve was torn by the physician when repositioning. there was no patient harm. | [5 empty] | The investigation of product damage (sterile sleeve damage) has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6a and d6b added f6/f8-should have been left blank on manufacturer device report 1220648-2023-05265. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32509 | 1220648-2023-05266 | 18270595 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18270595&pc=OZD | 11/07/2023 | 09/03/2024 |  | 12/05/2023 |  |  | 11/07/2023 | 12/05/2023 |  |  | 05/31/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024366762 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the doctor preformed percutaneous coronary intervention (pci) of left anterior descending artery and the patient had v-fib arrest, after return of spontaneous circulation, the impella cp was placed. the doctor reported that when they explanted the impella cp, they noticed what looked to be clot on the pigtail. patient was stable and no issues were noted from the impella. | [5 empty] | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32510 | 1220648-2023-05267 | 18270609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18270609&pc=OZD | 11/07/2023 | 04/16/2024 |  | 12/05/2023 |  |  | 11/07/2023 | 12/05/2023 |  |  | 08/31/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024395183 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | Pressure Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the 69-year-old male patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complaint reported that during impella cp support, there was a high purge pressure/purge system blocked alarm. the medical office recommended lysis therapy, however, staff decided against it. afterwards, the pump stopped. therefore, the pump was explanted and replaced with a new pump. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop and the high purge pressure issues has been completed since the original report was filed. the pump was returned, tested, and evaluated. the catheter was cut near the motor housing and fluid and biomaterial expelled from the catheter. the root cause of the pump stop was most likely due to biomaterial buildup contributing to the high purge pressure.
32511 | 1220648-2023-04987 | 18270637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18270637&pc=OZD | 10/21/2023 | 06/03/2024 |  | 12/05/2023 |  |  | 10/21/2023 | 12/05/2023 |  |  | 10/31/2024 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023283708 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the 67-year-old male patient was brought to the lab to reposition impella cp due to ongoing hemolysis throughout the night. it was felt that the impella was caught up under papillary muscle and was unable to be repositioned. echo in cath lab then revealed a thrombus in the left ventricle. some clot was dislodged and was pulled out attached to the pigtail. the physician felt strongly that the clot was there before impella was placed. the impella was therefore removed, and the patient was then placed on a balloon pump (iabp) and sent back to the unit. | [5 empty] | The investigation for the reported hemolysis and thrombus has been completed. the impella was not been returned for evaluation. according to the clinical, on the first day of impella cp support the patient experience suction events. in an attempt to resolve the issue the pump was repositioned. although it was accidentally pulled into the aorta it was successfully advanced back into the correct position. the next day the pump was repositioned a second time due to hemolysis concerns and the belief that the pump had become caught in the papillary muscles. after repositioning the device a thrombus was discovered in the lv. the decision was then made to explant the pump and replace it with another device. no product was returned for a visual inspection. data log analysis confirmed the presence of suction alarms and positioning issues. the most likely cause of the hemolysis was improper positioning. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿
32512 | 1220648-2023-04986 | 18270638 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18270638&pc=OZD | 10/21/2023 | 06/06/2024 |  | 12/05/2023 |  |  | 10/21/2023 | 12/05/2023 |  |  | 07/31/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024378111 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 40 YR |  | Male |  |  | White | Injury | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 40-year-old patient in the eu received hemodynamic support from an impella 5.5 heart pump during high-risk coronary artery bypass graft surgery and recovery thereafter. after three days, a purge fluid line leak occurred at the luer connector. the user was able to seal the leak with glue. the impella 5.5 then ran for another five days without any further problems until the patient was successfully weaned off the impella support. | [5 empty] | The investigation into the purge leak ¿ pump issue has been completed since the original report was submitted. the device was not returned for investigation. as a result, the root cause of the purge leak could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
32513 | 1220648-2023-05022 | 18270639 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18270639&pc=OZD | 10/24/2023 | 09/23/2025 |  | 12/05/2023 |  |  | 10/24/2023 | 12/05/2023 |  |  | 07/31/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024378320 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | Pumping Stopped | Device Sensing Problem |  |  |  |  | Stroke/CVA | No Clinical Signs, Symptoms or Conditions | [17 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 51-year-old female patient noted as high risk percutaneous cardiac insertion was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that a patient on impella support experienced leaking from the purge sidearm. there was no harm to the patient as a result of this incident. | Device demonstrated a pump stop on the evening of (b)(6) 2023. rn was advised to leave the device on surgical mode until explant. | The investigation uncovered a placement signal issue. |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation into the reported purge leak, pump stop and placement signal issue has been completed since the original report was filed. in the patient updates on 10/08 at 1:23pm all heparin as topped due to an intracranial bleed on 10/07. on 10/20 there was an update from csc that the patient was still not on anticoagulants due to the cva. there were no issues or concerns with the impella. on 11/13 there were still no systemic ac's due to intracranial hemorrhages. in the log below, heparin can be seen to be stopped om 10/07 at the time of the reported bleed. there are no spikes in the motor current. it is unknown if heparin was the cause of the hemorrhagic stroke. the root cause of the stroke/cva is most likely due to the patients condition. on 11/12 there purge pressure varies and there low purge flow at the reported time of sidearm leak in the patient updates. there was trouble shooting attempted by changing the bag and cassette but that did not resolve the issue. the purge flow remains to vary for the remainder of the case. the location of the reported leak is unknown. the root cause of the purge leak-pump is most likely due to a damaged sidearm. on 11/19 the placement signal spikes to 3000 and becomes unreliable. the 5s parameters are consistent with a break in the fiber line but the location is unknown and it is unknown if the device failed to meet specification. the root cause of the placement signal issue is most likely due to a fiber break. additionally, they happen simultaneously as there is an electrical open pump stop. the motor current spikes, and the pump stops with alarm 115 and 119 triggered. this is indicative of an open electrical line but the location is unknown and it is unknown if the device failed to meet specification. the root cause of the pump stop is most likely due to open electrical lines correction : section h6 : the medical device problem code in section h6, was inadvertently left out in the initial report which has been added to this report. correction: section b5 was updated in this report as some details were inadvertently left out in the initial report . | B5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2023-05022. b7 other relevant history, including preexisting medical conditions updated. d6a / d6b implant and explant updated. d10 concomitant medical products and therapy dates updated.
32514 | 1220648-2023-04930 | 18270643 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18270643&pc=OZD | 10/14/2023 | 03/04/2024 |  | 12/05/2023 |  |  | 10/18/2023 | 12/05/2023 |  |  | 11/30/2024 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024312156 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  |  |  | Malfunction | No | Yes | Improper Flow or Infusion | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/18/2023 |  |  |  |  | The user facility reported a 68-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, there were ongoing suction alarms. the physician attempted to reposition without success. the pigtail was caught around the papillary muscle. the device was swapped for another impella cp with no patient harm. | [5 empty] | The investigation into the positioning issues has been completed. the device was not returned for investigation. according to clinical findings, the doctor discovered the pump deep in the ventricle. data logs indicate that placement signal returned to normal range after a noted spike in motor current, but without any resolution in the low pump flow issue. the cause of the low pump flow issue was most likely biomaterial based on data log review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32515 | 1220648-2023-04907 | 18270645 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18270645&pc=OZD | 10/14/2023 | 02/06/2024 |  | 12/05/2023 |  |  | 10/14/2023 | 12/05/2023 |  |  | 02/28/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024320762 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant in japan reported the fourth day of impella cp management, the pump stopped and restarted three times in p-4 condition. the patient was hemodynamically stable, therefore the impella cp was removed. act was managed in about 130 seconds. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop issue has been completed since the original report was filed. upon investigation of the pump, a large clot was observed in the inflow cage and around the impeller. there was also blood ingress in the filter all the way to the yellow luer that likely resulted after the pump stop as there were no purge related issues prior. data logs show a high motor current spike followed by the pump stop, which is indicative of biomaterial ingestion. the root cause of the pump stop was determined to be ingested biomaterial.
32516 | 1220648-2023-04988 | 18270658 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18270658&pc=OZD | 10/19/2023 | 01/18/2024 |  | 12/05/2023 |  |  | 10/19/2023 | 12/05/2023 |  |  | 08/31/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024393358 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Improper Flow or Infusion | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/19/2023 |  |  |  |  | The user facility reported a 66-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient on impella support experienced hemolysis. the pump was repositioned via echo and fluid was given, however, the impella flow dropped to p-2. patient¿s family decided to go comfort care due to patient poor prognosis. as a result, impella was explanted. | [5 empty] | The investigation into the reported hemolysis has been completed since the original report was filed. no product was returned for investigation and no data logs were provided. the pump was confirmed to not be in proper position via echo. the root cause of the hemolysis issue was most likely improper positioning of the device. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32517 | 1220648-2023-05268 | 18270779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18270779&pc=OZD | 11/10/2023 | 12/03/2025 |  | 12/05/2023 |  |  | 11/10/2023 | 12/05/2023 |  |  |  |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1229946 |  | 0042-0000 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported 72-year-old male patient was attempted to be implanted with an impella cp smartassist heart pump for hemodynamic support. however, the automated impella controller (aic) skipped the last part of prep, and the pump was ultimately not implanted. it is unclear if the defect was caused by the pump or the aic. other devices were used to complete the procedure. there was no reported patient harm. | [5 empty] | The automated impella controller (aic) device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | Data analysis: console logs could not be fully recovered. the case was initiated on 10 november 2023. a few minutes after the pump was first connected and after case start wizard step 6 the aic displayed a placement signal not reliable alarm (#1036), which cleared after the pump was disconnected. the pump was then reconnected and calibration completed, but the pump was not started; the user navigated back to the previous step and then forward to step 7. finally, the pump was reconnected and, per the clinical note, both the aic and the pump were replaced. lt log showed 0 motor current for the entire recorded period. rt log showed lamp at 100% with almost 0 snr on the initial connection, with snr improving after the pump was reconnected with increasing purge flow ticks. device analysis: console (b)(4) was not returned to fs for further investigation. root cause: the root cause for the reported issue was not determined since aic was not returned. the root cause for the placement signal not reliable alarm (#1036) was not determined since aic was not returned. d2 common device name updated. d4 unique identifier (udi) # updated. d10 concomitant medical products and therapy dates updated. g1 reporting contact email and reporting contact fax number updated.
32518 | 1220648-2023-05269 | 18270793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18270793&pc=OZD | 11/09/2023 | 01/16/2024 |  | 12/05/2023 |  |  | 11/09/2023 | 12/05/2023 |  |  |  |  | 12/05/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1094547 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Mechanical Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The biomed/hospital team reported upon an automated impella controller (aic) with purge disc issues. the purge cassette would not seat/fit into the console. no patient use or harm/injury noted. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation for the reported controller purge disc issues has been completed. the controller and the data logs were returned for evaluation. the device was evaluated and the reported failure was unable to be reproduced. there were no signs of glucose leaks in or around the purge. data analysis performed found that on 10/27, the controller would not progress past start. if the purge motor shaft was unable to rotate, the purge cassette cannot be inserted if the tooling is not aligned. the root cause of the aic issue could not be determined since the issue could not be produced. however, there most common cause of the reported issue is glucose build up in the purge assembly since that is the most common cause for the purge motor to not be able to rotate. the purge motor was most likely blocked causing the purge cassette to be unable to be properly connected.
32519 | 1220648-2023-05270 | 18270834 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18270834&pc=OZD | 11/10/2023 | 01/16/2024 |  | 12/05/2023 |  |  | 11/10/2023 | 12/05/2023 |  |  | 09/30/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024401729 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that upon attempting the impella 5.5 delivery, the impella was unable to be implanted due to the anatomy of the patient. the implant was aborted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. as per clinical, the physician placed a graft tunneled supraclavicular on the innominate and later the impella was unable to pass due to anatomy. the cause of the delivery issue was most likely patient condition related with the pump unable to pass through innominate due to anatomy per clinical.
32520 | 1220648-2023-05271 | 18271440 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18271440&pc=OZD | 11/09/2023 | 02/16/2024 |  | 12/05/2023 |  |  | 11/09/2023 | 12/05/2023 |  |  | 03/31/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024342475 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following implant of an impella cp pump, the 60-year-old male patient developed a hematoma coinciding with uncontrolled bleeding at the access site. due to the noted condition, the pump was explanted and the patient was placed on inotropes until an impella 5.5 could be implanted the following morning. | [5 empty] | The investigation into access site bleeding ¿ major issue has been completed. the device was not returned for investigation. according to clinical information, the pump was removed due to hematoma and uncontrolled bleeding at access site. the cause of the access site bleeding issue was not determined since product was not returned and sufficient clinical information was not provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32521 | 1220648-2023-05272 | 18271462 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18271462&pc=OZD | 11/09/2023 | 04/18/2024 |  | 12/05/2023 |  |  | 11/09/2023 | 12/05/2023 |  |  | 07/31/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024386974 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the 78-year-old male patient decompensated and the impella cp was implanted. the staff noted that the impella showed ¿purge system kinked¿ twice during setup. after insertion, pump ran at 3.9l/min for the first minute of support. after 1-2 minutes, pump showed ¿impella in ventricle¿ and ¿suction¿ alarms and flows dropped to ~1 l/min. the doctor attempted to reposition, but suction persisted even at p-2. the staff noted that no thrombus had been visible on bedside echo prior to insertion. doctor decided to explant impella and implant intra-aortic balloon pump (intra-aortic balloon pump (iabp) for support. | [5 empty] | The investigation into the low pump flow has been completed since the original report was filed. the pump was returned, tested, and evaluated. visual inspection found a kink on the cannula. the root cause of low flow was due to kinked cannula. the cause of kinked cannula was likely due to the pump deep in the ventricle. patient had a small lv. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
32522 | 1220648-2023-05273 | 18271467 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18271467&pc=OZD | 11/09/2023 | 02/28/2024 |  | 12/05/2023 |  |  | 11/09/2023 | 12/05/2023 |  |  | 09/30/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024398352 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a 60-year-old male patient experienced a decrease in distal flow following impella implant. it was noted that the patient had required a high dose of levo during a scheduled high risk percutaneous coronary intervention procedure leading up to the loss in distal pulse. as a result of the loss of flow, an external fem-fem bypass was placed. support continued with the implanted pump until the patient eventually passed from their condition. there was no allegation or evidence of device association between the impella pump and the patient's passing. | [5 empty] | The investigation into the limb ischemia ¿ reversible has been completed. the device and data were not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.
32523 | 1220648-2023-05274 | 18271483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18271483&pc=OZD | 11/09/2023 | 10/04/2024 |  | 12/05/2023 |  |  | 11/09/2023 | 12/05/2023 |  |  | 08/31/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024390728 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 63-year-old male in st elevated myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that an implanted impella 5.5 pump began triggering false placement signal low alarms during patient support. the alarm was initially disabled; however, the decision was ultimately made to replace the pump with a new impella 5.5 device. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for optical signal issue has been completed. pump was not returned for investigation. the data logs were returned and investigated. all parameters were found to be within specification and suction was observed prior to placement signal low alarms. the placement signal was seen trending down. pump was explanted due to the placement signal low and also since the pump was ran for 35 days. therefore, per data logs review, the root cause of the optical signal issue was most likely patient condition related. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05274: d6a and d6b revised from the initial submission in accordance with updated procedures. g1 reporting contact fax number missing.
32524 | 1220648-2023-05275 | 18271515 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18271515&pc=OZD | 11/09/2023 | 01/16/2024 |  | 12/05/2023 |  |  | 11/09/2023 | 12/05/2023 |  |  | 07/31/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024375096 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 pump met resistance during attempted delivery. it was noted that the pump was unable to pass the anastomosis due to the sharp angle of the subclavian artery. as a result, the 5.5 implant was aborted and an impella cp pump was placed for planned support. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the cause of the delivery issue was patient condition related due to patient condition of small, diseased blood vessels.
32525 | 1220648-2023-05276 | 18271542 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18271542&pc=OZD | 11/10/2023 | 03/10/2025 |  | 12/05/2023 |  |  | 11/10/2023 | 12/05/2023 |  |  | 07/31/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024392479 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  | White | Injury | Yes | Yes | Difficult to Advance | Material Twisted/Bent | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that difficulty was encountered while attempting to implant an impella cp pump for patient support. after initially being unable to advance the pump due to the patient's tortuous anatomy, a 14 fr long introducer was placed and the pump was re-inserted. during this second attempt, resistance was met as the pump was advancing through the sheath due to a kink in the introducer. despite the noted issue, the pump was able to advance into the left ventricle. it was noted that the arteriotomy experienced persistent oozing following the implant for which the sutures were adjusted, manual pressure was held, and epi/lido was injected. support was maintained with the implanted pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Pump was not returned for investigation. clinical details mention imaging was used for impella advancement and the patient had tortuous vessels. difficulty was experienced while getting impella through the sheath due to a sheath kink which was due to the vessel tortuosity. therefore, the root cause of difficulty inserting pump through introducer issue was a sheath kink due to patients tortuous vessels. as per the clinical details, persistent oozing from the right arteriotomy occurred after inserting the guidewire repositioning unit, requiring manual pressure for successful hemostasis. therefore, the root cause of the access site bleeding issue was most likely due to lack of vessel recoil onto the sheath.
32526 | 1220648-2023-05277 | 18271593 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18271593&pc=OZD | 11/10/2023 | 01/16/2024 |  | 12/05/2023 |  |  | 11/10/2023 | 12/05/2023 |  |  | 09/30/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024398901 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/10/2023 |  |  |  |  | The user facility reported a 59-year-old male (unknown race) patient was to be implanted with an impella 5.5 device for mechanical circulatory support. the pump implant was attempted and the physician could not cross the aorta due to the patient's anatomy. an impella cp pump was subsequently implanted for patient support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported delivery issue -anatomy has been completed since the original report was filed. no product was returned. as per clinical first and second 5.5 pump would not cross aortic arch and after multiple attempts with torque and pressure on second pump the catheter caused to bend due to resistance at aortic arch. the cause of the delivery issue was most likely patient condition related with both pumps unable to pass aortic arch after multiple attempts per procedure.
32527 | 1220648-2023-05278 | 18271656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18271656&pc=OZD | 10/31/2023 | 09/02/2024 |  | 12/05/2023 |  |  | 10/31/2023 | 12/05/2023 |  |  | 06/30/2025 |  | 12/05/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024360219 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Retroperitoneal Hemorrhage | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/31/2023 |  |  |  |  | The user facility reported a 73-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. following impella implant, the patient developed a hematoma at the insertion site. manual pressure was applied to resolve the noted condition. roughly four days after implant, the patient also experienced a retroperitoneal bleed. anticoagulants were stopped as a result of the noted condition. support with the implanted pump was maintained throughout the event. | [5 empty] | The investigation for the vascular damage/gi bleed has been completed. the product was not returned for investigation. per the clinical information provided, the retroperitoneal bleeding occurred on the fourth day, but no other information regarding the incident was provided. the root cause of the vascular damage issue was not determined due to insufficient clinical information and unknown relation to impella. b.7 revised as information was inadvertently omitted from the initial manufacturer device report number 1220648-2023-05278. d.6a and d.6b revised from the initial submission of manufacturer device report number 1220648-2023-05278 in accordance with updated procedures. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2023-05278 and should not have been. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2023-05278 was submitted. h.6 code 925 was added since the initial manufacturer device report number 1220648-2023-05278 was submitted in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report number 1220648-2023-05278 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/pos, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32528 | 1220648-2023-05202 | 18255286 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18255286&pc=OZD | 11/09/2023 | 01/26/2024 |  | 12/04/2023 |  |  | 11/09/2023 | 12/04/2023 |  |  | 08/31/2025 |  | 12/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024398012 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Deformation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/09/2023 |  |  |  |  | The user facility reported an 87-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the impella pump kinked just below the outlet. patient is a small person. insertion was smooth. physician is an experienced impella user with excellent skills. the pump was replaced. no patient harm or injury was noted. | [5 empty] | The investigation into the product damage (cannula kink) has been completed. the device was returned for benchtop investigation. based on the clinical information provided, the pump was kinked following a smooth insertion. it was reported that the small patient had a right aortic arch. the returned product was visually inspected, and the cannula kink was confirmed just below the outlet cage. the cause of the cannula kink was use related due to difficult navigation of the aortic arch which likely required excessive force. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
32529 | 1220648-2023-05203 | 18255374 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18255374&pc=OZD | 11/29/2023 | 02/12/2024 |  | 12/04/2023 |  |  | 11/29/2023 | 12/04/2023 |  |  |  |  | 12/04/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 1748660 |  | 0042-0040 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/29/2023 |  |  |  |  | The user facility reported a 56-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support and the automated impella controller (aic) had a "controller error" alarm. the aic was replaced successfully. reportable due to potential patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - system self-check error has been completed since the original report was filed. the product was returned for investigation. data logs confirmed the reported failure mode given there was a controller error alarms triggered during the clinical procedure. the failure was unable to be reproduced throughout testing. the cause of the issue was not determined since the failure could not be reproduced.
32530 | 1220648-2023-05204 | 18255777 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18255777&pc=OZD | 11/12/2023 | 01/24/2024 |  | 12/04/2023 |  |  | 11/12/2023 | 12/04/2023 |  |  | 09/30/2025 |  | 12/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024393866 |  | 1000211 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/12/2023 |  |  |  |  | The user facility in japan reported a 62-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. after 29 days of support purge flow began to fluctuate on the impella 5.5 and they found that liquid was leaking from the purge line filter. gauze and pressure was applied and the team fixed the area where the purge pressure could be maintained to address the issue. patient was on support for additional 2 days post issue. no patient harm reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the purge sidearm leak has been completed. the device was returned for evaluation. upon visual inspection, the purge tubing between the sidearm filter neck and red impella plug was found to be damaged. the keyence imaging showed a white discoloration at the damaged end of sidearm filter joint. no damages to retainer or other physical abnormalities were found. the cause of the purge leak was damage to the filter to tubing joint specifically at the tubing to filter neck bond.
32531 | 1220648-2023-05205 | 18256105 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18256105&pc=OZD | 11/10/2023 | 08/22/2024 |  | 12/04/2023 |  |  | 11/10/2023 | 12/04/2023 |  |  | 09/30/2025 |  | 12/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024401021 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Mechanical Problem | Difficult to Remove | Defective Device | Patient-Device Incompatibility |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/10/2023 |  |  |  |  | The user facility reported a 69-year-old male patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. at the end of the case while explanting the cp through the 14fr introducer sheath, the pump would not come through the sheath despite great effort to pull it through. instead, md pulled the entire unit out as one without issue. there was no vessel damage per angiography. patient is stable post removal. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the introducer damage and delivery issues has been completed. no product or images were returned. there was difficulty inserting the 14 french peel away sheath. while explanting the impella cp through the 14 french introducer sheath, the pump would not come through the sheath. the physician pulled the entire unit out as one without issue. it was noted that the sheath had "buckled/folding in on peel-away. no product or images were returned. the root cause of the difficulty inserting/removing the introducer was not determined as no product was returned and limited clinical information was provided. the root cause of the removal issue was a introducer sheath kink. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ b.7 information was added as it was inadvertently omitted from the previously submitted manufacturer device report number 1220648-2023-05205. d.6a and d.6b revised from the initial submission in accordance with updated procedures since the initial manufacturer device report 1220648-2023-05205 was submitted. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2023-05205 and should not have been. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2023-05205 was submitted. h.6 codes 1384 and 1528 were reported incorrectly on the initial manufacturer device report number 1220648-2023-05205. new codes were added. h.10 additional manufacturer narrative instructions for use were inadvertently omitted from the previously submitted manufacturer device report number 1220648-2023-05205.
32532 | 1220648-2023-05206 | 18256212 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18256212&pc=OZD | 11/04/2023 | 08/13/2025 |  | 12/04/2023 |  |  | 11/04/2023 | 12/04/2023 |  |  |  |  | 12/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1264733 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller (aic) had a system self check failure message after powering on the aic. the aic was shutdown and restarted and produced the same error. there was no reported patient harm. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | A.2 age should have been left blank on the initial manufacturer device report number 1220648-2023-05206 as it is unknown. a.3 unknown should have been selected on the initial manufacturer device report number 1220648-2023-05206 as it is unknown. d.1 and d.2 brand name and common device name were revised in accordance with updated procedures since the initial manufacturer device report number 1220648-2023-05206 was submitted. d.4 model number, serial number and catalog number were revised in accordance with updated procedures since the initial manufacturer device report number 1220648-2023-05206 was submitted. e.2 and e.3 revised in accordance with updated procedures since the initial manufacturer device report number 1220648-2023-05206 was submitted. g.1 reporting contact email was revised since the initial manufacturer device report number 1220648-2023-05206 was submitted in accordance with updated procedures. reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2023-05206 and was added.
32533 | 1220648-2023-05207 | 18256356 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18256356&pc=OZD | 11/05/2023 | 04/30/2024 |  | 12/04/2023 |  |  | 11/05/2023 | 12/04/2023 |  |  | 07/31/2025 |  | 12/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024376931 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/05/2023 |  |  |  |  | The user facility reported a 66-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an impella cp pump experienced flow and positioning issues immediately following implant. the device was said to have gotten caught underneath the papillary muscle. repositioning was attempted, but the flow rate remained inadequate, and the device was pulled back across the aortic valve. after being unable to re-advance the pump over the valve, the decision was made to replace the pump. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation has been completed. the pump was not returned for investigation. the total run time was only six minutes. the data logs showed an impella position in aorta alarm during pump use. according to clinical notes, the pump was pulled back into the aorta, and the doctor was unable to place the pump in the proper position. the cause of the positioning issue was most likely related to use since the pump was accidentally pulled back into the aorta during repositioning. an improper position of the pump could lead to a low pump flow issue. d.4 revised model number as previously entered incorrectly. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 added code 4628 as it was inadvertently left off the previously submitted manufacture device report.
32534 | 1220648-2023-05208 | 18256416 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18256416&pc=OZD | 11/06/2023 | 02/28/2024 |  | 12/04/2023 |  |  | 11/06/2023 | 12/04/2023 |  |  | 09/30/2025 |  | 12/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024400702 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/06/2023 |  |  |  |  | The user facility reported a 68-year-old male patient was implanted with an impella device for mechanical circulatory support. the complainant reported that a patient experienced a loss in pulse distal to their impella access site following impella cp implant. it was noted that the limb was cool to the touch and that symptoms of pain and discoloration were present. the patient had a history of severe peripheral vascular disease and the condition did not improve by the following day. as a result, the pump was explanted after less than two days of support. | [5 empty] | The investigation into the limb ischemia ¿ reversible has been completed. the device and data were not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32535 | 1220648-2023-05209 | 18256472 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18256472&pc=OZD | 11/06/2023 | 09/23/2024 |  | 12/04/2023 |  |  | 11/06/2023 | 12/04/2023 |  |  | 08/31/2025 |  | 12/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024393805 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/06/2023 |  |  |  |  | The user facility reported a (b)(6) female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the sterile sleeve on the implanted impella cp pump was torn by the doctor during patient support. the sleeve was re-secured via tegaderm and the need for further repositioning was left up to the doctor. as support continued, it was noted that the patient experienced some bleeding from their groin when heparin was initiated. an additional suture at the access site allowed for hemostasis and support with the implanted pump was maintained. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for product damage (sterile sleeve damage) and access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the sterile sleeve damage was most likely due to use as the clinical states the physician tore the sleeve. the root cause of the major access site bleeding was most likely use related as adding an additional suture resolved the bleeding. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05209. d4: model number. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. h6 component code revised from the initial submission in accordance with updated procedures.
32536 | 1220648-2023-05210 | 18256511 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18256511&pc=OZD | 11/06/2023 | 03/14/2024 |  | 12/04/2023 |  |  | 11/06/2023 | 12/04/2023 |  |  | 09/30/2025 |  | 12/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 EU |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2024408918 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Female |  |  | White | Injury | Yes | Yes | Difficult to Advance | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 59-year-old female patient was scheduled to receive hemodynamic support from an impella 5.5 smart assist heart pump during high-risk coronary artery bypass surgery. several attempts to insert the impella 5.5 via the axillary artery were unsuccessful. it was reported that arteriosclerosis was present and vascular injury occurred. the vascular injury did not require intervention. the impella 5.5 was then introduced via a direct aortic approach. the impella procedure subsequently proceeded without any further difficulties. the patient was weaned from impella pump support and mobilized two days after the procedure. | [5 empty] | The investigation into the delivery issues- anatomy and the vascular damage - peripheral vessel issues has been completed since the original report was filed. the device was not returned for investigation. per the provided clinical information, the root cause of the delivery issue was patient condition related to small/ diseased vessels due to stenosis. the root cause of the vascular damage issue was use related due to improper technique of implanting the impella against the ifu. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32537 | 1220648-2023-05211 | 18256574 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18256574&pc=OZD | 11/06/2023 | 01/16/2024 |  | 12/04/2023 |  |  | 11/06/2023 | 12/04/2023 |  |  | 08/31/2025 |  | 12/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2024387756 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/06/2023 |  |  |  |  | The complainant reported that a patient of unknown age and gender was attempted to be implanted with an impella cp heart pump for hemodynamic support. however, the procedure had to be aborted due to patient anatomy. the patient had tortuous vascular anatomy in the groin artery. no patient harm reported due to the issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into delivery issues- anatomy has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the root cause of the delivery issue- anatomy was determined to be patient condition because of the tortuous vascular anatomy of the patient. the failure mode will be monitored and trended.
32538 | 1220648-2023-05212 | 18256648 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18256648&pc=OZD | 11/06/2023 | 04/29/2024 |  | 12/04/2023 |  |  | 11/06/2023 | 12/04/2023 |  |  | 08/31/2025 |  | 12/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024393346 |  | 0048-0045 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/06/2023 |  |  |  |  | The user facility reported a 74-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. following successful implant, the impella cp pump failed to obtain flows above .5 liters per minute. as a result, the pump was subsequently explanted and a clot was noted at the inflow. a new impella cp pump was implanted with no further issue. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow has been completed since the original report was submitted. the device was not returned for investigation. data logs reported suction with low flow during the case run without any motor current spike. according to clinical findings, the device was explanted due to low pump flow, and a clot was noted on the inflow cage. the cause of the low pump flow issue was most likely biomaterial, based on given clinical details. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
32539 | 1220648-2023-05213 | 18256667 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18256667&pc=OZD | 11/06/2023 | 04/23/2024 |  | 12/04/2023 |  |  | 11/06/2023 | 12/04/2023 |  |  | 08/31/2025 |  | 12/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024393070 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/06/2023 |  |  |  |  | The user facility reported a 78-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the cp pump experienced low flows following implant and subsequently triggered suction alarms. upon further observation, it was identified that the cannula of the pump was kinked. as a result, the decision was made to replace the device with a new impella cp pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow issue has been completed since the original report was submitted. the device was not returned for investigation. data logs show suction alarms throughout the one-hour case run along with low pump. based on given clinical details, the cause of the low pump flow issue was most likely kinked cannula.
32540 | 1220648-2023-05214 | 18256798 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18256798&pc=OZD | 11/06/2023 | 05/23/2024 |  | 12/04/2023 |  |  | 11/06/2023 | 12/04/2023 |  |  | 08/31/2025 |  | 12/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024395729 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/06/2023 |  |  |  |  | The user facility reported an (b)(6) male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the implanted impella cp pump experienced a spike in motor current coinciding with left ventricle waveform issues after cross clamp removal post coronary artery bypass graft surgery. despite troubleshooting, suction issues persisted and adequate flow levels could not be achieved. as a result, the decision was made to explant the pump. there was no harm to the patient as a result of the failure. | [5 empty] | The investigation into the low pump flow issue has been completed since the original report was submitted. the pump was not returned for investigation the cause of the low pump flow issue was most likely use related, based on given clinical details. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32541 | 1220648-2023-05215 | 18256855 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18256855&pc=OZD | 11/06/2023 | 12/04/2023 |  | 12/04/2023 |  |  | 11/06/2023 | 12/04/2023 |  |  |  |  | 12/04/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1347481 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | Charging Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an automated impella controller (aic) experienced a battery failure alarm. the aic was removed from service and a loaner was sent to the customer site for use as needed. there was no patient harm. | [5 empty] | The investigation into the reported "aic - battery failure (red alarm)" has been completed. the product was returned for investigation. the device was inspected and the failure mode was reproduced on an engineering bench. battery testing confirmed battery 1 cell 1 was outputting 683 mv, significantly under the 4v the rest of the cells were outputting. the battery 1 pack voltage was measured to be 0 v rather than the expected 16 v, and the battery lcd indicator was blank (fully discharged) due to the cell imbalance. during data analysis, there was a battery failure alarm (transport connection blown/ missing) due to low pack voltage. additionally, the lt logs showed that battery 1 had a cell imbalance on cell 1 of 600 mv. upon conclusion of the investigation, the battery failure alarm was due to a defective battery. the root cause of the defective battery was determined to be a single cell failure.
32542 | 1220648-2023-05216 | 18257197 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18257197&pc=OZD | 11/06/2023 | 03/06/2024 |  | 12/04/2023 |  |  | 11/06/2023 | 12/04/2023 |  |  | 03/31/2025 |  | 12/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024324163 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that impella 5.5 was inserted into the left ventricle without complication. the device started at p2 and when it was increased to p6 flows then dropped drastically. troubleshoot was initiated with no improvement in impella's flows. the impella was exchanged with no patient harm. | [5 empty] | The investigation into the low pump flow issue has been completed since the original report was filed. the device was returned and tested after arriving in fs. visual inspection found no issues on the pump. the failure mode could not be reproduced as the pump ran without issue under bench test. the root cause of low flow was most likely due to patient condition as no problem was found on the pump. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
32543 | 1220648-2023-05217 | 18257309 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18257309&pc=PYX | 11/06/2023 | 09/25/2025 |  | 12/04/2023 |  |  | 11/06/2023 | 12/04/2023 |  |  | 03/31/2025 |  | 12/04/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024394946 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a 52-year-old male patient experienced active oozing from their femoral access site following impella rp implant. additional sutures were placed to treat the condition and support continued as planned. no further intervention was required. | Additional information was provided that stated "on the last day of support, high purge pressure was reported. the purge flow decreased from 18 ml/hr to 2.5 ml/hr with purge pressure in the 900s mmhg. no kinks were noted and the purge cassette was changed, but that did not resolve the issue. tissue plasminogen activator (tpa) was ran to increase purge flow. no leaks were observed. the pump was replaced with a new pump. | One day after implant, it was then noted that the placement signal became unreliable as the signal was lost and the flow calculation was no longer available. the staff was advised on how to monitor the functionality of the pump without the failed application. the coinciding alarm was disabled and support was maintained with the implanted pump. |  |  |  | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ | The investigation for the reported access site bleeding and high purge pressure has been completed. the impella device was returned for evaluation. visual inspection of the pump found no biomaterial was observed around the impeller. the pump was purged in the lab and high purge pressure was not able to be reproduced. the data logs were reviewed and confirmed high purge pressure and low purge flows on the last day of support. there are suction alarms immediately before and during the event. the root cause of the high purge pressure was most likely biomaterial; tpa was ran and the high purge pressure gradually resolved indicating that a clot was flushed out. the root cause of the access site bleeding was not determined due to a lack of clinical details. b1 checked product problem h6-device code-added code 925 h6-problem code updated to 1491. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data logs confirm high purge pressure and low purge flows on the last day of support. there are suction alarms immediately before and during the event. the purge cassette was not returned. the high purge pressure was unable to be reproduced in the lab. the data logs also confirm placement signal not reliable alarms and the placement signal out of range at the same time as flows are displayed to be 0l/min. data logs show dp sensor drift which was confirmed when run in a water bucket in the lab. the root cause of the high purge pressure was most likely biomaterial; tpa was ran and the high purge pressure gradually resolved indicating that a clot was flushed out. the root cause of the access site bleeding - major was not determined due to a lack of clinical details. the root cause of the placement signal issue was determined to be dp sensor drift. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
32544 | 1220648-2023-05218 | 18257365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18257365&pc=OZD | 11/07/2023 | 01/16/2024 |  | 12/04/2023 |  |  | 11/07/2023 | 12/04/2023 |  |  | 08/31/2025 |  | 12/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC |  |  | US |  | 2024386642 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/07/2023 |  |  |  |  | The user facility reported a 76-year-old female undergoing cardiac surgery was scheduled for impella cp implant for mechanical circulatory support. it was noted that the impella cp optical pump met resistance during attempted implant. during the delivery, it was documented that the catheter would not advance past the patient's right iliac. due to the resistance associated with the patient's anatomy, the decision was made to abort the implant. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported delivery issue -anatomy has been completed since the original report was filed. the root cause of the delivery issue was determined to be patient condition because of patient vascular anatomy.
32545 | 1220648-2023-05219 | 18257394 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18257394&pc=OZD | 11/05/2023 | 03/14/2024 |  | 12/04/2023 |  |  | 11/07/2023 | 12/04/2023 |  |  | 08/31/2025 |  | 12/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024391481 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/07/2023 |  |  |  |  | The user facility reported a 80-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that a patient suffered an aortic dissection during impella cp support. it was noted via contrast imaging during a supported percutaneous coronary intervention, that a dissection was present in the ascension to the aortic arch. it was unknown when the dissection occurred. a stent was placed to treat the dissection and the patient's hemodynamics quickly stabilized. the pump was subsequently explanted following the intervention. | [5 empty] | The investigation into the vascular damage - great vessel issue has been completed since the original report was filed. the device was not returned for investigation. the root cause of the vascular damage - great vessel was not determined since the relationship to the impella is unknown and there are limited clinical details about the dissection. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
32546 | 1220648-2023-05220 | 18257530 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18257530&pc=OZD | 11/07/2023 | 04/30/2024 |  | 12/04/2023 |  |  | 11/07/2023 | 12/04/2023 |  |  | 07/31/2025 |  | 12/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024391327 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/07/2023 |  |  |  |  | The user facility reported a 74-year-old male patient was implanted with an impella cp device for mechanical circulatory support. the complainant reported that upon attempted implant of the impella cp pump, the cannula prolapsed on itself and the pigtail became entangled in the outflow of the device. it was noted that the issue occurred while attempting to advance the device across the aortic valve. as a result of the noted issue, the decision was made to abort the implant. there was no patient harm. | [5 empty] | The impella device was received from the customer and an evaluation of the device is pending. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues- use has been completed since the original report was filed. the pump was returned and evaluated. there was a kink in the cannula and the guidewire was twisted it was reported that while attempting to cross the av, wire access to the lv was lost. the root cause of the delivery issues is most likely due to use as the cannula was kinked and guidewire was twisted/unraveling. in accordance with updated procedures, sections d6, d9, and h10 have been revised from the initial submission.
32547 | 1220648-2023-05221 | 18257762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18257762&pc=OZD | 11/06/2023 | 09/23/2024 |  | 12/04/2023 |  |  | 11/06/2023 | 12/04/2023 |  |  | 09/30/2025 |  | 12/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024402402 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that upon insertion the 58-year-old male patient had a vessel dissection from the long 14fr peel away sheath/introducer, to be utitlized for the impella cp access. the staff removed the impella and perclosed the arteriotomy and patient had good hemostasis in less than 10 minutes. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported vascular damage has been completed. according to the clinical, while the long peel away sheath was being inserted the patient began complaining of pain. upon further investigation, a dissection was discovered. the long peel away sheath was then removed and replaced with the short peel away sheath. after the patient was stable, the physician considered leaving the impella device in place. ultimately, the decision was made to remove the pump and post close the arteriotomy with two perclose. hemostasis was then achieved after less than ten minutes of pressure. no product was returned for an investigation. data log analysis was not necessary for this complaint. the most likely cause of the peripheral vessel vascular damage was use related. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk pci states the following: section: warnings & cautions: cautions. ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ added- h6 impact code 4624, h6 component code 925.
32548 | 1220648-2023-05222 | 18257781 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18257781&pc=OZD | 11/06/2023 | 02/02/2024 |  | 12/04/2023 |  |  | 11/06/2023 | 12/04/2023 |  |  | 07/31/2025 |  | 12/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024380696 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was further reported that a hematoma was identified on the left groin site at the insertion of the impella sheath. manual pressure was been applied to site along with markers made to outside edges of hematoma. heparin was held due to elevated partial thromboplastin time (ptt)and patient was given1 unit of blood products given. patient received 2 more units of blood products and hemodynamics improved with no bleeding issues. | The user facility reported a 66-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant had a support ongoing with a cp when after five days the pump was explanted. the pump support had been successful despite bleeding issues that prompted the team to hold anticoagulation. the pump was explanted and at the time of explant there was thrombosis seen. the clot was evacuated with vascular surgery performing penumbra and endarterectomy. |  |  |  |  | The investigation for the reported thrombus and the access site bleed has been completed. no product was returned. clinical information was sufficient to determine the root cause of the thrombus. the cause of the thrombus issue was patient condition related per clinical review. the cause of the access site bleeding issue was not determined due to insufficient clinical information and since product was not returned. abiomed does not recommend evaluating based on ptt values. section h6: the health effects clinical code and impact code were inadvertently left out in the initial report which has been added to this report. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32549 | 1220648-2023-05223 | 18257818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18257818&pc=OZD | 11/06/2023 | 12/03/2025 |  | 12/04/2023 |  |  | 11/06/2023 | 12/04/2023 |  |  |  |  | 12/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1146764 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during training, the automated impella controller (aic) had a yellow ¿controller error¿ alarm that occurred upon turning on. the aic was turned off and back on and the yellow ¿controller error¿ alarm persisted. no patient was involved, and this error did not occur in a clinical setting, only in training purposes. | [5 empty] | Data analysis: on the complaint date (nov 06, 2023), the console displayed ¿controller error (loss of communication with keypad or between the kbd and i2c for kbd hw revision 1) ¿ alarm,¿ event #24. this confirms the reported failure mode. device analysis: the console was tested upon arrival at fs, and no issue was found with the ribbon cable between the 4-in-1 and the glass keyboard. upon performing the aic 200 power cycle, the issue could not be reproduced. root cause: the root cause for the controller error could not be determined due to the inability to reproduce it. d2 common device name updated. g1 reporting contact email and reporting contact fax number updated. | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
32550 | 1220648-2023-05224 | 18257869 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18257869&pc=OZD | 11/28/2023 | 03/06/2024 |  | 12/04/2023 |  |  | 11/28/2023 | 12/04/2023 |  |  | 09/30/2025 |  | 12/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024401338 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/28/2023 |  |  |  |  | The user facility reported a 36-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a 5.5 delivered for support of a younger patient being treated for cgs (cardiogenic shock) and cardiomyopathy. the pump was placed but after 2 weeks the pump was removed due to saturated flows. the pump was removed and patient placed on bi-v centrimag. | [5 empty] | The investigation into the saturated flow issue has been completed since the original report was filed. the pump was returned and tested after arriving in fs. visually inspected and biomaterial ingestion over the purge gap was observed. the cause of the saturated flow issue was due to gradual biomaterial ingestion over purge gap per field investigation and log review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32551 | 1220648-2023-05226 | 18257885 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18257885&pc=OZD | 11/11/2023 | 07/02/2024 |  | 12/04/2023 |  |  | 11/11/2023 | 12/04/2023 |  |  | 08/31/2025 |  | 12/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024395739 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/11/2023 |  |  |  |  | Additional information: care was withdrawn and the patient expired. per adiomed safety office, abiomed does not have enough information to associate use of device with outcome. | The us complainant had placed an impella cp for support of the patient post cabg (coronary artery bypass grafting). the patient had the cabg and on the icu floor the patient arrested and team made choice to place a cp. the cp was placed but the access site had persistent bloody ooze. the team troubleshot by rbc (red blood cell) infusion x3 units, manual pressure, and a stich. the pump remained on for a total of 82 hours when care was withdrawn and patient expired. the death is not attributed to impella. |  |  |  |  | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of access site bleeding was determined to be use issue because the issue was resolved upon addition of sutures at the bleeding site. b5 additional information added. d6a, d6b added d4 model number corrected b7 added h5 corrected h6 component code added f6 and f8 should have been blank in the initial mdr. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32552 | 1220648-2023-05225 | 18257897 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18257897&pc=OZD | 11/29/2023 | 01/17/2024 |  | 12/04/2023 |  |  | 11/29/2023 | 12/04/2023 |  |  | 09/30/2025 |  | 12/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024401720 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/29/2023 |  |  |  |  | The user facility reported a 73-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella 5.5 procedure was aborted due to the patient¿s anatomy. the impella was not able to pass the aortic arch. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the cause of the delivery issue was most likely due to patient condition.
32553 | 1220648-2023-05227 | 18257915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18257915&pc=OZD | 11/09/2023 | 01/16/2024 |  | 12/04/2023 |  |  | 11/09/2023 | 12/04/2023 |  |  | 08/31/2025 |  | 12/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024399165 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/09/2023 |  |  |  |  | The user facility reported an 82-year-old male post cardiotomy cardiogenic shock / low cardiac output syndrome was to be implanted with an impella 5.5 for mechanical circulatory support. the delivery was not possible due to the native anatomy. despite all the pump attempts and wire exchanges, the 5.5 failed and so the team instead placed a cp. no harm or injury. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported delivery issue -anatomy has been completed since the original report was filed. the cause of the delivery issue was patient condition related due to disease or geometry issues of the proximal innominate artery.
32554 | 1220648-2023-05228 | 18258073 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18258073&pc=OZD | 11/13/2023 | 01/24/2024 |  | 12/04/2023 |  |  | 11/13/2023 | 12/04/2023 |  |  | 09/30/2025 |  | 12/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024401783 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/13/2023 |  |  |  |  | The user facility reported an unknown aged male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. upon impella insertion the guidewire kinked, and md was unable to advance. a new device was successfully implanted instead. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues- use has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the most likely root cause of the delivery issue was most likely due to use issue as the wire was kinked and the pump was attempted to be delivered over the kinked wire.
32555 | 1220648-2023-05229 | 18260453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18260453&pc=OZD | 11/05/2023 | 11/17/2025 |  | 12/04/2023 |  |  | 11/06/2023 | 12/04/2023 |  |  |  |  | 12/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1557944 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Battery Problem | Computer Operating System Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 11/06/2023 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant in japan had a cp support in which when preparing the case/starting support the (automated impella controller) aic powered up with a red battery failure alarm. the case was completed with the cp and there was no noted harm or injury from any delay. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | A.2 revised age as it was reported incorrectly on initial manufacturer device report number 1220648-2023-05229. a.3 and a.5 added gender and ethnicity as they were inadvertently omitted from the initial manufacturer device report number 1220648-2023-05229. b.3 revised date of event as it was reported incorrectly on initial manufacturer device report number 1220648-2023-05229. d.1 revised brand name as it was reported incorrectly on the initial manufacturer device report number 1220648-2023-05229. d.4 revised model number and catalog number from the initial submission of manufacturer device report number 1220648-2023-05229 in accordance with updated procedures. d.9 device returned to manufacturer date was added as it was inadvertently omitted from the initial manufacturer device report number 1220648-2023-05229. e.1 added adrress line 2 as it was inadvertently omitted from the initial manufacturer device report number 1220648-2023-05229. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2023-05229 and should not have been. g.1 reporting contact email was revised since initial manufacturer device report number 1220648-2023-05229 was submitted in accordance with updated procedures. reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2023-05229 and was added. g.4 revised pma/510(k) number as it was reported incorrectly on initial manufacturer device report number 1220648-2023-05229. h.3 device not returned should not of been selected on initial manufacturer device report number 1220648-2023-05229. was the device evaluated should not have been selected on initial manufacturer device report number 1220648-2023-05229 as it has not yet been. if the device was not evaluated selection was incorrect on initial manufacturer device report number 1220648-2023-05229 as the device was returned and the investigation of the device was anticipated. h.5 revised label for single use as it was erroneously entered on the initial manufacturer device report number 1220648-2023-05229. h.6 medical device problem code was revised and component code was added since the initial submission of manufacturer device report number 1220648-2023-05229 in accordance with updated procedures. h.8 revised usage of device as it was erroneously entered on the initial manufacturer device report number 1220648-2023-05229.
32556 | 1220648-2023-05230 | 18260802 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18260802&pc=OZD | 11/05/2023 | 01/11/2024 |  | 12/04/2023 |  |  | 11/05/2023 | 12/04/2023 |  |  | 08/31/2025 |  | 12/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2024394334 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/05/2023 |  |  |  |  | The user facility reported a 56-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported bleeding from the impella site, despite advancing the repositioning unit into place. the pump was removed as a result. no further hemodynamic support was warranted. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major has been completed. the product was not returned. clinical notes detail an access site bleed upon insertion of impella cp. the patient had high act. the cause of the access site bleeding was high patient act values.
32557 | 1220648-2023-05175 | 18243929 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18243929&pc=OZD | 11/03/2023 | 02/15/2024 |  | 12/01/2023 |  |  | 11/03/2023 | 12/01/2023 |  |  | 09/30/2025 |  | 12/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024399136 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/03/2023 |  |  |  |  | The user facility reported a 46-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the patient developed a right chest wall hematoma roughly three days after impella 5.5 implant via right axillary access. the hematoma was noted to be accompanied by significant pain for with medication was provided. although the staff was unwilling to evacuate the hematoma due to infection risks, the decision was made to halt heparin administration. about three days later, it was reported that the impella site was still swollen; however, the pain had notable decreased. support with the implanted pump remained ongoing. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed. the device was not returned for investigation. as per the clinical details, doctor reported patient having hematoma at impella site. the cause of the access site bleeding issue was not determined since product was not returned and sufficient clinical information was not provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32558 | 1220648-2023-05176 | 18243952 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18243952&pc=OZD | 11/03/2023 | 04/16/2024 |  | 12/01/2023 |  |  | 11/03/2023 | 12/01/2023 |  |  | 08/31/2025 |  | 12/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2024395234 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/03/2023 |  |  |  |  | The user facility reported a patient of unknown age/gender in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. one day after implant, the patient had access site bleeding which resulted in low hemoglobin levels. as a result, the patient was given 12 units of blood and fem stop was used to address the bleeding. additionally, on 3rd day patient had blue leg which the physician suspected of thrombus, which resolved next day. groin surgically treated, and patient is currently stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major and the limb ischemia ¿ reversible has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia and the access site bleeding issues was not determined.
32559 | 1220648-2023-05177 | 18243984 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18243984&pc=OZD | 11/03/2023 | 08/14/2024 |  | 12/01/2023 |  |  | 11/03/2023 | 12/01/2023 |  |  | 08/31/2025 |  | 12/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024390218 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Female |  |  | White | Malfunction | No | Yes | Failure to Advance | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during the procedure for an impella cp placement for a 65-year-old female patient, insertion was not possible. there was too much kinking due to patient anatomy and calcification. there was no reported patient harm. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the delivery issues and the product kink has been completed. the impella was not returned for evaluation. per the clinical information provided, the patient had extremely calcified vessels. the root cause of the delivery issue and product kink was patient condition related to diseased vessels. d3-corrected mfr name and address added-d6a/d6b.
32560 | 1220648-2023-05178 | 18244004 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18244004&pc=OZD | 11/03/2023 | 08/21/2024 |  | 12/01/2023 |  |  | 11/03/2023 | 12/01/2023 |  |  | 08/31/2025 |  | 12/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024386722 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Failure to Advance | Material Twisted/Bent | Migration |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/03/2023 |  |  |  |  | The user facility reported a 75-year-old female patient was implanted with an impella device for mechanical circulatory support. the us complainant had a cp placed for support of a (non-st segment elevation myocardial infarction ) nstemi patient. the pump came across the valve and out of position. the team observed the pump to be kinked at pump repositioning attempts. the pump was subsequently explanted and replaced. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the positioning and product damage has been completed. product was not returned for investigation. per the clinical information provided, the pump was tried implanting without a wire which is against the impella ifu. the root cause of the positioning issue was use related to improper technique. d4: corrected model number added-d6a/d6b. f6/f8 should have been left blank in the initial report.
32561 | 1220648-2023-05179 | 18244051 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18244051&pc=OZD | 11/02/2023 | 04/17/2024 |  | 12/01/2023 |  |  | 11/02/2023 | 12/01/2023 |  |  | 08/31/2025 |  | 12/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024395645 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/02/2023 |  |  |  |  | The user facility reported a 67-year-old male patient was implanted with an impella cp device for mechanical circulatory support. during support, the patient developed a hematoma and the pump was subsequently explanted. the patient was considered stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
32562 | 1220648-2023-05180 | 18244101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18244101&pc=OZD | 11/02/2023 | 02/15/2024 |  | 12/01/2023 |  |  | 11/02/2023 | 12/01/2023 |  |  | 04/30/2025 |  | 12/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024350404 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  |  | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/02/2023 |  |  |  |  | The user facility reported a 46-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. there was impella in aorta alarms. impella was not able to be seen on echo. given that the patient had not been on support since position was an issue overnight with stable mean arterial pressure( map) and index, decision was made to remove at bedside. removal was successful. patient is stable post intervention. | [5 empty] | The investigation into positioning issues has been completed. the device was not returned for investigation. according to clinical information, the impella was pulled out into the aorta. the cause of the positioning issue was use related, based on provided clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32563 | 1220648-2023-05181 | 18244125 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18244125&pc=OZD | 11/02/2023 | 04/05/2024 |  | 12/01/2023 |  |  | 11/02/2023 | 12/01/2023 |  |  | 03/31/2025 |  | 12/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024319114 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/02/2023 |  |  |  |  | The user facility reported a 44-year-old male patient was implanted with an impella 5.5 device for mechanical circulatory support. the impella had high purge pressure and the purge system blocked alarm was going off. the team started tissue plasminogen activator (tpa) to resolve the issue. no patient harm reported due to the issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the high purge pressure issue has been completed since the initial report was submitted. the device was not returned for investigation. the cause of the high purge pressure issue was most likely biomaterial, based on data log review.
32564 | 1220648-2023-05182 | 18244144 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18244144&pc=OZD | 11/02/2023 | 09/03/2024 |  | 12/01/2023 |  |  | 11/02/2023 | 12/01/2023 |  |  | 09/30/2025 |  | 12/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024399107 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/02/2023 |  |  |  |  | The user facility reported a 67-year-old male patient was implanted with an impella cp device for mechanical circulatory support. the us complainant had a cp support for a patient with cgs (cardiogenic shock) that was transferred to a tertiary medical center for continued monitoring. the patient was found to need care escalation to the axillary placed 5.5 pump. when the cp was explanted there was limb thrombosis observed that prompted the team to perform an embolectomy. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." d.4 revised model number from the initial submission of manufacturer device report number 1220648-2023-05182 in accordance with updated procedures. d.6a and d.6b revised from the initial submission of manufacturer device report number 1220648-2023-05182 in accordance with updated procedures. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2023-05182 and should not have been. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2023-05182 was submitted. h.6 code 925 was added since the initial manufacturer device report number 1220648-2023-05182 was submitted in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2023-05182 in accordance with updated procedures.
32565 | 1220648-2023-05183 | 18244154 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18244154&pc=OZD | 11/01/2023 | 01/12/2024 |  | 12/01/2023 |  |  | 11/01/2023 | 12/01/2023 |  |  | 07/31/2025 |  | 12/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024385840 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/01/2023 |  |  |  |  | The user facility reported a 70-year-old male in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. due to the calcification of the patient's arteries, the physician was unable to deliver the impella cp. there was no noted damage to the impella and no patient injury sustained. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical information provided was sufficient to determine the root cause of the issue. clinical details states that the patient iliac arteries were completely calcified and pump was unable to advance into aorta. the cause of the delivery issue was patient condition related due to pump unable to pass through vasculature with patient having calcified vessels.
32566 | 1220648-2023-05184 | 18244160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18244160&pc=OZD | 11/01/2023 | 08/26/2024 |  | 12/01/2023 |  |  | 11/01/2023 | 12/01/2023 |  |  | 02/28/2025 |  | 12/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024325818 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Migration | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/01/2023 |  |  |  |  | The user facility reported a 69-year-old male post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. while the patient was on support, they pulled their device out of their groin. manual pressure was applied to the site and an impella 5.5 pump was later implanted for ongoing support. there was no patient harm due to this event. | The patient did not survive the procedure. per medical safety officer review use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the root cause of the positioning issue was use related. f6/f8-should have been left blank on manufacturer device report 1220648-2023-05184.
32567 | 1220648-2023-05185 | 18244166 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18244166&pc=OZD | 11/01/2023 | 03/26/2024 |  | 12/01/2023 |  |  | 11/01/2023 | 12/01/2023 |  |  | 03/31/2025 |  | 12/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024346536 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following implant of an impella cp pump, the 79-year-old female patient developed a hematoma at the access site. manual pressure was initially utilized at the site followed by application of a sand bag to maintain said pressure. due to the condition, the utilization of heparin was halted and the site was monitored until the hematoma improved. support was maintained throughout the event. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the initial report was submitted. the product was not returned for investigation. per the clinical information provided, the root cause of the access site bleeding was most likely due to anticoagulation management. | The impella device was discarded and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32568 | 1220648-2023-05186 | 18244174 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18244174&pc=OZD | 11/01/2023 | 01/26/2024 |  | 12/01/2023 |  |  | 11/01/2023 | 12/01/2023 |  |  | 09/30/2025 |  | 12/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024398260 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/01/2023 |  |  |  |  | The user facility reported a 63-year-old male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the patient experienced some bleeding from their impella access site during support. additional sutures were placed at the site and a unit of blood was given to the patient to treat the condition. support continued with the implanted pump following the intervention. | [5 empty] | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for benchtop investigation. based on the clinical information provided, patient had some bleeding at impella surgical access site and additional sutures were placed at surgical site resolving bleeding. the cause of the access site bleeding issue was most likely use related per clinical review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32569 | 1220648-2023-05187 | 18245574 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18245574&pc=OZD | 11/01/2023 | 01/16/2024 |  | 12/01/2023 |  |  | 11/01/2023 | 12/01/2023 |  |  | 07/31/2025 |  | 12/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2024389243 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/01/2023 |  |  |  |  | The user facility reported a 78-year-old female patient was scheduled for impella cp implant for mechanical circulatory support. during the attempted delivery, the impella cp pump was said to have met resistance. it was noted that the device would not track over the aortic arch due to severe calcification and existing aortic stenosis. due to the noted issue, the decision was made to abort the implant. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported delivery issue -anatomy has been completed since the original report was filed. the root cause of the delivery issue was determined to be patient condition because the patient anatomy and impella was unable to pass over the aortic arch.
32570 | 1220648-2023-05188 | 18245665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18245665&pc=OZD | 11/01/2023 | 02/07/2024 |  | 12/01/2023 |  |  | 11/01/2023 | 12/01/2023 |  |  | 06/30/2025 |  | 12/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024378415 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/01/2023 |  |  |  |  | The user facility reported an unknown age female with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. hematoma noted on arrival to unit. there was suction on the impella, manual pressure was held, and fluids and blood products were administrated. physician was unable to stop bleeding/hematoma development to groin. patient was taken back to cardiac cath lab and impella was removed. positive distal pulses and peripheral injection shows no active bleeding and positive distal flow. patient is stable. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. due to the lack of clinical details and no product returned, the cause of the access site bleeding issue was not determined since product was not returned and due to insufficient clinical information. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32571 | 1220648-2023-05189 | 18245719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18245719&pc=OZD | 11/01/2023 | 12/01/2023 |  | 12/01/2023 |  |  | 11/01/2023 | 12/01/2023 |  |  | 05/31/2025 |  | 12/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024359256 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/01/2023 |  |  |  |  | The user facility reported a 70-year-old male patient was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient went into ventricular fibrillation and had to be defibrillated twice. as the physician believed the cause was related to the positioning of the pump, the decision was made to explant the pump. a hematoma later developed following explant, but no intervention was required and the condition was said to be resolving. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. a review of the device logs showed no positioning issues and no abnormalities in functionality were noted. upon conclusion of the investigation, a definitive relation between the event and the impella could not be established. the vf issue was subsequently attributed to the patient condition. should additional related information become available, a supplemental report will be filed.
32572 | 1220648-2023-05190 | 18245743 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18245743&pc=OZD | 11/01/2023 | 04/30/2024 |  | 12/01/2023 |  |  | 11/01/2023 | 12/01/2023 |  |  | 09/30/2025 |  | 12/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024400536 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/01/2023 |  |  |  |  | The user facility reported a 61-year-old female patient was implanted with an impella device for mechanical circulatory support. us complainant reported that there was impella pump saturation and the impella had to be replaced. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation has been completed since the original report was submitted. the device was not returned for investigation. data logs showed saturated pump flow while running on p8 without any reported motor current spike. the pump was replaced without issue. the cause of the saturated pump flow issue was not determined since the product was not returned and sufficient clinical information was not provided. a.5 and d.5 revised ethnicity and model number as previously entered incorrectly. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 codes 2199 and 4627 were reported incorrectly on the previously submission report. the correct code has been added to health effect - impact codes. | H.11 model number (d.4) was revised and not d.5 as previously reported on the last manufacture device report.
32573 | 1220648-2023-05191 | 18246195 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18246195&pc=OZD | 11/04/2023 | 01/12/2024 |  | 12/01/2023 |  |  | 11/04/2023 | 12/01/2023 |  |  | 09/30/2025 |  | 12/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024394370 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/04/2023 |  |  |  |  | The user facility reported a 59-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the patient turned in bed causing damage to the impella catheter where it meets the red plug. the pump stopped with an unsuccessful restart. the impella will be replaced. no patient harm reported due to the issue. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the reported pump stop has been completed since the original report was filed. upon inspection of the product, the catheter was pulled out from the kink protector, and the wires were damaged. there is evidence of excessive force as it was also kinked, and the clinical confirms excessive force. however, when inspecting the bond between the catheter and kink protector there is an insufficient amount of the adhesive ¿loctite 4062¿ and the bond is weak. in the logs "alarm 115" occurs and the motor current spikes to 30,000, signifying a break of the electrical lines. these lines likely broke due to the excessive force being applied to the weak bond. a device history record review was done and there were no reworks/qn's associated with the lot or any other product malfunctions. the adhesive that was used passed inspection. the root cause of the pump stop was due to open electrical lines.
32574 | 1220648-2023-05192 | 18246208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18246208&pc=OZD | 11/04/2023 | 01/12/2024 |  | 12/01/2023 |  |  | 11/04/2023 | 12/01/2023 |  |  | 05/31/2025 |  | 12/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024359286 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | Ischemia | Vascular Dissection | Thrombocytopenia | [8 empty] | Disability | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/04/2023 |  |  |  |  | The user facility reported a 63-year-old female was implanted with an impella device for mechanical circulatory support. the us complainant had placed a cp via the axillary insertion site for a patient awaiting surgery, covid+ and with other cardiac comorbidities. the pump support was for just 2 days and at explant the site developed a hematoma. additionally there was loss of sensation and possible nerve/brachial plexus injury. the harms prompt fda reporting. the patient condition is noted to be "not thriving". final patient outcome is not known. | A notification received via abiomed¿s long-term outcome and quality indicator impella (loqi) registry, indicating, "concern for hemolysis on 11/3, with plasma free hgb 199, ldh of ~1500." it was noted that the pump was explanted as a result. | A notification received via abiomed¿s long-term outcome and quality indicator impella (loqi) registry, indicating, thrombocytopenia, acute ischemic lue, and a large "rent" in axillary artery. thrombocytopenia was noted as "likely due to consumption and hemolysis and not due to heparin-induced thrombocytopenia." the reported acute ischemic lue was noted as "completely occluded by impella" and lead to removal of the pump. lastly, a "large rent" was noted in the axillary artery, which was "associated with impella insertion." |  |  |  | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Our investigation of this event is still in progress. upon closure, a supplemental mdr will be filed. | Our investigation of this event is still in progress. upon closure, a supplemental mdr will be filed.
32575 | 1220648-2023-05193 | 18246296 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18246296&pc=OZD | 11/04/2023 | 04/01/2024 |  | 12/01/2023 |  |  | 11/04/2023 | 12/01/2023 |  |  | 09/30/2025 |  | 12/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024397163 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/04/2023 |  |  |  |  | The user facility reported a 63-year-old male post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. the patient experienced some minor oozing from their access site during impella cp optical support. the patient was given fresh frozen plasma and packed red blood cells to counteract any blood loss. support continued with the implanted pump following the intervention. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
32576 | 1220648-2023-05194 | 18246301 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18246301&pc=OZD | 11/04/2023 | 01/26/2024 |  | 12/01/2023 |  |  | 11/04/2023 | 12/01/2023 |  |  | 08/31/2025 |  | 12/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024395904 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/04/2023 |  |  |  |  | The user facility reported a 78-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed an access site bleed which was treated by angle matching and two units of blood infused. the bleed had occurred with patient being rolled. the pump remained on for just under 2 weeks. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for benchtop investigation. based on clinical description, bleeding at impella access site observed when patient rolled over and blood transfusion, angle matching done to resolve bleeding. the cause of the access site bleeding issue was patient movement due to patient rolling over during support per clinical.
32577 | 1220648-2023-05195 | 18246334 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18246334&pc=OZD | 11/08/2023 | 09/03/2024 |  | 12/01/2023 |  |  | 11/07/2023 | 12/01/2023 |  |  | 08/31/2025 |  | 12/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2024391653 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/07/2023 |  |  |  |  | The user facility in japan reported a male (age unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. a thrombus was observed in the patient and in the impella when the impella was removed. the patient was stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6a and d6b added h6 component code added the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05195: f6/f8-should have been left blank h.6 code 4627 was reported incorrectly. new code was added to health effect - impact code.
32578 | 1220648-2023-05196 | 18249029 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18249029&pc=OZD | 11/02/2023 | 01/17/2024 |  | 12/01/2023 |  |  | 11/02/2023 | 12/01/2023 |  |  | 09/30/2025 |  | 12/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024400639 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/02/2023 |  |  |  |  | The user facility reported a 65-year-old male in acute myocardial infarction/cardiogenic shock was scheduled for impella cp implant for mechanical circulatory support. despite multiple attempts at implanting the impella cp pump, the device ultimately was unable to traverse the patient's tortuous aorta and aortic valve. as a result, the implant was aborted. there was no patient harm. | [5 empty] | The impella device was discarded from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the root cause of the delivery issue was most likely patient condition related.
32579 | 1220648-2023-05197 | 18249927 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18249927&pc=OZD | 11/12/2023 | 09/22/2024 |  | 12/01/2023 |  |  | 11/12/2023 | 12/01/2023 |  |  | 07/31/2025 |  | 12/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024381087 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/12/2023 |  |  |  |  | The user facility reported a 82-year-old male post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. on day 3 of the support the purge pressure was high and purge system blocked alarms were triggered. as such the team adjusted the purge and used the med affairs information on tissue plasminogen activator protocol. several hours later, purge flow and pressure returned to earlier levels. no patient harm reported due to the issue. | Family and patient decided to withdraw care. the procedure outcome was that patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The investigation into the reported high purge pressure (hpp) has been completed since the original report was filed. the purge pressure began to steadily rise over the course of a day before the hpp alarm was triggered and they were in suction conditions. during this time period, there was no additive in the purge fluid. the solution was changed to d5w with 50unit/ml of heparin and the purge pressure was slightly reduced. tissue plasminogen activator (tpa) was then added into the purge solution and the pressure began to relieve and purge flow increased. the root cause of the high purge pressure is most likely due to a biomaterial occlusion. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-05197 was provided in sections b1, b5, d6, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
32580 | 1220648-2023-05198 | 18251020 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18251020&pc=OZD | 11/26/2023 | 02/22/2024 |  | 12/01/2023 |  |  | 11/26/2023 | 12/01/2023 |  |  | 04/30/2025 |  | 12/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024347700 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/26/2023 |  |  |  |  | The user facility reported a male patient hwo had ventricular fibrillation and arrest in the field was implanted with an impella device for mechanical circulatory support. additionally, the patient was supported by va-ecmo. the complainant reported that the patient had a cold leg with poor distal perfusion. in order to improve flow left femoral external bypass with two 6fr sheaths procedure was done. | [5 empty] | The investigation into the reported limb ischemia ¿ reversible been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32581 | 1220648-2023-05199 | 18252108 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18252108&pc=OZD | 11/07/2023 | 01/24/2024 |  | 12/01/2023 |  |  | 11/07/2023 | 12/01/2023 |  |  | 08/31/2025 |  | 12/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024387842 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Deformation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/07/2023 |  |  |  |  | The user facility reported a 37-year-old male with cardiomyopathy was to be implanted with an impella 5.5 for mechanical circulatory support. the team advanced the pump and a kink in the pump was observed. the pump was explanted and replaced by a new 5.5 pump with no further issues. | [5 empty] | The investigation into the reported delivery issue has been completed since the original report was filed. the root cause of the delivery issue was most likely due to use issue as the second 5.5 was able to pass through without issue. the failure modes will be monitored and trended. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32582 | 1220648-2023-05200 | 18252135 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18252135&pc=OZD | 11/02/2023 | 04/17/2024 |  | 12/01/2023 |  |  | 11/02/2023 | 12/01/2023 |  |  | 10/31/2024 |  | 12/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024292268 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/02/2023 |  |  |  |  | The user facility reported a 82-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was oozing from the right groin. hematoma was expanding needing manual compression. the hematoma was worsening, and the patient was taken back to the cath lab and the impella was removed. patient is stable post explant. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32583 | 1220648-2023-05201 | 18253012 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18253012&pc=OZD | 05/03/2023 | 04/17/2024 |  | 12/01/2023 |  |  | 11/02/2023 | 12/01/2023 |  |  | 10/31/2024 |  | 12/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024287675 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/02/2023 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry (loqi). according to the report, a 76 year old male patient (presenting with acute myocardial infarction and cardiogenic shock), had the impella device removed and underwent a fogarty embolectomy within the shared femoral artery and popliteal artery. a cfa endarterectomy with bovine patch angioplasty was performed and placement of a negative pressure wound vac. this allegation was made with a possible relationship to the impella device. | [5 empty] | The investigation for the reported thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombus could not be determined. the instructions for use states ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ d4-updated model number. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32584 | 1035166-2023-00129 | 18237297 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18237297&pc=DYB | 10/03/2023 | 09/04/2024 |  | 11/30/2023 |  |  | 11/02/2023 | 11/30/2023 |  |  | 12/01/2025 |  | 11/30/2023 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3021 |  |  | OSCOR INC. |  |  | US |  | DP-17331 |  | 0052-3021 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | 64 YR |  | Female |  |  |  | Injury | Yes | Yes | Break | Fluid/Blood Leak | [5 empty] | Appropriate Term / Code Not Available | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | We have been notified of a complaint involving an oscor 23 fr introducer (material # 0052-3021, lot # dp17331). on (b)(6) 2023, a 64-year-old female patient was presented to physician with right heart failure. it was reported an impella rp device was placed via right femoral vein. technique used for insertion was followed according to impella rp instructions for use (ifu). when removing the introducer, the introducer sheath cracked, resulting in oozing from sheath during procedure. before leaving cath lab, the patient required a blood transfusion, (b)(4) unit of packed red blood cell (prbc). the patient was brought to surgical intensive care unit (sicu). care was withdrawn the next day (b)(6) 2023, and patient expired. per end user, cause of death, care was noted to be withdrawn from patient. the device was discarded. | [5 empty] | Conclusion not yet available, evaluation in process. a follow-up will be submitted as soon as the investigation is complete. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | The following sections were updated in follow-up 1: b4, g3, g6, h2, h6, and h11. no product was provided for analysis. there was blood loss reported and the patient required a blood transfusion, 1 unit of packed red blood cell (prbc). the patient was brought to surgical intensive care unit (sicu). care was withdrawn the next day. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. no product was returned to oscor inc. for evaluation; however, a complaint notification was provided. without the return of the actual device in question for evaluation, oscor inc. is unable to determine the exact cause for this incident. at this time, it is not possible to assign a definitive root cause for the event as reported. the complaint is non-verifiable as the product was not returned for evaluation. no corrections or corrective actions will be taken. device was not returned for analysis and review of the manufacturing documents did not reveal any discrepancies that would have contributed to this event. this complaint is non-verifiable. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
32585 | 1220648-2023-05150 | 18238836 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18238836&pc=OZD | 11/06/2023 | 08/20/2024 |  | 11/30/2023 |  |  | 11/06/2023 | 11/30/2023 |  |  | 09/30/2025 |  | 11/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024399830 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/06/2023 |  |  |  |  | The user facility reported a 76-year-old female in acute myocardial infarction/cardiogenic shock was to be implanted with an impella 5.5 device for mechanical circulatory support. the physician used a lot of force upon inserting the impella 5.5 via axillary which is off-label usage. as a result, the catheter kinked above the motor housing. a new impella was attempted with no success. the team placed the patient on an intra-aortic balloon pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the product damage has been completed. according to the clinical details, the physician attempted to place the pump perc ax. excessive force was applied, and they were unable to pass the pump. as a result of the excessive force the catheter kinked above the motor housing. the decision was then made to use a second pump. no product was returned for an investigation. data log analysis was not necessary for this complaint. the most likely cause of the delivery issue was use related. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock states the following: section: warnings & cautions: cautions. ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ added- h6 impact code 4627 added-d6a/d6b. f6/f8 should have been left blank in the initial report.
32586 | 1220648-2023-05151 | 18242162 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18242162&pc=OZD | 11/09/2023 | 04/09/2024 |  | 11/30/2023 |  |  |  | 11/30/2023 |  |  | 06/30/2025 |  | 11/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED INC |  |  | US |  | 2024379657 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/09/2023 |  |  |  |  | The user facility reported a 71-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the cp pump inserted and had immediate low flow and suction alarms. the malfunction prompted the team to check the sheath/aspirate for clot and remove the pump. the pump was removed and clot was then observed at the outlet of the pump. there was no lasting harm or injury. the pump was not replaced by another cp but rather an iabp. | [5 empty] | The investigation into the low pump flow issue has been completed since the original report was filed. the pump was returned, tested, and evaluated. the returned complaint pump was visually inspected. no broken blades found. cannula kink near outlet observed. biomaterial ingested in the impeller purge gap after dried dextrose dissolved. clinical data confirmed clot visualized near outlet. the cannula kink may occur while the pump explanted. the cause of the low pump flow was biomaterial ingestion. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
32587 | 1220648-2023-05152 | 18243211 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18243211&pc=OZD | 11/01/2023 | 02/28/2024 |  | 11/30/2023 |  |  |  | 11/30/2023 |  |  | 09/30/2025 |  | 11/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024398680 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/01/2023 |  |  |  |  | The user facility reported a xx-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after placing impella cp, no flow was observed past impella sheath and it is believed that impella repo sheath was occlusive. before leaving cath lab an external fem-fem was placed to provide flow to distal limb. a doppler pulse was successfully obtained thereafter. patient was on device support for 1 more day post issue. | [5 empty] | The investigation into the limb ischemia ¿ irreversible has been completed. the device and data were not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32588 | 1220648-2023-05153 | 18243212 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18243212&pc=OZD | 11/01/2023 | 04/25/2024 |  | 11/30/2023 |  |  |  | 11/30/2023 |  |  | 07/31/2025 |  | 11/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024388869 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 68-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported the staff were unable to increase the impella past p3 due to continuous suction alarms. echo was showing correct placement. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation for the low pump flow has been completed. the pump was not returned for investigation. the data log shows several suction alarms during the case. the doctor took a stepwise decrease in the p-level after seeing the suction alarms. there were no motor current spikes or placement signal trends noted during the suction event. the cause of the low pump flow issue was not determined since product was not returned and sufficient clinical information was not provided.
32589 | 1220648-2023-05155 | 18243224 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18243224&pc=OZD | 11/01/2023 | 01/26/2024 |  | 11/30/2023 |  |  |  | 11/30/2023 |  |  | 09/30/2025 |  | 11/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024398918 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/01/2023 |  |  |  |  | The user facility reported a 75-year-old male post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. after five days of support, the team explanted the pump due to limb ischemia which was diagnosed by lost pulses. the team had used fem-fem bypass to support and perfuse the limb. the patient was stable after the explant. | [5 empty] | The investigation into the limb ischemia ¿ irreversible has been completed. the device was not returned for investigation. per the provided clinical information, the patient lost pulses over impella leg and external fem to fem bypass resolved ischemia on left femoral leg. the cause of the limb ischemia issue was patient condition related with fem-fem bypass resolving ischemia on impella leg per clinical. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32590 | 1220648-2023-05158 | 18243225 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18243225&pc=OZD | 11/02/2023 | 03/19/2024 |  | 11/30/2023 |  |  |  | 11/30/2023 |  |  | 05/31/2025 |  | 11/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024358925 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A patient of unknown age and gender should receive hemodynamic support with an impella cp smartassist heart pump. the automated impella controller (aic) displayed an ¿impella defective¿ alarm that could not be resolved by replacing the aic. a new impella cp smartassist heart pump was successfully used. the patient suffered no harm from the event. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation for the reported pump stop has been completed. the impella device was returned for evaluation. visual and functional evaluations found no abnormalities. data logs were reviewed which revealed impella defective alarm was observed after pump plug connect. therefore, the root cause of the pump did not start was unable to be determined, as the issue was unable to be reproduced. d4-updated model number
32591 | 1220648-2023-05159 | 18243226 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18243226&pc=OZD | 11/02/2023 | 10/10/2024 |  | 11/30/2023 |  |  |  | 11/30/2023 |  |  |  |  | 11/30/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | XXXXXX |  | 0042-0000 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/02/2023 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the (automated impella controller) aic had a failure while on patient use. it was further reported that it was a catheter issue, but the hospital would like to verify if the aic is functioning as expected. the aic will be returned for evaluation. | [5 empty] | The investigation findings for manufacturer device report (mdr) (b)(4)will be found in mdr (b)(4). | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32592 | 1220648-2023-05154 | 18243227 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18243227&pc=OZD | 11/01/2023 | 01/23/2024 |  | 11/30/2023 |  |  |  | 11/30/2023 |  |  | 03/31/2024 |  | 11/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2023214480 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/01/2023 |  |  |  |  | The user facility reported a 61-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp support for a patient admitted with angina and found to be in (acute myocardial infarction/ cardiogenic shock) ami/cgs. the pump was explanted after just over 10 hours due to entrained clot in the lv (left ventricle). the thrombus did not cause any harm or injury. no noted embolization. pump was explanted and not replaced. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the thrombus has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. data logs were reviewed, and no motor current spikes or any other abnormalities related to biomaterial injection were observed. the cause of the thrombus issue was most likely patient condition related.
32593 | 1220648-2023-05161 | 18243229 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18243229&pc=OZD | 11/02/2023 | 11/18/2025 |  | 11/30/2023 |  |  |  | 11/30/2023 |  |  |  |  | 11/30/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1299521 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller (aic) had an air purge system alarm. the aic was removed and replaced. the aic has been sent back for investigation. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
32594 | 1220648-2023-05162 | 18243246 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18243246&pc=OZD | 11/02/2023 | 02/16/2024 |  | 11/30/2023 |  |  |  | 11/30/2023 |  |  | 04/30/2025 |  | 11/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024365539 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/02/2023 |  |  |  |  | The user facility reported a 77-year-old female/male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that upon removal of sheath and insertion of repo sheath; patient noted to have hematoma to right groin. scrub tech continued to suture sheath in place, at which point the doctor noted that hematoma was growing and decision made to slowly wean and remove impella in order to remove repositioning sheath. the impella was successfully weaned and removed. the doctor attempted to place angioseal which failed. upon attempting to tie down perclose sutures, the doctor noted that there was a tear in the femoral artery. manual pressure held. hemoglobin dropped to 8.4,(b)(6) units of packed red blood cells were transfused. surgical cutdown and repair performed. patient is stable post intervention. | [5 empty] | The investigation into the access site bleeding ¿ major and the vascular damage - peripheral vessel issues has been completed. the device was not returned for investigation. as per the clinical details, during a procedure involving the placement of a repositioning sheath, the patient, with a bmi of 32, experienced a femoral artery dissection. the tear was likely exacerbated by the patient's body habitus. the cause of the vascular damage issue was most likely patient condition related, based on given clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
32595 | 1220648-2023-05163 | 18243248 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18243248&pc=OZD | 11/02/2023 | 09/10/2024 |  | 11/30/2023 |  |  |  | 11/30/2023 |  |  | 09/30/2025 |  | 11/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024394998 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the 47 year-old male patient returned from the operating room for hematoma washout at the impella 5.5 access site. patient's tingling and numbness sensation resolved post washout per patient/rn. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (unites states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿ added- h6 component code 925 added- d6a/d6b.
32596 | 1220648-2023-05164 | 18243249 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18243249&pc=OZD | 11/03/2023 | 01/15/2024 |  | 11/30/2023 |  |  |  | 11/30/2023 |  |  | 07/31/2025 |  | 11/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024380529 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/03/2023 |  |  |  |  | The user facility reported a 44-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported there were on going placement signal not reliable issues. additionally, there was a pump stop that occurred. tolerated off impella relatively well in lieu of milrinone infusion, however, upon arriving to the or their pa pressures shot up to 100 and was given lasix and escalated vasopressor/inotropic support. there was a subsequent impella exchange to a different impella. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation for the pump stop has been completed. the pump was not returned for evaluation. however, data logs were reviewed which showed towards the end of the case, a motor current spike of 30,000 and alarm 115 & alarm 119 occurred. this is indicative of all the motor lines being open. the root cause of the pump stop is most likely due to open electrical lines.
32597 | 1220648-2023-05167 | 18243257 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18243257&pc=OZD | 11/03/2023 | 02/29/2024 |  | 11/30/2023 |  |  |  | 11/30/2023 |  |  | 08/31/2025 |  | 11/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024395055 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/03/2023 |  |  |  |  | The user facility reported a 64-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support after coding in the cath lab. the patient had a large hematoma that would require surgical closure and device removal. two blood products were required. the impella 5.5 was placed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported access site bleeding -major has been completed since the original report was filed. pump was not returned for investigation. as per clinical details, the cause of the access site bleeding issue was not determined since product was not returned and sufficient clinical information was not provided.
32598 | 1220648-2023-05165 | 18243260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18243260&pc=OZD | 11/03/2023 | 08/20/2024 |  | 11/30/2023 |  |  |  | 11/30/2023 |  |  | 05/31/2025 |  | 11/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024352115 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/03/2023 |  |  |  |  | The user facility reported a 68-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had placed a cp for support of patient with known cad (coronary artery disease) and cabgx3 (coronary artery bypass grafting). the pump was placed and when transferring patient to the air transport the pump came out of position. the pump was not repositioned across the valve and the team was explanted and replaced. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. per the clinical information provided, while transferring the patient from aircraft litter to the stretcher, the pump got displaced into the aorta. therefore, the root cause of the positioning issue was patient movement. f6/f8 should have been left blank in the initial report.
32599 | 1220648-2023-05156 | 18243446 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18243446&pc=OZD | 11/01/2023 | 06/11/2024 |  | 11/30/2023 |  |  | 11/01/2023 | 11/30/2023 |  |  | 08/31/2025 |  | 11/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024396623 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 53-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient had hemolysis and volume was given while on impella support. additionally, the patient had oozing at the insertion site, manual pressure was held and the staff confirmed the blue hub at skin level and angle matched properly. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the reported hemolysis has been completed. no product was returned for analysis. the root cause of the access site bleeding - minor cannot be determined due to insufficient clinical information provided. data log analysis was conducted. there were instances of suction but no motor current spikes or a trend in the placement signal. it cannot be determined if the hemolysis resolved after the rn gave volume. the root cause of the hemolysis cannot be determined. the instructions for use referenced in the initial report is for the impella cp with smartassist system in sections: potential adverse events (united states), which now includes the statement "cardiac or vascular injury (including ventricular perforation).¿, use of echocardiography for positioning of the impella catheter, hemolysis, and suction. impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions including catheter position, pre-existing medical conditions, and small left ventricular volumes may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
32600 | 1220648-2023-05160 | 18243447 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18243447&pc=OZD | 11/02/2023 | 03/19/2024 |  | 11/30/2023 |  |  | 11/02/2023 | 11/30/2023 |  |  | 08/31/2025 |  | 11/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024392976 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/02/2023 |  |  |  |  | The user facility reported a 58-year-old male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support, the impella pump stopped. troubleshooting to restart the device was unsuccessful. the pump was removed and replaced with a new impella. | [5 empty] | The investigation for the pump stop has been completed. the impella was returned for evaluation. upon evaluation of the pump, a large impeller gap was identified. data logs confirm the high motor current pump stop with motor current spikes. the root cause of the pump stop was determined to be bearing wear as seen by an enlarged impeller purge gap. added- d9 device return date, h6 impact code 4629. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
32601 | 1220648-2023-05157 | 18243546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18243546&pc=OZD | 11/01/2023 | 03/22/2024 |  | 11/30/2023 |  |  | 11/01/2023 | 11/30/2023 |  |  | 04/30/2025 |  | 11/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024349712 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 23 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increased Pump Speed | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/01/2023 |  |  |  |  | The user facility reported a 23-year-old female with cardiomyopathy was implanted with an impella cp for mechanical circulatory support while awaiting left ventricle assist device (lvad) implant. following more than 21days of support, axillary impella cp began to show signs of impending pump failure, including elevated motor current (mc) (average mc ~600 mamp @ p-5). impella connect was used to show md and nps prior day's mc at same p-level which was ~100 mamp lower. patient was planned for lvad procedure first-case next morning. md strategized with care team for actions should pump fail. pump explanted successfully the following day after successful lvad placement. | [5 empty] | The investigation into the pump stop issue has been completed since the initial report was submitted. the device was not returned for investigation. according to the clinical details, the doctor decided to explant the pump after observing elevated motor current. the cause of the saturated pump flow issue was most likely biomaterial, based on data log review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32602 | 1220648-2023-05166 | 18243550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18243550&pc=OZD | 11/03/2023 | 02/28/2024 |  | 11/30/2023 |  |  | 11/03/2023 | 11/30/2023 |  |  | 09/30/2025 |  | 11/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024397317 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Other |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/03/2023 |  |  |  |  | The user facility reported a 64-year-old female post cardiac surgery was implanted with an impella cp device for mechanical circulatory support after the patient coded and received three rounds of cpr. the impella cp placed via the femoral artery insertion site. the patient had pulse loss issues in the impella limb. the team made choice to explant the impella due to the limb ischemia. the team observed thrombus burden when removed. | [5 empty] | The investigation into the limb ischemia ¿ reversible has been completed. the device and data were not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32603 | 1220648-2023-05168 | 18243841 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18243841&pc=OZD | 11/03/2023 | 01/31/2024 |  | 11/30/2023 |  |  | 11/03/2023 | 11/30/2023 |  |  | 08/31/2025 |  | 11/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024385749 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/03/2023 |  |  |  |  | The user facility reported a 68-year-old female patient was implanted with an impella device for mechanical circulatory support. the complainant reported that an impella cp pump stopped unexpectedly during patient support. attempts to restart the pump were unsuccessful and swapping to a new aic (automated impella controller) did not resolve the noted issue. the decision was subsequently made to replace the pump with a new impella cp device. upon explant, a thrombus was noted in the cannula and outlet of the pump. there was no lasting patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop has been completed. the pump and data logs were returned for evaluation. upon investigation of the pump, a large impeller purge gap was identified. a clot was observed in the enlarged purge gap. the data logs confirmed the high motor current pump stop as well as alarms with motor current spikes to 1511ma. the pump stop occurred beyond the eight day design life. the root cause of the pump stop was determined to be bearing wear as seen by an enlarged impeller purge gap.
32604 | 1220648-2023-05169 | 18243848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18243848&pc=OZD | 11/03/2023 | 02/02/2024 |  | 11/30/2023 |  |  | 11/03/2023 | 11/30/2023 |  |  | 06/30/2025 |  | 11/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024373018 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/03/2023 |  |  |  |  | The user facility reported a 43-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed ischemia in their impella leg during support. it was noted that the patient began to experience severe pain in their right leg during a scheduled percutaneous coronary intervention. an angiogram of the right iliac was performed and showed an occlusion with no flow distal to the sheath. as a result of the noted condition, the decision was made to explant the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible issue has been completed. the product was not returned for review. as per clinical, the patient developed severe pain in the right leg and angiogram of the right iliac revealed an occlusion with no flow distal to the sheath. later the pump was explanted. the cause of the limb ischemia issue was patient condition related with the right iliac occluded and the device was explanted to resolve the ischemia.
32605 | 1220648-2023-05170 | 18243865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18243865&pc=OZD | 11/03/2023 | 03/14/2024 |  | 11/30/2023 |  |  | 11/03/2023 | 11/30/2023 |  |  | 04/30/2025 |  | 11/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024350456 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/03/2023 |  |  |  |  | The user facility reported a 63-year-old female (race unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support during emergent percutaneous cardiac intervention. the implanted impella cp pump was unable to obtain adequate flows during patient support. as a result, the pump was explanted and the patient was switched to a tandem heart device with oxygenation for ongoing support. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the low pump flow and access site bleeding has been completed. the impella device was not returned for evaluation. data logs were reviewed which showed low pump flow with several suction alarms during the case run. there were no reported events related to biomaterial interaction or improper positioning. the cause of the issue was not determined since product was not returned and sufficient clinical information was not provided.
32606 | 1220648-2023-05171 | 18243897 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18243897&pc=OZD | 11/03/2023 | 11/30/2023 |  | 11/30/2023 |  |  | 11/03/2023 | 11/30/2023 |  |  | 08/31/2025 |  | 11/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024395935 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the 66-year-old male patient with a history of coronary artery disease was implanted with an impella cp pump for mechanical circulatory support. following implant, the pump was unable to ramp up and was subsequently explanted. upon examination, it was noted that a clot was on the outflow track of the device. there was no patient harm. | [5 empty] | The investigation into the reported low pump flow issue has been completed. the impella cp pump was returned for analysis. during inspection, heavy biomaterial clot was found in the outflow track. the root cause of the low pump flow issue was determined to be the biomaterial found in the outflow track. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32607 | 1220648-2023-05172 | 18243903 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18243903&pc=OZD | 11/03/2023 | 03/13/2024 |  | 11/30/2023 |  |  | 11/03/2023 | 11/30/2023 |  |  | 06/30/2025 |  | 11/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024373446 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported 66-year-old male patient on impella cp experienced lower than expected flow, with suction alarms, upon attempted ramp-up of the device. the pump was therefore removed and replaced. there was no reported patient harm. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation for the reported low pump flow has been completed. the impella device was returned for evaluation and visually inspected. biomaterial was observed around impeller and purge gap. the cause of the low pump flow issue was most likely biomaterial.
32608 | 1220648-2023-05173 | 18243913 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18243913&pc=OZD | 11/03/2023 | 09/11/2024 |  | 11/30/2023 |  |  | 11/03/2023 | 11/30/2023 |  |  | 07/31/2025 |  | 11/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024385752 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 81-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient experienced persistent oozing from their impella access site following implant. forward tension sutures were placed with angle matching gauze, but the condition remained. despite multiple troubleshooting attempts, the decision was made to explant the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely patient condition related, the impella was removed as the patient had a classified vessel condition as per the clinical description. added- h6 component code 925 added- d6a/d6b.
32609 | 1220648-2023-05174 | 18243921 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18243921&pc=OZD | 11/03/2023 | 03/18/2024 |  | 11/30/2023 |  |  | 11/03/2023 | 11/30/2023 |  |  | 01/31/2025 |  | 11/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024297941 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/03/2023 |  |  |  |  | The user facility reported a 79-year-old male in acute myocardial infarction/cardiogenic shock was to be implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 was unable to pass through the subclavian artery due to the calcification of the patient¿s vessel. the catheter was kinked and, after several unsuccessful attempts, the case was aborted. | [5 empty] | The investigation into the product damage (cannula kink) issue has been completed since the original report was filed. the device was returned and tested. no issues were found on the pump. the root cause of delivery issue was due to patient condition. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32610 | 1220648-2023-05126 | 18229793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18229793&pc=PYX | 11/07/2023 | 09/26/2024 |  | 11/29/2023 |  |  | 11/07/2023 | 11/29/2023 |  |  | 03/31/2024 |  | 11/29/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023217102 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 94 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Advance | Patient Device Interaction Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/07/2023 |  |  |  |  | The user facility reported a 94-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. the impella rp was unable to advance past the tricuspid valve. physician also noted that the rp was interfering with the temporary transvenous pacing (tvp) and could not be advanced. the impella rp was removed. | The patient did not survive the procedure. per medical safety officer review use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of delivery issues- anatomy has been completed. the impella device was not returned for evaluation. the cause of the delivery issue was use related due to interaction with tvp catheter that was unable to advance rp past tv after multiple attempts. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05126. b5 revised from the initial submission in accordance with updated procedures, including patient outcome and the mso statement. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact fax number missing. h6 code 2524 was reported incorrectly. new code was added to medical device problem code instructions for use: rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ rp flex with smart assist system with automated impella controller section: warnings and cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings and cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
32611 | 1220648-2023-05127 | 18229868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18229868&pc=OZD | 10/31/2023 | 03/19/2024 |  | 11/29/2023 |  |  | 10/31/2023 | 11/29/2023 |  |  | 06/30/2025 |  | 11/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024376127 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 62-year-old female noted as high risk coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the impella cp had low flows despite correct placement. the error message displayed as suction. the flow did not improve leading to pump explant. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation for the reported low pump flow has been completed. the impella device was returned for evaluation. the pump was visually inspected and biomaterial was observed in purge gap. no cannula kink was observed. no broken blades were noted. cp pump was connected to the aic and run for 5 minutes in p-9 in the investigation lab, average flow was 3.8 liters/minute which is above specification limit indicating saturated flow. data logs not available for review. the root cause of the low flow issue was most likely biomaterial ingestion.
32612 | 1220648-2023-05128 | 18229920 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18229920&pc=OZD | 11/03/2023 | 03/04/2024 |  | 11/29/2023 |  |  | 11/03/2023 | 11/29/2023 |  |  | 08/31/2025 |  | 11/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024394326 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/03/2023 |  |  |  |  | The user facility reported a 79-year-old male in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support after the physician aborted the case for inserting an impella 5.5. the impella cp was attempted to be inserted, but the case had to be aborted as well due to the inability to push across the patient's aortic valve. the physician was unable to make the turn out of left subclavian artery due to patient's anatomy. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported delivery issue- anatomy has been completed since the original report was filed. returned complaint cp pump visually inspected. no kink observed in catheter and cannula. no other abnormality observed. the cause of the delivery issue most likely was patient condition related due to patient anatomy.
32613 | 1220648-2023-05129 | 18232393 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18232393&pc=OZD | 11/01/2023 | 04/25/2024 |  | 11/29/2023 |  |  | 11/01/2023 | 11/29/2023 |  |  | 08/31/2025 |  | 11/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024385019 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an implanted impella cp pump experienced persistent suction events during patient support. despite position and volume management, the low flow persisted and the decision was made to replace the pump with another impella cp device. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation into the low pump flow has been completed since the original report was submitted. the device was not returned for investigation. data logs show low pump flow during the case run, with a noted trend in placement signal. there were no motor current spikes reported during the case. the cause of the low pump flow issue was most likely patient condition related, based on data log review.
32614 | 1220648-2023-05130 | 18232665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18232665&pc=OZD | 10/31/2023 | 04/09/2024 |  | 11/29/2023 |  |  | 10/31/2023 | 11/29/2023 |  |  | 05/31/2025 |  | 11/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2024350103 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient of unknown age and gender was implanted with an impella cp for mechanical circulatory support. a high pp was reported even though there were no kinks noted on the purge line. all mechanical causes were ruled out and lysis therapy was started. a few hours later, the purge arm started leaking and the pump was therefore explanted. there was no patient harm. | [5 empty] | The investigation into the high purge pressure and the purge leak ¿ pump issues has been completed since the initial report was submitted. the device was not returned for investigation. according to clinical records, the doctor reported high purge pressure without observing any kinks in the purge line. subsequently, the pump was explanted after a leak was observed from the yellow lure. the cause of the high purge pressure and the purge leak issues was not determined since product was not returned and limited clinical information was provided. | The impella device was not returned by the customer and therefore, evaluation of the device was not possible. upon conclusion of the investigation, a supplemental report will be submitted. per the ifu: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir. ¿clean the connector cable with 70% isopropyl alcohol.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
32615 | 1220648-2023-05131 | 18234659 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18234659&pc=OZD | 11/08/2023 | 09/26/2024 |  | 11/29/2023 |  |  | 11/08/2023 | 11/29/2023 |  |  | 07/31/2025 |  | 11/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | 08-NOV-2023 |  |  | ABIOMED, INC. |  |  | US |  | 2024376926 |  | 09-NOV-2023 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  |  | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/08/2023 |  |  |  |  | The user facility reported a patient admitted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the impella was trapped under papillary muscle. an attempt was made to reposition. during repositioning the impella was pulled into the aorta. the decision was made to pull the impella into descending aorta and place intra-aortic balloon pump in cardiac cath lab. the patient was stable. | [5 empty] | The investigation for the positioning issue has been completed. pump was not returned for investigation. data logs were returned and "impella in aorta" alarms were observed. clinical confirmed that the pump was stuck in papillary muscle and later pulled into aorta while repositioning. therefore, the root cause of the positioning issue was use related to improper technique. added- h6 impact code 4642 corrections to the initial mfr report: added-d6a/d6b. f6/f8 should have been left blank. h5 labeled for single use changed to yes. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32616 | 1220648-2023-05132 | 18234673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18234673&pc=OZD | 09/02/2023 | 11/06/2024 |  | 11/29/2023 |  |  | 09/02/2023 | 11/29/2023 |  |  | 07/31/2025 |  | 11/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024371649 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/02/2023 |  |  |  |  | The user facility reported a 73-year-old female in acute myocardial infarction/cardiogenic shock was to be implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 implant was aborted due to patient anatomy and possible stenosis. additionally, bleeding occurred during insertion. blood products were required. surgeon patched the area and confirmed it was not actively bleeding. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation of delivery issues and access site bleeding has been completed. the impella device was not returned for evaluation. the cause of the delivery issue was most likely patient condition related, the impella 5.5 implant was aborted due to patient anatomy and possible stenosis. the root cause of the access site bleeding issue was not determined since product was not returned and sufficient information is not provided in the clinical description. instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella 5.5 with smartassist for use during cardiogenic shock section: warnings and cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella 5.5 with smartassist section: warnings and cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella 5.5 with smartassist for use during cardiogenic shock with automated impella controller section: warnings and cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use during impella with smart assist system with automated impella controller section: warnings and cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings and cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05132 d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. h10 instructions for use missing.
32617 | 1220648-2023-05133 | 18234696 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18234696&pc=OZD | 11/05/2023 | 02/02/2024 |  | 11/29/2023 |  |  | 11/05/2023 | 11/29/2023 |  |  | 07/31/2025 |  | 11/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024386908 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/05/2023 |  |  |  |  | The user facility reported a 48-year-old male (race unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient lost pulsatility in the lower extremity. the patient was stable and the impella was removed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported, limb ischemia has been completed since the original report was filed. product was not returned for review. as per clinical the cause of the limb ischemia issue was most likely patient condition related with ischemia observed on impella leg .
32618 | 1220648-2023-05134 | 18234702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18234702&pc=OZD | 08/11/2023 | 11/29/2023 |  | 11/29/2023 |  |  | 08/11/2023 | 11/29/2023 |  |  | 05/31/2025 |  | 11/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024344645 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/11/2023 |  |  |  |  | The user facility reported a 60-year-old male in acute myocardial infarction/cardiogenic shock was to be implanted with an impella 5.5 device for mechanical circulatory support. the device was unable to fully advance out of introducer during insertion through tunneled bifurcating graft. new introducer and new impella 5.5 pump then inserted without issue. | [5 empty] | The investigation into "device unable to fully advance out of introducer during insertion" has been completed. product was not returned for review. the root cause of the difficulty inserting the pump through introducer was unable to be determined as limited clinical details were provided and no product was returned for review.
32619 | 1220648-2023-05135 | 18234812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18234812&pc=OZD | 10/21/2023 | 02/21/2024 |  | 11/29/2023 |  |  | 10/21/2023 | 11/29/2023 |  |  | 08/31/2025 |  | 11/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024397985 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a 55-year-old female patient on impella cp support experienced loss of pulses in the right lower extremity and as a consequence a distal perfusion catheter had to be placed. flow was able to be restored in the right foot. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the patient's limb ischemia has been completed. the product was not returned for evaluation. the clinical evaluation follows: any device contacting the heart wall in conjunction with a poor patient condition could result in ventricular fibrillation/ventricular tachycardia (vf/vt) but since fluoroscopic imaging and/or sufficient clinical details were not provided, a relation between impella and the vf/vt is unable to be determined.
32620 | 1220648-2023-05136 | 18234819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18234819&pc=OZD | 10/30/2023 | 01/26/2024 |  | 11/29/2023 |  |  | 10/30/2023 | 11/29/2023 |  |  | 06/30/2025 |  | 11/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2024374106 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Break | Crack | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/30/2023 |  |  |  |  | The user facility reported a 60-year-old white undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported a cracked side arm during impella support. it was unclear if the damage was sustained due to extended usage of bicarb in the purge system or unintentional force applied by the medical/hospital staff. the pump was removed and replaced as a result. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation into the reported purge leak has been completed since the original report was filed. the pump was not returned for investigation. the staff noted that there was high purge pressure and then the sidearm was cracked. the purge flow increased, and purge pressure decreased, so the pump was turned off right when the leak was noticed. it cannot be confirmed if the product failed to meet specification or where the leak is located. the root cause of the purge leak-pump is most likely due to a damaged sidearm.
32621 | 1220648-2023-05137 | 18234832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18234832&pc=OZD | 10/30/2023 | 09/10/2024 |  | 11/29/2023 |  |  | 10/30/2023 | 11/29/2023 |  |  | 11/30/2023 |  | 11/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023219510 |  | 005042 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | Activation, Positioning or Separation Problem | [5 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 64-year-old male noted as high risk coronary intervention was implanted with an impella 2.5 device for mechanical circulatory support. the complainant reported the patient had a right femoral artery dissection during dilation prior to sheath placement, using the 10 fr dilator. the dissection was managed by two preclose sutures that had been placed prior to dissection. these were deployed and manual pressure resulted in hemostasis. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the delivery issues and vascular damage has been completed. the impella was not returned for evaluation. however, clinical details were sufficient to determine the root cause. the root cause of the delivery issues was most likely due to use issue. the root cause of the vascular damage was most likely due to use issue. the instructions for use for the impella cp with smartassist system states the following: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b1-selected product problem added- d6a/d6b. h6 med dev prob code 2906 changed to 2682.
32622 | 1220648-2023-05138 | 18234839 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18234839&pc=OZD | 10/13/2023 | 02/28/2024 |  | 11/29/2023 |  |  | 10/13/2023 | 11/29/2023 |  |  | 08/31/2025 |  | 11/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024386830 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a 65-year-old female patient who was just placed on impella cp support had total occlusion of proximal left anterior descending (lad). femoral antegrade access obtained in left femoral artery and joined with retrograde right side femoral sheath to perfuse left leg. impella cp was removed and the patient was escalated to impella 5.5. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible has been completed. the device and data were not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
32623 | 1220648-2023-05139 | 18234841 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18234841&pc=OZD | 10/31/2023 | 01/31/2024 |  | 11/29/2023 |  |  | 10/31/2023 | 11/29/2023 |  |  | 03/31/2025 |  | 11/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024343525 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Fracture | [6 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a 14fr being placed for access for the impella cp when the team observed the 14fr short dilator to have fractured. the team made choice to utilize the 14fr long sheath and dilator. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation for the reported introducer damage has been completed. the introducer was returned for evaluation. during visual inspection, a split at the end of the long dilator was noted. there was no other damage noted. the cause of the introducer kit damage is dilator tip split that was most likely use-related as best practices were not followed and serial dilation was skipped. the cause of the access site bleeding is most likely use related as best practices were not followed and serial dilation was skipped. without serial dilation excessive force was required. the instructions for use states the following: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ added b1: adverse event. added b2: required intervention. added h6: health effect-clinical code. updated h1: serious injury.
32624 | 1220648-2023-05140 | 18234845 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18234845&pc=OZD | 10/31/2023 | 04/17/2024 |  | 11/29/2023 |  |  | 10/31/2023 | 11/29/2023 |  |  | 08/31/2025 |  | 11/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024385342 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/31/2023 |  |  |  |  | The user facility reported a 67-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was in cardiac care unit as a large hematoma was noted at the impella access site. two units of blood were given. the patient was taken to a hybrid operation room for repair of the access site. the right femoral artery was cut down and suture was placed. site continued to ooze but significantly slow. patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical details provided was sufficient to determine the root cause. the root cause of the access site bleeding was most likely due to patient movement as the bleeding occurred after transfer to the icu.
32625 | 1220648-2023-05141 | 18234846 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18234846&pc=OZD | 10/31/2023 | 04/12/2024 |  | 11/29/2023 |  |  | 10/31/2023 | 11/29/2023 |  |  | 05/31/2025 |  | 11/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024340899 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/31/2023 |  |  |  |  | The user facility reported a 60-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. on day seven of impella support the patient received tissue plasminogen activator (tpa) purge due to pf (purge flow rate) issues. on day ten the patient had suboptimal flows for perfusion and impella was removed to address the issue. a clot was found in the outflow. the patient has active mural thrombus. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the high purge pressure and the low pump flow issue has been completed since the original report was filed. the pump was returned, tested, and evaluated. pump was found cut off at the catheter and a big chunk of biomaterial was observed around purge gap and out flow cage. the cause of the low pump flow issues was biomaterial ingestion. the cause of the high purge pressure was biomaterial buildup in the purge gap since administering tpa in the purge resolved the issue.
32626 | 1220648-2023-05142 | 18234848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18234848&pc=OZD | 11/15/2023 | 04/29/2024 |  | 11/29/2023 |  |  | 11/15/2023 | 11/29/2023 |  |  | 07/31/2025 |  | 11/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024391366 |  | 0045-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/15/2023 |  |  |  |  | The user facility reported a 71-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp pump placed for support of a patient turned down for cabg (coronary artery bypass grafting. the pump was delivered but the flow was not optimal. the team troubleshot by fluid delivery but, instead found the need to replace the pump. the 2nd pump was placed for flows to be optimal. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow issue has been completed since the original report was filed. the pump was returned, tested, and evaluated. data confirmed the low flow when pump started & remained to the rest of the case that triggered auto suction response and suction alarms. however, visual inspection found no issues on the pump. the failure mode could not be reproduced as the pump ran without issue under bench test. the root cause of low flow was most likely due to patient condition as no problem was found on the pump. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
32627 | 1220648-2023-05143 | 18234851 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18234851&pc=OZD | 11/15/2023 | 09/19/2024 |  | 11/29/2023 |  |  | 11/15/2023 | 11/29/2023 |  |  | 04/30/2025 |  | 11/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2024339491 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Death | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/15/2023 |  |  |  |  | The user facility reported a 68-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant in japan had a cp support ongoing when on day 17 of the support the team performed neuro imaging in the ct (computed tomography) suite and diagnosed a cerebral bleed. the pump remained on for support for 3 more days when pump explanted as patient expired from being made a dnr. the cerebral harm is reportable, but the death will be closed as unrelated. the patient was made a dnr which is the cause of death. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported stroke/cva has been completed. however, as the necessary clinical information was not provided and the device was not returned, the root cause of the stroke/cva could not be determined. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. corrected information for the initial mfr 1220648-2023-05143 was provided in sections b2, b5, d6a, d6b, h1, h3, and h6.
32628 | 1220648-2023-05144 | 18234942 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18234942&pc=OZD | 11/02/2023 | 08/20/2024 |  | 11/29/2023 |  |  | 11/02/2023 | 11/29/2023 |  |  | 11/30/2024 |  | 11/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024313965 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Use of Device Problem | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/02/2023 |  |  |  |  | The user facility reported a 75-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient self-explanted the impella by pulling out the peel away sheath. patient was stable. echo and repeat mixed venous gas were ordered. no known patient harm or injury. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the root cause of the positioning issue was use related to improper technique. d6a and d6b added the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05144: f6/f8-should have been left blank. h.6 code 4642 and 1670 was reported incorrectly.
32629 | 1220648-2023-05145 | 18234959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18234959&pc=OZD | 11/02/2023 | 03/26/2024 |  | 11/29/2023 |  |  | 11/02/2023 | 11/29/2023 |  |  | 05/31/2025 |  | 11/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024345464 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/02/2023 |  |  |  |  | The user facility reported a 65-year-old male undergoing coronary artery bypass grafting was implanted with an impella 5.5 device for mechanical circulatory support. the initial orientation of the pump was towards the lateral wall. physician made decision to manually reorient the position in the heart while attempting to advance. the physician pushed the pump too hard on insertion and perforated the left ventricle. surgical repair was performed and blood products were given. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation for the reported ventricle perforation has been completed. the pump was not returned for investigation. as per the clinical details, doctor pushed pump too hard on insertion and perforated lv. the cause of the ventricle perforation issue was most likely use related, based on given clinical details.
32630 | 1220648-2023-05146 | 18234976 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18234976&pc=OZD | 11/20/2023 | 09/21/2024 |  | 11/29/2023 |  |  | 11/20/2023 | 11/29/2023 |  |  | 09/30/2025 |  | 11/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024402212 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/20/2023 |  |  |  |  | The user facility reported a 76-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the echo showed mural clot at apex and a small pericardial effusion was noted. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the thrombus issue has been completed since the initial report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-05146 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
32631 | 1220648-2023-05147 | 18234982 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18234982&pc=OZD | 11/02/2023 | 08/19/2024 |  | 11/29/2023 |  |  | 11/02/2023 | 11/29/2023 |  |  | 08/31/2025 |  | 11/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024385590 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/02/2023 |  |  |  |  | The user facility reported an 83-year-old male undergoing coronary artery bypass grafting with posterior graft and mitral valve repair was implanted with an impella 5.5 device for mechanical circulatory support. the plan was to pull back impella plus caution lifting heart. md did not wish to pull back impella and when md lifted heart defect noted in the left ventricle (lv) wall. md stated lv wall is very thin. the defect was repaired, and the case was completed successfully. the patient was stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation for the ventricle perforation has been completed. pump was not returned for investigation. the clinical information provided mentions that during lifting the heart the left ventricle wall was seen to be very thin leading to ventricle perforation. therefore, the root cause of the ventricle perforation issue was most likely patient condition related. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- d6a/d6b. f6/f8 should have been left blank in the initial report.
32632 | 1220648-2023-05148 | 18234985 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18234985&pc=OZD | 11/01/2023 | 08/20/2024 |  | 11/29/2023 |  |  | 11/01/2023 | 11/29/2023 |  |  | 03/31/2025 |  | 11/29/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024324193 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/01/2023 |  |  |  |  | The user facility reported a 71-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. during support, the patient was treated in unit with intra-venous fluids, staff recommended afterload management to assist with hemolysis. urinalysis was drawn and was positive for red blood cells. after administering fluid, urine became clear, however the impella was still removed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. the root cause of the hemolysis issue was most likely patient condition related since it was resolved with ivf administration. h6 component code added. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8-should have been left blank on manufacturer device report 1220648-2023-05148.
32633 | 1220648-2023-05149 | 18234987 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18234987&pc=OZD | 11/01/2023 | 11/29/2023 |  | 11/29/2023 |  |  | 11/01/2023 | 11/29/2023 |  |  |  |  | 11/29/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1417971 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 11/01/2023 |  |  |  |  | The user facility reported a 49-year-old male post cardiotomy cardiogenic shock / low cardiac output syndrome was to be implanted with an impella 5.5 for mechanical circulatory support. during set-up of the automated impella controller (aic), the team observed the socket was broken. the damage may have occurred with prior support. the team switched to another aic without issue. there was no patient harm. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
32634 | 1220648-2023-05118 | 18221499 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18221499&pc=OZD | 10/29/2023 | 05/01/2024 |  | 11/28/2023 |  |  | 10/30/2023 | 11/28/2023 |  |  | 06/30/2025 |  | 11/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2024375620 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  |  |  | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during implant of impella cp for a 50-year-ld male patient, the catheter was kinked by the physician. there was no reported patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the product damage (cannula kink) issue has been completed since the original report was filed. the pump was returned and evaluated. according to the clinical, during insertion of the impella cp pump the physician kinked the catheter. visually inspected the pump and found catheter kinked as well. the cause of the product damage was catheter kink due to improper technique after implant. to conform with updated procedures, section d6 has been revised with updated information.
32635 | 1220648-2023-05119 | 18221924 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18221924&pc=OZD | 10/30/2023 | 02/16/2024 |  | 11/28/2023 |  |  | 10/30/2023 | 11/28/2023 |  |  | 06/30/2025 |  | 11/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2024375770 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 89 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Advance | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during implant of an impella cp, there was a lot of friction when the nurse tried to insert the long dilator into the long sheath introducer. the nurse removed the dilator, flushed the sheath, and attempted to insert it again but was unsuccessful. the dilator kinked due to force used. a short introducer therefore had to be used due to the patient anatomy. there was no reported patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues- use and the product damage (cannula kink) issues has been completed. the device was not returned for investigation. based on clinical description, it was determined that this was an introducer issue. the root cause of the introducer issue was most likely patient condition related to small/ disease vessels.
32636 | 1220648-2023-05120 | 18222046 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18222046&pc=OZD | 10/30/2023 | 09/11/2024 |  | 11/28/2023 |  |  | 10/30/2023 | 11/28/2023 |  |  |  |  | 11/28/2023 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1165069 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 10/30/2023 |  |  |  |  | The user facility reported that during a case preparation, the automated impella controller, (aic) unit displayed a controller failure error. a replacement loaner was sent to the facility. no patient is involved. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - controller failure (red alarm) has been completed. after data log review and device analysis it was concluded the root cause was defective a purge pressure sensor. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05120.
32637 | 1220648-2023-05121 | 18222136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18222136&pc=OZD | 10/30/2023 | 04/29/2024 |  | 11/28/2023 |  |  | 10/30/2023 | 11/28/2023 |  |  | 08/31/2025 |  | 11/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024395914 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 54-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant had a cp support ongoing when there were signs of limb ischemia. the pain and loss of pulses prompted the team to discuss intervention. the bypass was not performed, rather the team explanted as patient was stable. limb ischemia resolved following explant. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ to conform to updated procedures, section d6 has revised with updated information.
32638 | 1220648-2023-05122 | 18222275 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18222275&pc=OZD | 10/30/2023 | 08/15/2024 |  | 11/28/2023 |  |  | 10/30/2023 | 11/28/2023 |  |  | 08/31/2025 |  | 11/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024395724 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/30/2023 |  |  |  |  | The user facility reported a 73-year-old male noted for high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient had hemolysis, and hematuria present. the impella was repositioned and eventually the device was removed within several hours of repositioning due to hemolysis without any complications. patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. per the ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported hemolysis has been completed. the impella device was not returned for evaluation. data logs were investigated and there was no positioning issue, motor current trends or suction alarms observed. clinical details provided are limited. therefore, the root cause of the hemolysis issue was not determined since product was not returned and data logs were inconclusive. the instructions for use for the impella cp with smartassist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions including catheter position, pre-existing medical conditions, and small left ventricular volumes may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ added-h6 impact code 4628, h6 component code 925 in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8 should have been left blank in the initial report.
32639 | 1220648-2023-05021 | 18222404 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18222404&pc=OZD | 10/30/2023 | 05/01/2024 |  | 11/28/2023 |  |  | 10/30/2023 | 11/28/2023 |  |  | 07/31/2025 |  | 11/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024391342 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Female |  | Non Hispanic | White | Malfunction | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/30/2023 |  |  |  |  | The user facility reported a 28 -year-old male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the pump was explanted after 4 days with signs of hemolysis. lab values were indicative of hemolysis. pump was explanted and the patient was taken for heart transplant. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the hemolysis issue has been completed since the original report was filed. the pump was returned, tested, and evaluated. a cut was observed at the catheter but there were no other abnormalities noted. the data logs indicate some suction and positioning alarms when they start the pump, but it immediately resolves, along with at the end when they are weaning. the root cause of the hemolysis cannot be determined due to limited information provided. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ in order to conform to updated procedures, sections d6 and h10 have been revised with updated information.
32640 | 1220648-2023-05123 | 18222411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18222411&pc=OZD | 10/30/2023 | 03/07/2024 |  | 11/28/2023 |  |  | 10/30/2023 | 11/28/2023 |  |  | 08/31/2025 |  | 11/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024381560 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 51-year-old male in st elevated myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient developed hemolysis during impella support coinciding with persistent suction issues. heparin was halted temporarily as a result of the hemolysis and the positioning and p-level were monitored to troubleshoot the suction. despite providing additional volume and verifying the pumps positioning, the suction remained unresolved. the pump was subsequently explanted and the patient was transferred to a centrimag acute circulatory support system for ongoing support. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the reported low pump flow and hemolysis has been completed. the impella device was not returned for evaluation. data logs were reviewed which indicated several suction alarm with little less than expected pump flow. there was no motor current spike or positioning alarms were recorded. however clinical details provided were sufficient to determine the root cause of the low pump flow. the cause of the issue low pump flow was patient condition related since the patient was on biventricular support, and the data log did not confirm any events related to biomaterial interaction or improper positioning of the pump. however, clinical details provided were insufficient to determine the root cause of the hemolysis. the cause of the hemolysis issue was not determined since product was not returned.
32641 | 1220648-2023-05124 | 18223095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18223095&pc=OZD | 10/29/2023 | 04/29/2024 |  | 11/28/2023 |  |  | 10/29/2023 | 11/28/2023 |  |  | 06/30/2025 |  | 11/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024398978 |  | 1000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported 72-year old male patient on impella 5.5 support had ongoing low flows/suction issues. the patient had been on pup support for 2 days when the suction occurred. the pump still is on for support and there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ | The investigation for the low pump flow has been completed. no product was returned for a visual inspection. data log analysis confirmed suction and low pump flow with no motor current spikes and a gradual decrease in placement signal. according to the clinical, on the fourth day of impella 5.5 support suction and low pump flow were recorded. the patient was intubated on va ecmo at the time of these alarms. no other clinical information was provided. the most likely cause of the low pump flow was the patient's condition. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
32642 | 1220648-2023-05125 | 18223311 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18223311&pc=OZD | 10/30/2023 | 09/24/2024 |  | 11/28/2023 |  |  | 10/30/2023 | 11/28/2023 |  |  | 04/30/2025 |  | 11/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2024345367 |  | 0048-0024 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 61 YR |  | Male |  |  |  | Death | Yes | Yes | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/30/2023 |  |  |  |  | The user facility reported a 61-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the pump stopped running on the 14th day. it did not restart and was removed. the patient subsequently died, and there was no causal relationship between impella and death, and the cause of death was acute myocardial infarction. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore,?an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation has been completed since the original report was filed. the purge cassette was returned for evaluation, but the cp pump was not returned. the purge cassette was connected to the console and primed. it was purged with green fluid. the purge cassette was connected to a good sample pump and purge flow. purge pressure was observed to be normal. no purge leak or any other purge related alarms were observed. the issue did not reproduce during the investigation. as per clinical review, the impella was in support for three hours and then air in the purge alarm appeared while operating the patient in intensive care unit. the alarm resolved only after change of the purge cassette. the pump stop was observed after 14 days of support. data logs were reviewed and air in purge alarm appeared after three hours in support after reduction to lower p-level for surgery and alarms resolved after changing the purge cassette. the returned device functioned to specification when primed, purged with a good pump and no issues were observed. the cause of the pump stop issue was not determined since the product was not returned and with alarm pump stop observed along with high motor current pump stop based on log review. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-05125 was provided in sections b1, b2, b5, d6a, d6b, h1, and h6.
32643 | 1220648-2023-05097 | 18208660 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18208660&pc=OZD | 11/10/2023 | 01/17/2024 |  | 11/27/2023 |  |  | 11/10/2023 | 11/27/2023 |  |  | 10/31/2025 |  | 11/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 SMARTASSIST SET, JP |  |  | ABIOMED, INC. |  |  | US |  | 2024407801 |  | 1000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/10/2023 |  |  |  |  | The user facility in japan reported a 59-year-old female admitted with multivessel coronary artery disease and in heart failure, reported an impella 5.5 was attempted to be implanted. however, after exposing the blood vessel and checking the diameter, the implant was aborted. the physician made the choice to implant an impella cp. no known patient harm or injury. reportable due to potential patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the cause of the delivery issue was patient condition related due to patient condition of small blood vessels.
32644 | 1220648-2023-05098 | 18210641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18210641&pc=OZD | 11/06/2023 | 04/09/2024 |  | 11/27/2023 |  |  | 11/06/2023 | 11/27/2023 |  |  |  |  | 11/27/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1143273 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the 63-year-old male patient was on an impella cp for mechanical support. the automated impella controller (aic) displayed the alarm "controller defective" for approx. 5min. the pump continued to run without any problems; however, the aic was exchanged for another console. there was no harm to the patient. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation into the aic - controller failure (red alarm) issue has been completed since the initial report was submitted. the console was returned, visually inspected, and tested. the reported failure was confirmed in the data logs; however, the failure was unable to be reproduced during testing. the cause of the issue was not determined since the issue could not be reproduced.
32645 | 1220648-2023-05099 | 18210941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18210941&pc=OZD | 11/03/2023 | 04/15/2024 |  | 11/27/2023 |  |  | 11/03/2023 | 11/27/2023 |  |  | 02/29/2024 |  | 11/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023188970 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during index percutaneous coronary intervention (pci), with impella support, of the left anterior descending artery (lad), the patient of unknown age and gender on impella cp support, developed distal thrombus embolization with no-reflow. the patient on the impella cp went into ventricular fibrillation (vf), which was fixed with defibrillation (1x). the informaiton of the vf was shared via the study team after the procedure. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported ventricular fibrillation (vf) and thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of vf and thrombus could not be determined as the device was not returned.
32646 | 1220648-2023-05100 | 18211256 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18211256&pc=OZD | 11/22/2023 | 02/09/2024 |  | 11/27/2023 |  |  | 11/22/2023 | 11/27/2023 |  |  | 12/31/2024 |  | 11/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024315879 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/22/2023 |  |  |  |  | The user facility reported a 71-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella support and experienced low flows once the impella started. two blood product units were given to the patient and the low flow resolved. additionally, there was a high purge pressure alarm and purge system blocked alarm. tpa was given and this resolved the alarms. reportable due to high purge pressure and low pump flows and necessary intervention. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation into the reported low pump flow and high purge pressure has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. according to the clinical, a low flow suction alarm was initially recorded. volume was added and the issue was not resolved. the logs confirm low flow and suction alarm. the cause of the low flow is undetermined because no product was returned and not enough clinical information was provided. according to the clinical, a purge system blocked alarm was recorded. tpa was administered immediately, and the alarm was resolved. the logs confirm a purge system blocked alarm. tpa was added to the purge and the pressure decreased to a normal range. the cause of the high purge pressure is biomaterial as tpa resolved the issue.
32647 | 1220648-2023-05102 | 18211276 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18211276&pc=OZD | 11/17/2023 | 05/29/2024 |  | 11/27/2023 |  |  | 11/17/2023 | 11/27/2023 |  |  | 12/31/2024 |  | 11/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024291885 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/17/2023 |  |  |  |  | The user facility reported a 64-year-old malewas implanted with an impella device for mechanical circulatory support. the us complainant had placed a 5.5 pump to support a patient admitted with cardiomyopathy while awaiting heart transplant. the 5.5 pump was placed via the axillary artery graft and the iabp was removed. the pump is still on support to date but, there has been access site bleeding. the team treated the bleed by evacuation and blood products. | The patient's outcome was reviewed by abiomed's medical safety officer. it was determined that there was not enough information to associate use of device with outcome. |  |  |  |  | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. the patient was stable after the evacuation and the issue was resolved. the instructions for use referenced in the initial report is from the ifu for impella 5.5 with smart assist for use during cardiogenic shock, section: potential adverse events (united states), which now includes statement "cardiac or vascular injury (including ventricular perforation)¿ added-b5 medical safety officer's review d4-updated model number added- d6a/d6b.
32648 | 1220648-2023-05101 | 18211287 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18211287&pc=OZD | 11/03/2023 | 01/16/2024 |  | 11/27/2023 |  |  | 11/03/2023 | 11/27/2023 |  |  | 07/31/2025 |  | 11/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024391580 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/03/2023 |  |  |  |  | The user facility reported a 71-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the physician attempted to advance impella though the iliac, but it was too occluded. the case was aborted. no patient harm was reported due to this issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into delivery issues- anatomy has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the cause of the issue was patient condition related because of the complex vasculature anatomy of the patient.
32649 | 1220648-2023-05105 | 18211855 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18211855&pc=PYX | 10/27/2023 | 04/30/2024 |  | 11/27/2023 |  |  | 10/27/2023 | 11/27/2023 |  |  | 12/31/2023 |  | 11/27/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023158174 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/27/2023 |  |  |  |  | The user facility reported 79-year-old female with cardiomyopathy was implanted with an impella rp for mechanical circulatory support. the patient developed a hematoma at the impella access site. manual pressure was used and the physician applied additional sutures. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical details provided was sufficient to determine the root cause. the root cause of the access site bleeding was most likely issue was use related as manual pressure was held and the physician applied more sutures to resolve the issue as per the clinical description. added- d6a/d6b in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
32650 | 1220648-2023-05103 | 18212125 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18212125&pc=OZD | 11/22/2023 | 11/12/2024 |  | 11/27/2023 |  |  | 11/22/2023 | 11/27/2023 |  |  | 06/30/2025 |  | 11/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024357263 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Material Integrity Problem | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/22/2023 |  |  |  |  | The user facility reported a 51-year-old male was implanted with an impella device for mechanical circulatory support. the us complainant had a cp placed via the 14fr sheath to support a patient admitted post vf (ventricular fibrillation) arrest and in need of cabg (coronary artery bypass grafting) for multivessel coronary artery disease. upon removal of the 14fr sheath the pump was advance and bent back upon itself. the team needed to use a snare from contralateral leg to remove the wire and pump from the anatomy. the team placed a new pump. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the product kink has been completed. pump was not returned for investigation. as per the clinical information provided, the cannula kink happened when removing the introducer for swapping the repo unit. therefore, the root cause of the product damage cannula kink issue was use related to improper technique.
32651 | 1220648-2023-05104 | 18212179 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18212179&pc=OZD | 10/27/2023 | 08/20/2024 |  | 11/27/2023 |  |  | 10/27/2023 | 11/27/2023 |  |  | 09/30/2025 |  | 11/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024399163 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  |  | Malfunction | No | Yes | Material Twisted/Bent | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/27/2023 |  |  |  |  | The user facility reported a 72-year-old male (race unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella cp was repositioned in icu with echo and cannula "folded on itself" in 2 different spots. the patient was brought to ccl and md had to snare pigtail from the aorta and straighten out the cannula in order to remove the device. a new cp was placed successfully. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. the pump was not returned for investigation. according to the clinical report, the cannula was found to be folded during repositioning. the doctor tried several times but was unable to position it correctly. new device was placed. the cause of the positioning issue was not determined since product was not returned and sufficient clinical information was not provided. added- d6a/d6b. f6/f8 should have been left blank in the initial report.
32652 | 1220648-2023-05106 | 18212200 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18212200&pc=OZD | 10/28/2023 | 05/01/2024 |  | 11/27/2023 |  |  | 10/28/2023 | 11/27/2023 |  |  | 07/31/2025 |  | 11/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024390520 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/28/2023 |  |  |  |  | The user facility reported a 71-year-old female (race unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced persistent hemolysis throughout the course of support. despite successful repositioning during instances of suction or deep positioning, the hemolysis remained ongoing and the decision was eventually made to explant the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the reported hemolysis has been completed. no product was returned for a visual inspection. data logs confirmed suction alarms followed by positioning issues in the aorta and ventricle. according to the clinical, shortly after beginning impella cp support hemolysis was detected. hemolysis remained present throughout support, and ultimately resulted in removal of the device on the second day of support. suction alarms were also recorded on the second day of support. first the pump was found in the aorta and was successfully repositioned. then two hours later the pump was found in the ventricle and was successfully repositioned again. the most likely cause of the hemolysis was improper positioning. potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ to conform with updated procedures, sections d6 & h10 have been revised with updated information.
32653 | 1220648-2023-05107 | 18212206 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18212206&pc=OZD | 10/27/2023 | 09/10/2024 |  | 11/27/2023 |  |  | 10/28/2023 | 11/27/2023 |  |  | 06/30/2025 |  | 11/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - EU |  |  | ABIOMED, INC. |  |  | US |  | 2024403295 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/28/2023 |  |  |  |  | The user facility reported a patient of unknown age, race and gender in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, a hematoma formed in the patient's groin. hemoglobin dropped to 7.5, therefore three units of red blood cells were administered. | It was further reported that the patient expired while on support. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was anticoagulation management due to high patient act values as per the clinical details provided. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- b5 mso review, h6 component code 925 in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8 should have been left blank in the initial report. g1-corrected mfr site fax number.
32654 | 1220648-2023-05108 | 18212235 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18212235&pc=OZD | 10/28/2023 | 02/28/2024 |  | 11/27/2023 |  |  | 10/28/2023 | 11/27/2023 |  |  | 04/30/2025 |  | 11/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP JP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024339551 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient with existing lower limb ischemia had to be escalated from an impella cp to an impella 5.5 pump in response to the condition. it was noted that the ischemia developed following prior insertion of an intra-aortic balloon pump. a fem-fem bypass had been completed prior to impella implant, but the decision was made to escalate to impella 5.5 support to avoid further complication in the limb. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible has been completed. the device and data were not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
32655 | 1220648-2023-05109 | 18212252 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18212252&pc=OZD | 11/05/2023 | 09/21/2024 |  | 11/27/2023 |  |  | 11/05/2023 | 11/27/2023 |  |  | 06/30/2025 |  | 11/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024360916 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Off-Label Use | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/05/2023 |  |  |  |  | The user facility in japan reported a 82-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after close to a month of support the pump stopped. the stoppage caused no harm or injury. the pump had been used beyond indication. the patient later expired from underlying causes of ami. the device was used for more than the designed number of days, doctors determined that the device was not related to the death. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: the impella cp® system has been approved for = 4 days. however, weaning could be delayed beyond the normal use for temporary support as an unintended consequence of continued instability of the patient¿s hemodynamics. inability to wean the patient from the device within a reasonable time frame should result in consideration of a more durable form of left ventricular support. | The investigation into the pump stop issue has been completed since the original report was filed. the device was returned and visually inspected. analysis of the pump revealed a large purge gap indicating bearing wear. the root cause of the pump stop was a wearing bear. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-05109 was provided in sections b1, b2, b5, d6a, d6b, h1, and h6.
32656 | 1220648-2023-05110 | 18212277 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18212277&pc=OZD | 10/28/2023 | 02/06/2024 |  | 11/27/2023 |  |  | 10/28/2023 | 11/27/2023 |  |  | 09/30/2025 |  | 11/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 EU |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2024399669 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that 65-year-old male patient was implanted with an impella 5.5 for mechanical circulatory support. however, purge fluid was noted to be leaking between purge reservoir and check valve. purge flow 12ml/h, pp 400mmhg. it is unknown if intervention was required for the leak. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation into the purge leak - pump issue has been completed. the device was not returned for benchtop investigation. according to the clinical information provided, purge fluid was leaking between purge reservoir and check valve. the cause of the purge leak issue was not determined since product was not returned and due to insufficient clinical information. | The investigation into the purge leak ¿ pump issue has been completed. the device was not returned for investigation. as per clinical, purge fluid was leaking between purge reservoir and check valve. the cause of the purge leak issue was not determined since product was not returned and due to insufficient clinical information.
32657 | 1220648-2023-05111 | 18212311 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18212311&pc=OZD | 10/29/2023 | 04/30/2024 |  | 11/27/2023 |  |  | 10/29/2023 | 11/27/2023 |  |  | 08/31/2025 |  | 11/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024398018 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 40-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that following implant of an impella cp pump, pulsatile bleeding was noted at the insertion site. after some discussion, the decision was made to explant the pump. hemostasis was successfully achieved via angiography following explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. added- d6a/d6b. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
32658 | 1220648-2023-05112 | 18212390 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18212390&pc=OZD | 10/29/2023 | 01/17/2024 |  | 11/27/2023 |  |  | 10/29/2023 | 11/27/2023 |  |  | 09/30/2025 |  | 11/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024394986 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/29/2023 |  |  |  |  | The user facility reported a 70-year-old female with cardiomyopathy was to be implanted with an impella 5.5 for mechanical circulatory support. the impella 5.5 was unable to pass through patient¿s native axillary artery. decision was made to place an impella cp via axillary access. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the cause of the delivery issue was patient condition related due to patient¿s native axillary artery unable to pass 5.5 pump.
32659 | 1220648-2023-05113 | 18212446 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18212446&pc=OZD | 10/29/2023 | 02/07/2024 |  | 11/27/2023 |  |  | 10/29/2023 | 11/27/2023 |  |  | 08/31/2025 |  | 11/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024392145 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Improper Flow or Infusion | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/29/2023 |  |  |  |  | The user facility reported a 55-year-old male patient was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient developed heparin induced thrombocytopenia roughly five days after implant. the patient was switched from heparin to bival as a result of the condition and support continued. however, a couple days later, the pump experienced a sudden loss of flow coinciding with a loss of motor current pulsatility. despite troubleshooting, the flow issues persisted and the decision was made to explant the pump. | [5 empty] | The impella device was received from the customer and an evaluation of the device is pending. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported thrombocytopenia and saturated flow has been completed since the original report was filed. the logs were analyzed and a motor current spike was observed on the seventh day of impella 5.5 support, and followed by a drop in motor current and pump flow. this data is consistent with biomaterial ingestion. the impella 5.5 was returned for analysis and no biomaterial or other abnormalities were found. however, clinical details specify that biomaterial was observed upon explant. the most likely root cause of the low pump flow was determined to be biomaterial ingestion. the most likely root cause of the thrombocytopenia is patient condition and the relationship to impella is unknown.
32660 | 1220648-2023-05114 | 18212455 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18212455&pc=OZD | 10/28/2023 | 02/28/2024 |  | 11/27/2023 |  |  | 10/28/2023 | 11/27/2023 |  |  | 09/30/2025 |  | 11/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024397005 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/28/2023 |  |  |  |  | The user facility reported a 35-year-old female with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, the patient developed ischemia post placement of the abiomed 14 fr introducer. the sheath was peeled away and repositioning sheath was advanced, however, a reperfusion catheter could not be placed. the pump was removed, and a new impella was placed via 14fr gore sheath (physician elected to leave introducer in place during support). | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible has been completed. the device and data were not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
32661 | 1220648-2023-05115 | 18212504 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18212504&pc=OZD | 10/28/2023 | 01/16/2024 |  | 11/27/2023 |  |  | 10/28/2023 | 11/27/2023 |  |  | 08/31/2025 |  | 11/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024382967 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported the surgeon gained access to subclavian artery of a 65-year-old female patient, via surgical cut down and inserted axillary kit. impella 5.5 was inserted over abiomed .018 wire, but was unable to be fully delivered. the device reportedly got stuck near ascending aorta due to patient anatomy. surgeon opted to remove 5.5 and tried with impella cp. there was no known patient harm or injury. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation into the reported delivery issue -anatomy has been completed since the original report was filed. the root cause of the delivery issue- anatomy was most likely due to patient condition.
32662 | 1220648-2023-05116 | 18212515 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18212515&pc=OZD | 11/11/2023 | 09/06/2024 |  | 11/27/2023 |  |  | 11/12/2023 | 11/27/2023 |  |  | 12/31/2024 |  | 11/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024321373 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/12/2023 |  |  |  |  | The user facility reported a 33-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient had a pre existing genetic clotting disorder. patient was hypercoagulable and per the nurse "is throwing clots everywhere". the pump was removed and a 3 foot long clot was evacuated from the site. the patient developed compartment syndrome in the impella leg requiring a fasciotomy. staff stated the patient will most likely need a below the knee amputation. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of thrombus and limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05116. d3 manufacturer name, city, state, and fax number. d4 model number. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 manufacturing site fax number missing. h6 health effect - clinical code 1942 was inadvertently not included. h6 code 2199 was reported incorrectly. component code was added. instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings and cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
32663 | 1220648-2023-05117 | 18212547 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18212547&pc=OZD | 10/29/2023 | 09/22/2024 |  | 11/27/2023 |  |  | 10/29/2023 | 11/27/2023 |  |  | 07/31/2025 |  | 11/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024390534 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  |  |  | Death | No | Yes | Activation, Positioning or Separation Problem | Device Handling Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/29/2023 |  |  |  |  | The user facility reported a 76-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp being delivered for support of a patient who arrested in the community. when placed via the 14fr long sheath the introducer kinked and was replaced. the team successfully placed the cp via the short sheath. the team additionally placed ecmo (extracorporeal membrane oxygenation) and reperfusion sheaths. the patient condition was poor despite all measures and care was withdrawn. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. as noted in the impella ifu: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrected information for the initial mfr 1220648-2023-05117 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding. | The investigation into the introducer damage ¿ mechanical issue has been completed since the original report was filed. the device was returned for investigation. the 14fr x25cm sheath and 14fr dilator were returned for evaluation. multiple kinks were found along the sheath of the introducer. in addition, the dilator was found to be scratched and deformed at the tip. the cause of the sheath kink and dilator damage was most likely due to patient's calcified iliac artery.
32664 | 1220648-2023-05093 | 18207737 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18207737&pc=OZD | 10/30/2023 | 08/20/2024 |  | 11/26/2023 |  |  | 10/30/2023 | 11/26/2023 |  |  | 05/31/2025 |  | 11/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024344534 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Material Puncture/Hole | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 76-year-old female noted as high risk percutaneous cardiac intervention (hrpci) was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported there was blood noticed in the sterile sleeve during patient support in the icu. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation for the sterile sleeve damage has been completed. pump was not returned for investigation. as per the clinical details, blood noticed in the sterile sleeve during patient support. the cause of the product damage issue was not determined since product was not returned and sufficient clinical information was not provided. the instructions for use referenced in the initial report is for the impella 5.5 with smartassist for use during cardiogenic shock in the following section: section: warnings & cautions: cautions added- d6a/d6b.
32665 | 1220648-2023-05094 | 18207741 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18207741&pc=OZD | 11/01/2023 | 01/16/2024 |  | 11/26/2023 |  |  | 11/01/2023 | 11/26/2023 |  |  | 08/31/2025 |  | 11/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024393031 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Hispanic |  | Injury | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/01/2023 |  |  |  |  | The user facility reported a 67-year-old male post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. the impella 5.5 could not pass the axillary artery due to the patient¿s vascular anatomy. an impella cp was placed instead. no patient harm reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into delivery issues- anatomy has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the cause of the delivery issue was patient condition related due to vascular anatomy.
32666 | 1220648-2023-05095 | 18207761 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18207761&pc=OZD | 11/01/2023 | 01/26/2024 |  | 11/26/2023 |  |  | 11/01/2023 | 11/26/2023 |  |  | 09/30/2025 |  | 11/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024401113 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/01/2023 |  |  |  |  | The user facility reported a 75-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. during placement of the impella 5.5 there was a placement signal unreliable yellow alarm on the aic. the aic was switched out and the alarm appeared again. it was discussed with the surgeon that the optical sensor may have been damaged. the impella was replaced and a new 5.5 was successfully inserted. no patient harm reported due to the issue. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the optical signal issue has been completed. product and data logs were returned and investigated. there was a major kink in the pump catheter near the motor end of the pump which was confirmed by light continuity test and optical bench parameters through logs. per the investigation results, the root cause of the optical signal issue was a damaged optical fiber line due to excessive force. note that the placement signal failure occurred prior to pump activation and during delivery.
32667 | 1220648-2023-05096 | 18207762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18207762&pc=OZD | 11/04/2023 | 01/17/2024 |  | 11/26/2023 |  |  | 11/04/2023 | 11/26/2023 |  |  | 09/30/2025 |  | 11/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024399765 |  | 1000100 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/04/2023 |  |  |  |  | The user facility reported a 78-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an impella cp pump met resistance during attempted implant. it was noted that the device would not track over the aortic arch due to severe calcification and existing aortic stenosis. due to the noted issue, the decision was made to abort the implant. there was no patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the cause of the delivery issue was patient condition related due to patient anatomy. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32668 | 1220648-2023-05053 | 18190312 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18190312&pc=OZD | 10/25/2023 | 03/13/2024 |  | 11/22/2023 |  |  | 10/25/2023 | 11/22/2023 |  |  | 06/30/2025 |  | 11/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024379305 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 60-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that an implanted impella cp pump experienced an unspecified issue during support attributed to a clot. it was stated that the device had been running without heparin for the entirety of the patient's support due to the patient's bleed risk. the pump was explanted and, upon running the pump in a clean saline bowl, a fibrous clot was observed. | [5 empty] | The investigation into the access site bleeding ¿ minor and the thrombus issues has been completed since the original report was filed. the device was not returned for investigation. the cause of the access site bleeding was most likely lack of vessel recoil onto the sheath since the bleeding was noted after insertion of guidewire reposition unit and it was resolved with placement of femstop. the cause of the thrombus issue was most likely patient condition related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32669 | 1220648-2023-05054 | 18190357 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18190357&pc=PYX | 10/26/2023 | 01/18/2024 |  | 11/22/2023 |  |  | 10/26/2023 | 11/22/2023 |  |  | 05/31/2025 |  | 11/22/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024355147 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/26/2023 |  |  |  |  | The user facility reported a 70-year-old male undergoing coronary artery bypass grafting and mitral valve repair was implanted with an impella rp device for mechanical circulatory support. resistance was met during attempted impella rp implant. it was noted that the pump was unable to advance into place and the shaft of the catheter subsequently kinked. the pump was removed and an impella 5.5 pump was placed for patient support. there was no noted patient harm. | [5 empty] | The investigation into the delivery issues- use has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the root cause of the delivery issue was most likely due to patient condition. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32670 | 1220648-2023-05055 | 18190369 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18190369&pc=OZD | 10/26/2023 | 01/11/2024 |  | 11/22/2023 |  |  | 10/26/2023 | 11/22/2023 |  |  |  |  | 11/22/2023 |  | AUTOMATED IMPELLA CONTROLLER |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1371823 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after loading the purge cassette and disc into the automated impella controller (aic), the console screen would not advance to the next step. the aic was therefore removed from service and a loaner was sent to the customer site. there was no known patient harm resulting from the reported issue. | [5 empty] | The investigation into the reported aic - purge disc/flag issues has been completed since the original report was filed. after reviewing the imc logs from the reported occurrence date, the issues with not being able to get past step 2 of case start wizard was confirmed in multiple instances. for every case start that the user attempted, a purge disc not detected alarm event #402 would occur during step 2. in order to continue past step 2, the purge cassette must be inserted and a purge disc must be detected. the consoles inability to detect the purge disc was the reason that the user was unable to advance past step 2. there were also a couple of purge system open alarms (event #407) found in the logs. because the purge retainer was broken and not flush with the purge assembly, when purge pressure was applied to the purge disc it would record very low pressure or none at all thus resulting in this purge system open alarm. there were also cases of impella stopped (motor current high -- detected by the epc) (event #13) and impella stopped (mechanical/electrical issue) (event #115). these 2 alarms were suspected to be caused by a worn out pump. pump 415209 is suspected to be a demo pump because it had been used on 8 different consoles.device analysis: console was tested after arriving in fs. upon examining aic for any visible defects, it was evident that the purge retainer was broken. the portion of the purge retainer that was fastened to the purge assembly was no longer intact. if the retainer is not flush with the console, there can be an inconsistent detection of the purge disc depending on how it sits in tandem with the purge flag. this may be the reason why the purge disc was not able to be detected which would result in the console not being able to advance past step 2 of the case start wizard. | The automated impella controller (aic) was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
32671 | 1220648-2023-05056 | 18190391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18190391&pc=OZD | 10/26/2023 | 02/16/2024 |  | 11/22/2023 |  |  | 10/26/2023 | 11/22/2023 |  |  | 07/31/2025 |  | 11/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024379806 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the 65-year-old male patient developed a hematoma at their access site during impella 5.5 support. the heparin drip was halted as a result of the noted condition. support with the implanted pump continued following the event. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed. the device was not returned for investigation. as per the clinical details, the cause of the access site bleeding issue was not determined since product was not returned and sufficient clinical information was not provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32672 | 1220648-2023-05057 | 18190409 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18190409&pc=OZD | 10/27/2023 | 05/13/2025 |  | 11/22/2023 |  |  | 10/27/2023 | 11/22/2023 |  |  |  |  | 11/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1393502 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a male patient noted as high risk percutaneous cardiac insertion was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that a patient on impella support experienced an automated impella controller (aic) error alarm occurred in the middle of the night. the alarm resolved within a few seconds and support with the device continued until the aic could be replaced roughly two days later. there was no harm to the patient as a result of this incident. | [5 empty] | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. the logs confirmed the reported failure mode given there was a controller error alarm. during the investigation the root cause was determined to be a software issue. | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | D2 common device name revised on manufacturer device report 1220648-2023-05057 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2023-05057 in accordance with updated procedures.
32673 | 1220648-2023-05058 | 18190422 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18190422&pc=PYX | 10/27/2023 | 08/19/2024 |  | 11/22/2023 |  |  | 10/27/2023 | 11/22/2023 |  |  | 02/28/2025 |  | 11/22/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024348480 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/27/2023 |  |  |  |  | The user facility reported a 52-year-old male with a heart transplant was implanted with an impella rp for mechanical circulatory support. the placement signal for their device became unreliable during patient support. as a result of the noted issue, the decision was made to explant the pump. there was no patient harm. | [5 empty] | The investigation for the optical signal issue has been completed. data logs show that 8 hours into the case the placement signal (ps) begins to drift downwards and eventually placement signal not reliable (psnr) dp out of range alarm is triggered with flow calculations disabled. flows remain disabled for 1.5 hrs before returning. cvp is also trending downwards at this time. on the returned product, both sensors found to be working. optical 5s parameters in spec. dp drift was not reproduced in a pressurized flow loop. psnr alarm not reproduced. pump passed electrical and voltage/signal testing. the root cause of the placement signal issues was most likely patient condition related; both sensors were reacting to the pump's environment, both sensors found to be within spec, and the ps issue was not reproduced in the lab. added-d6a/d6bf6/f8 should have been left blank in the initial report. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
32674 | 1220648-2023-05059 | 18190438 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18190438&pc=OZD | 10/27/2023 | 08/15/2024 |  | 11/22/2023 |  |  | 10/27/2023 | 11/22/2023 |  |  | 05/31/2025 |  | 11/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024354656 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fracture | Difficult to Remove | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/27/2023 |  |  |  |  | The user facility reported a 61-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that seven centimeters of the utilized guidewire broke off and remained in the peel away sheath during attempted removal. the sheath was subsequently pulled back and peeled away, at which point the wire was able to be pulled out. the repositioning sheath was reinserted for ongoing support. there was no patient harm resulting from the noted issue. | [5 empty] | The investigation for the product damage has been completed. no product was returned for evaluation. according to the clinical details, 7 cm of the .018 wire broke off and remained in the peel-away sheath. the wire was pulled out during the peel-away introducer. the doctor kinked the wire several times and had to apply excessive force. the cause of the product damage issue was not determined since product was not returned and sufficient clinical information was not provided. d4-corrected model number added- d6a/d6b. f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32675 | 1220648-2023-05060 | 18190465 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18190465&pc=OZD | 10/27/2023 | 06/18/2024 |  | 11/22/2023 |  |  | 10/27/2023 | 11/22/2023 |  |  | 05/31/2025 |  | 11/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024365474 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient expired while on support. upon review by abiomed's medical safety officer, there is not enough information to associate the use of device with the patient's outcome. | The us complainant had impella cp pump support ongoing when the patient had a run of atrial fibrillation that prompted 5x defibrillation. the defib did return the patient to normal sinus rhythm (nsr). |  |  |  |  | The investigation into the atrial fibrillation with rapid ventricular response issues has been completed since the original report was submitted. the device was not returned for investigation. as the necessary information was not provided, the root cause was not determined. b.7 revised as this information was inadvertently omitted from the previously submitted report. d.6a and d.6b revised from the initial submission in accordance with updated procedures. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted report incorrectly. no ifus are needed to be reported for this report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32676 | 1220648-2023-05061 | 18190709 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18190709&pc=OZD | 10/27/2023 | 06/13/2024 |  | 11/22/2023 |  |  | 10/27/2023 | 11/22/2023 |  |  | 07/31/2025 |  | 11/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024387385 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/27/2023 |  |  |  |  | The user facility reported a 71-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed hemolysis during impella cp support. as a result of the condition, the doctor attempted to reposition the pump: however, the device was pulled back across the aortic valve. the pump was subsequently explanted and a new impella cp pump was placed for ongoing support. | [5 empty] | The investigation into the positioning issues and the hemolysis issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the positioning issue is most likely due to use as the pump was accidentally pulled into the aortic valve during repositioning of the device. the root cause of the hemolysis issue cannot be determined. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. per the ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32677 | 1220648-2023-05062 | 18190757 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18190757&pc=OZD | 10/24/2023 | 01/12/2024 |  | 11/22/2023 |  |  | 10/24/2023 | 11/22/2023 |  |  | 04/30/2025 |  | 11/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 380 SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | GM |  | 2024344579 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/24/2023 |  |  |  |  | The user facility reported an unknown gender patient was brought in to be implanted with an impella cp for unknown reason. the physician had already tried to position the pump, but then stopped because the patient got ventricular fibrillation and the physician decided not to use the impella device. patient is stable. | [5 empty] | The investigation into the ventricular fibrillation (vf) has been completed. the device was not returned for benchtop evaluation and investigation. per the clinical information provided, the cause of the vf/vt (ventricular tachycardia) was most likely patient condition related because the timing of the event correlated with placement/repositioning and any device contacting the heart wall in conjunction with a poor patient condition could result in vt/vf. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32678 | 1220648-2023-05063 | 18190848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18190848&pc=OZD | 10/28/2023 | 10/31/2025 |  | 11/22/2023 |  |  | 10/28/2023 | 11/22/2023 |  |  |  |  | 11/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1428325 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 | [18 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 10/28/2023 |  |  |  |  | The user facility reported a 59-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a 5.5 support ongoing when there was position signal loss. the loss of signal prompted the team to replace the console. the pump had signal when on the 2nd console. | [5 empty] | The impella device was received from the customer and an evaluation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: ¿ the case associated with the reported complaint was initiated on oct 26, 2023, for the pump sn: (b)(4). ¿ on the complaint date, oct 28, 2023, during support, the console displayed ¿placement signal not reliable (opm invalid signal, optical pump connected)¿ event #(b)(6). ¿ the optical parameters were abnormal (light value: 0.53v and snr value: 0.5). the rt log showed the contrast oscillating between -3276 and 3276. the ps spiked to 3000, confirming the reported failure mode. ¿ subsequently, the console displayed the alarm ¿controller error (defective optical sensor lamp)¿ with event #(b)(6). device analysis: the console was tested upon arrival in fs, and the reported failure mode (psnr) was reproduced. although the average input voltage to the lamp was 4.8v, the lamp remained off. a hairline crack was found in the purge retainer and another on the top left side of the console; however, these were unrelated to the reported failure." root cause: the root cause of the loss of placement signal was a lamp malfunction. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b7 other relevant history, including preexisting medical conditions updated. d2 common device name updated. d4 model number updated. d10 concomitant medical products and therapy dates updated. f6 and f8 should have been left blank on the initial report. g1 reporting contact email and reporting contact fax number updated.
32679 | 1220648-2023-05064 | 18190897 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18190897&pc=OZD | 10/28/2023 | 01/25/2024 |  | 11/22/2023 |  |  | 10/28/2023 | 11/22/2023 |  |  | 06/30/2025 |  | 11/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024370698 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/28/2023 |  |  |  |  | The user facility reported a 55-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella access site began to bleed and a hematoma formed. dressing was removed and pressure was held but hemostasis was not achieved. physician decided to remove the impella to not risk groin complications. pump removed and perclose sutures tightened. site is dry with no further hematoma. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed. the device was not returned for benchtop investigation. according to the clinical information provided, patient had hematoma formed around the access site after insertion of gwru in cath lab with all best practices being followed. the cause of the access site bleeding issue was most likely due to lack of vessel recoil onto the sheath per clinical review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. per the ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32680 | 1220648-2023-05068 | 18192025 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18192025&pc=OZD | 10/27/2023 | 03/20/2024 |  | 11/22/2023 |  |  | 10/27/2023 | 11/22/2023 |  |  | 08/31/2025 |  | 11/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024391657 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | Pumping Stopped | Improper Flow or Infusion |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/27/2023 |  |  |  |  | The user facility reported a 66-year-old female (race unknown) post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. the pump was supporting for over two days when it was explanted after purge leak and high purge pressure observations. follow up with team determined the pump stopped. the impella 5.5 was explanted and the patient remained stable. | [5 empty] | The investigation for the reported high purge pressure, pump stop, and purge leak issues has been completed. the pump was returned for investigation. as per the clinical information, a purge leak was noted from the red handle, and the doctor explanted the pump after a pump stop event. the returned product has a kink near the motor housing and significant biomaterial on the impeller and purge gap. the red handle came wrapped with white tape, which was opened to see the origin of the leak. the purge leak was found coming from the catheter down the motor line. upon further investigation of the purge line, a crack was observed near the motor housing, where the kink was observed on the catheter. also, blood was observed inside purge line and catheter near motor housing. data logs show a purge leak trend from the beginning of the case. they also confirmed purge pressure, followed by saturated flow and a pump stop with several motor current high alarms. the cause of the purge leak was determined to be damaged purge tubing. a kink was noted near the motor housing, where the purge line was found damaged inside the catheter. this allowed purge to leak inside the catheter and reach the red handle. the cause of the pump stop issue was determined to be biomaterial. an open purge line allowed biomaterial and blood ingress into the purge gap, reaching the motor housing and causing clotting in the motor. h6 impact code updated to 4629. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use for the related events are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
32681 | 1220648-2023-05069 | 18192077 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18192077&pc=OZD | 10/26/2023 | 09/19/2024 |  | 11/22/2023 |  |  | 10/26/2023 | 11/22/2023 |  |  | 05/31/2025 |  | 11/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024353389 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Material Separation | Adverse Event Without Identified Device or Use Problem | [5 empty] | Air Embolism | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/26/2023 |  |  |  |  | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 49-year-old male post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. the patient was being transferred to the tertiary center upmc via helicopter. during the transport, the team noted the cable had "bumped into something from the side" and the cable became damaged and separated. additionally there was damage to the purge which allowed for air to embolize to the heart. the embolization was seen on transesophageal echocardiography in the or during pump replacement. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-05069 was provided in sections b2, b5, d6a, d6b, h1, and h6. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the reported product damage and embolism has been completed. the impella device was returned for evaluation and found that that cable plugs (gold pins) were separated from the cord suggesting that excessive force was used. the two halves of the white cable connectors showed evidence of sufficient potting indicating that they were bonded properly and that excessive force resulted in the damaged connector. as a result, the root cause of the cable separation issue was use-related as excessive force was used, which subsequently caused a pump stop. according to clinical data provided, it appeared that the purge tubing was severed and not clamped which lead to air noted in heart on trans-esophageal echo in operating room. therefore, the root cause of the embolism issue was use related to improper technique.
32682 | 1220648-2023-05070 | 18192175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18192175&pc=OZD | 10/26/2023 | 08/15/2024 |  | 11/22/2023 |  |  | 10/26/2023 | 11/22/2023 |  |  | 05/31/2025 |  | 11/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024352412 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Difficult to Remove | Difficult to Advance | Material Twisted/Bent |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/26/2023 |  |  |  |  | The user facility reported a 45-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. difficulty was met during attempted impella 5.5 implant. due to the patient's anatomy, resistance was met when attempting to pass the cannula past the anastomosis. further observation noted that the complications were attributed to a kink in the guidewire. after some troubleshooting, the impella was advanced into the left ventricle; however, the guidewire became stuck in the pump and could not be removed. both components were subsequently removed and some damage was seen on the catheter above the motor hub. a new wire and pump were subsequently used for implant without further complication. there was no patient harm. | [5 empty] | The investigation of product damage (guidewire damage - peripheral vessels) and delivery issues- use has been completed. the pump was not returned for investigation. according to the clinical details, the pump was unable to pass after the innominate artery due to observed guide wire kink. the pump and guide wire were removed together because of the kink. additionally, the catheter above the motor housing was found to be damaged. the doctor used a new pump. the patient was obese, with a reported vessel size of less than 7 mm. the cause of the delivery issue was most likely related to the patient condition, based on the provided clinical details. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ f6/f8-should have been left blank on manufacturer device report 1220648-2023-05070. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32683 | 1220648-2023-05066 | 18192190 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18192190&pc=OZD | 10/26/2023 | 01/12/2024 |  | 11/22/2023 |  |  | 10/26/2023 | 11/22/2023 |  |  | 05/31/2025 |  | 11/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024352360 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported being unable to pass the impella 5.5 device through the graft/anastomosis site of the 54-year-old patient due to significant resistance. despite multiple attempts, insertion was aborted. there was no reported patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical information provided was sufficient to determine the root cause. clinical information states that multiple attempts to advance were unsuccessful and the 5.5 implant was aborted. the root cause of the delivery anatomy issue was most likely patient condition related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
32684 | 1220648-2023-05072 | 18192265 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18192265&pc=OZD | 10/26/2023 | 09/20/2024 |  | 11/22/2023 |  |  | 10/26/2023 | 11/22/2023 |  |  | 05/31/2025 |  | 11/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024356549 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/26/2023 |  |  |  |  | The patient was successfully weaned from the impella device, however did not survive. per medical safety officer review use of the device cannot conclusively be associated with the outcome. | The user facility reported a 75-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. there was bleeding present at groin impella access site. dressing was taken down and manual pressure for 30 minutes, fem stop back in place. systemic anticoagulant was on hold until bleeding was better controlled. patient is stable and was on support for additional 5 days post the issue. |  |  |  |  | The investigation for access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b5 added patient outcome and mso statement the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05072. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact fax number missing. h6 component code revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32685 | 1220648-2023-05073 | 18192326 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18192326&pc=OZD | 10/23/2023 | 01/31/2024 |  | 11/22/2023 |  |  | 10/24/2023 | 11/22/2023 |  |  | 06/30/2025 |  | 11/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024377956 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 75-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that a patient on impella support had to be given cpr and as a result the patient experienced bleeding from multiple sticks to the artery. impella had to be explanted in order to repair the damaged artery and stop the bleeding. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. clinical details states that there were multiple sticks to the artery from attempting arterial access while active cpr was being done is the culprit of the bleed. therefore, the root cause of the access site bleed was most likely due to improper technique was used to puncture the access site. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32686 | 1220648-2023-05074 | 18192344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18192344&pc=OZD | 10/24/2023 | 01/11/2024 |  | 11/22/2023 |  |  | 10/24/2023 | 11/22/2023 |  |  | 06/30/2025 |  | 11/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024360203 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a 74-years-old male patient with multiple cardiac comorbidities being worked up for surgery. the team attempted to place access and the impella cp via the right femoral but due to calcium the team instead placed the impella via the left femoral and the percutaneous coronary intervention (pci) was done via the right femoral access. the long sheath was damaged/cracked in the access and was found to be allowing blood loss. the team still proceeded with the placement of the cp and the pci. | [5 empty] | The investigation for the reported introducer damage has been completed. the device was not returned for evaluation. clinical details states the long sheath was damaged/cracked in the access and was found to be allowing blood loss. however, without the device to determine the location of the damage, we cannot determine the root cause of the reported introducer sheath damage. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
32687 | 1220648-2023-05077 | 18192548 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18192548&pc=OZD | 10/11/2023 | 06/06/2024 |  | 11/22/2023 |  |  | 10/11/2023 | 11/22/2023 |  |  | 06/30/2025 |  | 11/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024372517 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/11/2023 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. complainant in japan reported the patient was moving a lot and the impella slipped out of the left ventricle and was reinserted, but it was not able to be placed in the right position. the patient was weaned off of the impella. | [5 empty] | The investigation for the reported positioning issue has been completed. according to the clinical, on the second day of impella cp support the patient began to move and the cannula slipped out of the lv. repositioning was attempted but an optimal position could not be found. the decision was then made to remove the pump. no product was returned for a visual inspection. data log analysis was not necessary for this complaint. the most likely cause of the positioning issues was patient movement. d4-updated model number added- d6a/d6b. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32688 | 1220648-2023-05076 | 18192553 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18192553&pc=OZD | 10/25/2023 | 09/03/2024 |  | 11/22/2023 |  |  | 10/25/2023 | 11/22/2023 |  |  | 08/31/2025 |  | 11/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2024391755 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/25/2023 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. md reports he cannot see a placement signal. the device's motor current was pulsatile and impella flow seen (no details values obtained). aic (automated impella controller) alert requested checking impella cannula for kinks. on fluoroscopy the impella cannula is well positioned in the lv. due to these problems, the doctor stopped the impella and then called the ccs. | [5 empty] | The investigation for the optical signal issue has been completed. upon investigation of the pump, optical parameters were found to be within specification and the pump passed the light continuity test and pressure test. the root cause of the optical signal issue was not determined since data logs were not returned. d6a and d6b added d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05076 d3-corrected mfr name, city, and state d4-corrected model number d6a and d6b added f6/f8 should have been left blank in the initial report.
32689 | 1220648-2023-05078 | 18192745 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18192745&pc=OZD | 10/26/2023 | 09/26/2024 |  | 11/22/2023 |  |  | 10/26/2023 | 11/22/2023 |  |  | 07/31/2025 |  | 11/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024392394 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/26/2023 |  |  |  |  | It was further reported that the patient expired sometime after the procedure. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that use of the device cannot conclusively be associated with the outcome. | The user facility reported a 58-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient had significant bleeding from around the impella repositioning sheath in the right femoral artery. physician performed cut down to area and had vascular physician/team obtain hemostasis around sheath with umbilical tape placed around artery and sheath. |  |  |  |  | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- b5 mso review, h6 component code 925 corrections to the initial mfr report: d3-corrected mfr name, city, and state. d4-corrected model number. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8 should have been left blank. g1-corrected mfr site name and address. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32690 | 1220648-2023-05079 | 18193001 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18193001&pc=PYX | 10/26/2023 | 04/16/2024 |  | 11/22/2023 |  |  | 10/26/2023 | 11/22/2023 |  |  | 02/28/2025 |  | 11/22/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024313906 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/26/2023 |  |  |  |  | The user facility reported a 52-year-old male noted as heart transplant rejection was implanted with an impella rp device for mechanical circulatory support. the medical team was able to insert the impella pump several times however the pump would not maintain the proper position. after a few failed attempts the medical team decided to explant the pump. the impella rp was replaced with and impella rp flex via right internal jugular vein. | [5 empty] | The investigation into the positioning issues has been completed since the original report was filed. the pump was returned, tested, and evaluated. the pigtail showed slight deformation, but no other issues were found with the returned product. the patient had a condition of obesity. the cause of the positioning issue was most likely patient condition related. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
32691 | 1220648-2023-05080 | 18193126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18193126&pc=OZD | 10/26/2023 | 01/12/2024 |  | 11/22/2023 |  |  | 10/26/2023 | 11/22/2023 |  |  |  |  | 11/22/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1292317 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 10/26/2023 |  |  |  |  | The user facility reported a 79-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. at the end of the case, the blue purge disc holder broke off the automated impella controller (aic) while removing the purge cassette. there was no patient harm. | [5 empty] | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were not relevant to the reported issue. the console was evaluated, and it was confirmed that the purge disc retainer had broken off on the console. the root cause of this issue was a broken purge retainer. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
32692 | 1220648-2023-05081 | 18193317 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18193317&pc=OZD | 10/26/2023 | 03/07/2024 |  | 11/22/2023 |  |  | 10/26/2023 | 11/22/2023 |  |  | 07/31/2025 |  | 11/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024370697 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the 58-year-old female was attempted to be implanted with an impella 5.5 but the pump would not start. there was a high motor current. therefore, the case was aborted. there was no reported patient harm. | [5 empty] | The investigation into the pump did not start issue has been completed since the original report was filed. the device was returned. product was returned for investigation. biomaterial was observed on the impeller, and a cannula kink was noted during the physical examination. the root cause of the pump did not start issue was an open electrical line, due to excessive force, found in red handle after monkey fist glue. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
32693 | 1220648-2023-05082 | 18193364 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18193364&pc=OZD | 10/03/2023 | 09/20/2024 |  | 11/22/2023 |  |  | 10/26/2023 | 11/22/2023 |  |  | 04/30/2025 |  | 11/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024345680 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/26/2023 |  |  |  |  | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 74-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient had access site bleeding which was resolved by adjusting the angle and refixing. blood products were given to address the issue as well. additionally, the patient was on impella support for 78.50 hours before the patient expired. the cause of death was determined to be acute heart failure. there was no relationship between the cause of death and the impella. |  |  |  |  | The investigation into the reported access site bleeding ¿ major has been completed since the original report was filed. per the clinical notes, the angle of insertion had to be adjusted to resolve access site bleeding. the cause of the access site bleeding was use related because the adjusted angle of insertion resolved the issue. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: bleeding access site ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32694 | 1220648-2023-05083 | 18193387 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18193387&pc=OZD | 10/26/2023 | 01/18/2024 |  | 11/22/2023 |  |  | 10/26/2023 | 11/22/2023 |  |  |  |  | 11/22/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1123952 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 10/26/2023 |  |  |  |  | The user facility reported that during the preparation the automated impella controller (aic) purge disc sensor was not registering that the purge disk was inserted. troubleshooting did not resolve the issue. the aic was exchanged and issue was resolved. no patient was involved. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the malfunction of the automated impella controller has been completed. the console and the data logs were returned for evaluation. the data logs were reviewed from the reported occurrence date and confirmed that the purge disc not detected alarm occurred over 15 times throughout the reported occurrence date. the console was visually inspected and it was evident that the purge flag was broken. further examination revealed that there was also a significant amount of buildup/residue around the purge assembly and behind the purge assembly on the battery bracket. the root cause of this issue is a broken purge flag. d9 updated.
32695 | 1220648-2023-05085 | 18193417 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18193417&pc=OZD | 10/26/2023 | 09/26/2024 |  | 11/22/2023 |  |  | 10/26/2023 | 11/22/2023 |  |  | 06/30/2025 |  | 11/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024370639 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/26/2023 |  |  |  |  | It was further reported that the patient expired while on support. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that use of the device cannot conclusively be associated with the outcome. | The user facility reported a xx-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. on day one of impella support the right groin was oozing. manual pressure was held for 20 minutes, and pressure dressing was applied. the second day pressure dressing was applied again. on the third day since the groin was still oozing fem stop was applied overnight and blood transfusion started. patient passed away later. the death is unrelated to the device. |  |  |  |  | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- h6 component code 925 corrections to the initial mfr report: added-b5 mso review: in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8 should have been left blank. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32696 | 1220648-2023-05084 | 18193442 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18193442&pc=OZD | 10/26/2023 | 09/11/2024 |  | 11/22/2023 |  |  | 10/26/2023 | 11/22/2023 |  |  | 06/30/2025 |  | 11/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024373479 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/26/2023 |  |  |  |  | The user facility reported a 74-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. femoral access was made and a leg immobilizer used on limb. the right limb was known to have a deep vein thrombosis and the impella access was on the left limb. after a matter of hours the pump was explanted due to the access site bleed and hematoma. additionally, manual compression held for 40 min and fem stop applied and left in place for over an hour to address the issue. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b.7 revised as information was inadvertently omitted on the initial manufacturer device report number 1220648-2023-05084. d.6a and d.6b revised from the initial submission of manufacturer device report number 1220648-2023-05084 in accordance with updated procedures. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2023-05084 and should not have been. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2023-05084 was submitted. h.6 component code was added since the initial manufacturer device report number 1220648-2023-05084 in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2023-05084 in accordance with updated procedures. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: bleeding access site ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32697 | 1220648-2023-05087 | 18194123 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18194123&pc=PYX | 10/27/2023 | 01/11/2024 |  | 11/22/2023 |  |  | 10/27/2023 | 11/22/2023 |  |  | 01/31/2025 |  | 11/22/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024305852 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/27/2023 |  |  |  |  | The user facility reported a 64-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. upon insertion and after pump turned on there was alarms of purge issues. high pressure ultimately escalated to purge block with less then 1 ml an hour of flow through the purge tubing troubleshooting did not help resolve the issue so that physician decided to proceed with tissue plasminogen activator (tpa). a couple of hours after there was impella stop alarm. all steps were taken at that point to try and restart but ultimately rp was explanted. patient was stable. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ | The investigation for the reported high purge pressures and a pump stop has been completed. the device was not returned for evaluation, however, the data logs were reviewed. review of the data logs revealed there were no indicative motor current spikes, as it was flowing normally for 3 minutes before the pump stopped. before the high motor current pump stop, the purge pressure was dropping. after the high motor current pump stop, it was able to restart. purge pressure began to rise and 9 minutes later and the high purge pressure alarm was triggered and purge system blocked. there are no indicative signs of ingestion or evidence of a long bag change. while the high purge pressure and pump stops were confirmed via the data logs, without the device or sufficient clinical details, the root cause of the high purge pressure and pump stop cannot be determined.
32698 | 1220648-2023-05071 | 18194138 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18194138&pc=OZD | 10/02/2023 | 10/18/2024 |  | 11/22/2023 |  |  | 10/06/2023 | 11/22/2023 |  |  |  |  | 11/22/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, JP |  |  | ABIOMED, INC. |  |  | US |  | 1383039 |  | 0042-0000-JP |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Failure to Prime | [6 empty] | Multiple Organ Dysfunction Syndrome | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 10/24/2023 |  |  |  |  | As reported on mdr 1220648-2023-04787: the user facility reported a 67-year-old male in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. during the priming step, the automated impella controller (aic) blacked out. priming appears to have continued during the blackout. a reboot was attempted but a controller error alarm appeared on the screen. the aic was replaced. on (b)(6) 2023, the user facility reported the patient was implanted with an impella cp once the new aic was in place. during support, the patient expired due to multiple organ failure. it is unknown if the aic screen black out during priming causing a delay in procedure caused or contributed to the patient¿s death. there is no allegation against the impella cp and the patient's death. there were no problems with the operation of the impella cp. | [5 empty] | The investigation for manufacturer device report 1220648-2023-05071 will be submitted in manufacturer device report 1220648-2023-04787. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32699 | 1220648-2023-05065 | 18194507 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18194507&pc=OZD | 10/27/2023 | 03/14/2024 |  | 11/22/2023 |  |  | 10/27/2023 | 11/22/2023 |  |  | 08/31/2025 |  | 11/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  |  |  | 2024380542 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/27/2023 |  |  |  |  | The user facility reported a 76-year-old male undergoing coronary artery bypass grafting surgery was implanted with an impella 5.5 device for mechanical circulatory support. on day three of support, there was a high purge pressure alarm. the impella 5.5 was explanted and the patient was alert and doing well. follow up with the team determined the cannula was kinked. | [5 empty] | The investigation for the reported high purge pressure issue and the cannula kink issue has been completed. the impella device was not returned for evaluation. however, data logs and clinical information provided was sufficient to determine the root cause. data logs were reviewed finding a 7-minute gradual rise in purge pressure followed by saturated pump flow. this condition was resolved after 30 minutes. the cause of the high purge pressure issue was most likely biomaterial based on data log review. the instructions for use states ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ added health effect impact codes: 4644, 4627. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32700 | 1220648-2023-05075 | 18194592 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18194592&pc=OZD | 10/24/2023 | 03/06/2024 |  | 11/22/2023 |  |  | 10/24/2023 | 11/22/2023 |  |  | 07/31/2025 |  | 11/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024390111 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 71-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant had a cp support ongoing with signs of hemolysis in a patient admitted for percutaneous coronary intervention (pci). the pci was multi-vessel with the cp supporting the patient's hemodynamics. it was noted that providers were aware and working to wean the patient off impella support. | [5 empty] | The investigation into the hemolysis issue has been completed since the original report was filed. the device was not returned for investigation. according to the clinical details, hemolysis was resolved with pump reposition at the start of the case. the cause of the hemolysis issue was improper positioning of the device based on data log review and provided clinical information. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
32701 | 1220648-2023-05086 | 18194625 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18194625&pc=OZD | 10/27/2023 | 01/26/2024 |  | 11/22/2023 |  |  | 10/27/2023 | 11/22/2023 |  |  | 07/31/2025 |  | 11/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024377020 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported the 52-year-old male patient was agitated in bed and ripped the red impella 5.5 plug and broke the impella in half. the grey was broken on the proximal portion of the red impella plug. team brought to bedside and emergently explanted at bedside. patient was stable and the device was not re-implanted. | [5 empty] | The investigation into the reported product damage has been completed since the original report was filed. product was returned for investigation. the pump was damaged from the kink protector area inside the red handle. the bond joining the catheter to the the kink protector was intact. data logs were returned and there was a motor current spike along with impella disconnected while running alarm further confirming instant damage. per the given clinical information, the root cause of the pump stop issue was an open electrical line due to excessive force used. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
32702 | 1220648-2023-05092 | 18194877 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18194877&pc=OZD | 10/25/2023 | 11/19/2025 |  | 11/22/2023 |  |  | 10/27/2023 | 11/22/2023 |  |  |  |  | 11/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1352449 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Power Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 10/27/2023 |  |  |  |  | The user facility in japan reported that after a procedure was complete, the team tried to disconnect the connection cable and turn off the automated impella controller (aic) power, but the aic had to be shut down instead. when the aic was restarted to try to troubleshoot, an alarm stating "controller error" appeared. the aic was to be returned for evaluation and/or repair. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
32703 | 1220648-2023-05088 | 18194967 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18194967&pc=OZD | 10/28/2023 | 01/12/2024 |  | 11/22/2023 |  |  | 10/28/2023 | 11/22/2023 |  |  | 11/30/2024 |  | 11/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023277474 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/28/2023 |  |  |  |  | The user facility reported a 53-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. while prepping the impella 5.5 device alarm ¿purge blocked¿ was active. no purge fluid was seen exiting the pump, but the physician wanted to proceed with the use of the pump. after the pump was inserted within 3 seconds there was a pump stop alarm. the pump was removed and a new 5.5 was inserted without any issue. | [5 empty] | The investigation for the reported high purge pressure and pump stop has been completed. the impella device was not returned for evaluation. however, data logs were reviewed and it can be seen that there was the "purge system blocked and high purge pressure" alarm. additionally the pump was attempted to be started and it would not start as they were getting "impella stopped - motor current high (hmc)." the logs also revealed that priming was able to be completed with no issues and the purge fluid used was recommended. the root cause of the pump did not start and high purge pressure cannot be determined due to lack of product returned. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella
32704 | MW5148465 | 18204930 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18204930&pc=OZD | 11/11/2023 | 11/20/2023 |  | 11/22/2023 |  |  |  | 11/22/2023 |  |  |  |  | 11/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | ABIOMED IMPELLA DEVICE |  |  | OZD |  |  | IPX1 |  |  | ABIOMED INC. | [8 empty] | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Voluntary report | Yes | Risk Manager | [6 empty] | 69 YR |  | Male | 73 KG | Non Hispanic | White | Malfunction | No | Yes | [7 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | [9 empty] | Impella console malfunctioning. device was given back to abiomed impella device rep named (b)(4) who can be reached at (b)(4).
32705 | MW5148470 | 18205045 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18205045&pc=OZD | 03/28/2023 | 11/20/2023 |  | 11/22/2023 |  |  |  | 11/22/2023 |  |  |  |  | 11/22/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 WITH SMARTASSIST |  |  | OZD | [5 empty] | ABIOMED, INC. | [8 empty] | No | [6 empty] | Health Professional |  |  | Voluntary report | Not Available | Patient | [6 empty] | 52 YR |  | Male | 100 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | [7 empty] | Exposure to Body Fluids | Pain | Anxiety | Numbness | [9 empty] | Disability | Other | Life Threatening | Hospitalization | Required Intervention |  | AMLODIPINE 2.5 MG TABLET | ATORVASTATIN 10 MG TABLET | B-D U/F PEN NEEDLE 31G X 8 MM MISC | CALCIUM CARBONATE-VITAMIN D 500-200 MG-UNIT TABS PER TABLET | CLONAZEPAM 0.5 MG TABLET | DEXCOM G7 SENSOR | INSULIN LISPRO (1 UNIT DIAL) 100 UNIT/ML PEN-INJECTOR | LANSOPRAZOLE 30 MG DR CAPSULE | LANTUS SOLOSTAR 100 UNIT/ML PEN INJECTION | MAGNESIUM OXIDE 400 MG TABLET | METFORMIN 500MG | MULTIPLE VITAMIN (MULTIVITAMIN) TABLET | MYCOPHENOLATE MOFETIL 500 MG TABLET | OLMESARTAN 40 MG TABLET | OZEMPIC (1 MG/DOSE) 4 MG/3ML SOPN | SILDENAFIL 20 MG TABLET | TACROLIMUS 1 MG CAPSULE | TACROLIMUS 5 MG CAPSULE | [5 empty] | Initial submission | [9 empty] | Abiomed - impella 5.5 with smartassist purge fluid leak after spending approximately two months in the hospital at the (b)(6) hospital for heart failure, i was placed on the abiomed impella 5.5 with smartassist. this medical device was implanted because i was awaiting a heart transplant, and my prognosis was poor. on approximately (b)(6) 2022, the abiomed impella 5.5 with smartassist purge fluid began to leak from the purge sidearm of the pump. the purge leak impacted the impella 5.5 system, and low purge pressures occurred. alarms went off and i notified the medical staff. when the medical staff arrived, i informed them that the purge fluid was leaking from the machine, and it made my gown and sheets wet. i also informed them that the purge pressures had significantly dropped. they tried to repair the leak by rapping it with a maxi-pad. unfortunately, this did not solve the problem and they continued to try to repair the machine for approximately 2-weeks. the machine continued leaking for two weeks causing me a great deal of stress, knowing this machine was keeping me alive. i was very upset and did not understand why the medical device was failing. the low pressures continued for about two weeks and the hospital contacted the abiomed representative, and they could not resolve the leakage. they informed me that it was too risky to swap the impella 5.5 machine for another machine, so they kept changing the gauze they eventually used. they also used bone wax to try to remedy the leakage to increase the purge pressures. this helped initially, but eventually did not solve the problem. the purge pressures kept going up and down, putting me at great risk for injury and death. it was clear that the impella 5.5 had defective components and device malfunctions that caused a life-threatening situation. on (b)(6) 2022, a donor was identified, and i went through a heart transplant. following the heart transplant, i began to experience complications, and i was placed on an ecmo machine (extracorporeal membrane oxygenation) for approximately 5 days. i soon regained consciousness and i realized that my left foot was significantly numb, and i was in a lot of pain. the medical team treated the numbness and pain with gabapentin, but this did not help symptoms or relieve the pain. i continued to experience severe numbness in my foot and pain for another 1 1/2 month, until i was discharged from the hospital. following discharge, i was treated by a physical therapist, neurologist and a pediatrist. my numbness continues, and my pain is getting worse in my left foot. currently, the pediatrist is treating the numbness and pain in my foot with cortisone injections, but this has not helped. i realize that after numerous treatments over the past 1 1/2 years, my numbness and pain is getting worse. it has caused a substantial disruption in my ability to conduct normal life functions (i.e., walking, running & sleeping). it has also caused me a great deal of anxiety, which i am being treated for by my primary care provider. the entire incident with the impella 5.5 leakage has been traumatic and could have been prevented. after conducting research, i discovered that other patients that experienced purge fluid leaks with impella 5.5, are experiencing similar problems, disability, and permanent damage. i am reporting this problem because my condition is getting worse, and i need help treating this problem. i also hope this report will lead to product improvements that will prevent disability and permanent damage in other patients. dr. (b)(6).
32706 | 1220648-2023-05024 | 18181199 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18181199&pc=OZD | 10/26/2023 | 02/08/2024 |  | 11/21/2023 |  |  | 10/26/2023 | 11/21/2023 |  |  | 08/31/2025 |  | 11/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024388744 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that impella 5.5 had two pump stops. the clinician was able to restart the pump both times. there was no reported patient harm. | [5 empty] | The investigation for the reported pump stop has been completed. the impella device was returned for evaluation finding it had been cut open close to the motor housing. the device was visually inspected, and biomaterial was found over the outflow cage. the catheter was observed to be punctured near 30 cm marking with hole in the catheter. blood ingress was observed in the catheter, but clear in the purge line. no other abnormalities were found. data logs were reviewed, and several 119 pump stop alarms were observed at the end. pump was able to restart after several occurrences of 119 pump stop. no motor current spike was observed to align with biomaterial ingestion. the cause of the pump stop issue was most likely blood ingress due to a hole in the catheter resulting from improper use per field investigation and log review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
32707 | 1220648-2023-05025 | 18185097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18185097&pc=OZD | 10/31/2023 | 11/21/2023 |  | 11/21/2023 |  |  | 10/31/2023 | 11/21/2023 |  |  | 05/31/2025 |  | 11/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024369992 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/31/2023 |  |  |  |  | The user facility reported a 66-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was bleeding from the impella access site during support. the physician was required to re-suture the catheter to the patient and adjust the angle by using gauze and the issue was resolved. | [5 empty] | The investigation into the reported access site bleeding has been completed. the device was not returned by the medical facility. as per clinical, repositioning sheath was not sutured close enough to the skin at time of bleeding. bleeding issue was resolved after re-suturing and angle matched the dressing. the cause of the access site bleeding issue was related to the use of suturing of the repositioning sheath, based on given clinical information. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32708 | 1220648-2023-05026 | 18185219 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18185219&pc=OZD | 11/17/2023 | 04/15/2024 |  | 11/21/2023 |  |  | 11/17/2023 | 11/21/2023 |  |  | 06/30/2025 |  | 11/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024371779 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/17/2023 |  |  |  |  | The user facility reported a 84-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp support ongoing with the peel-away sheath retained in the femoral. the right leg was dusky, cold, and numb with lost pulses. retention of the sheath goes against the ifu recommendations. the team placed distal antegrade sheaths and observed the sheath sidearm to be thrombosed. the sheath was retained and patient transported to tertiary center. limb maintained and support continues. | [5 empty] | The investigation for the reported limb ischemia and thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia and thrombus could not be determined as the device was not returned nor sufficient clinical details. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32709 | 1220648-2023-05027 | 18185241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18185241&pc=OZD | 10/31/2023 | 02/12/2024 |  | 11/21/2023 |  |  | 10/31/2023 | 11/21/2023 |  |  |  |  | 11/21/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1186350 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure of Device to Self-Test | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 10/31/2023 |  |  |  |  | The user facility reported that during a case preparation, the automated impella controller, (aic) had a error message that system self check failed - change console immediately. troubleshooting was performed with no resolution. a replacement loaner was sent to the facility. no patient is involved. | [5 empty] | The investigation into the aic - system self-check error has been completed since the original report was filed. the product was returned for investigation. no console logs were available at the time of the issue occurrence because the self check failure prevented the console from booting up. therefore, console logs were not used in this investigation. the console was tested and physically inspected after being returned to field service. a firmware check was performed, which revealed that there was a lack of communication between the 4-in-1 and the battery 1 circuitry on the electronic control board (pbm). the cause of the aic issue was contamination. there was corrosion of the pbm traces from leakage of electrolyte of the supercapacitors. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
32710 | 1220648-2023-05028 | 18185401 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18185401&pc=OZD | 10/31/2023 | 08/20/2024 |  | 11/21/2023 |  |  | 10/31/2023 | 11/21/2023 |  |  | 02/28/2025 |  | 11/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024325764 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/31/2023 |  |  |  |  | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant in the us had a cp come out of position while in the ct suite for imaging. the pump was explanted and replaced. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the root cause of the positioning issue was patient movement. d6a and d6b added the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05028: d4 model number, f6/f8-should have been left blank, g1 reporting contact fax number missing. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32711 | 1220648-2023-05029 | 18185518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18185518&pc=OZD | 10/25/2023 | 05/07/2024 |  | 11/21/2023 |  |  | 10/25/2023 | 11/21/2023 |  |  | 07/31/2025 |  | 11/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024387378 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/25/2023 |  |  |  |  | The user facility reported a 58-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella cp was placed during cardiac catheterization laboratory percutaneous cardiac intervention and a coronary artery perforation occurred. patient was taken for emergent coronary artery bypass grafting (cabg)/left anterior descending artery perforation containment. impella catheter fractured and malfunctioned when cross clamped with soft aortic clamp. the decision was made to place a second impella post cabg procedure. | [5 empty] | The investigation into the product damage (hole in catheter) has been completed since the original report was filed. the pump was returned, evaluated, and visually inspected and no catheter kink, biomaterial or any damage to impeller outflow or blades were found. according to the clinical, shortly after beginning impella cp support a catheter fracture was discovered. immediately after cross-clamping the aorta with soft clamps a pump stop occurred. the cause of the pump stop issue was use related per the returned product investigation, the data log, and the clinical review. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
32712 | 1220648-2023-05030 | 18185539 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18185539&pc=OZD | 10/30/2023 | 09/10/2024 |  | 11/21/2023 |  |  | 10/30/2023 | 11/21/2023 |  |  | 09/30/2025 |  | 11/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024394994 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/30/2023 |  |  |  |  | The user facility reported a 47-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the pump was placed and after 6 days the team observed access site bleeding. the team made decision to hold heparin/pause anticoagulation. the patient remains on support while awaiting surgery. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was anticoagulation management due to high patient act values; heparin was stopped to achieve hemostasis as per the clinical description. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (unites states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿ added- h6 component code added- d6a/d6b f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32713 | 1220648-2023-05031 | 18185601 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18185601&pc=OZD | 11/18/2023 | 10/30/2024 |  | 11/21/2023 |  |  | 11/18/2023 | 11/21/2023 |  |  | 03/31/2024 |  | 11/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023199493 |  | 005048 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 13 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/18/2023 |  |  |  |  | The user facility in japan reported a 13-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant in japan had a 2.5 placed for support along with ecmo (extracorporeal membrane oxygenation). the team had the pump come out of support after just over 3 hours of support. the pump could not be re-delivered across to the lv (left ventricle¿). the pump was explanted. the ecmo supported the patient without harm or injury. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05031. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact fax number missing. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32714 | 1220648-2023-05032 | 18185760 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18185760&pc=OZD | 10/27/2023 | 01/22/2024 |  | 11/21/2023 |  |  | 10/27/2023 | 11/21/2023 |  |  | 08/31/2025 |  | 11/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024397336 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had impella cp support ongoing for a female patient admitted for an ami/cgs and in need of hemodynamic support. the patient suffered limb ischemia in the impella accessed limb. the pulses were lost and so the team placed a fem-fem bypass. | [5 empty] | The investigation for the report limb ischemia has been completed. no device was returned for evaluation. as per clinical information, the patient had distal limb ischemia and fem-fem external bypass resolved the ischemia. the cause of the limb ischemia issue was patient condition related with fem-fem bypass resolving ischemia on the impella leg per clinical. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32715 | 1220648-2023-05033 | 18187519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18187519&pc=OZD | 10/26/2023 | 01/11/2024 |  | 11/21/2023 |  |  | 10/26/2023 | 11/21/2023 |  |  | 08/31/2025 |  | 11/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024392514 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/26/2023 |  |  |  |  | The user facility reported a 54-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. during the process of advancing the impella through the subclavian artery, significant resistance was at the bend of the subclavian. multiple attempts to advance past the bend were unsuccessful and the 5.5 implant was aborted due to fear of damaging the artery. | [5 empty] | The investigation into the delivery issues- anatomy has been completed. the device was not returned for benchtop evaluation and investigation. per the clinical information provided, the root cause of the delivery issue was patient¿s condition related due to stenosis in the vessels at the bend of the subclavian artery. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32716 | 1220648-2023-05036 | 18187685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18187685&pc=OZD | 10/23/2023 | 07/24/2024 |  | 11/21/2023 |  |  | 10/23/2023 | 11/21/2023 |  |  | 05/31/2025 |  | 11/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024346434 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the 65-year-old male patient's access site was swollen and painful following placement of impella 5.5 as a transplant workup. the site was drained of old clot/hematoma, did not look infectious, and was sutured shut after hematoma evacuation. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use referenced in the initial report is for the impella 5.5 with smart assist for use during cardiogenic shock in the section: potential adverse events (united states), which now includes the statement "cardiac or vascular injury (including ventricular perforation)¿. added- h6 component code 925 in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32717 | 1220648-2023-05035 | 18187686 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18187686&pc=OZD | 10/23/2023 | 02/28/2024 |  | 11/21/2023 |  |  | 10/23/2023 | 11/21/2023 |  |  | 04/30/2025 |  | 11/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024349496 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that in icu the 71-year-old female patient¿s bilateral feet were purple. impella cp access side (right lower extremity) had absent pedal pulse but present pt pulse via doppler. the lower left extremities (lle) (non-pump side) also had absent pedal and pt pulses. the plan is to use a warming blanket and conduct frequent pulse checks. | [5 empty] | The investigation into the reported limb ischemia -reversible has been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32718 | 1220648-2023-05034 | 18187762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18187762&pc=OZD | 10/22/2023 | 02/28/2024 |  | 11/21/2023 |  |  | 10/22/2023 | 11/21/2023 |  |  | 06/30/2025 |  | 11/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024379717 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/22/2023 |  |  |  |  | The user facility reported a 33-year-old female with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support after the patient began to decompensate post trachial dilation. the patient developed limb ischemia secondary to percutaneous right axillary approach. the medical team decided to explant the pump and later replaced it with another impella cp with left axillary access. the patient was in critical condition and was being closely monitored. | [5 empty] | The investigation into the reported limb ischemia -reversible has been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32719 | 1220648-2023-05037 | 18187921 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18187921&pc=OZD | 10/23/2023 | 08/26/2024 |  | 11/21/2023 |  |  | 10/23/2023 | 11/21/2023 |  |  | 07/31/2025 |  | 11/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024379206 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Improper Flow or Infusion | Positioning Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/23/2023 |  |  |  |  | The user facility reported a 66-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, the impella relayed suction alarms. transesophageal echocardiography showed inlet interfering with mitral valve. the following day the physician was able to successfully reposition, but the pump started displaying signs of flow saturation. decision was made to remove impella 5.5 at bedside. it was noted by the cardiac surgeon that the patient's heart had evidence of cardiac tissue erosion from where the impella was mispositioned. the patient is stable. | [5 empty] | The investigation for the positioning issue and the saturated flow issues has been completed. according to the clinical description, the patient began experiencing position related suction alarms. transthoracic echocardiogram confirmed the inlet was interfering with the mitral valve. multiple unsuccessful attempts were then made to reposition the pump. however, following one more attempt the next day the physician was able to successfully reposition the pump. shortly after the pump was repositioned, the pump began displaying signs of flow saturation. the decision was then made to remove the pump. it was later noted that the patient's heart had evidence of cardiac tissue erosion from where the pump was mispositioned. product evaluation confirmed biomaterial on the impeller and no other abnormalities. additionally, when attempting to reproduce the saturated flow in the lab the pump stopped three times. data log analysis confirmed multiple positioning issues throughout support. high purge pressure and purge system blocked alarms were also recorded on the second day of support. the cause of the saturated flow was biomaterial. the cause of the positioning issues was not determined due to lack of clinical details. b.1 adverse event was selected incorrectly on the initial manufacturer device report number 1220648-2023-05037. b.7 revised as information was inadvertently omitted from the initial manufacturer device report number 1220648-2023-05037. d.4 revised model number from the initial submission of manufacturer device report number 1220648-2023-05037 in accordance with updated procedures. d.6a and d.6b revised from the initial submission of manufacturer device report number 1220648-2023-05037 in accordance with updated procedures. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2023-05037 and should not have been. h.6 code 4628 was added as it was inadvertently omitted from the initial manufacturer device report number 1220648-2023-05037. | The impella device was received from the customer and an evaluation is pending. upon investigation closure, a supplemental mdr will be filed.
32720 | 1220648-2023-05038 | 18187927 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18187927&pc=OZD | 10/23/2023 | 01/23/2024 |  | 11/21/2023 |  |  | 10/23/2023 | 11/21/2023 |  |  | 08/31/2025 |  | 11/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024394661 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/23/2023 |  |  |  |  | The user facility reported a 50-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during implant and after the impella went around the aortic arch, it became kinked on the cannula. staff attempted to move to unkink and was unable to straighten. device was removed, cleaned and re-implanted and the kink occurred again. the team was unable to go above p3 without suction event. patient placed on ecmo and a new impella cp was placed without any complications. there was no patient harm. | [5 empty] | The investigation into the product damage (cannula kink) has been completed since the original report was filed. the device was returned for investigation. per the investigation and the clinical information provided, the root cause of the low pump flow issue was a kinked cannula due to small/ diseased vessel. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
32721 | 1220648-2023-05039 | 18187974 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18187974&pc=OZD | 10/23/2023 | 07/31/2024 |  | 11/21/2023 |  |  | 10/23/2023 | 11/21/2023 |  |  | 08/31/2025 |  | 11/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024387871 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/23/2023 |  |  |  |  | The user facility reported a 45-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support as a bridge to transplant. during support, heparin was stopped due to the patient experiencing lots of oozing from the impella access site. the patient remained on impella support for 19 days and was weaned successfully. | [5 empty] | The investigation for access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding is determined to be anti-coagulation management. the heparin was stopped, and the patient healed later. h6 component code added the following have been reported incorrectly or missing from manufacturer device report: in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8-should have been left blank. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32722 | 1220648-2023-05040 | 18188004 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18188004&pc=OZD | 10/24/2023 | 01/26/2024 |  | 11/21/2023 |  |  | 10/24/2023 | 11/21/2023 |  |  | 09/30/2025 |  | 11/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 JP |  |  | OZD |  |  | IMPELLA 5.5 JP |  |  | ABIOMED, INC. |  |  | US |  | 2024393870 |  | 1000211 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Difficult to Advance | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that insertion of the impella 5.5 was difficult due to the 60-year-old male patient¿s vascular anatomy. the impella was therefore removed and there was a kink in the proximal part of the motor. a second try of reinserting the impella and a purge related alarm occurred, and even after troubleshooting, there was no improvement. a new impella was used. there was no harm to the patient. | [5 empty] | The investigation into the delivery issues- use has been completed. the device was returned for benchtop investigation. the pump was visually inspected, and catheter kink near motor housing was observed. the root cause of the delivery issue was patient condition related due to patient vessels condition. | The impella device was returned by the customer and an evaluation is ongoing. upon investigation closure, a supplemental mdr will be filed.
32723 | 1220648-2023-05041 | 18188021 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18188021&pc=OZD | 10/24/2023 | 01/26/2024 |  | 11/21/2023 |  |  | 10/24/2023 | 11/21/2023 |  |  | 07/31/2025 |  | 11/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024388296 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/24/2023 |  |  |  |  | The user facility reported a 66-year-old female in inferior myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the placement of the impella cp required the patient to go for a cutdown with vascular surgery for arterial repair and with blood product replacement required. the patient was escalated to impella 5.5. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation into the vascular damage - peripheral vessel has been completed. the device was not returned for investigation. per the provided clinical information, the placement of the cp required the patient to go for a cutdown with vascular surgery for arterial repair with blood product replacement due to the patient having a very high bmi and difficult anatomy. the root cause of the vascular damage - peripheral vessels is most likely due to the patient's condition.
32724 | 1220648-2023-05042 | 18188052 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18188052&pc=PYX | 10/24/2023 | 03/25/2024 |  | 11/21/2023 |  |  | 10/24/2023 | 11/21/2023 |  |  | 02/28/2024 |  | 11/21/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023188628 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Migration | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 69 year old male patient was implanted with an impella rp pump for mechanical circulatory support. it was documented that following implant, flows were unable to rise above 1 lpm. repositioning was completed; however, each time the pump was released following troubleshooting, the device would retract back into poor positioning. manual manipulations were attempted for over two hours, after which the decision was made to explant the pump. the patient was subsequently placed on a right ventricular assist device via centrimag. there was no patient harm as a result of the noted issue. | [5 empty] | The investigation into the positioning issues and the low pump flow issue has been completed since the initial report was submitted. the product was not returned for investigation. as per the clinical details, the doctor found the pump in a poor position every time after starting. the cause of the positioning issue was most likely due to use as the pump was able to be placed in the proper position after some time. the root cause of the low pump flow was undetermined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32725 | 1220648-2023-05043 | 18188094 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18188094&pc=OZD | 10/24/2023 | 02/16/2024 |  | 11/21/2023 |  |  | 10/24/2023 | 11/21/2023 |  |  | 08/31/2025 |  | 11/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024395133 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 70-year-old female noted as high risk coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the impella pump had to be swapped out due to unspecified positioning issues. there was no patient harm. | [5 empty] | The investigation into the positioning issues has been completed. the device was not returned for investigation. the cause of the positioning issue was not determined since product was not returned and sufficient clinical information was not provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32726 | 1220648-2023-05044 | 18188100 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18188100&pc=OZD | 10/24/2023 | 06/19/2024 |  | 11/21/2023 |  |  | 10/24/2023 | 11/21/2023 |  |  | 05/31/2025 |  | 11/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024356493 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Migration | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/24/2023 |  |  |  |  | The user facility reported a 60-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that a patient became agitated after waking up and pulled their implanted impella cp pump into their aorta. as the patient was deemed to be hemodynamically stable, the decision was made to explant the pump. there was no patient harm as a result of the noted issue. | [5 empty] | The investigation for the positioning issue has been completed. according to the clinical, on the second day of impella cp support the patient woke up agitated and accidentally pulled the pump into the aorta. no attempt was made to reposition the pump. the patient was then gradually weaned off of support and the pump was removed. no product was returned for a visual inspection. data log analysis was not necessary for this complaint. the most likely cause of the positioning issues was patient movement. added- d6a/d6b. d4-corrected model number. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32727 | 1220648-2023-05045 | 18188155 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18188155&pc=OZD | 10/24/2023 | 08/01/2024 |  | 11/21/2023 |  |  | 10/24/2023 | 11/21/2023 |  |  | 08/31/2025 |  | 11/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024393210 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/24/2023 |  |  |  |  | It was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that use of the device cannot conclusively be associated with the outcome. | The user facility reported a 55-year-old male presented with acute myocardial infarction/cardiogenic shock and with acute stroke symptoms, hemiparesis with chest paint with significant waves. the patient was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced three episodes that prompted defib x3. the family made the decision to make the patient a dnr due to neuro status. the patient subsequently expired with no allegation or evidence of device association. |  |  |  |  | The investigation for the reported ventricular fibrillation/ventricular tachycardia (vf/vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf/vt could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- b5 mso review, h6 component code 925 in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. per the ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32728 | 1220648-2023-05046 | 18188173 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18188173&pc=OZD | 10/24/2023 | 09/19/2024 |  | 11/21/2023 |  |  | 10/24/2023 | 11/21/2023 |  |  | 08/31/2024 |  | 11/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023262334 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 65-year-old female patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that an implanted impella cp optical pump encountered persistent suction issues during patient support. this resulted in limited support levels and pump escalation was recommended. however, due to the patient overall condition, the decision was made to withdraw care. there was no patient harm as a result of the noted issue. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported low pump flow has been completed since the original report was filed. the pump was not returned for investigation. data logs show low pump flow throughout the entire case with an active suction alarm. two motor current spikes were noted during the case. the cause of the low pump flow issue was most likely biomaterial based on data log review. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-05046 was provided.
32729 | 1220648-2023-05047 | 18188188 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18188188&pc=OZD | 10/24/2023 | 10/24/2025 |  | 11/21/2023 |  |  | 10/24/2023 | 11/21/2023 |  |  |  |  | 11/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1126094 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Activation Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The hospital's biomed department reported that their aic triggered an unspecified error message that called for a replacement. there was no reported patient harm. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | A.2 age should have been left blank on the initial manufacturer device report number 1220648-2023-05047 as it is unknown. a.3 unknown should have been selected on the initial manufacturer device report number 1220648-2023-05047 as it is unknown. d.1 and d.2 brand name and common device name were revised in accordance with updated procedures since the initial manufacturer device report number 1220648-2023-05047 was submitted. d.4 model number were revised in accordance with updated procedures since the initial manufacturer device report number 1220648-2023-05047 was submitted. serial number and catalog number were reported incorrectly on the initial manufacturer device report number 1220648-2023-05047 and were revised. g.1 mdr reporting contact email was revised from the initial submission of manufacturer device report number 1220648-2023-05047 in accordance with updated procedures. reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2023-05047 and was added. h.6 code 2199 was reported incorrectly on the initial manufacturer device report number 1220648-2023-05047.
32730 | 1220648-2023-05048 | 18188286 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18188286&pc=OZD | 10/24/2023 | 06/06/2024 |  | 11/21/2023 |  |  | 10/24/2023 | 11/21/2023 |  |  | 05/31/2025 |  | 11/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024348705 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Migration | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 54 year old male patient was implanted with an impella cp pump for mechanical circulatory support. it was documented that while changing the patient's position, the impella cp came out of the aorta. the impella was therefore removed and an intra-aortic balloon pump (iabp) was inserted. there was no reported patient injury. | [5 empty] | The investigation into the positioning issues has been completed since the original report was submitted. the device was not returned for investigation. based on the data analysis, the most likely cause of the positioning issues was patient movement. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32731 | 1220648-2023-05049 | 18188334 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18188334&pc=OZD | 10/25/2023 | 01/12/2024 |  | 11/21/2023 |  |  | 10/25/2023 | 11/21/2023 |  |  | 05/31/2025 |  | 11/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024353422 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/25/2023 |  |  |  |  | The user facility reported a 60-year-old male with cardiomyopathy was implanted with an impella 5.5 as a bridge to transplant. during support, the impella was displaying high purge pressure and purge pressure alarms. upon inspection of the lines there were no kinks noted. the patient remained on impella support until 18-nov-2023 when the patient successfully received a heart transplant. there was no harm to the patient as a result of the pressure alarms. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported high purge pressure issue has been completed. the impella device was not returned for evaluation. clinical details states that high purge pressure alarms were seen with purge pressure rise above 1000 mmhg and purge flow reduced less than 2 ml/hr. data logs were reviewed finding a purge pressure rise immediately within 2 minutes and purge flow ticks stalling were observed due to likely kink in the purge line. the cause of the high purge pressure issue was most likely a kinked purge line in the catheter per data log review.
32732 | 1220648-2023-05050 | 18188342 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18188342&pc=OZD | 10/25/2023 | 05/15/2024 |  | 11/21/2023 |  |  | 10/25/2023 | 11/21/2023 |  |  | 04/30/2025 |  | 11/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024348228 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/25/2023 |  |  |  |  | The user facility reported a 59-year-old male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support, the pump was found to be in the papillary. an unsuccessful attempt was made to reposition the pump. the patient's lactate dehydrogenase was greater than 3000 indicative of hemolysis. the pump was on p-2 setting for 45 minutes with no changes in baseline hemodynamics. the physician was unable to free the pigtail portion of the pump from papillary with forward advancement and clockwise torque. the pump was pulled back until freed from structure, but pump was noted to be fully in the aorta with inability to cross the aortic valve. the physician made the decision to fully remove pump and monitor the patient's hemodynamics closely. | [5 empty] | The investigation into the hemolysis and the positioning issues has been completed since the original report was submitted. the device was not returned for investigation. data log analysis confirmed suction alarms and positioning issues were present. the most likely cause of the hemolysis was improper positioning. the most likely cause of the positioning issues was use related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
32733 | 1220648-2023-05051 | 18188348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18188348&pc=OZD | 10/25/2023 | 06/07/2024 |  | 11/21/2023 |  |  | 10/25/2023 | 11/21/2023 |  |  | 09/30/2025 |  | 11/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024394606 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/25/2023 |  |  |  |  | The user facility reported a 73-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. patient went into torsades multiple times and had to be shocked out of it. the reason for the torsades is unknown. the patient was attempted to be weaned off the impella and it was unsuccessful because the ejection fraction (ef) became worse. patient is still on support and is stable. | [5 empty] | The investigation into the vf/vt/af/at (torsades) issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32734 | 1220648-2023-05052 | 18188356 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18188356&pc=OZD | 10/25/2023 | 05/13/2024 |  | 11/21/2023 |  |  | 10/25/2023 | 11/21/2023 |  |  | 08/31/2025 |  | 11/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024393181 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  |  |  | Injury | No | Yes | Migration | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/25/2023 |  |  |  |  | The user facility reported a 69-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the implanted impella cp pump was accidentally pulled out of the ventricle during repositioning. the pump was subsequently removed as a result of the noted issue. the patient was placed on ecmo for ongoing support. there was no patient harm. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the positioning issue was use related. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32735 | 1220648-2023-05006 | 18169571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18169571&pc=OZD | 10/20/2023 | 11/17/2025 |  | 11/20/2023 |  |  | 11/20/2023 | 11/20/2023 |  |  |  |  | 11/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1691800 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 39 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | Followup |  | 10/20/2023 |  |  |  |  | The user facility reported a 39-year-old male (race unknown) post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. during support, the automated impella controller (aic) displayed a controller failure (red alarm). the team swapped out the aic with another aic with no issues. there was no patient harm related to this event. | [5 empty] | Correction to d4: lot number and serial number corrected. upon investigation closure, a supplemental report will be submitted. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | D.2 revised common device name in accordance with updated procedures since the initial manufacturer device report 1220648-2023-05006 was submitted. d.4 revised model number in accordance with updated procedures since the initial manufacturer device report 1220648-2023-05006 was submitted. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2023-05006 and should not have been. g.1 revised mdr reporting contact name and address and manufacturing site name and address as they were entered incorrectly on the initial manufacturer device report 1220648-2023-05006. reporting contact email was revised in accordance with updated procedures since the initial manufacturer device report 1220648-2023-05006 was submitted. g.6 30 days was inadvertently omitted on the first follow up manufacturer device report 1220648-2023-05006. h.6 code 2199 was reported incorrectly on initial manufacturer device report 1220648-2023-05006.
32736 | 1220648-2023-05007 | 18173003 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18173003&pc=OZD | 10/24/2023 | 10/02/2024 |  | 11/20/2023 |  |  | 10/24/2023 | 11/20/2023 |  |  |  |  | 11/20/2023 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1554466 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Power Up | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | A patient of unknown age and gender in the eu received hemodynamic support with an impella cp heart pump. it was reported that the automated impella controller (aic) took an unusually long time to turn on and detect the purge cassette during preparation of the impella cp heart pump. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation of aic - power on issues/blank screen has been completed. data logs reviewed and device analysis performed and the root cause could not be determined due to insufficient information and not being able to reproduce. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05007.
32737 | 1220648-2023-05008 | 18173070 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18173070&pc=OZD | 08/07/2023 | 11/20/2023 |  | 11/20/2023 |  |  | 08/07/2023 | 11/20/2023 |  |  | 04/30/2025 |  | 11/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024350779 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/07/2023 |  |  |  |  | The user facility reported a 79-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella kinked upon insertion distal to the outlet. they were unable to flow above 1.3l, the device was removed and a new impella was inserted without any complications. no patient harm reported due to the issue. | [5 empty] | The investigation into the reported low pump flow has been completed. product was returned for evaluation. there was a kink found on cannula near the distal land of the outlet. per the clinical details provided, the root cause of the low flow issue was a kinked cannula related to tortuous aorta. instructions for use for the related event are as follows ¿avoid manual compression of the inlet and outlet areas of the cannula assembly.¿ ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
32738 | 1220648-2023-05009 | 18173102 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18173102&pc=OZD | 10/24/2023 | 06/24/2024 |  | 11/20/2023 |  |  | 10/31/2023 | 11/20/2023 |  |  | 04/30/2025 |  | 11/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024337973 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/31/2023 |  |  |  |  | The user facility in japan reported a 74-year-old male post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. a thrombus was confirmed in the abdomen at the time of the impella removal. thrombus was removed from the patient successfully. the patient is stable. | [5 empty] | The investigation into the thrombus issue has been completed since the original report was submitted. the product was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32739 | 1220648-2023-05010 | 18173546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18173546&pc=OZD | 10/24/2023 | 08/15/2024 |  | 11/20/2023 |  |  | 10/24/2023 | 11/20/2023 |  |  | 01/31/2025 |  | 11/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024332122 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/24/2023 |  |  |  |  | The user facility reported a 77-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. when explanting impella sheath, there was not sufficient bleed back before tightening down perclose closure. unable to find doralis pedis or posterior tibialis pulse. patient became increasingly restless and was over sedated, neo started for hypotension and was given narcan in the ccu and patient became alert and complained of back pain. patient continued to deteriorated overnight. was taken to cath lab intubated on neo, epi, levo, and dop for abdominal femoral runoff that found dissection from cfa to popliteal artery. a vascular surgeon was consulted. | [5 empty] | The investigation of limb ischemia and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the limb ischemia could not be determined. the cause of the vascular damage issue was most likely patient condition related to diseased vessels. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05010: d4 model number in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8-should have been left blank h.6 code 4643 was reported incorrectly. new codes were added to health effect - impact code. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32740 | 1220648-2023-05011 | 18173586 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18173586&pc=OZD | 10/31/2023 | 03/13/2024 |  | 11/20/2023 |  |  | 10/31/2023 | 11/20/2023 |  |  | 04/30/2025 |  | 11/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024336821 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/31/2023 |  |  |  |  | The user facility reported a 64-year-old female post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. the team had difficulty advancing to left ventricle position with only transesophageal echocardiography and not fluoro imaging. the pump advanced and caused a left ventricle perforation. the team patched the perforation with suture and patch. patient survived perforation. the patient remains on support and was transferred to tertiary center for continued monitoring/care. | [5 empty] | The investigation into the ventricle perforation issue has been completed since the original report was filed. the device was not returned for investigation. based on the clinical account, the root cause of the perforation was due to improper positioning and the pump being inserted too deep. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
32741 | 1220648-2023-05012 | 18173797 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18173797&pc=OZD | 10/24/2023 | 03/26/2024 |  | 11/20/2023 |  |  | 10/24/2023 | 11/20/2023 |  |  | 05/31/2025 |  | 11/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024349346 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/24/2023 |  |  |  |  | The user facility reported a 70-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. while on support, there was an ¿impella in the aorta¿ alarm during surgery. this was resolved by increasing the p-level. the pump was cross-clamped and experienced a spike in motor current and stopped. the stopped pump was removed at the end of the coronary artery bypass procedure procedure. there was no patient harm due to these events. | [5 empty] | The investigation into the pump stop, the positioning issues, and the low pump flow has been completed since the initial report was submitted. the product was not returned for investigation. the cause of the pump stop issue was most likely use related since the doctor mentioned the cross-clamp event during the pump stop. the cause of the positioning issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the low pump flow issue was most likely patient condition related since it was resolved after giving blood volume. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32742 | 1220648-2023-05013 | 18174206 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18174206&pc=OZD | 10/30/2023 | 08/20/2024 |  | 11/20/2023 |  |  | 10/30/2023 | 11/20/2023 |  |  | 04/30/2025 |  | 11/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024355002 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/30/2023 |  |  |  |  | The user facility reported a 70-year-old female with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the patient had elevated plasma free hemoglobin, so the physician was concerned about hemolysis. as a result, the impella cp was explanted and replaced with a new one to address the issue. | [5 empty] | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. instruction for use as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05013: in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8-should have been left blank. g1 reporting contact fax number missing. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32743 | 1220648-2023-05014 | 18174904 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18174904&pc=OZD | 10/21/2023 | 08/12/2024 |  | 11/20/2023 |  |  | 10/21/2023 | 11/20/2023 |  |  | 08/31/2025 |  | 11/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024386888 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/21/2023 |  |  |  |  | The user facility reported a 58-year-old female with a heart transplant rejection was implanted with an impella cp device for mechanical circulatory support. upon impella insertion, there was a kink in the catheter at the cannula and outlet junction. staff was unable to resolve the issue and the impella was removed. the patient was on ecmo at the time of implant. after the impella was removed, the patient was becoming more pulsatile and the team left patient on ecmo with flows at 2.8. | [5 empty] | The investigation for the product damage has been completed. the impella device was not returned for evaluation. therefore, without sufficient clinical details, the root cause of the reported issues product damage could not be determined. the instructions for use for the impella cp with smartassist system states the following: section: warnings & cautions: cautions : ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ added- d6a/d6b f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32744 | 1220648-2023-05015 | 18175318 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18175318&pc=PYX | 10/21/2023 | 01/12/2024 |  | 11/20/2023 |  |  | 10/21/2023 | 11/20/2023 |  |  | 03/31/2025 |  | 11/20/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024394959 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  |  | White | Malfunction | No | Yes | Increase in Pressure | Migration | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the physician adjusted the impella rp pump after reviewing the x-ray. the physician believed the pump might have been deep and pulled the pump back. after repeat x-ray the physician decided the pump needed to be advanced a cm. high purge pressure alarms started after final adjustment. during this time the 73-year-old female patient developed hemolysis symptoms. a purge cassette change rectified the alarms, and the patient was placed on p-7. after the final adjustment and p-level lowered, the patient¿s hemolysis symptoms were resolved. physician and staff instructed to monitor motor current for upwards trend. the team had suspected clot entrainment and removed the impella rp. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed since the original report was filed. the impella pump was returned for investigation the device was reviewed and there was no indication of biomaterial or kinks noticed in the cassette tubing or on the catheter. however, due to the log signature of the purge pressure rising, it looks like the purge line was most likely kinked causing the high purge pressure. the high purge pressure was not able to be reduced. the root cause of the high purge pressure is most likely due to a kinked purge line.
32745 | 1220648-2023-05016 | 18175800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18175800&pc=OZD | 10/21/2023 | 01/25/2024 |  | 11/20/2023 |  |  | 10/21/2023 | 11/20/2023 |  |  | 02/28/2025 |  | 11/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024321395 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/21/2023 |  |  |  |  | The user facility reported a 67-year-old female was implanted with an impella cp device for mechanical circulatory support after an aortic valve replacement. during support, the patient had no flow around the impella sheath. an external fem-fem was placed and support continued. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible issue has been completed. the device was not returned for benchtop investigation. according to the clinical information provided, patient had ischemia with no flow around the impella sheath and external fem-fem bypass resolved ischemia. the cause of the limb ischemia issue was most likely patient condition related with fem-fem bypass done to resolve ischemia.
32746 | 1220648-2023-05018 | 18175828 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18175828&pc=OZD | 10/25/2023 | 04/18/2024 |  | 11/20/2023 |  |  | 10/25/2023 | 11/20/2023 |  |  | 09/30/2025 |  | 11/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024397218 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Power Up | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/25/2023 |  |  |  |  | The user facility reported a 54-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the patient had a very challenging anatomy. the physician removed the impella twice from the body to re-wire or try a new approach. once the impella was in the proper position and ready to start, an impella stop alarm occurred. after trying to start the pump several times unsuccessfully, a new pump was used. the patient is stable, and the new pump was successfully implanted. | [5 empty] | The investigation into the pump did not start issue has been completed since the initial report was submitted. the pump was returned and tested. during testing and investigation, the pump was unable to be started in the lab and three open electrical lines were found. the root cause of the pump did not start issue was determined to be three open electrical lines in the catheter by the motor housing. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
32747 | 1220648-2023-05017 | 18175836 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18175836&pc=OZD | 10/21/2023 | 08/01/2024 |  | 11/20/2023 |  |  | 10/21/2023 | 11/20/2023 |  |  | 05/31/2025 |  | 11/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024362437 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso)and it was determined that use of the device cannot conclusively be associated with the outcome. | The us complainant reported the 69-year-old female patient had a hematoma during impella cp support. manual pressure was held and blood products were given to treat the noted condition. support with the implanted pump was maintained throughout the event. |  |  |  |  | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added b5 -mso review, h6 component code 925 added- d6a/d6b. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿
32748 | 1220648-2023-05019 | 18176127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18176127&pc=OZD | 10/22/2023 | 03/04/2024 |  | 11/20/2023 |  |  | 10/22/2023 | 11/20/2023 |  |  | 05/31/2025 |  | 11/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024352027 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that 79-year-old female patient experienced continuous suction on impella cp despite adequate preload and right ventricle position. suction was unable to be resolved so a decision was made to exchange for new impella cp. upon removal clot noted at base of inlet. there was no reported patient harm. | [5 empty] | The investigation into the reported low pump flow has been completed since the original report was filed. returned complaint cp pump (b)(6) visually inspected. no cannula kink observed. no broken blades found. no biomaterial observed in inlet/ outlet flow cages or around impeller. dried dextrose around the purge gap dissolved. pump run for 10 minutes in investigation lab on p-9. no low pump flow reproduced. pump flow was 4l/m on p-9 which indicated saturated pump flow. the cause of the low pump flow issue most likely was biomaterial (clot) ingestion base on the clinical data they noted clot at inlet cage that may washed of before return for investigation. data logs reviewed: no motor current spikes observed. suction alarms occurred. low pump flow dropped below 2l/m at p-4. patient had calcified/stenotic valve condition. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32749 | 1220648-2023-05020 | 18176714 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18176714&pc=OZD | 10/22/2023 | 06/26/2024 |  | 11/20/2023 |  |  | 10/22/2023 | 11/20/2023 |  |  | 01/31/2025 |  | 11/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024306649 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was further reported that the patient expired while on support. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that there was not enough information to associate use of the device with the outcome. | The complainant reported a 69-year-old male was implanted with an impella cp for mechanical circulatory support after being on balloon pump. impella cp was noted to be placed over the pap muscle. the patient had to be taken back to the cath lab for the pump to be repositioned. the physician pulled back across the atrioventricular valve and was unable to advance back into the left ventricle. therefore, the pump was removed, and a new cp was placed. there was no reported patient harm. mdr required. |  |  |  |  | The investigation for the positioning issue has been completed. according to the clinical, on the first day of impella cp support the physician noted that the pump was in contact with the papillary muscle. in an attempt to resolve the issue the pump was repositioned, but the pump was accidentally pulled across the aortic valve and was unable to be advanced back into the left ventricle. the decision was then made to explant the pump and replace it with another device. no product was returned for a visual inspection. data log analysis was not necessary for this complaint. the most likely cause of the positioning issues was use related. added- b5 mso review, h6 impact code 4628 added- d6a/d6b. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32750 | 1220648-2023-05023 | 18179373 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18179373&pc=OZD | 10/25/2023 | 06/13/2024 |  | 11/20/2023 |  |  | 10/25/2023 | 11/20/2023 |  |  | 08/31/2025 |  | 11/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024387278 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a 58-year-old male patient on impella cp support experienced hemolysis that occurred during p-9 support. there were no positioning alarms. hemoglobin was administered and the cp was explanted, and replaced with an impella 5.5. | [5 empty] | The investigation has been completed. according to the clinical information, on the second day of impella cp support, hemolysis symptoms were discovered. three days later, the decision was made to remove the pump and replace it with an impella 5.5 device. three days after this swap, haptoglobin was administered and the hemolysis resolved. no product was returned for a visual inspection. data log analysis did not confirm the presence of any abnormalities. the cause of the hemolysis was not determined because no product was returned and not enough clinical information was given. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ d.6a, d.6b, e.1 address line 2 and g.1 reporting contact fax number were revised from the initial submission in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. additional manufacturer narrative on the previously submitted report stated that the device was not returned but the device was discarded. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿
32751 | 1220648-2023-05004 | 18167199 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18167199&pc=OZD | 10/23/2023 | 03/22/2024 |  | 11/19/2023 |  |  | 10/23/2023 | 11/19/2023 |  |  | 07/31/2025 |  | 11/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024374354 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant in japan reported a 40-year-old female patient that presented with acute myocardial infarction and cardiogenic shock was placed on impella cp support. approximately 19 days after initiation of the device, the pump stopped. the impella was removed and catecholamine support was administered. | [5 empty] | The investigation for the reported pump stop has been completed. the pump nor the data logs were not returned for evaluation. however, without the device nor sufficient clinical information for evaluation, the exact root cause of the pump stop could not be determined. | The impella device was not received from the customer as it was discarded and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
32752 | 1220648-2023-05005 | 18167200 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18167200&pc=OZD | 10/24/2023 | 02/28/2024 |  | 11/19/2023 |  |  | 10/24/2023 | 11/19/2023 |  |  | 04/30/2025 |  | 11/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024353615 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/24/2023 |  |  |  |  | The user facility reported a 67-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient had been on intra-aortic balloon pump (iabp) and began to decompensate. the decision was made to remove iabp and escalate to the impella cp device. during support, the patient lost pulses in the right leg prior to being transferred to an outside hospital. the md was not comfortable placing antegrade sheath for distal perfusion. upon arrival to accepting hospital, impella assisted percutaneous cardiac intervention (pci) was performed and impella was removed following pci. | [5 empty] | The investigation into the reported limb ischemia -reversible has been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32753 | 1220648-2023-04989 | 18159198 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18159198&pc=OZD | 10/20/2023 | 02/07/2024 |  | 11/17/2023 |  |  | 10/20/2023 | 11/17/2023 |  |  | 04/30/2025 |  | 11/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024351026 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella support of a 70-year old female patient a high purge pressure alarm triggered when decreasing from p3 to p2, pf <1. there were no kinks in purge tubing. the physician acknowledged the patient¿s anticoagulation was sub therapeutic. heparin assay <0.1, on day 2 of pump support for approximately 24 hours. team did not want to utilize tpa as the plan was to explant. the team attempted cassette change which did not resolve the alarm. therefore, the alarm audio was disabled and pump was explanted successfully to plan bedside with manual pressure to achieve hemostasis. | [5 empty] | The investigation for the high purge pressure has been completed. the impella cp pump was returned for evaluation. visual inspection was performed, and biomaterial ingestion was observed at the purge gap around the impeller. no catheter kink or other abnormalities were found. high purge pressure alarm was reproduced within five minutes of purging with purge pressure >1050 mmhg and purge flow <2 ml/hr. no other abnormalities were found. the root cause of the high purge pressure issue was due to biomaterial ingestion in purge gap based on field investigation and clinical review. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
32754 | 1220648-2023-04990 | 18159265 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18159265&pc=OZD | 10/20/2023 | 12/01/2025 |  | 11/17/2023 |  |  | 10/20/2023 | 11/17/2023 |  |  |  |  | 11/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1691800 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 39 YR |  | Male |  |  |  | Malfunction | No | Yes | Activation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | A 39-year-old male patient received hemodynamic support from an impella rp smartassist heart pump and an impella 5.5 smartassist heart pump during high-risk cardiac surgery and as bridge to heart transplant. the automated impella controller (aic) running the impella 5.5 device showed a controller failure red alarm and had to be exchanged. the patient suffered no harm from the incident. | [5 empty] | Data analysis: on the complaint date (oct 20, 2023), during support, the console displayed ¿controller error (loss of communication with keypad or between the kbd and i2c for kbd hw revision 1) ¿ alarm,¿ event #24. the issue was never resolved for the rest of the case. this confirms the reported failure mode. device analysis: the console was tested upon arrival at fs, and the ¿controller error,¿ event #24 was reproduced intermittently. no issue was found with the ribbon cable between the 4-in-1 and the glass keyboard. root cause: the root cause for the controller error could not be determined due to the intermittent failure mode. there is no evidence in the log to confirm a controller failure (red alarm). the controller error (yellow alarm) was misreported as a controller failure (red alarm). h3 updated. | A.1 revised patient identifier (in confidence) as it was reported incorrectly on initial manufacturer device report number 1220648-2023-04990. d.1 and d.2 brand name and common device name were revised in accordance with updated procedures since the initial manufacturer device report 1220648-2023-04990 was submitted. d.4 model number, serial number and catalog number were revised in accordance with updated procedures since the initial manufacturer device report 1220648-2023-04990 was submitted. g.1 revised mdr reporting contact name and address and manufacturing site name and address as they were reported incorrectly on initial manufacturer device report 1220648-2023-04990. reporting contact email was revised in accordance with updated procedures since the initial manufacturer device report 1220648-2023-04990 was submitted. h.6 code 2199 was reported incorrectly on initial manufacturer device report 1220648-2023-04990. | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
32755 | 1220648-2023-04991 | 18159353 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18159353&pc=OZD | 10/20/2023 | 04/12/2024 |  | 11/17/2023 |  |  | 10/20/2023 | 11/17/2023 |  |  | 05/31/2025 |  | 11/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024354572 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 71-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that a hematoma was observed an hour after being returned to icu. the impella was also having suction events and pulsatility was getting narrower. after re-evaluation of the patient, the groin was noticed to have a large hematoma around the impella site. manual pressure and femstop was applied and hematoma was reduced in size and the site returned to normal. two units of packed red blood cells (prbcs) were given to replace the volume that was lost. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. clinical details states that the bleeding was resolved when heparin levels dropped. therefore, the root cause of the access site bleed was most likely due to anticoagulation management. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32756 | 1220648-2023-04992 | 18159485 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18159485&pc=OZD | 10/20/2023 | 04/16/2024 |  | 11/17/2023 |  |  | 10/20/2023 | 11/17/2023 |  |  | 03/31/2025 |  | 11/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024341948 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported the 82-year-old male patient had access site bleeding and manual pressure was held. the repositioning sheath was re-sutured and a c-clamp was applied to the groin. additionally, the staff were having issues obtaining distal pulses on the access site leg. external bypass was given as an option to help obtain distal flow to access leg if perfusion to limb became a concern. | [5 empty] | The investigation for the reported access site bleeding and limb ischemia has been completed. the impella device was not returned for evaluation. however, clinical information provided was sufficient to determine the root cause. the root cause of access site bleeding was determined to be use issue because the bleeding was resolved after repositioning sheath was re-sutured. hence, the root cause of the issue is determined to be use issue due to lack of proper technique in closing the access site. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32757 | 1220648-2023-04993 | 18159610 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18159610&pc=OZD | 10/20/2023 | 06/26/2024 |  | 11/17/2023 |  |  | 10/20/2023 | 11/17/2023 |  |  | 08/31/2025 |  | 11/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024394933 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information: case wrap-up: the impella device was explanted as care is being withdrawn. patient remains on vent and pharmacological support. the patient expired. abiomed medical safety office review determined abiomed does not have enough information to associate use of the device with the outcome. | The us complainant reported that while on impella cp support, the 42-year-old patient was hemolyzing. the doctor therefore repositioned the impella, withdrawing it 2 cm. it is unknown if repositioning would have resolved the condition as the pump was explanted the following day. |  |  |  |  | The investigation into the hemolysis and the positioning issues has been completed since the original report was submitted. the device was not returned for investigation. according to the clinical, on the second day of impella cp support the patient began hemolyzing with pink/red urine. after the pump was repositioned, the hemolysis remained. shortly afterward the patient was withdrawn from support. data log analysis confirmed positioning issues and a low placement signal trend. the most likely cause of the hemolysis was improper positioning. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ additional information b5. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿patient conditions, including catheter position, pre-existing medical conditions, and small left ventricular volumes, may also play a role in patient susceptibility to hemolysis.¿
32758 | 1220648-2023-04994 | 18159780 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18159780&pc=OZD | 10/20/2023 | 02/05/2024 |  | 11/17/2023 |  |  | 10/20/2023 | 11/17/2023 |  |  | 07/31/2025 |  | 11/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024391691 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | Renal Impairment | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a 56-year-old male patient had an impella 5.5 as bridge to left ventricular assist device (lvad) implant. patient¿s creatinine and ldh began to significantly uptrend over the course of four days. during this change in lab results a swan was inserted (eventually clotted off) showing that the patient was hypervolemic. after lvad placement, patient¿s kidney function, hemolysis markers and symptoms improved quickly. the physician believes there was clot formation in the device causing the clinical deterioration. | [5 empty] | The investigation into the into the hemolysis and renal issues has been completed since the original report was filed. product was returned and investigated. per the clinical information, there were concerns for hemolysis with creatinine and lactate dehydrogenase (ldh) beginning to significantly uptrend over the course of four days and peaking during impella 5.5 support. impella 5.5 pump was returned in a cut open state at catheter 35 cm marking. visually inspected and keyence imaging utilized to confirm huge thrombus/biomaterial being stuck over the outside of the cannula and significant biomaterial was wrapped around the impeller at outflow. data logs were reviewed, and motor current spike was observed. no suction alarms or position issue alarms were found. the cause of the hemolysis with subsequent renal impairment (as evidenced by rising creatinine) was due to significant biomaterial ingestion (thrombus). | The device was returned and an evaluation is pending. upon conclusion of the investigation, a supplemental report will be submitted with a summary of the results. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32759 | 1220648-2023-04995 | 18159883 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18159883&pc=OZD | 10/20/2023 | 11/03/2025 |  | 11/17/2023 |  |  | 10/20/2023 | 11/17/2023 |  |  | 06/30/2025 |  | 11/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024374184 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported the 81-year-old male patient had a small hematoma at the access site and heparin was temporarily paused. support continued; however, four days later a ct scan showed a small intracranial hemorrhage. heparin was again discontinued and the pump was explanted the following day. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32760 | 1220648-2023-04996 | 18162426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18162426&pc=OZD | 10/23/2023 | 02/16/2024 |  | 11/17/2023 |  |  | 10/23/2023 | 11/17/2023 |  |  | 07/31/2025 |  | 11/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024392392 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/23/2023 |  |  |  |  | The user facility reported a 64-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the physician peeled away and repositioned the repositioning sheath, noticed bleeding around the blue suture hub. pump removed due to this bleeding concern and replaced with an intra-aortic balloon pump . | [5 empty] | The investigation into the access site bleeding ¿ major has been completed since the original report was filed. the device was returned for investigation. it was reported that the root cause of the bleeding was use related issue due to improper use technique to insert the blue hub into the patient. the blue hub is not intended to be in patient. | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for investigation. based on the clinical information provided, the cause of the bleeding was use related issue due to improper use technique to insert the blue hub in to the patient. the blue hub is not intended to be in patient. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32761 | 1220648-2023-04997 | 18162586 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18162586&pc=OZD | 10/22/2023 | 09/19/2024 |  | 11/17/2023 |  |  | 10/22/2023 | 11/17/2023 |  |  | 07/31/2025 |  | 11/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024390702 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The us complainant reported the 65-year-old male patient had minimal distal right leg flows. the patient had a distal perfusion sheath inserted last night. however, due to lack of heaprinization the sheath clotted off and was subsequently removed. staff was discussing escalating to an impella 5.5 due to high lactic and increased requirement for pressors/inotropes. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the limb ischemia ¿ reversible has been completed. the device and data were not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-04997 was provided in sections b2, b5, d6a, d6b, h1, and h6.
32762 | 1220648-2023-04998 | 18162621 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18162621&pc=OZD | 11/03/2023 | 09/26/2024 |  | 11/17/2023 |  |  | 11/03/2023 | 11/17/2023 |  |  | 06/30/2025 |  | 11/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024360023 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Activation, Positioning or Separation Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/03/2023 |  |  |  |  | The user facility reported a 48-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported that upon impella insertion the left axillary was dilated and the impella sheath was inserted. when the dilator was removed from the sheath, the sheath kinked. when the impella cp was inserted the over wire, they were not able to pass through sheath. the impella insertion was aborted. | Additional information noted that the procedure outcome was that patient expired due to cardiovascular insufficiency. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The investigation into the difficulty inserting/removing pump through introducer has been completed since the initial report was submitted. the product was not returned for investigation. as per the clinical details, axillary artery dilation and impella sheath insertion were complicated by sheath kinking, hindering the passage of wires and impella pump. the cause of the issue difficulty inserting pump through introducer was an introducer sheath kink, based on provided clinical details. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-04998 was provided in sections b1, b5, d6, h1, and h6.
32763 | 1220648-2023-04999 | 18162652 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18162652&pc=OZD | 09/20/2023 | 02/19/2024 |  | 11/17/2023 |  |  | 09/20/2023 | 11/17/2023 |  |  | 02/28/2025 |  | 11/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024321041 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/20/2023 |  |  |  |  | The user facility reported a 77-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported a failure of a closure device resulting in the need for surgical intervention to close the impella access site. patient required surgical intervention to close groin post procedure. | [5 empty] | The investigation into the vascular damage - peripheral vessel issue has been completed. the device was not returned for investigation. the root cause of the vascular damage - peripheral vessel issue was not determined since product was not returned and there is a lack of clinical details. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
32764 | 1220648-2023-05000 | 18162704 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18162704&pc=OZD | 11/14/2023 | 09/17/2024 |  | 11/17/2023 |  |  | 11/14/2023 | 11/17/2023 |  |  | 09/30/2025 |  | 11/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024401750 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/14/2023 |  |  |  |  | The user facility reported a 67-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported that the patient was on impella 5.5 support. upon explant, a clot was found at the outlet of the device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for thrombus has been completed. as the necessary clinical information was not provided, the root cause of the thrombus was not determined. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05000. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8- should have been left blank. g1 reporting contact fax number missing.
32765 | 1220648-2023-05001 | 18162724 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18162724&pc=OZD | 10/21/2023 | 12/05/2023 |  | 11/17/2023 |  |  | 10/22/2023 | 11/17/2023 |  |  | 03/31/2024 |  | 11/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2023202190 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 16 YR |  | Female |  | Non Hispanic | Asian | Injury | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/22/2023 |  |  |  |  | The user facility reported a 16-year-old female with cardiomyopathy was implanted with an impella cp for mechanical circulatory support after being placed on ecmo due to viral myocarditis. during support, the device experienced a purge flow alarm. the purge system was replaced resulting in a temporary improvement. however, the purge flow decreased and increased again. it was determined that there was an abnormality in the pump. the pump was explanted and replaced with another impella cp. there was no health hazard to the patient. | [5 empty] | The investigation into the high purge pressure has been completed since the original report was filed. the device was returned for investigation. the data logs were viewed and there were motor current spikes followed by high purge pressure and low purge flow. additionally, pump flow became saturated. there was biomaterial found in the purge gap of the returned device. the root cause of the high purge pressure is due to the biomaterial. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
32766 | 1220648-2023-05002 | 18162883 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18162883&pc=OZD | 10/22/2023 | 01/16/2024 |  | 11/17/2023 |  |  | 10/22/2023 | 11/17/2023 |  |  | 09/30/2025 |  | 11/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024397208 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/22/2023 |  |  |  |  | The user facility reported a male (age unknown) with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. while on support, the patient experienced a stroke and had received iv haldol earlier over-night. the following morning the patient was noted to be more somnolent and unable to respond to questions with morning nursing assessment. no ischemia noted on ct scan. at this time, it is unknown what course of action was taken. the patient was also on va ecmo. the patient remains on impella support, to date the pump has supported for over 29 days. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported stroke/ cva has been completed since the original report was filed. no product was returned for investigation. data logs do not show evidence of biomaterial. the root cause of the stroke/cva was most likely patient condition related and the relation to the pump cannot be determined.
32767 | 1220648-2023-05003 | 18163002 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18163002&pc=OZD | 10/21/2023 | 07/23/2024 |  | 11/17/2023 |  |  | 10/21/2023 | 11/17/2023 |  |  | 08/31/2025 |  | 11/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024391681 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/21/2023 |  |  |  |  | The user facility reported a 69-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support as a bridge to transplant. while on support, the patient developed a hematoma at the access site. a washout was performed and heparin was paused. | [5 empty] | The investigation for access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6a/d6b added h6 component coded added 925 the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-05003. f6/f8-should have been left blank. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32768 | 1220648-2023-04957 | 18148426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18148426&pc=OZD | 10/18/2023 | 07/15/2024 |  | 11/16/2023 |  |  | 10/18/2023 | 11/16/2023 |  |  | 06/30/2025 |  | 11/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024367788 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/18/2023 |  |  |  |  | The user facility reported a 61-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during the support the patient developed a hematoma at the access site. as a result, the impella was removed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: bleeding access site: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8-should have been left blank in the initial report.
32769 | 1220648-2023-04958 | 18148614 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18148614&pc=OZD | 10/18/2023 | 08/02/2024 |  | 11/16/2023 |  |  | 10/18/2023 | 11/16/2023 |  |  | 08/31/2025 |  | 11/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024395816 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/18/2023 |  |  |  |  | The user facility reported an 83-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient had a hematoma noted following sheath exchange to repositioning sheath for which the patient was given one unit of blood and manual pressure was applied. additionally, the patient had massive bleeding noted out of the nasogastric tube and oropharyngeal area. patient further deteriorated, resulting in mass transfusion. unfortunately this was unsuccessful and the patient expired after 3.16 hours of impella support. no allegations were made towards the impella for cause of death. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was most likely patient movement, the patient kicked his legs, displacing the impella cp and causing a hematoma as per the clinical description. per medical safety officer review use of the device cannot conclusively be associated with the outcome. h6 component code added. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-04958: in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8-should have been left blank.
32770 | 1220648-2023-04959 | 18148813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18148813&pc=OZD | 10/18/2023 | 05/15/2024 |  | 11/16/2023 |  |  | 10/18/2023 | 11/16/2023 |  |  | 06/30/2025 |  | 11/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024379381 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Hispanic |  | Malfunction | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an 60-year-old female patient presenting for acute myocardial infarction and cardiogenic shock was implanted with an impella cp pump for mechanical circulatory support. it was documented that upon escalation from the impella cp pump to an impella 5.5, the surgeon identified that the entry point of the impella cp was high, and felt uncomfortable removing it and repairing arteriotomy. therefore, the surgeon consulted vascular. vascular came in, was able to get to the point of insertion of the impella, removed the impella cp, and repaired the arteriotomy. jackson pratt drain (jp) drain was placed prior to closure. | [5 empty] | The investigation into the removal issues has been completed since the original report was submitted. the device was not returned for investigation. based on the clinical details, the most likely cause of the removal issues was use related. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer as it was discarded and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32771 | 1220648-2023-04960 | 18148896 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18148896&pc=OZD | 10/18/2023 | 09/19/2024 |  | 11/16/2023 |  |  | 10/18/2023 | 11/16/2023 |  |  | 07/30/2025 |  | 11/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024386875 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  |  | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/18/2023 |  |  |  |  | The user facility reported a 80-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. there was high purge pressure. the cassette was replaced with no decrease in purge pressure. support continued with the implanted pump despite the noted issue. there was no reported patient harm. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event for the patient's demise noted: there is not enough evidence to exclude impella as an associated factor in the patient's outcome. patient's peri-procedural condition was poor. the patient's peri-procedural condition was critical. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | B2: the date of patient death is unknown. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-04960 was provided. note: corrected information provided in h6 is an addition to previous coding. | The investigation into the high purge pressure issue has been completed since the initial report was submitted. the device was not returned for investigation. the cause of the high purge pressure issue was most likely biomaterial, based on data log review.
32772 | 1220648-2023-04961 | 18149217 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18149217&pc=OZD | 10/18/2023 | 08/20/2025 |  | 11/16/2023 |  |  | 10/18/2023 | 11/16/2023 |  |  |  |  | 11/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1103720 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Intermittent Loss of Power | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) was turned on to start a new case for training with a demo impella when the aic screen turned black for a couple seconds. troubleshooting was initiated but the aic was unable to be used and an "unexpected controller shutdown" alarm appeared. there was no reported harm. | [5 empty] | The impella device was not returned by the customer and therefore, evaluation of the device was not possible. upon investigation completion, a supplemental mdr will be filed. | Corrections to the initial mfr report: g1 mdr reporting contact email. d2 device name updated. d9 device return date added. updated h3 to device evaluation anticipated.
32773 | 1220648-2023-04962 | 18149439 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18149439&pc=OZD | 10/18/2023 | 02/21/2024 |  | 11/16/2023 |  |  | 10/18/2023 | 11/16/2023 |  |  | 08/31/2025 |  | 11/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 380 SMART ASSIST SET, EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2024390807 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aneurysm | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/18/2023 |  |  |  |  | The user facility reported that a patient of unknown age and gender was implanted with an impella cp pump for mechanical circulatory support. during implant, an artery aneurysm of the right femoral was detected. the pump was subsequently explanted and a new impella pump was inserted for patient support. | [5 empty] | The impella device was not returned by the customer and therefore, evaluation of the device was not possible. upon conclusion of the investigation, a supplemental report will be submitted with a summary of the results. | The investigation into the vascular damage - peripheral vessel issue has been completed. the device was not returned for investigation. the root cause of the vascular damage - peripheral vessel issue was most likely patient condition related. reported unknown relation to impella.
32774 | 1220648-2023-04963 | 18149734 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18149734&pc=OZD | 10/18/2023 | 07/17/2024 |  | 11/16/2023 |  |  | 10/18/2023 | 11/16/2023 |  |  | 05/31/2025 |  | 11/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024352062 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a 65 year old female patient was implanted with an impella cp pump for mechanical circulatory support. while preparing for explant, the patient failed attempts to wean off the pump and the p level was increased back up to p-7. at that point in time, clots were noted in the patient and the decision was made to withdraw care. | It was further reported that care was withdrawn and the patient subsequently expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there was not enough information to associate use of the device with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- b5 mso review in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
32775 | 1220648-2023-04964 | 18150424 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18150424&pc=OZD | 10/06/2023 | 01/24/2024 |  | 11/16/2023 |  |  | 10/19/2023 | 11/16/2023 |  |  | 08/31/2025 |  | 11/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024388547 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Hospitalization | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/19/2023 |  |  |  |  | The user facility reported a patient of unknown age, race, gender in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient developed a thrombotic occlusion of the right femoral artery due to decreased activated clotting time. this event required prolonged hospitalization. this event occurred during/after repositioning of the unit. the device was explanted successfully. | [5 empty] | The investigation into the reported thrombus has been completed since the original report was filed. no product was returned. as per clinical patient had ischemia due to occluded right femoral artery and ischemia resolved after explant. the cause of the limb ischemia issue was patient condition related due to patient having thrombotic occlusion of the right femoral artery per clinical review. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32776 | 1220648-2023-04966 | 18151089 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18151089&pc=PYX | 10/20/2023 | 03/26/2024 |  | 11/16/2023 |  |  | 10/20/2023 | 11/16/2023 |  |  | 07/31/2025 |  | 11/16/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024392296 |  | 0046-0035 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/20/2023 |  |  |  |  | The user facility reported a 76-year-old male in inferior st elevated myocardial infarction was implanted with an impella rp device for mechanical circulatory support. during support, a blood clot became entangled will the pump and caused the pump to stop. the patient was also implanted with an impella cp for bipella support. the impella rp was explanted and the patient was supported with the impella cp and medically managed. | [5 empty] | The investigation into the pump stop and the thrombus issues has been completed since the initial report was submitted. the product was not returned for investigation. as per the clinical, impella rp was removed due to clot entrainment, resulting in pump stop. the cause of the pump stop issue was most likely biomaterial, based on data log review and provided clinical information. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
32777 | 1220648-2023-04965 | 18151223 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18151223&pc=OZD | 10/19/2023 | 11/22/2024 |  | 11/16/2023 |  |  | 10/19/2023 | 11/16/2023 |  |  |  |  | 11/16/2023 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1420348 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Failure to Analyze Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 64-year-old female patient had been implanted with an impella 5.5 for bridge to ventricular assist device (vad)/transplant. when the nurse attempted to change the purge cassette, the automated impella controller (aic) would not detect the new purge disc. therefore, the aic was exchanged, and the problem resolved. there was no reported patient harm. | [5 empty] | The device was returned and an evaluation is pending. upon conclusion of the investigation, a supplemental report will be submitted with a summary of the results. | The investigation into automated impella controller (aic) - purge disc/flag issues has been completed. the data logs from the reported occurrence date confirmed that the console issued the purge disk not detected alarm numerous times. the console was tested and the issue with properly detecting the purge disk was reproduced and the purge disk flag was confirmed to be broken. the root cause of this issue is a broken/missing purge disc detect flag. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-04965.
32778 | 1220648-2023-04967 | 18151535 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18151535&pc=PYX | 10/20/2023 | 05/03/2024 |  | 11/16/2023 |  |  | 10/20/2023 | 11/16/2023 |  |  | 07/31/2025 |  | 11/16/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024388940 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Deformation | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/20/2023 |  |  |  |  | The user facility reported a 64-year-old male in acute inferior st elevated myocardial infarction was implanted with an impella rp device for mechanical circulatory support. during support and after turning the patient the repeat chest xray showed the rp prolapsed and was explanted. the patient was stable and transferred to icu for continued care. | [5 empty] | The investigation into the positioning issues and the product damage (cannula kink) has been completed since the original report was submitted. the device was not returned by the user facility for investigation. as per clinical details, impella rp was pulled out after big turn. the cause of the positioning issue was most likely use related. d6a, d6b. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32779 | 1220648-2023-04968 | 18151682 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18151682&pc=OZD | 10/18/2023 | 07/15/2024 |  | 11/16/2023 |  |  | 10/18/2023 | 11/16/2023 |  |  | 07/31/2025 |  | 11/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024388631 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/18/2023 |  |  |  |  | The user facility reported a 65-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient had loss of pulses in the leg after inserting the introducer sheath. the sheath was peeled away due to lose of pulses and the patient coded after that. soon after, return of spontaneous circulation was achieved but the pulses did not come back in leg. external bypass was done at bedside in the morning. the plan is to keep impella a couple more days for stabilization. three days later, the patient was converted to central ecmo cannulation. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. the instructions for use for the impella cp with smartassist system states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ added- d6a/d6b. f6/f8 should have been left blank in the initial report.
32780 | 1220648-2023-04970 | 18151852 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18151852&pc=OZD | 10/13/2023 | 06/08/2024 |  | 11/16/2023 |  |  | 10/20/2023 | 11/16/2023 |  |  | 06/30/2025 |  | 11/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024378426 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 44-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that while on impella support the patient experienced hemolysis. | The device was explanted the same day it was implanted. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation has been completed. according to the clinical information, shortly after beginning impella cp support, the patient began to develop hemolysis. no attempts were made to resolve the issue and the pump was explanted later the same day. no product was returned for a visual inspection. the data log analysis confirmed no abnormalities. the cause of the hemolysis was not determined because no product was returned and not enough clinical information was given. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ b.3 date of event revised as previously entered incorrectly. b.5 revised as information was inadvertently omitted from the previously submitted manufacture device report. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 added code 925 as it was inadvertently left off the previously submitted manufacture device report. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures.
32781 | 1220648-2023-04971 | 18151998 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18151998&pc=OZD | 10/20/2023 | 03/08/2024 |  | 11/16/2023 |  |  | 10/20/2023 | 11/16/2023 |  |  | 01/31/2025 |  | 11/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024334600 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a 65-year-old male patient on impella cp support experienced profuse access site bleeding. the site was not properly angle match and the patient required 2 units packed red blood cells (prbc), 1 l normal saline. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was filed. the device was not returned for investigation. according to clinical details, access site was not properly angle match and bleeding was resolved after applying a proper angle match. the cause of the access site bleeding issue was use related.
32782 | 1220648-2023-04969 | 18152630 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18152630&pc=OZD | 04/24/2023 | 03/25/2024 |  | 11/16/2023 |  |  | 10/18/2023 | 11/16/2023 |  |  | 10/31/2024 |  | 11/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024297629 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Heart Block | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/18/2023 |  |  |  |  | A notification received on 2023-10-18 via the long-term outcome and quality indicator (loqi) impella registry alleged a 78 year old male patient had gone into complete heart block with a "probable" relationship alleged with the impella device, as it was believed the catheter was bumping into the av node. the pump was repositioned in the cath lab and intervention was required via insertion of a temporary pacer. | New information received via abiomed¿s long-term outcome and quality indicator impella registry indicating a patient had endured post procedural bleeding with a "definite" relationship to the impella device. as a result, a wound vac was applied and multiple units of prbc were delivered. "good hemostasis" was noted after applying additional sutures. |  |  |  |  | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the access site bleeding major and the vf/vt/af/at issue has been completed since the initial report was submitted. the product was not returned for investigation. the root cause of access site bleeding was determined to be use issue since oozing that was noticed at the access site had been rectified after applying the additional sutures at the anastomosis. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined.
32783 | 1220648-2023-04972 | 18154133 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18154133&pc=OZD | 10/21/2023 | 09/20/2024 |  | 11/16/2023 |  |  | 10/21/2023 | 11/16/2023 |  |  | 07/31/2025 |  | 11/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024392241 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Unknown |  |  |  | Death | Yes | Yes | Improper Flow or Infusion | [6 empty] | Obstruction/Occlusion | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/21/2023 |  |  |  |  | The user facility reported a 53-year-old patient (gender and race unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella flows were very low and the left ventricle systolic numbers were negative. the physician noted the patient's vessels were 8mm and the hospital did not have stents that big. the patient's right coronary artery occluded. medications were given to the patient, but they were not effective in keeping the artery open. the patient expired on support. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the low pump flow issue has been completed since the original report was filed. the pump was returned, tested, and evaluated. biomaterial was observed in the purge gap. dried dextrose dissolved. the cause of the low pump flow most likely was biomaterial ingestion. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b1, b2, b5, h1, and h6.
32784 | 1220648-2023-04973 | 18154422 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18154422&pc=OZD | 10/18/2023 | 05/29/2024 |  | 11/16/2023 |  |  | 10/18/2023 | 11/16/2023 |  |  | 06/30/2025 |  | 11/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024373640 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/18/2023 |  |  |  |  | There is not enough information to associate use of device with outcome. | The user facility reported a 56-year-old female with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. patient developed reddish brown urine after pump becoming pulled back across the valve and then returned to position via echo guided repositioning. the next day, they were no longer able to assess a pulse in the right leg. nurse reports leg has a mottled appearance at this time. the physician came to bedside to speak with patient wife about possible options including fem-fem bypass and moving impella to axillary site; however, wife declined both options and chose to make patient comfort measures only and withdraw care. |  |  |  |  | The investigation into the hemolysis and the limb ischemia ¿ irreversible issues has been completed since the original report was submitted. the device was not returned for investigation. based on clinical description & data analysis, the root cause of hemolysis was most likely due to pump was not in the proper position & the root cause of limb ischemia was most likely patient condition related. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
32785 | 1220648-2023-04974 | 18154544 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18154544&pc=OZD | 10/18/2023 | 03/06/2024 |  | 11/16/2023 |  |  | 10/18/2023 | 11/16/2023 |  |  | 08/31/2025 |  | 11/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024386702 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Problem | Unintended Movement | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/18/2023 |  |  |  |  | The user facility reported a 74-year-old female in inferior st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient was taken to icu for recovery. upon arrival to icu, left groin access site presented hematoma and active bleeding. the patient was brought to or for site treatment. impella cp was pulled back into aorta while transferring patient to or table. cardiology deemed patient hemodynamically stable enough so device could be explanted in or. patient sent to icu for recovery. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding: major and positioning issues has been completed since the original report was filed. the device was not returned for investigation. the root cause of the access site bleeding - major issue was most likely patient movement since it was reported to have started after transfer to the icu. the root cause of the positioning issues was most likely patient movement as the pump was pulled back into the aorta while transferring the patient to the or table.
32786 | 1220648-2023-04975 | 18155774 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18155774&pc=OZD | 10/19/2023 | 06/03/2024 |  | 11/16/2023 |  |  | 10/19/2023 | 11/16/2023 |  |  | 08/31/2025 |  | 11/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024385532 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported 62-year-old female was implanted with impella 5.5 for mechanical circulatory support. during support, hematuria in the foley bag. echo was ordered and volume bolus was given. | [5 empty] | The investigation into the hemolysis issue has been completed since the original report was submitted. the device was not returned for investigation. data log confirmed the presence of suction alarms and no other abnormalities. the cause of the hemolysis was not determined because no product was returned, and limited clinical information was given. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
32787 | 1220648-2023-04976 | 18155831 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18155831&pc=OZD | 10/19/2023 | 01/11/2024 |  | 11/16/2023 |  |  | 10/19/2023 | 11/16/2023 |  |  | 05/31/2025 |  | 11/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024345969 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Failure to Advance | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/19/2023 |  |  |  |  | The user facility reported a 64-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the medical team had difficulty inserting the impella pump and while trying to pass the pump through the catheter along the wire it kinked and placement signal also became erratic. the physician performing the procedure decided to remove the pump and replace it with a new one. no patient harm reported due to the issue. | [5 empty] | The investigation into the low pump flow has been completed. it is reported that the pump was unable to be delivered as the patient had diseased arteries/vessels. the product was returned, and it was confirmed that it was kinked by the motor housing. the placement signal was erratic because the pump was not in position yet and trying to be delivered. the root cause of the delivery issues is most likely due to the patient¿s condition. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32788 | 1220648-2023-04977 | 18155859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18155859&pc=OZD | 10/19/2023 | 02/02/2024 |  | 11/16/2023 |  |  | 10/19/2023 | 11/16/2023 |  |  | 08/31/2025 |  | 11/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024394601 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/19/2023 |  |  |  |  | The user facility reported a 84-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient on impella support experienced low pump flows. despite aggressive anticoagulation therapy, the medical team noticed a clot in the pigtail. md removed pump for washout and decided to replace with a new pump. after new pump was inserted proceeded with the case with no complications. | [5 empty] | The investigation into the reported low pump flow has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. according to the clinical, low flow was recorded despite aggressive anticoagulation therapy, correct positioning, and sufficient volume. the product was not returned for a visual inspection. the logs confirm low flow and suction alarms. the cause of the low flow is not determined because no product was returned. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32789 | 1220648-2023-04978 | 18155871 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18155871&pc=OZD | 10/19/2023 | 02/07/2025 |  | 11/16/2023 |  |  | 10/19/2023 | 11/16/2023 |  |  |  |  | 11/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 110031 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 10/19/2023 |  |  |  |  | The user facility's biomedical engineer reported that the automated impella controller (aic) is displaying a controller error message (yellow alarm). the aic will be returned for evaluation. there was no patient involvement. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the controller failure _yellow alarm has been completed. console data logs confirm that controller error was sent. real-time logs from the confirmed the failure. the console was returned for evaluation and confirmed for the reported failure mode of a controller error. the failure has been reproduced using a an optical bench simulator and the root cause investigated to be a deficiency of the communication protocol of the optical bench, resulting in misalignment of data communication packets received by epc. console sw operated as designed. the root cause of the controller error (opm comm crc invalid) [optical pump connected] is still undetermined but is known pattern failure of transient loss of optical data.
32790 | 1220648-2023-04979 | 18155891 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18155891&pc=OZD | 10/20/2023 | 03/08/2024 |  | 11/16/2023 |  |  | 10/20/2023 | 11/16/2023 |  |  | 05/31/2025 |  | 11/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024363102 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A patient of unknown age received hemodynamic support from an impella cp smartassist heart pump during high-risk percutaneous coronary intervention (hrpci). the impella remained in the heart after the coronary procedure for further support. one day after inserting the impella cp, the purge filter was damaged and purge solution was leaking. the user repaired the damage with a purge filter bypass and the impella cp ran for another seven days without any problems. the patient suffered no harm from the incident. | [5 empty] | The investigation into the purge sidearm leak has been completed. neither the device nor logs were returned for evaluation. however, the clinical report has been reviewed. the root cause of the purge leak was due to a broken purge filter on the yellow luer lock. however, the exact location of the leak could not be identified. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
32791 | 1220648-2023-04980 | 18155892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18155892&pc=OZD | 10/20/2023 | 11/16/2023 |  | 11/16/2023 |  |  | 10/20/2023 | 11/16/2023 |  |  | 05/31/2025 |  | 11/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024365495 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the medical team had the 18" guidewire inserted then backloaded and delivered the impella cp pump across the hemostatic valve of the introducer without issue. the impella cp was then placed into position across the aortic valve under fluoroscopy. the 18" wire was then removed, and the impella cp was started in auto, with flows only reaching 2.5l/min. shortly after the impella reached 2.5l/min of flow in auto, the blood pressure of the patient began to drop on the automated impella controller (aic) and on the monitor. impella positioning was confirmed on fluoroscopy, and transducer was re-zeroed, the patient remained hypotensive. the physician called for a stat echo which revealed a large pericardial effusion with tamponade. pericardiocentesis was performed and in the next hour the patient coded five times which required the patient to be administered multiple inotropes and five units of blood. at this point an attempt was made to take the patient in the or for evaluation of possible left ventricle (lv) perforation, but unfortunately the patient expired. | [5 empty] | The investigation into the reported "ventricle perforation" has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the root cause of the ventricle perforation is most likely due to the guidewire being pushed in too deeply upon insertion of the pump. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿
32792 | 1220648-2023-04981 | 18155913 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18155913&pc=OZD | 10/20/2023 | 04/17/2024 |  | 11/16/2023 |  |  | 10/20/2023 | 11/16/2023 |  |  | 07/31/2025 |  | 11/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024391686 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 24 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Limb Pain | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the 24-year-old male patient on impella 5.5 support, complained of right arm and hand pain. a doppler study was performed and it¿s negative for deep vein thrombosis (dvt). the icu physician stated that the pain is probably from nerve damage from the cutdown site. | It was reported that the impella was explanted with surgical closure method. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
32793 | 1220648-2023-04983 | 18155922 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18155922&pc=OZD | 10/20/2023 | 04/17/2024 |  | 11/16/2023 |  |  | 10/20/2023 | 11/16/2023 |  |  | 08/31/2025 |  | 11/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024381474 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported the impella cp was placed successfully with ideal positioning on a 65-year-old male patient. as soon as the impella was positioning in the left ventricle (lv), the patient went into ventricular tachycardia (vt) which turned into ventricular fibrillation (vf). the cp was removed and intra-aortic balloon pump (iabp) was placed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. | The investigation for the reported ventricular fibrillation/ventricular tachycardia (vf/vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf/vt could not be determined.
32794 | 1220648-2023-04982 | 18155924 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18155924&pc=OZD | 10/20/2023 | 01/31/2024 |  | 11/16/2023 |  |  | 10/20/2023 | 11/16/2023 |  |  | 08/31/2025 |  | 11/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024390887 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported 59-year-old male on impella 5.5 support, was shocked out of ventricular fibrillation (vf) and amio bolused with drip was started with no repeat episodes after point of care ultrasonography (pocus) was completed. the impella was also noted to have been pulled back 2cm. according to cardiology fellow, increasing p level to p7 preceded the vf episode. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. | The investigation for the reported ventricular fibrillation has been completed. the impella pump was not returned for investigation. the data log showed a trend in placement signal and several impella unknown position alarms. according to clinical information, ventricular fibrillation was resolved with medication and repositioning of the pump. the cause of the ventricular fibrillation issue was use related, as the pump was found in improper position during the incident.
32795 | 1220648-2023-04984 | 18155940 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18155940&pc=OZD | 10/21/2023 | 08/09/2024 |  | 11/16/2023 |  |  | 10/21/2023 | 11/16/2023 |  |  | 08/31/2025 |  | 11/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024394612 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/21/2023 |  |  |  |  | The user facility reported a 63-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. during support, the patient's right leg became ischemic after impella placement, however, imaging showed significant collateralization suggested of pre-existing disease. passive fem-fem bypass placed which resolved it initially but the medical team decided that the patient would need escalation to axillary impella 5.5. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6a/d6b added f6/f8-should have been left blank on manufacturer device report.
32796 | 1220648-2023-04985 | 18155969 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18155969&pc=OZD | 10/23/2023 | 02/05/2024 |  | 11/16/2023 |  |  | 10/23/2023 | 11/16/2023 |  |  | 06/30/2025 |  | 11/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2024390577 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/23/2023 |  |  |  |  | The user facility reported a 81-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient had right pedal pulses still absent, foot cool and mottled. clot burden was noted upon removal of the impella in cath lab. reportable due to limb ischemia and thrombus. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported limb ischemia has been completed. the impella device was not returned for evaluation. however, clinical details stated that after the pump was removed, there was clot burden on the right femoral. therefore, the root cause of the limb ischemia was likely due to patient condition.
32797 | MW5148237 | 18159410 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18159410&pc=OZD | 07/21/2023 | 11/14/2023 |  | 11/16/2023 |  |  |  | 11/16/2023 |  |  |  |  | 11/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED INC. | [8 empty] | No | [9 empty] | Voluntary report | Not Available | Unknown | [6 empty] | 73 YR |  | Male | 71 KG | Non Hispanic | White | Death | Yes | Yes | [7 empty] | Cardiogenic Shock | [12 empty] | Death | [28 empty] | Initial submission | [9 empty] | On (b)(6) 2023, (b)(6) was admitted to the (b)(6) hospital system due to cardiogenic shock. on (b)(6) 2023, an impella 5.5 pump was implanted. his condition worsened and on (b)(6) 2023 the impella was removed. on (b)(6) 2023 mr. (b)(6) died.
32798 | 1220648-2023-04931 | 18139933 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18139933&pc=OZD | 10/18/2023 | 06/06/2024 |  | 11/15/2023 |  |  | 10/18/2023 | 11/15/2023 |  |  | 07/31/2025 |  | 11/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024389254 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/18/2023 |  |  |  |  | The user facility reported a 66-year-old male in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient experienced ventricular fibrillation on day four of impella support. the impella was repositioned at bedside. several hours later the patient was taken to the cath lab and the impella was repositioned under fluoroscopy and no additional coronary blockage was noted. it was further reported that the patient had red urine but did not have hemolyzed labs. the physician did not want to reposition the impella again due to difficulty with positioning previously. the patient is slowly being weaned for removal and is being closely monitored for worsening blood presence in urine. | There is not enough information to associate use of device with outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the vf/vt/af/at (vfib) and the positioning issues has been completed since the original report was submitted. the device was not returned for investigation. the cause of the positioning issue was not determined because no product was returned, and not enough clinical information was given. as the necessary information was not provided, the root cause of the vt/vf was not determined. d6a, d6b
32799 | 1220648-2023-04932 | 18140129 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18140129&pc=OZD | 10/16/2023 | 02/21/2024 |  | 11/15/2023 |  |  | 10/16/2023 | 11/15/2023 |  |  | 07/31/2025 |  | 11/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2024381724 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/16/2023 |  |  |  |  | The user facility reported a patient (unknown age and gender) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the impella purge suddenly became blocked. it was confirmed that there were no kinks. the purge cassette was replaced, but the purge pressure remained unchanged. the impella was replaced. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the high purge pressure issue has been completed since the original report was filed. the product was returned and the tested. no physical abnormalities were observed. saturated pump flow was observed during testing. data logs showed a motor current spike and saturated flow during the event of high purge pressure, which is sign of biomaterial interaction. the cause of the high purge pressure issue was biomaterial based on product evaluation and data log review.
32800 | 1220648-2023-04933 | 18140212 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18140212&pc=OZD | 10/16/2023 | 08/08/2024 |  | 11/15/2023 |  |  | 10/16/2023 | 11/15/2023 |  |  | 06/30/2025 |  | 11/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024375663 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/16/2023 |  |  |  |  | The user facility reported a 69-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had a hematoma and was oozing from the access site. two units of blood and platelets were transfused. a figure eight suture was placed by the physician to control the bleeding. the patient was stable. the decision was eventually made to explant the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was most likely use related since figure 8 suture were placed by the physician to control bleeding as per the clinical description. h6 component code added section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8-should have been left blank on manufacturer device report 1220648-2023-04933.
32801 | 1220648-2023-04934 | 18140312 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18140312&pc=OZD | 10/16/2023 | 09/20/2024 |  | 11/15/2023 |  |  | 10/16/2023 | 11/15/2023 |  |  | 07/31/2025 |  | 11/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024387276 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Positioning Problem | [5 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/16/2023 |  |  |  |  | The user facility reported a 34-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the staff were unable to increase p level for increased flow due to large clot in lv and active suction alarms. additionally, the patient had hemolysis and the pump position was greater than 6cm from annulus, but no alarms present on console. pump remeasured and was noted to be 4.9cm, but free floating in lv. the catheter was eventually repositioned and the issue was resolved. | Additional information noted the patient expired on support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can catheter blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6.
32802 | 1220648-2023-04935 | 18140419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18140419&pc=OZD | 10/07/2023 | 03/01/2024 |  | 11/15/2023 |  |  | 10/19/2023 | 11/15/2023 |  |  | 06/30/2025 |  | 11/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024368155 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility in japan reported a 64-year-old male patient was implanted with an impella 5.5 pump for mechanical circulatory support. prior to switching to impella support, the patient was supported via intra-aortic balloon pump (iabp) and ecmo for an emergency coronary artery bypass graft (cabg). following the operation, the patient was implanted with an impella 5.5 pump; however, it was also noted via ct evaluation that the patient had suffered a cerebral infarction. the exact onset of the infarction is unknown. due to the condition, the patient was given a dnr status. the patient¿s blood pressure subsequently decreased and the patient passed away. a definitive relationship between the impella and the patient¿s passing could not be established. | [5 empty] | The impella device was not received from the customer as it was discarded and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the stroke/cva event has been completed. the device was not returned for investigation. as the clinical information was not provided, the true root cause of the stroke/cva could not be determined.
32803 | 1220648-2023-04936 | 18140484 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18140484&pc=OZD | 10/17/2023 | 07/26/2024 |  | 11/15/2023 |  |  | 10/17/2023 | 11/15/2023 |  |  | 07/31/2025 |  | 11/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024390796 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that oozing was noted from impella cp site of the 87-year-old male patient. dressing was changed, pressure dressing applied, and groin oozing improved over night with additional stitch placement. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access site bleeding is determined to be use issue because groin oozing improved over night with additional stitch placement. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- h6 component code 925 in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
32804 | 1220648-2023-04937 | 18140510 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18140510&pc=OZD | 10/16/2023 | 04/22/2024 |  | 11/15/2023 |  |  | 10/16/2023 | 11/15/2023 |  |  | 08/31/2025 |  | 11/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024392966 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Detachment of Device or Device Component | Positioning Problem |  |  |  |  | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/16/2023 |  |  |  |  | The user facility reported a 74-year-old male post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. the patient had impella suction events due to the impella's position. patient had device repositioned under echo guidance post episode of vt requiring multiple defibrillations. device migration is believed to have triggered event, patient had been re-intubated for comfort and was on additional pressor support. additionally, during the attempt to reposition, the device was pulled back in the aorta and sterile sleeve detached from proximal connection. patient is hemodynamically stable at this time and has regained improvement of cardiac function. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation for the reported arrhythmia, low pump flow, sleeve damage, and repositioning issues has been completed. the pump was not returned for investigation. data logs suggest an active suction alarm and an impella position in the aorta alarm during the experience of low pump flow. according to clinical findings, the pump was found in an improper position and accidentally pulled back into the aorta during repositioning. this condition of low pump flow and improper position was resolved after following the proper repositioning steps. damage was noted on the sterile sleeve during the repositioning of the device. the cause of the positioning issue was most likely use related, as clinical information suggests the device was pulled back into the aorta during repositioning, leading to a low pump flow issue. the cause of the sterile sleeve damage issue was most likely use-related, as it occurred during repositioning of the device. as the necessary information and returned device were not provided, the root cause of the vt/vf was not determined.
32805 | 1220648-2023-04938 | 18140703 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18140703&pc=OZD | 10/19/2023 | 01/25/2024 |  | 11/15/2023 |  |  | 10/19/2023 | 11/15/2023 |  |  | 07/31/2025 |  | 11/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024390232 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/19/2023 |  |  |  |  | The user facility reported a 52-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. placement signal was unreliable with loss of placement signals when switching the automated impella controller, plus they were unable to maintain support without manual pressure on impella cable connection to maintain connection. so the decision was made to remove pump and replace it to address that issue. no patient harm reported due to the issue. | [5 empty] | The investigation into the optical signal issue has been completed since the original report was filed. the product was returned and upon visually inspection, a piece of the inner patient cable had broken off and was in the optical fiber port. before removing the blocker, the pump was plugged into a benchtop aic. no signal was detected, the pump was unable to start, and it was unable to fully seat in the connector. after removal of the blocker, the pump was plugged back into the benchtop aic. a signal was detected, the pump was able to start, and it was fully seated in the connector. the field logs show the pump running on the second console without placement signal. this is likely after the tape was applied to increase the connection. the cause of the pump stop was determined to be a damaged patient cable connector which was preventing connection of the pump to the console. this damage most likely occurred during the transfer of the pump from the rex to the duke console. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32806 | 1220648-2023-04939 | 18140806 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18140806&pc=OZD | 10/16/2023 | 06/04/2024 |  | 11/15/2023 |  |  | 10/16/2023 | 11/15/2023 |  |  | 09/30/2025 |  | 11/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024395639 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Swelling/ Edema | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/16/2023 |  |  |  |  | The user facility reported a 65 year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the patient had a hematoma develop at the impella site coinciding with right arm swelling due to a partial thromboplastin time(ptt) up to 81.5. heparin paused temporarily and ptt is now normal at 48.5 and hematoma stabilized and is dissipating. the staff felt that the swelling was due to congestion from hematoma and will start decreasing now that bleeding had stopped. hematoma resolved and patient was considered stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access site bleeding was determined to be anticoagulation management because the ptt values were too high and heparin was stopped resulted in stabilizing the hematoma. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
32807 | 1220648-2023-04940 | 18141105 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18141105&pc=OZD | 10/17/2023 | 04/03/2024 |  | 11/15/2023 |  |  | 10/17/2023 | 11/15/2023 |  |  |  |  | 11/15/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1627246 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 10/17/2023 |  |  |  |  | The user facility reported a 76-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, the automated impella controller (aic) displayed a placement signal error. the aic was replaced and the placement signal error was resolved. there was no patient harm related to this event. | [5 empty] | The investigation into the aic - pump issue has been completed since the initial report was submitted. the console was returned, visually inspected, and tested. the cause of the aic issue was a defective optical bench. the optical bench displayed noise that disrupted signal-to-noise ratio (snr) measurement. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
32808 | 1220648-2023-04941 | 18141944 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18141944&pc=OZD | 10/19/2023 | 04/24/2024 |  | 11/15/2023 |  |  | 10/19/2023 | 11/15/2023 |  |  | 04/30/2025 |  | 11/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024348261 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Restricted Flow rate | Activation, Positioning or Separation Problem | [5 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/19/2023 |  |  |  |  | The user facility reported a male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient had red urine output/hemolysis. additionally, there were suction events on the impella and the impella was repositioned under transesophageal echocardiography (tee) guidance, inlet cage appeared to be getting caught in the mitral apparatus. patient was on impella support for 5 additional days post the issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the hemolysis, the low pump flow and the positioning issues has been completed since the original report was submitted. the device was not returned for investigation. the cause of the hemolysis issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the low pump flow issue was an interaction with crrt since the flow management on crrt successfully resolved the issue. the cause of the positioning issue was not determined since product was not returned and sufficient clinical information was not provided. clinical details also indicated a limb ischemia issue. because all the relevant information was not provided and the device was not returned, the root cause of the limb ischemia was undetermined.
32809 | 1220648-2023-04942 | 18142051 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18142051&pc=OZD | 10/16/2023 | 09/26/2025 |  | 11/15/2023 |  |  | 10/17/2023 | 11/15/2023 |  |  | 01/31/2025 |  | 11/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024329046 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/17/2023 |  |  |  |  | The user facility reported a 58-year-old male in anterior st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. during support, the cannula was found to have a kink between the white catheter shaft and the red plug. after numerous troubleshooting and echo images the pump is found be in good position and team was educated on use of motor current and flows for patient. there was no harm to the patient because of this incident. | The complainant also reported loss of optical signal. the placement signal and left ventricle signal waveforms were not available on the automated impella controller (aic) after multiple troubleshooting attempts. changing the aic did not resolve the issue. imaging confirmed the pump was in the correct position. after becoming aware of the loss of the optical sensor, the decision was made to leave the patient on support. |  |  |  |  | The investigation has been completed since the original report was submitted. according to the clinical information, on the first day of impella cp support, the patient arrived to the intensive care unit (icu) with a kink in the catheter. after a loss of the placement signal and left ventricle signal waveforms, the team attempted to resolve the issue with no success. although it was determined that they had lost the optical sensor, the team elected to leave the patient on support. no product was returned for a visual inspection. data log analysis confirmed the optical signal issue was during support and that the 5s parameters aligned with a kink in fiber line. the signal-to-noise ratio, light and contrast were decreasing while the lamp and gain increased with placement signal being lost. the most likely cause of the optical signal issue was a damaged fiber line with kink in catheter occurring after patient transfer to icu. b.7 revised as this information was inadvertently omitted from the previously submitted report. d.6a and d.6b revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | B5 updated to include placement signal issue description. h6 updated to include placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2023-04942. b7 other relevant history, including preexisting medical conditions updated. d4 expiration date updated. d10 concomitant medical products and therapy dates updated. f6 and f8 should have been left blank. g1 updated with correct mdr reporting contact email.
32810 | 1220648-2023-04943 | 18142074 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18142074&pc=OZD | 10/17/2023 | 01/24/2024 |  | 11/15/2023 |  |  | 10/17/2023 | 11/15/2023 |  |  | 07/31/2025 |  | 11/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024392211 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/17/2023 |  |  |  |  | The user facility reported a 74-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during implant, the impella sheath was found to have kinked and wires were no longer able to pass. a new long peel away sheath was placed slowly without issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues- use has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based on the clinical information provided, the long peel away introducer sheath was inserted and kinked due to tortuous vessels with wire unable to pass. second long introducer sheath placed slowly over same access and pump placed successfully. the cause of the introducer damage issue was introducer sheath kink due to tortuous vessel condition during insertion. the failure mode will be monitored and trended.
32811 | 1220648-2023-04944 | 18142119 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18142119&pc=OZD | 10/17/2023 | 02/21/2024 |  | 11/15/2023 |  |  | 10/17/2023 | 11/15/2023 |  |  | 07/31/2025 |  | 11/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024378414 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Unspecified Infection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/17/2023 |  |  |  |  | The user facility reported a 45-year-old female with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support while awaiting heart transplant. while on impella support, the patient had purulent drainage coinciding with an increase in white blood cells count indicating a potential infection. the patient was started on antibiotics to treat the noted condition. furthermore, it also identified that the patient developed a hematoma prompting a temporary halt to heparin administration. support with the implanted pump was maintained throughout the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related events are as follows: the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported infection and access site bleeding has been completed since the original report was filed. the root cause of the infection was not determined. the device is sterilized under validated conditions and there was no evidence to to indicate a break in the sterile barrier or pump contamination. unknown relation to impella. the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical information provided.
32812 | 1220648-2023-04945 | 18142155 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18142155&pc=OZD | 11/09/2023 | 03/08/2024 |  | 11/15/2023 |  |  | 11/09/2023 | 11/15/2023 |  |  |  |  | 11/15/2023 |  | IMPELLA AUTOMATED CONTROLLER (AIC) |  |  | IMPELLA AUTOMATED CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1108711 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/09/2023 |  |  |  |  | The us complainant had an in-service ongoing when the console (aic) battery alarmed. there was no patient use or harm/injury possible. the in-service revealed the red alarm. the aic was removed from service for analysis. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console (aic) battery failure alarm has been completed. the aic was returned for evaluation. during testing of the aic, the failure mode was reproduced on an engineering lab bench. the battery failure alarm was due to a defective battery. the aic logs were reviewed and confirmed the reported failure. the root cause of the defective battery was due to single cell failure, a known pattern failure.
32813 | 1220648-2023-04946 | 18142211 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18142211&pc=OZD | 10/17/2023 | 05/16/2024 |  | 11/15/2023 |  |  | 10/17/2023 | 11/15/2023 |  |  | 08/31/2025 |  | 11/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024394573 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/17/2023 |  |  |  |  | The user facility reported a 64-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. patient¿s urine /labs showed signs of hemolysis which was resolved with repositioning of the device. patient is stable. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore the investigation for the reported hemolysis has been completed., the root cause of the stroke could not be determined. added-the instructions for use referenced in the initial are from the impella cp with smart assist system section stated potential adverse events (united states). in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
32814 | 1220648-2023-04947 | 18142232 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18142232&pc=OZD | 10/18/2023 | 08/21/2024 |  | 11/15/2023 |  |  | 10/18/2023 | 11/15/2023 |  |  | 05/31/2025 |  | 11/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024362273 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Impairment | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/18/2023 |  |  |  |  | The user facility reported a 49-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced hemolysis which was diagnosed by tinged foly/urine. contributing factors small left ventricle, organ damage. continuous renal replacement therapy was initiated. | It was further reported that the patient expired while on support. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions including catheter position, pre-existing medical conditions, and small left ventricular volumes may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the hemolysis has been completed. data logs were reviewed which showed about 1.2515 minutes into the case, pump was in unknown position with low flow corresponding with the time the clinical team found pump was shallow. pump then ran without issue to the rest of the case, product was not returned for review. based on clinical description and data analysis, the root cause of hemolysis was most likely due to pump was not in the proper position. the instructions for use for the impella cp with smartassist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added-b5 mso review, h6 component code 925 in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8 should have been left blank in the initial report.
32815 | 1220648-2023-04948 | 18142765 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18142765&pc=OZD | 11/13/2023 | 09/22/2024 |  | 11/15/2023 |  |  | 11/13/2023 | 11/15/2023 |  |  | 08/31/2025 |  | 11/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024396384 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Separation Problem | [6 empty] | Hematoma | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/13/2023 |  |  |  |  | The user facility reported a 52-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella was implanted with no issues but upon attempted removal of wire most came out, but a portion got stuck, frayed, and broke off. still in patient as the physician did not want to remove as this time, patient has bilateral hematomas to both groin sites and ij site. pt received 7 units of blood and 1 unit ffp. | The procedural outcome noted that the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-04948 was provided in sections b1 and h1. note: corrected information provided in h6 is an addition to previous coding.
32816 | 1220648-2023-04949 | 18143627 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18143627&pc=OZD | 10/16/2023 | 01/24/2024 |  | 11/15/2023 |  |  | 10/16/2023 | 11/15/2023 |  |  | 04/30/2025 |  | 11/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024337479 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 9 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Separation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/16/2023 |  |  |  |  | The user facility reported a 9-year-old male with heart failure due to gene mutation was implanted with an impella 5.5 as a bridge to implant. during explant, the impella motor housing was pulled away from catheter due to cross clamp across motor housing. there was no patient harm related to this event. patient successfully transitioned to transplant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the product damage (detached inflow/outflow - peripheral vessel) issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based on the clinical information provided, the root cause of the product damage is mostly due to a use issue as they had a cross clamp placed across the motor housing.
32817 | 1220648-2023-04950 | 18144371 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18144371&pc=OZD | 10/17/2023 | 04/22/2024 |  | 11/15/2023 |  |  |  | 11/15/2023 |  |  | 05/31/2025 |  | 11/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024359195 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/17/2023 |  |  |  |  | The user facility reported unknown age male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. pump insertion was reported to be difficult due to patient anatomy. after the impella was placed in the left ventricle, it delivered only low blood flow. repositioning the impella did not bring any improvement. the impella cp was removed and biomaterial was visible on the pump after removal. it was decided to continue the pci procedure without impella support. the patient suffered no harm as a result of the incident. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow and the delivery issues- anatomy has been completed since the original report was filed. the pump was returned, tested, and evaluated. the root cause of the delivery issues is due to the patient¿s condition. the returned pump confirmed that there was biomaterial wrapped around the impeller. the root cause of the low pump flow is most likely due to ingested biomaterial.
32818 | 1220648-2023-04951 | 18144386 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18144386&pc=OZD | 10/17/2023 | 08/26/2024 |  | 11/15/2023 |  |  | 10/17/2023 | 11/15/2023 |  |  | 03/31/2025 |  | 11/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024347290 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/17/2023 |  |  |  |  | The user facility reported a 69-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. patient bleeding from the access site in the right groin. area was redressed and femstop placed with bleeding controlled. patient had a large hematoma and the staff made the decision to take the impella out and place 16fr cook sheath and place impella with fem-fem external for flow to the leg. good distal pulses. all sheaths secured with no oozing or bleeding. | The patient was made do not resuscitate (dnr) and expired while on support. per medical safety officer review use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-04951. g1 reporting contact fax number missing. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
32819 | 1220648-2023-04952 | 18144396 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18144396&pc=OZD | 10/17/2023 | 05/14/2024 |  | 11/15/2023 |  |  |  | 11/15/2023 |  |  | 04/30/2025 |  | 11/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024352814 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Migration | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was an "impella in aorta" alarm during a planned percutaneous coronary intervention (pci). the impella cannula shifted frequently during the case needing positional adjustments but patient remained stable just with positioning of the interventional equipment. the impella cp was moved back into optimal position. there was no patient harm. post pci, the patient remained hemodynamically stable and the impella weaned and explanted without issue. | [5 empty] | The investigation into the positioning issues has been completed since the original report was submitted. the device was not returned for investigation. according to the clinical, shortly after beginning impella cp support, an impella in aorta alarm was recorded. this occurred during a planned pci with positioning of interventional equipment utilizing shockwave. to resolve the issue, the pump was repositioned back into optimal position successfully. the most likely cause of the positioning issues was use related. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿
32820 | 1220648-2023-04953 | 18144512 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18144512&pc=OZD | 10/18/2023 | 05/16/2024 |  | 11/15/2023 |  |  |  | 11/15/2023 |  |  | 05/31/2025 |  | 11/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024353906 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/18/2023 |  |  |  |  | The user facility reported a 78-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the purge fluid was leaking from the side arm at the purge reservoir site. as a result, decision was made to wean and explant the impella. no patient harm reported due to the issue. | [5 empty] | The investigation for the purge leak has been completed. no product was returned for a visual inspection. data log analysis was not necessary for this complaint. according to the clinical, on the seventh day of impella cp support purge fluid was observed leaking from the sidearm. the cause of the pump purge leak was most likely a damaged sidearm, but the exact location of the leak could not be determined without return product. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32821 | 1220648-2023-04954 | 18144524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18144524&pc=OZD | 10/17/2023 | 05/07/2024 |  | 11/15/2023 |  |  |  | 11/15/2023 |  |  | 08/31/2025 |  | 11/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024393634 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Improper Flow or Infusion | [6 empty] | Stroke/CVA | Tachycardia | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a 57-year-old male patient on impella 5.5 support experienced ventricular tachycardia and suction events. the impella positioning was checked and appeared to be in the correct positioning. three days after these events, the physician removed the impella from direct placement and the patient had a severe middle cerebral artery (mca) stroke with bleeding and the patient ultimately expired. the physician believed that impella was the cause of stroke upon removal. | [5 empty] | The investigation into the reported issues has been completed since the original report was submitted. the device was not returned by the user facility for investigation. the cause of the low pump flow issue was most likely patient condition related, based on data log review and provided clinical details. as relevant clinical information related to the event was not provided, the root cause of the stroke/cva could not be determined. the root cause of the vf/vt/af/at was not determined because the necessary information was not provided, and the device was not returned. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32822 | 1220648-2023-04955 | 18144525 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18144525&pc=OZD | 10/18/2023 | 02/07/2024 |  | 11/15/2023 |  |  | 10/18/2023 | 11/15/2023 |  |  | 07/31/2025 |  | 11/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024370550 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Remove | Detachment of Device or Device Component | [5 empty] | Cardiac Perforation | Aortic Valve Insufficiency/ Regurgitation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/18/2023 |  |  |  |  | The user facility reported a 68-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support post coronary artery bypass grafting and mitral valve repair. when the patient was weaned from the impella 5.5 and brought to the or for pump removal, a great deal of resistance was met. the surgeon checked transesophageal echocardiography for location. after multiple attempts to remove the device, the impella 5.5 broke in the graft just above the motor. the patient's chest was opened to remove the graft and pump. post chest closure, the anesthesiologist noted the patient had moderate aortic insufficiency (ai). it was noted that the team did not believe the patient had ai prior to this procedure. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the reported removal issues, heart valve damage and cannula /pigtail detachment has been completed since the original report was filed. the product was returned and it was confirmed that the cannula was detached as the catheter broke and there was biomaterial found in the outflow/inflow cage. per the clinical details, the physician attempted to pull through the clot that was present using excessive force without assessing the resistance that was experienced. the root cause of the product damage (detached inflow/outflow-great vessel) and the heart valve damage is most likely due to use as the cannula became detached in the av node due to the excessive force used.
32823 | 1220648-2023-04956 | 18144545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18144545&pc=OZD | 10/16/2023 | 08/25/2024 |  | 11/15/2023 |  |  | 10/16/2023 | 11/15/2023 |  |  | 07/31/2025 |  | 11/15/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024389432 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/16/2023 |  |  |  |  | The user facility reported a 56-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. during support, the patient developed lower limb ischemia. the patient was on four pressors. the decision was made to escalate the patient to an impella 5.5. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Section h6 investigation coding has been updated, as the investigation has been completed. according to clinical, the patient was on four pressers and had developed lower limb ischemia. the patient arrived at cardiovascular intensive care unit and right lower extremity was cooler compared to left; no pedal pulse present. patient has a history of peripheral vascular disease. discussion made regarding escalation to an impella 5.5. the impella device was not return, and therefore an evaluation of the device was not possible. limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ·patients pre-morbid condition and peri-procedural condition ·any concomitant medication ·vascular size or variations thereof ·detailed description of the symptoms experienced by the patient. ·physical examination findings, including pulse assessment and visual examination of the affected limb. ·diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ·laboratory test results, including blood tests for markers of inflammation or clotting disorders. ·any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ·patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. correction was made to the following section(s). refer to said section for corrected data. b7: other relevant history. d6a implant date. d6b explant date. f6: date captured on the initial report should be disregard; section not applicable. f8: date captured on the initial report should be disregard; section not applicable.
32824 | 1220648-2023-04922 | 18131322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18131322&pc=OZD | 10/17/2023 | 02/06/2024 |  | 11/14/2023 |  |  | 10/17/2023 | 11/14/2023 |  |  | 07/31/2025 |  | 11/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024389251 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/17/2023 |  |  |  |  | The user facility reported a 57-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, there was a pump stop with multiple unsuccessful restarts. multiple conversations had with team in previous weeks about longevity of pump and its run time. md wanted to wait until decision was made by accepting hospital before performing change out to a new pump. cardiology was called after pump stoppage and pump was pulled into descending aorta. the patient was taken to cath lab and impella replaced with new impella cp. no issues noted on removal. there was no patient harm related to this event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation for the reported pump stop has been completed. the impella device was not returned for evaluation. however, data log was reviewed and it revealed 119 pump stops on day twenty. multiple attempts were made to restart the device with no success. without additional clinical evaluation and device not be returned, the root cause of the pump stop could not be determined.
32825 | 1220648-2023-04923 | 18136918 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18136918&pc=OZD | 10/30/2023 | 09/09/2024 |  | 11/14/2023 |  |  | 10/30/2023 | 11/14/2023 |  |  | 08/31/2025 |  | 11/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024396477 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/30/2023 |  |  |  |  | The user facility reported a 77-year-old male post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. the impella pump was difficult to pass though anastomosis and aortic valve. due to continuous pressure and force, the catheter on the pump kinked about 6 cm above the motor. the impella insertion was aborted and a different impella 5.5 was placed without issue. no patient harm was reported due to the issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was use related. h6 additional medical device problem code added instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella 5.5 with smartassist for use during cardiogenic shock section: warnings and cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings and cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings and cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-04923 in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8-should have been left blank g1 reporting contact fax number missing
32826 | 1220648-2023-04924 | 18136978 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18136978&pc=OZD | 10/16/2023 | 05/22/2024 |  | 11/14/2023 |  |  | 10/16/2023 | 11/14/2023 |  |  | 04/30/2025 |  | 11/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024366169 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported there was a kink in the impella cp outflow. the impella was therefore replaced. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer as it was discarded and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the product damage (cannula kink) issue has been completed since the original report was submitted. according to the clinical information, during impella cp insertion the cannula was kinked close to outflow resulting in flows of only 1.6-1.8. liter per minute. after no success with repositioning the pump, the decision was made to explant the device. no product was returned for a visual inspection. data log analysis confirmed low flows and no other abnormalities. the most likely cause of the low pump flow was a cannula kink. b.7 revised as this information was inadvertently omitted from the previously submitted report. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 code 2199 was reported incorrectly on the previously submitted report.
32827 | 1220648-2023-04925 | 18136979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18136979&pc=OZD | 10/16/2023 | 01/22/2024 |  | 11/14/2023 |  |  | 10/16/2023 | 11/14/2023 |  |  | 01/31/2025 |  | 11/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024316070 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female |  |  |  | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/16/2023 |  |  |  |  | The user facility reported a 44-year-old female (race unknown) with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support, the purge sidearm plastic filter cracked and was leaking. the team performed a purge sidearm bypass. the pump was not explanted. there was no patient harm related to this event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported purge sidearm leak has been completed. since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the purge leak-pump cannot be determined due to the product not being returned to confirm a leak or the location. the data logs were also not returned.
32828 | 1220648-2023-04926 | 18136982 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18136982&pc=OZD | 10/16/2023 | 01/12/2024 |  | 11/14/2023 |  |  | 10/16/2023 | 11/14/2023 |  |  | 07/31/2025 |  | 11/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024388784 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/16/2023 |  |  |  |  | The user facility reported a 74-year-old female post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. the pump had persistent suction problems and, despite the device being repositioned, the issue persisted. the decision was made to escalate the patient to an impella 5.5. there was no patient harm related to this event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow has been completed. the product was not returned. as per the clinical information provided, the placement signal is very erratic and there are low flows with suction alarms. the flow then increases abnormally with no change in p-level. the root cause of the low pump flow is most likely due to the patient¿s condition.
32829 | 1220648-2023-04927 | 18136983 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18136983&pc=OZD | 10/16/2023 | 02/21/2024 |  | 11/14/2023 |  |  | 10/16/2023 | 11/14/2023 |  |  | 07/31/2025 |  | 11/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024377024 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | No | Other | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/16/2023 |  |  |  |  | The user facility reported a 60-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, there were multiple "impella in aorta" alarms. the patient was unable to tolerate the impella and the choice was made to explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported positioning issue has been completed since the original report was filed. data review: data confirmed the failure mode. logs showed about 20.5 hours into the case, pump was in aorta that triggered impella position in aorta alarms. about 4 hours later, pump was disconnected from the console & the support ended. product was not returned for review. based on clinical description & data review, the root cause of positioning issue was most likely use related.
32830 | 1220648-2023-04928 | 18137021 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18137021&pc=OZD | 10/19/2023 | 05/07/2024 |  | 11/14/2023 |  |  | 10/19/2023 | 11/14/2023 |  |  | 07/31/2025 |  | 11/14/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024392354 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Improper Flow or Infusion | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that when inserting the impella cp on the 63-year-old male patient, unresolved suction alarms occurred and the impella could not go above p6. the physician looked at the right ventricle (rv) and said it was struggling. arterial blood gas (abg) was drawn and the patient was not saturating well. the decision was made to place ecmo and impella was brought down to p2. however, the impella kept getting suction alarms, but ecmo team said the patient had good volume. it was further reported that just before taking the patient off the table, the patient¿s left foot, same side as impella, was blue. the patient was brought to ccu and ct surgeon implanted the external perfusion catheter on the impella side. the surgeon reported that the left foot color looked better. the patient also had a little oozing at the groin. the patient has been moving a lot, which may have caused the oozing. leg immobilizer placed, dressing changed, and oozing seems better. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ | The investigation for the reported limb ischemia and low pump flow has been completed. the device was not returned for evaluation. clinical details were not sufficient to determine the root cause of the limb ischemia. without additional clinical details, the root cause of the limb ischemia could not determined. data logs were reviewed which showed a trend in placement signal with reported suction alarms and low pump flow during the case run. according to clinical observations, the right ventricle was struggling, and blood volume was administered during the case run. the cause of the low pump flow issue was most likely patient condition related, based on data log review. added-b1 selected product problem. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
32831 | 1220648-2023-04929 | 18137029 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18137029&pc=OZD | 10/19/2023 | 11/14/2023 |  | 11/14/2023 |  |  | 10/19/2023 | 11/14/2023 |  |  |  |  | 11/14/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1116075 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Injury | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a broken purge clip on the automated impella controller (aic). there was no reported patient harm. | [5 empty] | The investigation into the reported automated impella controller (aic) purge disc/flag issues has been completed. review of data logs is not relevant to this complaint as the clinical staff only reported that the aic had a broken purge clip. the returned console was examined arrival and the purge disk retainer on this console was confirmed to be broken. the root cause of the reported event was determined to be the broken purge retainer.
32832 | 1220648-2023-04908 | 18121300 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18121300&pc=OZD | 10/23/2024 | 03/13/2024 |  | 11/13/2023 |  |  | 10/14/2023 | 11/13/2023 |  |  | 08/31/2025 |  | 11/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024390858 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Unspecified Infection | Ventricular Fibrillation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The us complainant reported the 63-year-old male patient was showing concern for abscess around impella 5.5 access site. the patient was given antibiotics and support with the implanted pump was maintained. | Additional information received thru a notification received via abiomed¿s long-term outcome and quality indicator impella registry indicating this patient to have endured ventricular tachycardia / premature ventricular contractions (vt/pvc) with a "possible" relationship to the impella device. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | Instructions for use for the related event are as follows: potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation for the reported infection and ventricular tachycardia (vt) has been completed. the impella device was not returned for evaluation. without additional clinical information, the root cause of the infection nor the vt cannot be determined.
32833 | 1220648-2023-04909 | 18121500 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18121500&pc=OZD | 10/15/2023 | 05/14/2024 |  | 11/13/2023 |  |  | 10/15/2023 | 11/13/2023 |  |  | 07/31/2025 |  | 11/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024387983 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the 62-year-old female patient on impella cp support was experiencing suction alarms and started bleeding from somewhere, the urine was amber in color, and the patient was found to have hemolysis. the patient was required to receive 2 units of blood. when the patient showed to have slightly cooler and more pale leg decision was made to explant the impella. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. | The investigation for the reported hemolysis has been completed. no product was returned for a visual inspection. data log analysis confirmed the presence of suction alarms. no motor current spike or positioning issues found. according to the clinical, during cp pump support bleeding concerns, reduced preload was noted and were believed to have resulted in hemolysis. multiple blood products and volume were given, but the decision was made to explant the device due to limb ischemia concerns. also, placement signal decreasing close to end of case relative to hemolysis event. the most likely cause of the hemolysis was patient condition. potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
32834 | 1220648-2023-04910 | 18121701 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18121701&pc=OZD | 11/03/2023 | 04/23/2024 |  | 11/13/2023 |  |  | 11/03/2023 | 11/13/2023 |  |  | 08/31/2025 |  | 11/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024385750 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/03/2023 |  |  |  |  | The user facility reported a 68-year-old patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported a drop in pump flow during use of the device. troubleshooting attempts were made to resolve the issue, to no avail. the pump was removed and a thrombus was found in the cannula. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported low pump flow has been completed. the impella device was returned for evaluation. visual inspection found no issues on the pump. the failure mode could not be reproduced as the pump ran without issue under bench test. data logs show a spike in motor current and active suction alarm, followed by low pump flow during the start of the case. the root cause of low flow was most likely biomaterial ingestion as the thrombus was visible when pump was removed. the instructions for use states ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ added- d9 device return date, h6 impact code 4628 d4-updated model number.
32835 | 1220648-2023-04911 | 18121754 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18121754&pc=OZD | 10/16/2023 | 01/18/2024 |  | 11/13/2023 |  |  | 10/16/2023 | 11/13/2023 |  |  | 08/31/2025 |  | 11/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024394552 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/16/2023 |  |  |  |  | The user facility reported a 54-year-old male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the patient did not have an ischemic limb prior to impella insertion. the patient was on multiple high dose pressors and quickly declining. when the physician assessed the femoral artery he knew it was likely to be ischemic when impella was placed, but patient needed rapid support. at end of case the limb was found to have no pulse and putting in an antegrade perfusion sheath was discussed. the physician made the decision to transfer for 5.5 placement due to a need for escalation in support and to remove the cp and help re-perfuse the affected limb. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported limb ischemia has been completed. the impella device was not returned for evaluation. however, clinical information provided was sufficient to determine the root cause. as per clinical, the patient had ischemia on the impella leg and the pump was explanted later to resolve complication. the cause of the limb ischemia issue was most likely patient condition.
32836 | 1220648-2023-04912 | 18121859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18121859&pc=OZD | 10/16/2023 | 04/23/2024 |  | 11/13/2023 |  |  | 10/16/2023 | 11/13/2023 |  |  |  |  | 11/13/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1376304 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) had a controller error alarm. there was no reported patient harm. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the initial report was submitted. the console was returned and tested. the issue with the purge disc not consistently being detected was reproduced. the root cause of the issue was determined to be that the purge disc retainer was loose due to a broken internal mounting post.
32837 | 1220648-2023-04913 | 18122064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18122064&pc=OZD | 10/16/2023 | 11/13/2023 |  | 11/13/2023 |  |  | 10/16/2023 | 11/13/2023 |  |  |  |  | 11/13/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1445594 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported 72-year-old male patient on impella 5.5 support experienced purge alarms overnight. the rn discovered during cassette change that blue purge disc holder was loose and coming off the automated impella controller (aic). the aic was therefore exchanged. there was no reported patient harm. | [5 empty] | The investigation into the reported automated impella controller (aic) purge disc/flag issues has been completed. there were no logs associated with the reported occurrence date. this could be potentially due to the clinical staff misreporting the occurrence date, as the logs suggest that the reported issues occurred on (b)(6) 2023. during testing, the purge retainer was confirmed to be broken. the root cause of the reported event was determined to be a broken purge disk retainer.
32838 | 1220648-2023-04914 | 18122567 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18122567&pc=OZD | 10/15/2023 | 01/12/2024 |  | 11/13/2023 |  |  | 10/15/2023 | 11/13/2023 |  |  |  |  | 11/13/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1084349 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 10/15/2023 |  |  |  |  | The user facility reported a 57-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, the automated impella controller (aic) displayed "purge not detected" alarm. the aic was replaced successfully. there was no patient harm related to this event. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the console was returned for investigation, and it was visually inspected under a digital microscope. cracks on the lateral side of the purge retainer as well as a missing purge flag were observed. the root cause for the purge disc not detected alarms was a defective purge pressure sensor.
32839 | 1220648-2023-04915 | 18122941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18122941&pc=OZD | 10/15/2023 | 01/31/2024 |  | 11/13/2023 |  |  | 10/15/2023 | 11/13/2023 |  |  | 07/31/2025 |  | 11/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024374119 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/15/2023 |  |  |  |  | The user facility reported a 65-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the patient had elevated lactate dehydrogenase and the physician was unable to reposition the pump off the mitral valve. the physician exchanged the pump to achieve optimal positioning to alleviate increased risk for hemolysis. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow issue has been completed. the pump was not returned for investigation. as per clinical, doctor found pump to be deep near mitral valve/apparatus. the cause of the positioning issue was most likely use related as the replacement 5.5 was able to be placed and positioned without issue.
32840 | 1220648-2023-04916 | 18123002 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18123002&pc=OZD | 10/28/2023 | 01/22/2025 |  | 11/13/2023 |  |  | 10/29/2023 | 11/13/2023 |  |  |  |  | 11/13/2023 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1325612 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 10/29/2023 |  |  |  |  | The user facility in japan reported a 65-year-old male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during the support a yellow controller error alarm was triggered on the automated impella controller (aic). it was noted that the alarm emitted 3 times within the 12 hour period .no patient harm reported and the device support was continued. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the console (aic) controller error alarm has been completed. data logs confirmed the reported failure mode given there were multiple controller error alarm (opm comm crc error ¿ event set # 1045) triggered during the clinical procedure. real time (rt) logs confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) are represented as -16384 values due to the crc error. the console was returned for evaluation and found that it operated as expected. the root cause for the controller error was due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. the optical bench was replaced as a precaution due to multiple instances of the issue. corrections to the initial mfr report: d2 common device name updated f6/f8 should have been left blank. g1 mdr reporting contact email.
32841 | 1220648-2023-04917 | 18123091 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18123091&pc=OZD | 10/17/2023 | 08/15/2024 |  | 11/13/2023 |  |  | 10/17/2023 | 11/13/2023 |  |  | 08/31/2025 |  | 11/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  |  |  | 2024393690 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/17/2023 |  |  |  |  | The user facility reported a 69-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the medical team experienced difficulty delivering the impella device. the sheath was partially in vessel and the decision was made to remove dilator and exchange wire to perform percutaneous transluminal coronary angioplasty. the physician lost wire access and upon reinserting wire a dissection of the vessel was noted at distal tip of sheath. impella was properly inserted and angioplasty was performed. post procedure, the impella was weaned, explanted and site successfully closed with proglides x 2. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. the most likely cause of the peripheral vessel vascular damage was use related. h6 component code added. instructions for use as follows: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings and cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings and cautions: warnings section: pre-support evaluation section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-04917: in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8-should have been left blank.
32842 | 1220648-2023-04918 | 18124050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18124050&pc=OZD | 10/30/2023 | 08/20/2024 |  | 11/13/2023 |  |  |  | 11/13/2023 |  |  | 01/31/2025 |  | 11/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024316682 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/30/2023 |  |  |  |  | The user facility reported a 69-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the pump came out of position in the lv when the patient moved/kicked legs. the malpositioned pump had the pump performance level lowered and pump was explanted without any complications. no lasting harm or injury reported due to the issue.. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. the pump was not returned for investigation. pump was pulled back into the aorta due to patient movement, which was caused by kicking the lag. the cause of the positioning issue was patient movement, based on given clinical details. added- h6 impact code 4628 added- d6a/d6b f6/f8 should have been left blank in the initial report.
32843 | 1220648-2023-04919 | 18127507 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18127507&pc=OZD | 10/15/2023 | 06/19/2024 |  | 11/13/2023 |  |  | 10/15/2023 | 11/13/2023 |  |  | 07/31/2025 |  | 11/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024390892 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/15/2023 |  |  |  |  | The user facility reported a 68-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. there was bleeding from the impella access site. impella was removed and vessel angiography was performed. a new impella was inserted without any issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding - major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was determined to be use issue because the peel away sheath was not removed and left at the access site. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
32844 | 1220648-2023-04920 | 18127569 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18127569&pc=OZD | 10/17/2023 | 02/16/2024 |  | 11/13/2023 |  |  | 10/17/2023 | 11/13/2023 |  |  | 07/31/2025 |  | 11/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024390550 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Improper Flow or Infusion | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a 66-year-old male patient who had an impella cp pump implant, experienced low pump flow and bleeding from the access site (hematoma). the medical staff repositioned the pump in order to improve the pump flows. the left ventricle appeared to be empty but suction continued. ultimately, the medical staff had to explant the impella cp and replaced it with a new pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major and the low pump flow issues has been completed. the device was not returned for investigation. according to the clinical, major access site bleeding occurred after repositioning the sheath, and suction and low pump flows were recorded after prepping the patient for the icu. the most likely cause of the major access site bleeding was anticoagulation management, and the cause of the low pump flow was most likely patient condition.
32845 | 1220648-2023-04921 | 18127603 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18127603&pc=OZD | 10/17/2023 | 02/02/2024 |  | 11/13/2023 |  |  | 10/17/2023 | 11/13/2023 |  |  | 07/31/2025 |  | 11/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024372435 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported the impella 5.5 pump providing support to 39-year-old male patient, was pulled back out of the ventricle. the provider stated the device was pulled back towards the subclavian artery under echo. the impella was eventually explanted. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer as it was discarded and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ | The investigation into the positioning issues has been completed. the device was not returned for investigation. as per clinical, the impella device was malpositioned due to series of patient movement after weaning sedation. the impella device explanted with repositioning not attempted due to the concern of position and risk. the root cause of the positioning issue was use related with the impella device pulled out and repositioning not attempted due to concern of position and risk per clinical review. h.10 the following information was included on the additional manufacturer narrative previously submitted and should not have been: instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿
32846 | 1220648-2023-04906 | 18117421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18117421&pc=OZD | 10/30/2023 | 09/20/2024 |  | 11/12/2023 |  |  |  | 11/12/2023 |  |  | 08/31/2025 |  | 11/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024386029 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/30/2023 |  |  |  |  | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 74-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the 5.5 was placed and team discussed appropriate placement for the 5.5 to avoid any perforation,. left ventricle perforation was then noted and quickly repaired using pledget sutures x2 and a pericardial patch. no blood products needed during repair and vitals stable. the perforation was thought to be caused by a surgeon lifting the heart. the patient expired in the icu due to the complications from the bypass. the death was not related to the device. |  |  |  |  | The investigation into the reported ventricle perforation has been completed since the original report was filed. based on the clinical account, the root cause of the ventricle perforation is a use issue. it is noted that the surgeon lifted the heart to perform a vsd closure, which caused the perforation. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.
32847 | 1220648-2023-04882 | 18110011 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18110011&pc=OZD | 10/11/2023 | 01/11/2024 |  | 11/10/2023 |  |  | 10/11/2023 | 11/10/2023 |  |  | 04/30/2025 |  | 11/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024350512 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/11/2023 |  |  |  |  | The user facility reported a 78-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support for high risk percutaneous coronary intervention. while on impella support, it was noted that the patient developed ischemia in their lower limb. the pump was explanted as a result of the condition. | [5 empty] | The investigation into the limb ischemia ¿ reversible has been completed. the product was not returned. as per the clinical information provided, the root cause of the limb ischemia was determined to be patient condition due to peripheral arterial disease in the iliacs. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32848 | 1220648-2023-04883 | 18110061 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18110061&pc=OZD | 10/11/2023 | 03/25/2024 |  | 11/10/2023 |  |  | 10/11/2023 | 11/10/2023 |  |  | 08/31/2025 |  | 11/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024385074 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Material Separation | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/11/2023 |  |  |  |  | The user facility reported a 65-year-old male patient noted for st elevated myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. the sterile sleeve was reported to have been compromised during a repositioning event. the sleeve was stretched out along the catheter and it ripped away from the proximal hub (blue hub - lock mechanism). the exposed catheter was secured after it was cleansed with antiseptic solution and support with the implanted pump continued. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the product damage (sterile sleeve damage) and the positioning issues has been completed since the initial report was submitted. the product was not returned for investigation. the cause of the product damage (sterile sleeve damage) issue and the positioning issues was use related, based on given clinical details.
32849 | 1220648-2023-04884 | 18110773 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18110773&pc=OZD | 09/28/2023 | 03/26/2024 |  | 11/10/2023 |  |  | 10/12/2023 | 11/10/2023 |  |  | 10/31/2024 |  | 11/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024286214 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Hispanic | White | Malfunction | Yes | Yes | Pumping Stopped | [6 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/12/2023 |  |  |  |  | The user facility reported a 38-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the impella cp had position alarms. the team additionally noted air in purge that was troubleshot by de-air process. later, (15minutes) there was a non-resolving pump stop. the pump stopped and the team performed cpr for 20 minutes and the patient expired. the user facility had no backup impella cp pumps and no backup automated impella controllers. | [5 empty] | The investigation into the pump stop and the positioning issues has been completed since the initial report was submitted. the product was not returned for investigation. the cause of the pump stop and the positioning issues was not determined since product was not returned and limited clinical information was provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
32850 | 1220648-2023-04885 | 18110858 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18110858&pc=OZD | 10/12/2023 | 11/10/2023 |  | 11/10/2023 |  |  | 10/12/2023 | 11/10/2023 |  |  |  |  | 11/10/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1576159 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that their automated impella controller (aic) had a broken knob. a new aic loaner was shipped to the user site for use as needed. there was no harm to the patient as a result of this incident. | [5 empty] | The investigation into the reported knob issues issue has been completed. data log analysis log recorded the pump in the last usage of this console at the customer site. after the pump was removed, the console received repeated shutdown requests from the user, which is most likely indicating that the user was unable to confirm/select the shutdown request due to a broken knob. due to this issue, the console has been powered off by holding down the power switch for longer than 30 seconds. the console was tested, and upon trying to navigate through the service menu, observed the inconsistent rotatory knob selection. upon checking the 5v output across the test points tp100.0 and tp101.0, no 5v output was received every time upon pressing the rotatory knob. also didn¿t hear the audio from aic that is supposed to be heard every time upon pressing the rotatory knob. the root cause for the reported event was determined to be due to a defective component.
32851 | 1220648-2023-04886 | 18111245 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18111245&pc=OZD | 10/12/2023 | 03/07/2024 |  | 11/10/2023 |  |  | 10/12/2023 | 11/10/2023 |  |  | 05/31/2025 |  | 11/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024368913 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 67-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. patient had saturated pump flows and high motor current. the impella cp was at end of life and patient needed escalation for full support. the impella cp was explanted and an impella 5.5 was inserted via axillary access for ongoing support. | [5 empty] | The investigation into the saturated flow issue has been completed since the original report was filed. the device was not returned for investigation. data logs show high motor current with motor current spike and suction alarm was reported during time of saturated flow event. the cause of the saturated pump flow issue was most likely biomaterial based on data log review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32852 | 1220648-2023-04887 | 18111411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18111411&pc=OZD | 10/12/2023 | 01/18/2024 |  | 11/10/2023 |  |  | 10/12/2023 | 11/10/2023 |  |  | 04/30/2025 |  | 11/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024352924 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/12/2023 |  |  |  |  | The user facility reported a 57-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was found to have a thrombus in the left ventricle and the decision was made to remove impella. patient is stable post explant. | [5 empty] | The investigation for the thrombus has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the thrombus issue was most likely patient condition related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32853 | 1220648-2023-04888 | 18111546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18111546&pc=OZD | 10/13/2023 | 08/20/2024 |  | 11/10/2023 |  |  | 10/13/2023 | 11/10/2023 |  |  | 04/30/2025 |  | 11/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024332749 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 62-year-old male patient was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported the patient had an infection from their axillary graft. patient was given antibiotics and was scheduled for device explant. following removal, the decision was made to allow the infection to clear before further action could be taken for possible left ventricular assist device (lvad) consideration. | [5 empty] | The investigation for the infection has been completed. the impella device was not returned for evaluation. therefore, without sufficient clinical details, the root cause of the reported infection could not be determined. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock & cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ b.7 revised as information was inadvertently omitted from the initial manufacturer device report number 1220648-2023-04888. d.6a and d.6b revised from the initial submission of manufacturer device report number 1220648-2023-04888 in accordance with updated procedures. h.6 added code 925 as it was inadvertently left off the previously submitted manufacture device report number 1220648-2023-04888. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2023-04888 in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32854 | 1220648-2023-04890 | 18114478 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18114478&pc=OZD | 07/22/2021 | 01/22/2024 |  | 11/10/2023 |  |  | 07/22/2021 | 11/10/2023 |  |  | 04/30/2023 |  | 11/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022076862 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/12/2023 |  |  |  |  | The user facility reported a 83-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that a patient on impella support experienced some placement signaling alarms. the md discussed repositioning the pump to resolve the alarms. there was no reported harm to the patient as a result of this incident. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. low pump flow - the cause of the low pump flow cannot be determined since the product was not returned and enough clinical information is not available. this report is being filed as part of a retrospective review of historical records. | The investigation into the optical signal issue and the low pump flow issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. per the clinical information provided, the root cause of the optical signal issue and the low pump flow issue could not be determined since the product was not returned and enough clinical information is not available.
32855 | 1220648-2023-04894 | 18115082 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18115082&pc=OZD | 10/12/2023 | 10/17/2025 |  | 11/10/2023 |  |  | 10/12/2023 | 11/10/2023 |  |  | 07/31/2025 |  | 11/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024374304 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 10/12/2023 |  |  |  |  | The user facility reported a 52-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support as a bridge to transplant. during support, there were some suction events on the impella that were related to low central venous pressure after diuresis. there was no patient harm. the patient remained on support for five additional days and was successfully weaned. | Us complainant reported there were some suction events on the impella that were related to low cvp after diuresis. additionally, there were placement signal not reliable alarms. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow has been completed. there was no product returned. the data logs were reviewed. throughout the whole case there were three instances of suction conditions. low flow was seen on 10/2014 in the last image of the logs but there were no spikes in motor current. the root cause of the low pump flow could not be determined due to no product returned and insignificant clinical information. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. there was no product returned. the data logs were reviewed and the optical signal issue was found to be a problem with the automated impella controller (aic). the investigation of the placement signal not reliable (psnr) alarms is captured in 23-19146-2. throughout the whole case there were three instances of suction conditions. low flow was seen on 10/14 in the last image of the logs but there were no spikes in motor current. the root cause of the low pump flow cannot be determined due to no product returned and insignificant clinical information. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
32856 | 1220648-2023-04889 | 18115142 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18115142&pc=OZD | 10/11/2023 | 06/19/2024 |  | 11/10/2023 |  |  | 10/11/2023 | 11/10/2023 |  |  | 05/31/2025 |  | 11/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024352347 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information provided states that the patient had decompensated with increased vt overnight despite increase to p9 requiring levo and vaso so protek duo support added. | The user facility reported a 61-year-old male was implanted with an impella 5.5 device for mechanical circulatory support. during support, it was noted that the patient had a hematoma at the access site. the doctor evacuated the hematoma and the patient's discomfort improved. anticoagulants were on hold for 48 hours and support with the pump was maintained. |  |  |  |  | The investigation for the access site bleed and ventricular tachycardia (vt) has been completed. tthe device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed and vt could not be determined. the instructions for use referenced in the initial report is for the impella 5.5 with smart assist for use during cardiogenic shock in section: potential adverse events (united states), which now includes "cardiac or vascular injury (including ventricular perforation)¿ added- b5-additional event information, h6 clinical code 2095 in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32857 | 1220648-2023-04895 | 18115151 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18115151&pc=OZD | 10/12/2023 | 08/14/2024 |  | 11/10/2023 |  |  | 10/12/2023 | 11/10/2023 |  |  | 07/31/2025 |  | 11/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024389485 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 73-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported poor distal leg perfusion during use of the impella device. the pump was removed in response to this event. | [5 empty] | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6a and d6b added the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-04895: g1 reporting contact fax number missing. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32858 | 1220648-2023-04896 | 18115183 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18115183&pc=OZD | 10/12/2023 | 08/20/2025 |  | 11/10/2023 |  |  | 10/12/2023 | 11/10/2023 |  |  |  |  | 11/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1133152 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 10/12/2023 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) displayed a controller error. there was no patient involvement during this event. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Corrections to the initial mfr report: g1 mdr reporting contact email. d2 device name has been updated. f6 and f8 should have been left blank.
32859 | 1220648-2023-04897 | 18115266 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18115266&pc=OZD | 10/12/2023 | 08/05/2024 |  | 11/10/2023 |  |  | 10/12/2023 | 11/10/2023 |  |  | 05/31/2025 |  | 11/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024369015 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/12/2023 |  |  |  |  | The user facility reported a 44-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient woke up on the unit and pulled the sheath out of the right groin and impella out of the ventricle. rn decreased the impella to p2 and the patient was stable. the patient was taken to the operation room for right femoral artery repair. patient tolerated the procedure well. | [5 empty] | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. the root cause of the vascular damage issue was use related to excessive force applied by the patient. instructions for use as follows: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings and cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings and cautions: warnings. section: pre-support evaluation. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ d6a and d6b added. h6 component code added the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-04897: f6/f8- should have been left blank. g1 reporting contact fax number missing. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
32860 | 1220648-2023-04898 | 18115306 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18115306&pc=OZD | 10/13/2023 | 05/13/2024 |  | 11/10/2023 |  |  | 10/13/2023 | 11/10/2023 |  |  | 06/30/2025 |  | 11/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024370725 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Detachment of Device or Device Component | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/13/2023 |  |  |  |  | The user facility reported a 61-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. once the patient was weaned from the impella cp, the patient was brought to the cath lab for impella removal. during removal, the inflow cage became detached and remained in the vessel resulting in rapid blood loss. this was caused by the repositioning unit being hung on the cath lab drape. this was quickly rectified by the fellow and the impella device was pulled back and hemostasis quickly achieved with manual pressure. 14 french manta closure was advanced and while trying to secure the foot plate, the entire device came out of the artery. the patient was transfused with two units of packed red blood cells. | [5 empty] | The investigation into the access site bleeding ¿ major and the product damage (detached inflow/outflow - peripheral vessel) issues has been completed since the original report was submitted. the device was not returned for investigation. according to the clinical, while removing the impella device the inflow cage became detached and remained in the artery which resulted in rapid blood loss. the inflow cage was removed, manual pressure was applied, and two units of prbcs were transfused, and hemostasis achieved by manta closure. the cause of the detached inflow cage was not determined because no product was returned, and limited clinical information was given. potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32861 | 1220648-2023-04899 | 18115319 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18115319&pc=OZD | 10/13/2023 | 02/21/2024 |  | 11/10/2023 |  |  | 10/13/2023 | 11/10/2023 |  |  | 06/30/2025 |  | 11/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024390196 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 94 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | Material Deformation | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/13/2023 |  |  |  |  | The user facility reported a 94-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. after the impella implant the doctor could visualize a kink in the cannula. flows greater than 1.8 could not be achieved. physician decided to replace pump. when pump was removed, no evidence of defect was found. the replacement pump also appeared under fluoro in a similar manner and physician concluded it was patient anatomy. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported low pump flow and product damage has been completed since the original report was filed. no product was returned for investigation. the cannula kink was visualized under fluoro. data logs confirmed the low pump flow. the root cause of the low pump flow issue was most likely a kinked cannula, and the cannula kink most likely resulted from difficult patient anatomy.
32862 | 1220648-2023-04900 | 18115325 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18115325&pc=OZD | 10/13/2023 | 08/20/2024 |  | 11/10/2023 |  |  | 10/13/2023 | 11/10/2023 |  |  | 08/18/2025 |  | 11/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - EU |  |  | ABIOMED, INC. |  |  | US |  | 2024377188 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/13/2023 |  |  |  |  | Additional information received, that the patient expired while on support. per medical safety officer review use of the device cannot conclusively be associated with the outcome. | The user facility reported a patient of unknown age, race and gender was implanted with an impella cp for mechanical circulatory support. during support, a thrombus was found in the patient's left ventricle. no additional information is available at the time of this report. |  |  |  |  | The investigation of thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b5 updated with additional information: d6a and d6b added. instructions for use as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings and cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ f6/f8-should have been left blank on manufacturer device report 1220648-2023-04900. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32863 | 1220648-2023-04901 | 18115335 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18115335&pc=OZD | 10/13/2023 | 02/21/2024 |  | 11/10/2023 |  |  | 10/13/2023 | 11/10/2023 |  |  | 08/31/2025 |  | 11/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2024393177 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/13/2023 |  |  |  |  | The user facility reported a 74-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support via right femoral artery. a large hematoma had developed, during transport, over the impella sheath site. once the patient arrived to the critical care unit, the hematoma was worse and there was concern for distal limb ischemia. the decision was to exchange the impella device and place it in the left femoral artery. reportable due to patient's hematoma and limb ischemia and the removal of the device. | [5 empty] | The investigation into the reported access site bleeding ¿ major and limb ischemia - reversible has been completed since the original report was filed. product was not returned for review. based on clinical description, the root cause of access site bleeding was most likely due to patient movement. the root cause of limb ischemia was most likely due to patient condition related. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32864 | 1220648-2023-04902 | 18115339 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18115339&pc=OZD | 10/14/2023 | 07/10/2024 |  | 11/10/2023 |  |  | 10/14/2023 | 11/10/2023 |  |  | 08/31/2025 |  | 11/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2024387798 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/14/2023 |  |  |  |  | The user facility reported a 54-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. md calls from cathlab, needs assistance with implant, together successfully implanted pump, patient keeps getting ventricular fibrillation, has already received a total of 750 mg amiodarone, pump pulled back a little, pigtail lies free in the ventricle, pump parameters are good. | [5 empty] | The investigation for the reported ventricular fibrillation (vf) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf could not be determined. the instructions for use for the impella cp with smart assist in section: potential adverse events (united states). states ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32865 | 1220648-2023-04893 | 18115994 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18115994&pc=OZD | 10/12/2023 | 10/17/2025 |  | 11/10/2023 |  |  | 10/12/2023 | 11/10/2023 |  |  | 06/30/2025 |  | 11/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024373961 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Device Sensing Problem | Infusion or Flow Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 83-year-old female noted to be as high risk for coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported there were suction alarms on that resolved when volume was given. additionally, there was an impella position in aorta alarm that occurred while weaning the pump for explant. the pump was explanted. | [5 empty] | The investigation into the low pump flow and the positioning issues has been completed. the device was not returned for investigation. data logs confirmed the failure mode. logs showed low flow occurred when pump started that triggered suction response & suction alarms. based on clinical description and data review, the root cause of low flow was most likely due to patient condition as the low flow was resolved with volume given. the cause of the false position in aorta was patient condition related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
32866 | 1220648-2023-04892 | 18116030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18116030&pc=OZD | 10/12/2023 | 08/15/2025 |  | 11/10/2023 |  |  | 10/12/2023 | 11/10/2023 |  |  | 05/31/2025 |  | 11/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024356236 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/12/2023 |  |  |  |  | The user facility reported a 64-year-old female with non-ischemic cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. immediately upon initiation of support low purge pressure alarm with noted on the automated impella controller (aic) a alarm stating "low purge pressure" occurred. all connections were checked and were secure. purge solution dripped out of the purge disc areas. the decision was made to replace the impella 5.5 with another. there was no patient harm related to this event. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | D.6a and d.6b added since the initial submission of manufacturer device report number 1220648-2023-04892 in accordance with updated procedures. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2023-04892 and should not have been. g.1 revised mdr reporting contact email since the initial submission of manufacturer device report number 1220648-2023-04892 in accordance with updated procedures. h.3 revised as device evaluation is anticipated and was not selected on the initial manufacturer device report number 1220648-2023-04892. h.6 codes 2199 and 3221 were reported incorrectly on the initial manufacturer device report number 1220648-2023-04892 in accordance with updated procedures.
32867 | 1220648-2023-04891 | 18116173 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18116173&pc=OZD | 10/12/2023 | 01/19/2024 |  | 11/10/2023 |  |  | 10/12/2023 | 11/10/2023 |  |  | 07/31/2025 |  | 11/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024385955 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/12/2023 |  |  |  |  | The user facility reported an 83-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support post stent placement in the circumflex artery. the impella placement was verified under fluro per cath lab staff. the patient coded and, after achieving return of spontaneous circulation, and the plan was to go to or for coronary artery bypass grafting after the circumflex was perforated post stent placement. however, patient could not be stabilized and ct surgeons decided to open chest in cath lab. after opening chest, surgeons saw that the impella cp had perforated the left ventricle. ultimately under imaging, it was seen that there was no activity with the heart. staff all agreed there was nothing else that could be done. the patient deceased in cath lab and target organ damage was pronounced. | [5 empty] | The investigation for the ventricle perforation has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. tthe data logs confirmed that there were several "impella position in ventricle & unknown" alarms towards the end of support. additionally, motor current spike and motor speed dropped during these alarms indicating a potential interaction. the pump likely came out of position during cpr. the root cause of the ventricle perforation is most likely due to improper positioning of the device. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿
32868 | 1220648-2023-04903 | 18116596 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18116596&pc=OZD | 10/16/2023 | 02/21/2024 |  | 11/10/2023 |  |  | 10/16/2023 | 11/10/2023 |  |  | 04/30/2025 |  | 11/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024351129 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/16/2023 |  |  |  |  | The user facility reported a 77-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. a red alarm of ¿impella in ventricle¿ occurred during removal of peel away and immediately resolved with no issues or interruptions of patient care. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported positioning issue has been completed since the original report was filed. complaint device not returned for investigation. lt log reviewed and showed the impella position in ventricle alarm occurred at the beginning of the case. the cause of the positioning issue alarm most likely was use related due to improper technique as immediately resolved with no issues or interruptions of patient care.
32869 | 1220648-2023-04904 | 18116631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18116631&pc=OZD | 10/16/2023 | 10/24/2025 |  | 11/10/2023 |  |  | 10/16/2023 | 11/10/2023 |  |  |  |  | 11/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1399951 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 10/16/2023 |  |  |  |  | The user facility reported a 71-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the entire case performed without incident. patient moved to icu bed without issues. as rolling out of the cardiac cath. lab, console displayed 2 error messages. frist, yellow "controller error" and second, "controller failure" red alarm. messages told the team to switch to backup console and that purge system has stopped and to switch to backup controller. the team switched to backup automated impella controller (aic) without further issues. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Corrections to the initial mfr report: g1 mdr reporting contact email, d2 device name has been updated, f6 and f8 should have been left blank, updated h8 to reuse.
32870 | 1220648-2023-04905 | 18116693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18116693&pc=OZD | 10/19/2023 | 08/02/2024 |  | 11/10/2023 |  |  | 10/20/2023 | 11/10/2023 |  |  | 08/31/2025 |  | 11/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024392547 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 27 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/20/2023 |  |  |  |  | The user facility reported a 27-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. a hematoma was developed at the right axillary incision/pocket site post the impella 5.5 insertion. the device was tunneled, and the patient had excess tissue in that area. systemic anticoagulants were stopped. hemoglobin remained stable, not requiring transfusion and hematoma resolved. | [5 empty] | The investigation for access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was determined to be patient condition because of the excess tissue contributing for the bleed. b.2 revised as selection was missing from the initial submission of manufacturer device report number 1220648-2023-04905. d.6a and d.6b revised from the initial submission of initial manufacturer device report number 1220648-2023-04905 in accordance with updated procedures. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2023-04905 and should not have been. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report number 1220648-2023-04905 was submitted. h.6 code 925 was added since the initial submission of manufacturer device report number 1220648-2023-04905 in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report number 1220648-2023-04905 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32871 | 1220648-2023-04859 | 18104077 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18104077&pc=OZD | 07/13/2023 | 11/09/2023 |  | 11/09/2023 |  |  | 07/13/2023 | 11/09/2023 |  |  | 10/31/2024 |  | 11/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023283770 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | Unintended Movement | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/13/2023 |  |  |  |  | The user facility reported a 64-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the patient began moving violently causing pump to pull out of ventricle. staff emergently intubated patient and administered sedatives. the physician was unable to re-advance original impella and made the decision to place second impella cp. | [5 empty] | The investigation into the positioning issues has been completed. product was not returned for review. the root cause of the positioning issue was determined to be patient movement as the patient began moving violently forcing the pump out of the ventricle.
32872 | 1220648-2023-04860 | 18104441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18104441&pc=OZD | 11/08/2023 | 10/03/2024 |  | 11/09/2023 |  |  | 11/08/2023 | 11/09/2023 |  |  | 08/31/2025 |  | 11/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED INC. |  |  | US |  | 2024393137 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Activation, Positioning or Separation Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/08/2023 |  |  |  |  | The user facility reported a 38-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient had bleeding/oozing without hematoma out of the impella access site. the nurse redressed the site and performed angle matching and the bleeding resolved. additionally, the patient had red urine. echo confirmed the pigtail appeared to be in papillary and inlet measuring 2.1 from annulus. the impella was repositioned, and the hemolysis started to clear up. | [5 empty] | The investigation for the hemolysis and the positioning issue has been completed. pump was not returned for evaluation. data logs were investigated and "impella in aorta" alarms were observed. no suction or motor current spikes were observed. as per the provided clinical information, the pump was found to seated in the papillary muscle. due to the improper positioning of the pump, hemolysis was observed. after repositioning the pump, hemolysis resolved. therefore, the root cause of the hemolysis issue was improper positioning due to improper technique. data logs showed "impella in aorta" alarms. echo confirmed pump out of position as per clinical. therefore, the root cause of the positioning issue was most likely use related to improper technique. corrections to the initial mdr mfr report. d4: corrected model number. added: d6a/d6b. f6/f8 should have been left blank. g1: corrected mfr site name and address. h6 med dev prob code 2993 listed in error. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
32873 | 1220648-2023-04861 | 18106011 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18106011&pc=OZD | 10/11/2023 | 02/21/2024 |  | 11/09/2023 |  |  | 10/11/2023 | 11/09/2023 |  |  | 03/31/2025 |  | 11/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024345300 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/11/2023 |  |  |  |  | The user facility reported an 87 -year-old male noted for high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the impella was unable to pass across the aortic valve into optimal position. the impella was removed, and placement was tried a second time. again, the inflow of the device would cross the aortic valve but not the cannula. during second attempt, wire and impella kinked in aorta. the impella was removed and a kink on the catheter above the motor housing was observed. so a decision was made to place a new impella cp. no patient harm was reported due to the issue. | [5 empty] | The investigation into the reported product damage (cannula kink) and delivery issue has been completed since the original report was filed. product was returned for investigation. visual inspection found a kink on catheter below the end cap at mark 16. no other issues were found on the rest of the pump. the root cause of delivery issue was most likely due to stenotic valve patient condition. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32874 | 1220648-2023-04862 | 18106041 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18106041&pc=OZD | 11/01/2023 | 03/18/2024 |  | 11/09/2023 |  |  | 11/01/2023 | 11/09/2023 |  |  |  |  | 11/09/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, JP |  |  | ABIOMED, INC. |  |  | US |  | 1306449 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/01/2023 |  |  |  |  | The user facility in japan reported a 75-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. during support, there were repeated purge leaks and purge cassette replacement. the automated impella controller (aic) was replaced due to an alarm stating "purge system not detected". the purge system alarm prompted the team to follow the information for use and put a backup console into place. there was no patient harm related to this event. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the console was returned and tested after arriving in fs. the purge flag was confirmed to be broken. the cause of the issue was a defective component. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
32875 | 1220648-2023-04863 | 18106307 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18106307&pc=OZD | 02/21/2023 | 11/09/2023 |  | 11/09/2023 |  |  | 02/21/2023 | 11/09/2023 |  |  | 09/30/2024 |  | 11/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023275742 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/21/2023 |  |  |  |  | The user facility reported a 65-year-old male in anterior st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient was having intermittent ventricular tachycardia/ventricular fibrillation during implant of impella. all arrhythmias subsided once impella was in the left ventricle and turned on, during 30 min unloading time, and throughout intervention. | [5 empty] | The investigation into a patient ventricular tachycardia/ventricular fibrillation has been completed. no product was returned. data logs were reviewed, and suction alarms were found during the case. the cause of the ventricular tachycardia issue was not determined and unknown relation to impella based on clinical information and data log review.
32876 | 1220648-2023-04864 | 18106399 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18106399&pc=OZD | 11/07/2023 | 01/12/2024 |  | 11/09/2023 |  |  | 11/07/2023 | 11/09/2023 |  |  | 09/30/2025 |  | 11/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED INC. |  |  | US |  | 2024399853 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/07/2023 |  |  |  |  | The user facility reported a 66-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the physician attempted to place an impella cp, but had trouble advancing in the left ventricle due to interaction of the previously placed ebu guide catheter. physician removed the impella from body through the sheath and wanted to re-cross the aortic valve. after re-crossing, multiple attempts were made to advance impella cp back through the impella sheath over the 0.018 wire. the impella cp would not advance through the sheath. a new impella cp was opened and primed. the second impella cp was able to advance through the sheath over the 0.018 wire with no issues. | [5 empty] | The investigation into the reported difficulty inserting/removing pump through introducer has been completed since the original report was filed. the clinical stated the pump was able to advance in the lv due to the ebu catheter in place. the root cause of the positioning issue is most likely due to the interaction with another device. the pump and the introducer were returned for inspection and there were no deformities or damages found. the pump was able to be inserted into the introducer with no issues. the root cause of the difficulty inserting/removing the pump through the introducer is most likely due to use as it was able to pass through the introducer the first time and able to in the lab. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32877 | 1220648-2023-04865 | 18106415 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18106415&pc=OZD | 10/11/2023 | 02/21/2024 |  | 11/09/2023 |  |  | 11/09/2023 | 11/09/2023 |  |  | 05/31/2025 |  | 11/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024359771 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 61-year-old male noted for st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. during support, oozing was noted from the access site that prompted intervention. the team had explanted the pump and done contralateral ballooning, infused blood products, used topical tranexamic acid (txa), and closed the site with perclose and angioseal followed by manual hold. the perclose had failed. the access site was stable following the intervention. | [5 empty] | The investigation into the reported access site bleeding ¿ major has been completed since the original report was filed. product was not returned for review. based on clinical description, the root cause of access site bleeding was most likely due to patient movement. | Correction provided for h6 health effect - clinical code. the original report was coded for hemolysis inadvertently. the updated h6 code should reflect the reported access site bleed. upon investigation closure, a supplemental report will be submitted. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32878 | 1220648-2023-04866 | 18106589 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18106589&pc=OZD | 10/11/2023 | 12/04/2023 |  | 11/09/2023 |  |  | 10/11/2023 | 11/09/2023 |  |  | 04/30/2025 |  | 11/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024352885 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a 74-year-old male patient on impella cp support experienced bleeding from the access site and as a result, the medical team was required to replace the pump. the bleed was observed in the transfer of the patient to a tertiary center for continued monitoring. | [5 empty] | The investigation into the patient's access site bleeding has been completed. no product was returned for evaluation. clinical review determined the root cause of the access site bleeding-major is most likely due to a use issue as the plane seatbelt was suspected to have been placed over the access site. there was no bleeding at the time of implant. b1 adverse event was inadvertently checked with no. checked yes under b1 adverse event. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32879 | 1220648-2023-04867 | 18106620 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18106620&pc=OZD | 10/11/2023 | 01/12/2024 |  | 11/09/2023 |  |  | 10/11/2023 | 11/09/2023 |  |  | 05/31/2025 |  | 11/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024340814 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 58-year-old male noted for high risk percutaneous cardiac insertion was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the impella 5.5 was started and immediately the pump stopped with two different alarms occurring simultaneously. one alarm stated pump stopped due to controller and the other was an increased motor current pump stop alarm. the pump was attempted to be restarted three times with no success. the automated impella controller (aic) was switched out which did not resolve the pump stop. therefore, the pump was explanted. | [5 empty] | The investigation into the pump stop issue has been completed since the original report was filed. the impella pump was returned for investigation. visual inspection of the pump revealed that catheter was kinked close to red impella plug kink protector, but no biomaterial or blood ingress was observed. the pump stop issue did not reproduce. the root cause of the pump stop issue was not determined since the failure did not reproduce per field investigation. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
32880 | 1220648-2023-04868 | 18106750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18106750&pc=OZD | 04/14/2023 | 09/26/2024 |  | 11/09/2023 |  |  | 04/14/2023 | 11/09/2023 |  |  |  |  | 11/09/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1410103 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Connection Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The us complainant had a team attempting a high risk percutaneous coronary intervention (hrpci) with impella cp support for 75-year-old female patient. the team had done orbital atherectomy when the patient condition deteriorated and the cp was placed to the left ventricle while attempting to achieve return of spontaneous circulation (rosc). the automated impella controller (aic) however malfunctioned and during this malfunction the time of death (tod) was called. the death was not attributed to the attempted use of the cp. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-04868 was provided in sections b1, b2, b5, d6a, d6b, h1, and h6. | The investigation into the reported automated impella controller (aic) issues has been completed. logs were reviewed and analyzed with special attention to reported date of complaint. there was no evidence of any error or alarm other than ¿algs suspended opm signal invalid¿, which is expected behavior when a pump is disconnected. the logs further show that the subsequent pumps were run against the console without any non-conformance or issue. the console was visually inspected on the lab bench and there were no anomalies found. the aic operated within specification when a pump was then connected and run for more thirty minutes. the root cause of the reported issue was unable to be determined as there were no issues found during testing at the bench level. the device functioned/operated as per specification. this report is being filed as part of a retrospective review of historical records.
32881 | 1220648-2023-04869 | 18106941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18106941&pc=OZD | 10/11/2023 | 02/12/2024 |  | 11/09/2023 |  |  | 10/11/2023 | 11/09/2023 |  |  | 05/31/2025 |  | 11/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024350567 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/11/2023 |  |  |  |  | The user facility reported a 82-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was noted that the anchoring ring could not be secured. the doctor tried to secure the anchoring ring again after adjusting the position, but it could not be secured. various attempts were made such as pushing it while attempting to secure it, but the issue persisted. the patient was on impella support for 28.75 hours before the patient expired. the cause of death was determined to be multiple organ failure. no allegations were made towards the impella for the cause of death. | [5 empty] | The investigation into the positioning issues has been completed. the device was not returned for investigation. based on the clinical information provided, a pump position adjustment was made, and the anchor ring could not be secured afterwards. log analysis was not necessary for this complaint. the most likely cause of the positioning issue was touhy-borst related. the cause of the touhy-borst issue was unable to be determined because no product was returned. it was reported that the patient died from multiple organ failure, and there was no relationship between impella cp support and the cause of death. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32882 | 1220648-2023-04870 | 18107067 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18107067&pc=OZD | 10/11/2023 | 02/28/2024 |  | 11/09/2023 |  |  | 10/11/2023 | 11/09/2023 |  |  | 07/31/2025 |  | 11/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024387201 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/11/2023 |  |  |  |  | The user facility reported a 68-year-old female noted for st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. poor distal perfusion noted to right lower extremity (rle)(impella groin). patient with severe vascular disease at baseline and on high-dose vasopressors over last 24 hours per team. both extremities equally cold. no pedal pulses noted in rle with mottling to right thigh noted. the physician stated that they would monitor the patient closely and trend lactate. the nurse subsequently weaned vasopressors and support continued. | [5 empty] | The investigation into the reported limb ischemia -reversible has been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32883 | 1220648-2023-04871 | 18107138 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18107138&pc=OZD | 10/11/2023 | 06/06/2024 |  | 11/09/2023 |  |  | 10/11/2023 | 11/09/2023 |  |  | 07/31/2025 |  | 11/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024387161 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/11/2023 |  |  |  |  | The user facility reported a 63-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced impella access site bleeding. manual pressure held, fem-stop, and two units of packed red blood cells were given to treat the condition. support continued with the implanted pump. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely patient movement, the patient was restless & bleed as per the clinical description provided. the instructions for use referenced in the initial report is for the impella cp with smart assist system in the following sections general patient care considerations, entering cardiac output, and potential adverse events (united states), which now includes the statement "cardiac or vascular injury (including ventricular perforation).¿ | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32884 | 1220648-2023-04872 | 18107421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18107421&pc=OZD | 10/15/2023 | 05/23/2024 |  | 11/09/2023 |  |  | 10/15/2023 | 11/09/2023 |  |  | 07/31/2025 |  | 11/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024391364 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/15/2023 |  |  |  |  | The user facility reported a 71-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the team suspected the patient had hemolysis. the pump was repositioned and the issue was resolved. the patient remained on cp support for ten days and was escalated to an impella 5.5. | [5 empty] | The investigation into the hemolysis issue has been completed since the original report was submitted. according to the clinical information, on the first day of impella cp support, the patient's urine was pink in color. the following day the patient's urine turned red. the pump position was confirmed and the patient was placed on dialysis. no product was returned for a visual inspection. the data log analysis confirmed suction alarms after hemolysis occurred. the cause of the hemolysis was not determined because no product was returned and not enough clinical information was given. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 code 4641 was reported incorrectly on the previously submitted report. new code has been added to health effect - impact code. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
32885 | 1220648-2023-04873 | 18107821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18107821&pc=OZD | 11/06/2023 | 11/17/2025 |  | 11/09/2023 |  |  | 11/06/2023 | 11/09/2023 |  |  |  |  | 11/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1607645 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Run on Battery | Battery Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/06/2023 |  |  |  |  | The complainant reported that an impella controller (aic) being utilized for patient transport training was powered on, and a yellow alarm occurred stating ¿switch to back up controller¿. turned off and back on again and alarm had resolved. the aic was able to be used for the training but will be removed for analysis. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32886 | 1220648-2023-04874 | 18107835 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18107835&pc=OZD | 06/23/2023 | 11/10/2023 |  | 11/09/2023 |  |  |  | 11/09/2023 |  |  | 01/31/2025 |  | 11/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024313585 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Material Deformation | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/23/2023 |  |  |  |  | The user facility reported a 75-year-old male noted as high risk percutaneous coronary intervention was to be implanted with an impella cp device for mechanical circulatory support. the placement was aborted as the 14fr sheath kinked in the native anatomy and then the cp failed to deliver through the peripheral artery. there was no patient harm or injury from the attempt. | [5 empty] | H6 investigation conclusions has been corrected. | No product was returned for analysis. as per clinical, impella access gained through left femoral artery and patient had tortuous, calcified vessels. kinks were found at the distal end of introducer sheath during delivery and was aborted. no other damages were found relative to the pump. the cause of the delivery issue was patient condition related due to left iliac tortuosity complicating insertion and not advancing past the sheath and into left ventricle. the cause of the introducer damage issue was an introducer sheath kink due to diseased vessel patient condition.
32887 | 1220648-2023-04875 | 18107847 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18107847&pc=OZD | 11/03/2023 | 04/23/2024 |  | 11/09/2023 |  |  | 11/03/2023 | 11/09/2023 |  |  | 07/31/2025 |  | 11/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMART ASSIST SET, EU |  |  | ABIOMED, INC. |  |  | US |  | 2024378135 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/03/2023 |  |  |  |  | The user facility the eu reported a patient (gender, age unknown) was undergoing cardiac surgery, and was implanted with an impella 5.5 device for mechanical circulatory support. low pump flows occurred and team believe there is thrombus in cannula of the pump. the pump was explanted and replaced. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow and the thrombus issues has been completed since the original report was submitted. the device was not returned for investigation. according to clinical findings, the doctor discovered biomaterial in the cannula during the saturated pump flow issue. the cause of the saturated pump flow issue was most likely biomaterial found inside cannula.
32888 | 1220648-2023-04876 | 18107849 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18107849&pc=OZD | 10/14/2023 | 01/25/2024 |  | 11/09/2023 |  |  | 10/14/2023 | 11/09/2023 |  |  | 06/30/2025 |  | 11/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024377386 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/14/2023 |  |  |  |  | The user facility reported an 83-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. during support, the patient's hemoglobin dropped to 8.6 from 14, possibly from bleeding around the site. two units of blood products were given to the patient. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed. the device was not returned for benchtop investigation. according to the clinical information provided, patient was experiencing bleeding around impella sheath after insertion of guidewire repositioning unit (gwru). impella angle matched, and bleeding issue resolved. the cause of the access site bleeding issue was use related with patient bleeding from access site after gwru insertion and bleeding resolved after impella angle matched. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32889 | 1220648-2023-04877 | 18107885 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18107885&pc=OZD | 10/13/2023 | 02/26/2024 |  | 11/09/2023 |  |  | 10/13/2023 | 11/09/2023 |  |  | 02/28/2025 |  | 11/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024338545 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/13/2023 |  |  |  |  | The user facility reported an 81-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support after cardiac surgery. during support, the patient experienced access site bleeding. the bleeding was managed by mattress suture and fem-stop. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed. the device was not returned for investigation. the root cause of the access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32890 | 1220648-2023-04878 | 18108005 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18108005&pc=OZD | 10/11/2023 | 01/18/2024 |  | 11/09/2023 |  |  | 10/11/2023 | 11/09/2023 |  |  | 08/31/2025 |  | 11/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024391492 |  |  |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/11/2023 |  |  |  |  | The user facility reported a 62-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support for escalation of care from an impella cp pump. post impella 5.5 placement, the sterile sleeve was noted to be ripped. the patient was stable and remained on impella 5.5 support. it is unknown how the ripped sterile sleeve was addressed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported product damage (sterile sleeve damage) has been completed since the original report was filed. clinical notes detail that the sterile sleeve was found to be ripped after the insertion of the pump. no product was provided for evaluation. the cause of the sterile sleeve damage was not determined as the product was not returned for analysis.
32891 | 1220648-2023-04879 | 18108023 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18108023&pc=OZD | 10/11/2023 | 02/07/2024 |  | 11/09/2023 |  |  | 10/11/2023 | 11/09/2023 |  |  | 07/31/2025 |  | 11/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024388309 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation |  |  |  |  | 47 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Vascular Dissection | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/11/2023 |  |  |  |  | The user facility reported a 47-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that a patient on impella support had the pump stop and was unable to start it back due to allegedly a clot in the pump. the pump had to be explanted and replaced with a new one. upon explant, the artery was found perforated proximal to the 14fr sheath placement. a covered stent was applied. | [5 empty] | The investigation for the reported pump stop and vascular damage has been completed. the impella device was not returned for evaluation. however, review of clinical data and data logs were sufficient to determine a root cause. data logs revealed that about 15 minutes into the case, a motor current spike occurred and the pump stopped, which triggered an "impella stopped-motor current high" alarm. the patient's activated clotting time (act) was 133. the root cause of the pump stop was most likely due to biomaterial from lack of anticoagulation. the root cause of vascular damage was most likely due to use related as the vessel was perforated at the impella access site. updated g2: selected health professional updated h6 clinical code: 4440 correct: d4 model number. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
32892 | 1220648-2023-04880 | 18108029 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18108029&pc=OZD | 10/11/2023 | 01/11/2024 |  | 11/09/2023 |  |  | 10/11/2023 | 11/09/2023 |  |  | 06/30/2025 |  | 11/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024376501 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/11/2023 |  |  |  |  | The user facility reported a 63 -year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. there was an optical sensor warning appeared on the screen, purge system not working. staff reset the automated impella controller (aic) and replaced the aic to see if the pump would work properly. same issue occurred with both aics. the prompts to remove the wire and start the pump were not available. due to complex patient anatomy decision was made to replace the pump the new device was successfully replaced without any issue. no harm to patient reported due to the issue. . | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported optical signal issue has been completed since the original report was filed. the clinical details suggest that there was an ¿optical warning¿ alarm and that the ¿purge system was not flowing,¿ on both aic¿s. the logs from the first console cannot be viewed as it was a loaner console, and they were not returned. the logs were looked at from the second console (ic2833) and there were no optical alarms seen. the only alarm that was seen was ¿adjust lv.¿ the imc logs were viewed and there were no errors with the eeprom or instances of any optical issues. the optical parameters were taken in the lab and they were all within specification, insinuating there was no hard failure of the system. additionally, the logs show that the purge pressure is high at the start of the case but it does not trigger an alarm. the pressure can be seen decreasing at a steady rate. the pump was able to be purged in the lab and fluid was able to be seen exiting the gap. with the clinical information provided and the tests run in the lab, no problem was detected for the priming issue.
32893 | 1220648-2023-04881 | 18108031 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18108031&pc=OZD | 10/11/2023 | 03/22/2024 |  | 11/09/2023 |  |  | 10/11/2023 | 11/09/2023 |  |  | 07/31/2025 |  | 11/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024388075 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/11/2023 |  |  |  |  | The user facility reported a 61-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient's urine output was red. the impella cp was repositioned. the issue will be monitored. the patient remained on impella cp support for nine additional days and was weaned successfully. | [5 empty] | The investigation into the hemolysis issue has been completed since the initial report was submitted. the device was not returned for investigation. the data logs showed that the position of the pump was unknown along with suction and low pump flow. this was then resolved by repositioning the device. the root cause of the hemolysis is most likely due to improper positioning. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ . ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
32894 | 1220648-2023-04783 | 18091320 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18091320&pc=OZD | 10/30/2023 | 01/26/2024 |  | 11/08/2023 |  |  | 10/30/2023 | 11/08/2023 |  |  | 03/31/2025 |  | 11/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024339507 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/30/2023 |  |  |  |  | The us complainant had a cp pump supporting a patient when with patient movement/cough the pump dislodged out of the appropriate lv (left ventricle) position. as the patient was stable and being weaned off support the pump was explanted. no pump replacement took place. | [5 empty] | The investigation into the positioning issues has been completed. the device was not returned for investigation. as per clinical, patient was coughing and the impella dislodged and moved out of position alarming aorta and later impella was pulled out to explant. the cause of the positioning issue was due to patient movement with patient coughing leading to pump position in aorta per clinical review. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32895 | 1220648-2023-04743 | 18091357 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18091357&pc=OZD | 10/09/2023 | 06/06/2024 |  | 11/08/2023 |  |  | 10/09/2023 | 11/08/2023 |  |  | 04/30/2025 |  | 11/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024332874 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 23 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/09/2023 |  |  |  |  | The user facility reported a 23-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support as the patient was decompensating on intra-aortic balloon pump. on day four of the support, there was an access site bleed and small hematoma was also noted. the bleed was treated by turning off the heparin. the axillary site had an ultrasound performed. the hematoma was not further treated, but the team watched and questions if due to right internal jugular access for an impella rp or impella5.5 site as both access sites were in right torso. the patient was successfully bridged to heart transplant. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the pump was not returned for investigation. the root cause of the access site bleeding issue was most likely [low/high] patient activated clotting time values. the bleed was treated by turning off the heparin as per the clinical description. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 added code 925 as it was inadvertently left off the previously submitted manufacture device report. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted report incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32896 | 1220648-2023-04822 | 18091377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18091377&pc=OZD | 10/08/2023 | 06/03/2024 |  | 11/08/2023 |  |  | 10/08/2023 | 11/08/2023 |  |  | 08/31/2025 |  | 11/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024392374 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/08/2023 |  |  |  |  | The user facility reported a 65-year-old female (race unknown) undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. during support, the patient had bleeding at the impella access site. nurse stated that the patient required blood transfusion and pressure was held. the impella was explanted. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the pump was not returned for investigation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) . ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32897 | 1220648-2023-04782 | 18091380 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18091380&pc=OZD | 10/05/2023 | 08/21/2024 |  | 11/08/2023 |  |  | 10/05/2023 | 11/08/2023 |  |  | 07/31/2025 |  | 11/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024391208 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/05/2023 |  |  |  |  | It was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that use of the device cannot conclusively be associated with the outcome. | The user facility reported a 63year-old male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the patient was declining on current impella cp support. pulses were lost in the right lower extremities. an external fem bypass was placed temporarily. the patient was taken to the operation room and the impella cp was removed and the 5.5 was inserted without any issue. pulses were now present in both lower extremities post device placement. |  |  |  |  | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. the instructions for use for the impella cp with smartassist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ added-b5 mso review in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32898 | 1220648-2023-04832 | 18091384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18091384&pc=OZD | 10/17/2023 | 06/07/2024 |  | 11/08/2023 |  |  | 10/17/2023 | 11/08/2023 |  |  | 08/31/2025 |  | 11/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024395898 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Remove | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/17/2023 |  |  |  |  | The patient expired in the procedure area. however, the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there is no association between impella use and outcome. | The user facility reported a 53-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the impella was weaned and removed. sheath was pulled and angiogram performed of post closure, with pre-existing pro-glides revealed unsuccessful closure. multiple attempts were made but closure was unsuccessful. during placement of self-expanding stent proximally for dissection the patient became hypotensive and ventricular tachy-arrhythmias ensued. post angiogram showed successful hemostasis. after 40-50 minutes of advanced cardiovascular life support patient passed away. the death was not related to impella as per the physician. it was further reported that during the case a less than a golf ball sized hematoma was noted. one unit of blood was given to address the issue. |  |  |  |  | The investigation for the reported difficult to remove and subsequent bleeding issues has been completed. according to the clinical, the dissection/bleeding event occurred likely due to closure device failure. no product was returned for a visual inspection. data log analysis was not necessary for this complaint. the most likely cause of the vascular damage-peripheral vessels issue was use related. the instructions for use referenced in the initial report is for the impella cp with smart assist system in the general patient care considerations, entering cardiac output, potential adverse events (united states), which now includes the statement "cardiac or vascular injury (including ventricular perforation).¿, sections. added- b5 mso review in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32899 | 1220648-2023-04833 | 18091386 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18091386&pc=OZD | 11/03/2023 | 04/16/2024 |  | 11/08/2023 |  |  | 11/03/2023 | 11/08/2023 |  |  | 07/31/2025 |  | 11/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024391454 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/03/2023 |  |  |  |  | The user facility reported a 80-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the protect iv clinical trial case reported that the patient experienced atrial fibrillation with a left bundle branch block. it was further reported groin bleed at the impella access site. systemic heparin was held, manual pressure and fem stop required. | [5 empty] | The investigation for the reported access site bleed and atrial fibrillation has been completed. the device was not returned for evaluation. however, clinical details provided was sufficient to determine the root cause. the root cause of the access site bleeding was most likely due to use issue of patient movement since the bleed occurred in groin due to patient movement. the root cause of the the atrial fibrillation could not be determined as clinical details were not sufficient. d4-updated model number | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32900 | 1220648-2023-04834 | 18093192 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18093192&pc=OZD | 10/10/2023 | 01/24/2024 |  | 11/08/2023 |  |  | 10/10/2023 | 11/08/2023 |  |  | 04/30/2025 |  | 11/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024336981 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported a 61-year-old male patient was implanted with an impella 5.5 for care escalation from an impella cp. during the procedure, the impella 5.5 delivery failed due to anatomic limitations at the right axillary. the access site was therefore abandoned and the team planned for left axillary. the team later had to treat the access site as a vein was damaged during the access attempt. blood products and surgical repair was noted to have been needed. | [5 empty] | The investigation into the patient's access site bleeding and delivery issues has been completed. the product was not returned for evaluation. the root cause of the delivery issues - anatomy is most likely due to the patients diseased vessels, and the root cause of the access site bleeding - major is most likely due to use as the surgeon nicked the vein and intervention was required. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32901 | 1220648-2023-04835 | 18093277 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18093277&pc=OZD | 10/10/2023 | 01/23/2024 |  | 11/08/2023 |  |  | 10/10/2023 | 11/08/2023 |  |  |  |  | 11/08/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1205056 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Activation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported the automated impella controller (aic) had a purge disk error. the aic was therefore successfully replaced with another console. there was no reported patient harm. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the console was returned for investigation and was tested upon return. no logs exist from the reported occurrence date. during visual examination of this console, it appeared that the purge flag was still in place and purge disk retainer is secure. however, the purge retainer does have a visible hairline fracture in the portion of the purge retainer that is closer to the purge cassette. the hairline fracture would not have contributed to the consoles inability to detect the purge disc. the cause of the issue was not determined since the issue could not be reproduced.
32902 | 1220648-2023-04836 | 18093387 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18093387&pc=OZD | 10/10/2023 | 03/18/2024 |  | 11/08/2023 |  |  | 10/10/2023 | 11/08/2023 |  |  | 06/30/2025 |  | 11/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2024364386 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Nurse | Foreign | Health Professional | User facility | Company representation |  |  | 35 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user reported a 35-year-old patient was implanted with an impella cp pump for mechanical circulatory support. roughly five days after implant, it was documented that the pump stopped. the aic was swapped and the power was cycled, but the pump remained inoperative. the pump was subsequently replaced with another impella cp pump for ongoing support. there was no report of patient harm as a result of the failure. | [5 empty] | The investigation into the pump stop issue has been completed since the original report was filed. the device was not returned. as per the provided clinical information, additional to the pump stop, a purge leak was observed and resolved after the replacement of the purge cassette. the cause of the pump stop issue was not determined since product was not returned and sufficient clinical information was not returned. the cause of the purge leak issue was not determined since product was not returned. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
32903 | 1220648-2023-04837 | 18093538 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18093538&pc=OZD | 10/10/2023 | 01/12/2024 |  | 11/08/2023 |  |  | 10/10/2023 | 11/08/2023 |  |  |  |  | 11/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP OPTICAL |  |  | OZD |  |  | IMPELLA CP OPTICAL |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Migration | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user reported an impella heart pump was removed due to positioning issues. after correct placement, the heart pump slipped back into the aorta when the introducer sheath was removed. all attempts to place the pump correctly over the aortic valve were unsuccessful. the pump was therefore removed. there was no reported patient harm. | [5 empty] | The investigation into the reported positioning issues has been completed since the original report was filed. no product was returned. as per clinical review during the peel-away process, the impella was slipped back into the aorta and pump was unable to be correctly positioned again. the cause of the positioning issue was use related with pump position slipping into aorta during peel-away process. the failure modes will be monitored and trended. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32904 | 1220648-2023-04838 | 18093782 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18093782&pc=OZD | 10/10/2023 | 02/29/2024 |  | 11/08/2023 |  |  | 10/10/2023 | 11/08/2023 |  |  | 08/31/2025 |  | 11/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024393050 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/10/2023 |  |  |  |  | The user facility reported a 77-year-old male noted for high risk percutaneous coronary intervention was scheduled for impella cp implant for mechanical circulatory support. during attempted delivery, the impella cp would not pass the iliac. the impella cp was therefore removed from the sheath and clot was noted in the cannula in the distal end. a second pump was prepped and the short sheath was exchanged for the long impella sheath. the iliac was ballooned and the sheath was successfully placed. there was no reported patient harm. | [5 empty] | The investigation into the reported delivery issue -use, hematoma and vascular damage has been completed since the original report was filed. the pump was not returned for investigation. according to the clinical report, the patient was diagnosed with a calcified vessel condition. the cause of the delivery issue was patient condition related to calcified vessel condition since insertion was successful after placing long sheath and balloon angioplasty, based on given clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32905 | 1220648-2023-04839 | 18094348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18094348&pc=OZD | 10/10/2023 | 12/05/2025 |  | 11/08/2023 |  |  | 10/10/2023 | 11/08/2023 |  |  | 08/31/2025 |  | 11/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024389364 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Female |  |  |  | Malfunction | No | Yes | Air/Gas in Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/10/2023 |  |  |  |  | The user facility reported an 80-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the automated impella controller (aic) being used with an impella cp showed continuous alarm with air in system. the purge cassette was changed but the alarm continued and was not resolved until the pump was replaced. there was no known harm or injury to patient due to the issue. | [5 empty] | The impella pump was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The purge cassette was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. d.3 section revised as it was reported incorrectly on the initial manufacturer device report number 1220648-2023-04839. d.6a and d.6b added dates in accordance with updated procedures since the initial manufacturer device report number 1220648-2023-04839 was submitted. d.9 device return date was added as it was inadvertently omitted on initial manufacturer device report number 1220648-2023-04839. e.1 revised facility name and address line 1 as they were reported incorrectly on the initial manufacturer device report number 1220648-2023-04839. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2023-04839 and should not have been. g.1 revised mdr reporting contact email and manufacturing site fax number in accordance with updated procedures since the initial manufacturer device report number 1220648-2023-04839 was submitted. h.3 revised to reflect device was returned as it was reported incorrectly on the initial manufacturer device report number 1220648-2023-04839. h.6 codes 2199 and 3221 were reported incorrectly on the initial manufacturer device report number 1220648-2023-04839. h.10 additional manufacturer narrative was reported incorrectly on the initial manufacturer device report number 1220648-2023-04839 as a device was returned.
32906 | 1220648-2023-04840 | 18094388 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18094388&pc=OZD | 10/09/2023 | 04/22/2024 |  | 11/08/2023 |  |  | 10/09/2023 | 11/08/2023 |  |  | 08/31/2025 |  | 11/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024388739 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/09/2023 |  |  |  |  | The user facility reported a 63-year-old male presenting for acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. during support, suction alarms were triggered for which volume was given and continuous renal replacement therapy (crrt) was paused. issue was resolved and support continued. | [5 empty] | The investigation for the reported low pump flow has been completed. the pump was not returned for investigation. data logs show low pump flow without a corresponding placement signal trend and motor current spike, which was resolved after some time. as per the clinical information provided, suction alarms were reported and addressed by adjusting volume and temporarily pausing crrt. the cause of the low pump flow issue was most likely patient condition related, based on given clinical details. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
32907 | 1220648-2023-04841 | 18097588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18097588&pc=OZD | 10/09/2023 | 04/18/2024 |  | 11/08/2023 |  |  | 10/09/2023 | 11/08/2023 |  |  |  |  | 11/08/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1442220 |  | IC7536 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported the automated impella controller (aic) would not prime impella cassette and disc. the pump did not make a priming noise and fluid was coming out of the yellow port, but only at the gravity drip. staff attempted to reinstall the cassette twice and the aic would not prime. a new aic was brought in and the same purge cassette/purge disc was used and it was successful. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation into the aic - purge issue ¿ other issue has been completed since the initial report was submitted. the console was returned and tested. visual inspection of the purge assembly revealed excessive purge fluid residue, indicating leakage of purge fluid during clinical use. purge fluid residue was also found within the purge motor assembly, preventing the purge motor from spinning. the cause of the aic issue was contamination.
32908 | 1220648-2023-04842 | 18097670 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18097670&pc=OZD | 10/09/2023 | 02/28/2024 |  | 11/08/2023 |  |  | 10/09/2023 | 11/08/2023 |  |  | 07/31/2025 |  | 11/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | GM |  | 2024388108 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 64 YR |  | Male |  |  |  | Injury | No | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/09/2023 |  |  |  |  | The user facility reported a 64-year-old male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the patient was on bipella support with both a cp and rp that were both placed in the right groin. the leg was molted and cold. the family withdrew care on the same day this was noted. it's unclear which device caused or contributed to the molted/cold leg as both impella cp and rp were inserted in this limb. | [5 empty] | The investigation into the reported limb ischemia -reversible has been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/pos, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32909 | 1220648-2023-04843 | 18097671 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18097671&pc=PYX | 10/09/2023 | 01/24/2024 |  | 11/08/2023 |  |  | 10/09/2023 | 11/08/2023 |  |  | 06/30/2024 |  | 11/08/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | GM |  | 2023221294 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 64 YR |  | Male |  |  |  | Injury | No | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/09/2023 |  |  |  |  | The user facility reported a 64-year-old male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the patient was on bipella support with both a cp and rp that were both placed in the right groin. the leg was molted and cold. the family withdrew care on the same day this was noted. it's unclear which device caused or contributed to the molted/cold leg as both impella cp and rp were inserted in this limb. | [5 empty] | The investigation for the limb ischemia ¿ reversible issue has been completed. the impella device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the cause of the limb ischemia issue was most likely patient condition related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32910 | 1220648-2023-04844 | 18097673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18097673&pc=OZD | 10/09/2023 | 02/02/2024 |  | 11/08/2023 |  |  | 10/09/2023 | 11/08/2023 |  |  | 07/31/2025 |  | 11/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024380149 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/09/2023 |  |  |  |  | The user facility reported a 59-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient on impella support experienced low pump flow which the medical team was unable to resolve with repositioning of the pump. a decision was made to explant the pump. no patient harm reported due to the issue. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported low pump flow has been completed since the original report was filed. visual inspection of the device revealed biomaterial wrapped around the impeller. data log analysis revealed a motor current spike at the time of first suction alarm. the root cause of the low pump flow was due to biomaterial wrapped around the impeller.
32911 | 1220648-2023-04845 | 18097677 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18097677&pc=OZD | 10/09/2023 | 05/29/2024 |  | 11/08/2023 |  |  | 10/09/2023 | 11/08/2023 |  |  | 06/30/2025 |  | 11/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024372552 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/09/2023 |  |  |  |  | There's not enough information to associate use of device with outcome. | The user facility reported a 63-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient exhibited signs of hemolysis (hematuria and hemolyzed labs). the impella cp was explanted and an impella 5.5 was implanted as escalation of care. |  |  |  |  | The investigation into the hemolysis issue has been completed since the original report was submitted. the device was not returned for investigation. the cause of the hemolysis was not determined because no product was returned, and limited clinical information was given. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
32912 | 1220648-2023-04846 | 18097681 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18097681&pc=OZD | 10/10/2023 | 10/24/2025 |  | 11/08/2023 |  |  | 10/10/2023 | 11/08/2023 |  |  |  |  | 11/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1678477 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Electrical /Electronic Property Problem | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 10/10/2023 |  |  |  |  | The user facility reported a 68-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the automated impella controller (aic) had a controller error alarm. clinicians were advised not to switch the aic unless the pump has a failure. the patient is a status two heart transplant and is pump dependent. no harm reported to patient due to the issue. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. b.7 information was added as it was inadvertently omitted from initial manufacturer device report number 1220648-2023-04846. d.2 common device name was revised since the initial manufacturer device report number 1220648-2023-04846 was submitted in accordance with updated procedures. e.1 fax number was added as it was inadvertently omitted from initial manufacturer device report number 1220648-2023-04846. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2023-04846 and should not have been. g.1 revised mdr reporting contact email since the initial manufacturer device report number 1220648-2023-04846 was submitted in accordance with updated procedures. added fax number since the initial manufacturer device report number 1220648-2023-04846 was submitted in accordance with updated procedures. h.6 code 3221 was reported incorrectly on the initial manufacturer device report number 1220648-2023-04846 in accordance with updated procedures. | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. the impella device and data logs were returned for evaluation. the root cause of the controller error and loss of placement signal is due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed
32913 | 1220648-2023-04847 | 18097682 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18097682&pc=OZD | 10/10/2023 | 05/02/2024 |  | 11/08/2023 |  |  | 10/10/2023 | 11/08/2023 |  |  | 08/31/2025 |  | 11/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024393350 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/10/2023 |  |  |  |  | The user facility reported a 55-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the impella motor current unexpectedly dropped. suction occurred. the pump was removed and inspected for clot and then reinserted. a spike in motor current was observed. a new pump was inserted successfully. there was no patient harm or complications. | [5 empty] | The investigation for the reported low pump flow has been completed. the data logs were returned for investigation, and there was low flow and suction with no motor current spikes or placement signal trend before the pump was taken out. the pump had motor current spikes after it was inserted. the product was returned and there was a kink in the cannula. the root cause of the low pump flow is most likely due to the cannula kink. b.7 the patient's past medical history was inadvertently left off the original manufacture device report (mdr) and has been added. d.6a and d.6b revised from the initial submission in accordance with updated procedures. d.9 revised device returned to manufacturer? as previously entered incorrectly. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
32914 | 1220648-2023-04848 | 18097685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18097685&pc=OZD | 10/10/2023 | 03/04/2024 |  | 11/08/2023 |  |  | 10/10/2023 | 11/08/2023 |  |  | 06/30/2025 |  | 11/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024373020 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a 56-year-old male patient on impella cp support was found to have dark urine and lab specimens continued to result as hemolyzed. the pump was repositioned under echo and appears to be in good condition, but dark urine has continued. patient condition has improved dramatically, so the cardiology team started the wean process and the pump was eventually explanted. | [5 empty] | The pump was not returned for investigation. data logs show no signs related to motor current spikes or positioning issues. data logs show downward trending placement signal during pump support. as per the clinical details, patient's urine appears dark, and lab samples consistently show hemolysis. the cause of the hemolysis issue was most likely patient condition related since the data logs shows the downward placement signal trend. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
32915 | 1220648-2023-04849 | 18097712 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18097712&pc=OZD | 10/10/2023 | 02/05/2024 |  | 11/08/2023 |  |  | 10/10/2023 | 11/08/2023 |  |  | 07/31/2025 |  | 11/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024381085 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/10/2023 |  |  |  |  | The user facility reported a 37-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella cp was placed without any issues but when ramped up in auto, pump saturation was observed. the impella was removed and replaced. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the saturated flow issue has been completed. the device was returned for benchtop investigation. the returned pump visually was inspected and cannula kink near impella was observed. biomaterial around purge gap and impella was observed. the root cause of the saturated flow issue was biomaterial ingestion, and the cannula kink may have occurred during removal.
32916 | 1220648-2023-04850 | 18097713 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18097713&pc=OZD | 10/10/2023 | 01/31/2024 |  | 11/08/2023 |  |  | 10/10/2023 | 11/08/2023 |  |  | 06/30/2025 |  | 11/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024375523 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | A patient of unknown age and gender received hemodynamic support from an impella cp smartassist heart pump during high-risk percutaneous coronary intervention (hrpci). leg ischemia was reported below the impella pump access site. the impella cp heart pump was therefore removed and an impella 5.5 smartassist heart pump was inserted. the patient was subsequently stable. | [5 empty] | The investigation into the patient's limb ischemia has been completed. product was not returned for investigation. as per clinical records, the patient was escalated to a full support device due to limb ischemia. the cause of the limb ischemia issue was most likely patient condition related to small vessel. b1-selected serious injury b2 in the initial report had death incorrectly selected. there was no death reported. | The investigation into the patient's limb ischemia has been completed. product was not returned for investigation. as per clinical records, the patient was escalated to a full support device due to limb ischemia. without further information and considering the variety of potential factors such as the patient¿s clinical condition, history, risk factors and special investigation reports, it is not possible to determine the exact root cause of the limb ischemia and establish an association with impella use. revised h6 code/investigation conclusions from 50 to 4315 as we cannot definitively conclude the cause as attributed to patient's condition. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32917 | 1220648-2023-04851 | 18097718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18097718&pc=OZD | 10/25/2023 | 02/28/2024 |  | 11/08/2023 |  |  | 10/25/2023 | 11/08/2023 |  |  | 06/30/2025 |  | 11/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024370770 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/25/2023 |  |  |  |  | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported a patient presenting with poor perfusion in both lower extremities. the decision was made to remove the impella device due to diminished popliteal pulses that were found both pre and post impella insertion. after removal of the impella, however, a blood flow occlusion was detected due to a thrombus found caught in the patient's iliac stent. vascular surgery was required to restore blood flow. a determination on whether the issue arose due to impella explant process or closure device could not be determined. | [5 empty] | The investigation into the limb ischemia ¿ reversible has been completed. the device and data were not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32918 | 1220648-2023-04852 | 18097865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18097865&pc=OZD | 03/10/2023 | 11/08/2023 |  | 11/08/2023 |  |  | 03/10/2023 | 11/08/2023 |  |  |  |  | 11/08/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1239556 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Communication or Transmission Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The us complainant was setting up for patient to receive imeplla 5.5 support when the automated impella controller (aic) would not recognize the pump. troubleshooting measures were taken but, the team discovered the aic itself needed replacement. the instructions for use were followed and a backup contoller used. no harm or injury noted from delay. | [5 empty] | The investigation into the reported 'pump not detected' issue has been completed. console logs analysis from the reported day of event are mostly consistent with complaint and show that although pump connection was detected, the console was unable to read its eeprom, therefore unable to recognize the pump and its parameters. when console was received and tested with pump simulator and an engineering pump, the issue was able to be reproduced. the root cause of the automated impella controller (aic) issue was determined to be a defective impellatronics board. this report is being filed as part of a retrospective review of historical records.
32919 | 1220648-2023-04853 | 18097871 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18097871&pc=OZD | 10/14/2023 | 07/05/2024 |  | 11/08/2023 |  |  | 10/14/2023 | 11/08/2023 |  |  | 12/31/2024 |  | 11/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024305064 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  |  |  | Death | Yes | Yes | Pumping Stopped | [6 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/14/2023 |  |  |  |  | It was further reported that the patient's family made the decision to withdraw care. the patient expired. abiomed does not have enough information to exclude the device from the outcome. | The user facility reported a 76-year-old male (race unknown) post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. during support, the automated impella controller (aic) was displaying continual "air in purge" alarms causing the team to transfer to a new aic. once the pump was plugged into the new aic, an "impella defective" alarm was displayed and the p level was locked at p0. the team switched back to the original aic and the "impella defective" alarm persisted. the patient coded and return of spontaneous circulation was achieved approximately 15 minutes later. the impella 5.5 was explanted and patient sent to icu. |  |  |  |  | The investigation into the pump stop has been completed. data logs confirmed low purge and air in purge alarms on 10/14. the pump was unplugged, and upon plugging into a new console the impella defective alarm was triggered and the pump was unable to start. following a product evaluation of the pump, purge fluid was found in the red handle. the location of the leak was found to be a hole in the purge lumen within the catheter close to the motor housing. this leak allowed purge fluid to enter the red handle through the catheter, causing a corruption of the electronics within, resulting in the impella defective alarm. the cause of the pump stop was a purge leak in the purge line located close to the motor housing, resulting in corrupted electronics. this is considered a product malfunction. instructions for use for the related event: purge leak-pump impella 5.5 impella 5.5with smartassist for use during cardiogenic shock; section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ pump stop impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ b5 additional information b7 additional information f6 and f8 should have been blank in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. b2 added death and date of death h6 added health effect code. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
32920 | 1220648-2023-04854 | 18097902 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18097902&pc=OZD | 11/05/2023 | 01/15/2024 |  | 11/08/2023 |  |  | 11/05/2023 | 11/08/2023 |  |  | 04/30/2025 |  | 11/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024350753 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/05/2023 |  |  |  |  | The user facility reported a 70-year-old female was implanted with an impella device for mechanical circulatory support. the complainant reported that the pump was placed for support during angioplasty (ppci) and two hours into the case the impella stopped after an injection of ic nitro. the impella stopped due to high motor current and restarted on its own. the procedure was successful, and the patient is stable. | [5 empty] | The investigation for the reported pump stop has been completed. the impella device was returned for evaluation. upon visual inspection, there were no kinks noted. there was no biomaterial found on the pump. the pump passed electrical testing. when tested in the lab, the pump stop was unable to be reproduced at p-9. however, data logs were reviewed finding a high motor current spike immediately prior to the high motor current pump stop, which is indictive of biomaterial ingestion. the root cause of the pump stop was most likely biomaterial because the pump restarted on its own most likely as the clot passed through the pump. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
32921 | 1220648-2023-04855 | 18097908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18097908&pc=OZD | 10/13/2023 | 01/15/2024 |  | 11/08/2023 |  |  | 10/13/2023 | 11/08/2023 |  |  | 09/30/2024 |  | 11/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023282773 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  |  |  | Injury | Yes | Yes | Improper Flow or Infusion | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/13/2023 |  |  |  |  | The user facility reported a 78-year-old male (race unknown) in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. during support, the patient had had a hematoma around the access site. manual pressure was held. impella began to have suction events and the doctor noticed the hematoma was growing. the decision was made to remove and replace the impella. | [5 empty] | The investigation for the reported access site bleeding and the low pump flow has been completed. the impella device was returned for evaluation. the pump was visually inspected no damages were noted. the pump was ran in a basin of water and the low pump flow could not be reproduced. without the logs, the low flows could not be confirmed during the case. clinical information provided states that at the time of the bleed the activated coagulation time (act) was at 400 so the patient needed to receive 11,000 units of heparin. the root cause of the reported access site bleed was most likely due to anticoagulation management. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32922 | 1220648-2023-04856 | 18097911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18097911&pc=OZD | 10/12/2023 | 06/11/2024 |  | 11/08/2023 |  |  | 10/12/2023 | 11/08/2023 |  |  | 02/28/2025 |  | 11/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024313278 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/12/2023 |  |  |  |  | The user facility reported a 66-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support as a bridge to left ventricular assist device (lvad). while on impella support, the patient had a drop in hemoglobin and hematocrit and the patient had excess oozing at the access. the patient received multiple blood products. the patient was successfully implanted with lvad. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. a.5 revised ethnicity as previously entered incorrectly. b.7 revised as this information was inadvertently omitted from the previously submitted report. d.4 revised unique identifier (udi) # as previously entered incorrectly. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted report incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures. additional manufacturer narrative on the previously submitted report stated that the device was discarded but the device was not returned. type of investigation codes has been updated to reflect device not returned. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32923 | 1220648-2023-04857 | 18097916 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18097916&pc=OZD | 10/13/2023 | 12/04/2023 |  | 11/08/2023 |  |  | 10/13/2023 | 11/08/2023 |  |  | 07/31/2025 |  | 11/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024388346 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that while placing swan ganz, the impella cp went into the aorta and was unable to cross back in the left ventricle (lv). the impella was therefore explanted and a new impella cp was placed. | [5 empty] | The investigation into the malfunction of positioning issues has been completed. no product was returned. as per the clinical review while placing swan-ganz catheter, impella was pulled out into aorta and unable to cross back into the left ventricle (lv). no wireless re-access or other abnormalities were observed. the cause of the positioning issue was most likely use related with impella being pulled out into aorta and unable to reposition while accessing swan-ganz catheter per clinical review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32924 | 1220648-2023-04858 | 18097918 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18097918&pc=OZD | 10/13/2023 | 03/26/2024 |  | 11/08/2023 |  |  | 10/13/2023 | 11/08/2023 |  |  | 08/31/2025 |  | 11/08/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024392965 |  | 0048-0045 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 70-year-old male noted to be high risk for percutaneous cardiac intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient had hemolysis. echo completed with repositioning was done but hemolysis did not clear after repositioning. the impella was therefore weaned off and patient hemodynamics improved. | [5 empty] | The investigation for the reported hemolysis has been completed. the impella device was not returned for evaluation. data logs were reviewed which showed pump position was in aorta that triggered impella position in aorta alarms. based on clinical description and data analysis, the root cause of hemolysis was most likely due to pump was not in the proper position. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
32925 | 1220648-2023-04826 | 18084349 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18084349&pc=OZD | 10/08/2023 | 02/07/2024 |  | 11/07/2023 |  |  | 10/08/2023 | 11/07/2023 |  |  | 08/31/2025 |  | 11/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024389328 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a 57-year-old male patient on impella 5.5 support experienced a pump saturation problem which was resolved by using a tissue plasminogen activator (tpa) through the purge. there was no harm to the patient as a result of this incident. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow issue has been completed. the device was not returned for investigation. based on the clinical information provided, the patient had a high lv diastolic waveform, high purge pressure, and saturated flow. tpa was infused through the purge system and these issues were resolved. the logs confirm a gradual increase in motor current followed by saturated flow and spikes in purge pressure and placement signal. following tpa infusion, the flow, motor current, purge pressure, and placement signal decreased. the most likely cause of the saturated flow is biomaterial.
32926 | 1220648-2023-04827 | 18084413 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18084413&pc=OZD | 10/11/2023 | 02/02/2024 |  | 11/07/2023 |  |  | 10/11/2023 | 11/07/2023 |  |  | 01/31/2025 |  | 11/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024306257 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant in japan reported the patient had lower extremity ischemia on the impella cp side. the patient had a history of previous percutaneous transluminal angioplasty (pta), however the relationship between the ischemia and the impella cp is unknown. the patient suffered from crush syndrome due to lower limb ischemia and the patient outcome worsened. the patient was on impella cp support for 143.75 hours before the patient expired. the relationship between the patient death and the impella cp is unknown. | [5 empty] | The investigation for the reported limb ischemia has been completed. product was not returned for review. based on clinical description, the root cause of limb ischemia was most likely due to patient condition as the patient had a history of percutaneous transluminal angioplasty (pta) before inserting impella. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32927 | 1220648-2023-04828 | 18084498 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18084498&pc=OZD | 10/27/2021 | 01/31/2024 |  | 11/07/2023 |  |  | 10/11/2023 | 11/07/2023 |  |  | 08/31/2023 |  | 11/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022129424 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/11/2023 |  |  |  |  | The protect iv study team reported a patient of unknown age, gender or race experienced three "possible" relationships between impella cp support and the increase in patient's plasma free hemoglobin, cardiac enzymes and creatine levels. the patient was noted to be high risk percutaneous coronary intervention. it is unknown how this was managed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the reported hemolysis issue has been completed. the impella device was not returned for evaluation. the data logs did not record any events related to improper positioning of the pump. the pump was implanted after the hemolysis event occurred. the root cause of the hemolysis issue was not determined and determined to be unrelated to the impella.
32928 | 1220648-2023-04829 | 18084570 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18084570&pc=OZD | 10/12/2023 | 08/21/2024 |  | 11/07/2023 |  |  | 10/12/2023 | 11/07/2023 |  |  | 04/30/2025 |  | 11/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024348647 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 66-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient had an unidentifiable structure in left ventricle (lv) near tip of impella via bedside echo. team could not tell if structure was actually in contact with impella. no suction alarms were present. patient was on anticoagulants. the structure was not removed or analyzed as the decision was made to withdraw care. the family made choice and patient expired. | [5 empty] | The investigation of thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6a and d6b added. h6 component code added. mso statement: per medical safety officer review use of the device cannot conclusively be associated with the outcome of death. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-04829. b2 outcomes attributed to adverse event was missing. g1 reporting contact fax number missing. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32929 | 1220648-2023-04830 | 18087544 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18087544&pc=OZD | 10/09/2023 | 02/21/2024 |  | 11/07/2023 |  |  | 10/09/2023 | 11/07/2023 |  |  | 07/31/2025 |  | 11/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024390708 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported the 58-year-old male patient implanted with an impella had hemolysis that was confirmed with ldh >1000 and pfh elevated as high as 290. the patient was escalated to the impella 5.5 as a result of the noted condition. | [5 empty] | The investigation into the reported hemolysis has been completed since the original report was filed. pump was not returned for investigation, data logs does not indicated any abnormalities related to improper position or interaction with biomaterial. the cause of the hemolysis issue was not determined since product was not returned and sufficient clinical information was not provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis." ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
32930 | 1220648-2023-04831 | 18087993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18087993&pc=OZD | 10/12/2023 | 01/31/2024 |  | 11/07/2023 |  |  | 10/12/2023 | 11/07/2023 |  |  | 05/31/2025 |  | 11/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024360331 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/12/2023 |  |  |  |  | The user facility reported a 70-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, an alarm occurred stating "impella position in aorta". the impella was repositioned under echocardiogram. there was no patient harm. pump was weaned and explanted after 55 hours of support. | [5 empty] | The investigation into the low pump flow issue has been completed. the pump was not returned for investigation. according to clinical records, the pump was repositioned after triggering an impella position in aorta alarm. the cause of the positioning issue was not determined as insufficient clinical details about the event were provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32931 | 1220648-2023-04803 | 18073461 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18073461&pc=OZD | 10/09/2023 | 02/28/2024 |  | 11/06/2023 |  |  | 10/09/2023 | 11/06/2023 |  |  | 08/31/2025 |  | 11/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024393340 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 94 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Pain | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had an impella cp placed via the femoral artery for support during high risk percutaneous coronary intervention (hrpci) of an elderly patient the surgeons had turned down. the patient had limb pain due to the impella cp insertion. the team felt the pump/sheath access was causing limb ischemia. the team therefore made choice to explant the pump. | [5 empty] | The investigation into the limb ischemia ¿ reversible has been completed. the device and data were not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32932 | 1220648-2023-04804 | 18073495 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18073495&pc=OZD | 10/09/2023 | 02/09/2024 |  | 11/06/2023 |  |  | 10/09/2023 | 11/06/2023 |  |  | 08/31/2025 |  | 11/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024392980 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | No | Material Puncture/Hole | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that upon suturing impella cp peel away sheath for male patient, a hematoma was noted and pulsatile flow coming from around the sheath. the physician therefore decided to remove the impella due to stabilization of the patient's hemodynamics. once the sheath was removed, a small needle hole was seen on the peel away sheath. it was further reported that the hole was from the physician puncturing the sheath with a micro puncture needle. the patient remained stable after impella removal. | [5 empty] | The investigation into the access site bleeding ¿ major and the product damage (hole in catheter) issues has been completed. the device was not returned for investigation. based on the clinical information provided, the root cause of the access site bleeding cannot be determined due to insufficient clinical information. however, the root cause of the hole in introducer was due to use, as it was stated from the representative that the physician punctured a microneedle hole in the sheath about an inch from the hemostatic valve. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32933 | 1220648-2023-04805 | 18073555 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18073555&pc=OZD | 10/09/2023 | 04/12/2024 |  | 11/06/2023 |  |  | 10/09/2023 | 11/06/2023 |  |  | 04/30/2025 |  | 11/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024355094 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | Yes | Increase in Pressure | Activation, Positioning or Separation Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information provided states the patient had ischemia to the access site leg and ecmo leg was also reported. it is unknown when the limb ischemia began. it was diagnosed through a cold foot and a distal limb perfusion catheter was placed to try to resolve the issue; it was unsuccessful. the patient has peripheral vascular disease. the pump remained on support. | The complainant reported that a male patient on impella cp support experienced bleeding from the access site for which the patient had to receive blood products in order to make the patient hemodynamically stable. additionally, the patient also experienced an impella cp in the aorta alarm and a high pressure alarm. the medical staff therefore had to reposition the pump. the patient is still on impella support, however he is in a critical condition. |  |  |  |  | The investigation for the reported access site bleed, positioning issues, and high purge pressures has been completed. the impella device was not returned for investigation. however, clinical details provided were sufficient to determine the root cause of the many issues. bleeding was reported from multiple access sites. the cause of the access site bleeding issue was most likely patient condition related since patient had multiple access site bleeding. as far as positioning issues, data logs were reviewed which showed several "impella position in aorta" alarms with noted saturation in pump flow. this condition was happened after changing pump to different console (aic) and it was resolved after some time. the cause of the positioning issue was most likely patient movement, as the issue occurred after transferring the patient to a higher level of care and was resolved after repositioning. as far as high purge pressure, data logs shows the quick spike in purge pressure, which was resolved within 7 minute with one intermediate resolution in purge pressure. the cause of the issue high purge pressure was most likely kinked purge line, based on data log review. clinical details were not sufficient to determine the root cause of the limb ischemia. therefore the root cause of the limb ischemia could not be determined. the instructions for use states ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32934 | 1220648-2023-04806 | 18073565 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18073565&pc=OZD | 10/08/2023 | 06/24/2025 |  | 11/06/2023 |  |  | 10/08/2023 | 11/06/2023 |  |  |  |  | 11/06/2023 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1141484 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  |  |  | Death | Yes | Yes | Connection Problem | Priming Problem | [5 empty] | Cardiac Arrest | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [10 empty] | Required Intervention | Life Threatening | Death | [22 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the automated impella controller (aic) pressure sensor did not recognize impella cp pump after priming. the aic was therefore swapped for a different console but the patient expired prior to pump insertion. | The complainant reported a patient in cardiogenic shock from an acute myocardial infarction was being prepped for an impella cp device implant for mechanical circulatory support. after priming, the automated impella controller (aic) pressure sensor did not recognize the impella cp pump. consequently, the aic had to be exchanged for the backup aic. however, the patient would expire before the impella cp could be implanted for support. |  |  |  |  | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation for the console (aic) pump issue has been completed. data logs revealed that the console had experienced the epm crystal issue, and that the console had then automatically performed steps to mitigate it. when the epm crystal issue occurs, the console is unable to recognize the presence of the pump. this matches the reported complaint. the console was returned for evaluation. . the reported complaint could not be reproduced. the pump detection issue was confirmed to be due to the rarely occurring epm crystal issue. the impellatronic pca does not have permanent damage, but this intermittent pump detection failure occurs infrequently and already has a sw mitigation in place. d2-corrected common device name. | Medical safety officer reviewed the event and determined there is not enough information to exclude the automated impella controller as an associated factor in the patient's outcome. event was revisited and determined there is an association with the demise outcome. consequently, the following updates to respective sections of this follow-up 2 supplemental medwatch report were completed accordingly: section b1: to reflect adverse event and product problem. section b2: outcomes attributed to adverse event. section b5: event description details including adverse event/outcome. section h1: type of reportable event death. section h6: health effect/medical device/component codes. note: investigation outcome statement(s) and its h6 coding (type, findings and conclusions) remain unchanged as previously reported under follow-up 1 supplemental medwatch report 1 report.
32935 | 1220648-2023-04807 | 18073593 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18073593&pc=OZD | 10/07/2023 | 11/06/2023 |  | 11/06/2023 |  |  | 10/07/2023 | 11/06/2023 |  |  |  |  | 11/06/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1122151 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant in the eu reported patient was treated for acute myocardial infarction/cardiogenic shock and was scheduled to receive hemodynamic support with an impella cp pump. during the preparation of the impella cp heart pump, the automated impella controller (aic) did not recognize the purge cassette. the backup controller was therefore used and the problem was solved. there was no reported patient harm. | [5 empty] | The investigation into the reported aic - purge disc/flag issues has been completed. review of data logs confirmed that a purge disk not detected alarm occurred. the console was received for analysis and tested. upon examining the automated impella controller (aic) for any visible defects, it was evident that the purge disc detection flag was broken which would result in a ¿purge disc not detected¿ alarm. the root cause of this issue was determined to be a broken/missing purge disc detection flag.
32936 | 1220648-2023-04808 | 18075349 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18075349&pc=OZD | 10/22/2023 | 06/04/2024 |  | 11/06/2023 |  |  | 10/22/2023 | 11/06/2023 |  |  | 05/31/2025 |  | 11/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024357000 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/22/2023 |  |  |  |  | The user facility reported a 76-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. rn went to redress the site and claimed that the touhy borst valve was loose, the catheter slipped back to 93cm and there was increased bleeding at the site. md at bedside to reposition pump with echo guidance, and the pigtail remained across the aortic valve, md was able to re-advance the catheter back to the previous 96cm and locked the catheter in place. bleeding subsided with a brief period of manual pressure. | [5 empty] | The investigation into the positioning issues and the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. based on the data analysis, the most likely cause of the positioning issues was a touhy-borst issue. the most likely cause of the access site bleeding was use related. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32937 | 1220648-2023-04809 | 18075492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18075492&pc=OZD | 10/07/2023 | 02/21/2024 |  | 11/06/2023 |  |  | 10/07/2023 | 11/06/2023 |  |  | 05/31/2025 |  | 11/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024357072 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Migration | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after a 63-year-old male patient was transported to another facility, the impella cp pulled back across the aortic valve. therefore, the physician decided to remove the impella and continue full extracorporeal membrane oxygenation (ecmo). | [5 empty] | The investigation into the reported positioning issue has been completed since the original report was filed. pump was not returned for investigation. data confirmed pump was in aortic area (about 8.25 hours into the case). about 2.75 hours later, pump was stopped manually & disconnected from the console. product was not returned for review. based on clinical description, the root cause of positioning issue was most likely due to patient movement. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32938 | 1220648-2023-04810 | 18075514 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18075514&pc=OZD | 10/10/2023 | 09/21/2024 |  | 11/06/2023 |  |  | 10/10/2023 | 11/06/2023 |  |  | 12/31/2024 |  | 11/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024286495 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Female |  |  | White | Death | Yes | No | Device Damaged by Another Device | [6 empty] | Unspecified Vascular Problem | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | A 72-year-old female in the eu, with a previous medical history of cardiac surgery (coronary artery bypass graft), received hemodynamic support from an impella cp pump during a high-risk percutaneous coronary intervention (hrpci). during removal of the impella cp device after successful pci procedure, the subclavian/aorta stent that had been in place from a previous intervention was noted to have been deformed. it is unclear if the impella device caused the deformation, however it cannot be excluded. the stent had been placed in the subclavian artery to stabilize the patient. cpr was noted to have been required during the procedure. the same impella device was also reinserted to further support the patient, despite recommendation to use a new one. one day after the adverse event, the patient died from multi organ failure. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-04810 was provided in sections b2, b5, d6a, d6b, h1, and h6. | The investigation into the vascular damage - great vessel issue has been completed. the device was not returned for investigation. according to clinical details, during the device removal, a subclavian/aorta stent came out along with the pump. despite the recommendation for a new pump, the doctor chose to reuse the existing one, which operated effectively after cpr was administered. the cause of the vascular damage issue was most likely use related based on given clinical details.
32939 | 1220648-2023-04811 | 18075520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18075520&pc=PYX | 10/07/2023 | 04/12/2024 |  | 11/06/2023 |  |  | 10/07/2023 | 11/06/2023 |  |  | 01/31/2024 |  | 11/06/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023174031 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Migration | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/07/2023 |  |  |  |  | The user facility reported a 66-year-old male with bilateral pulmonary embolism was implanted with an impella rp device for mechanical circulatory support. the pump was initially delivered successfully; however, after sheath peeled away, the pump came out of appropriate position. the physician aborted impella rp usage for medications. | [5 empty] | The investigation into the positioning issues has been completed since the original report was submitted. the device was not returned for investigation. as per given clinical details, the cause of the positioning issue was use related since the impella was pulled out while removing the repositioning sheath. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32940 | 1220648-2023-04812 | 18075540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18075540&pc=OZD | 10/10/2023 | 01/17/2024 |  | 11/06/2023 |  |  | 10/10/2023 | 11/06/2023 |  |  | 07/31/2025 |  | 11/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024391253 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported that the impella cp was having continuous suction alarms. the staff therefore tried to reposition the device, but was unsuccessful. additionally, there was bleeding from the impella access site that had to be redressed. the impella was ultimately replaced due to low pump flow and concern for a clot ingestion. | [5 empty] | The impella devices were received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ | The investigation into the reported low pump flow and positioning issue has been completed since the original report was filed. cp pump returned for investigation. visually inspected the pump and no cannula kink, catheter kink or biomaterial ingestion were found. no abnormalities relative to impeller blades were observed. pump was connected to the console and ran at p-9. the pump was running at a flow of 3.8 l/min at p-9 without any issues. no mechanical or physical abnormalities were found, and low pump flow issue did not reproduce. data logs were reviewed, several suction alarms were found with p-level drop and repositioning at p-2 did not resolve suction alarms at end. no motor current spike observed relative to biomaterial ingestion. as per clinical impella access site was re-dressed, and the angle of insertion was not appreciated during the dressing change causing the access site to bleed. pump having low flows after p-level drop and continuously suctioning on p-2 even after repositioning device. second cp inserted had pump saturation and replaced with third cp pump started in auto but moved to p4 since increasing p-level consistently resulted in suction alarms. the cause of the access site bleeding issue was use related since the angle of insertion maintained was not appreciated during the dressing change causing the groin site to bleed and manual pressure held to resolve the issue. the cause of the low pump flow issue was patient condition related with low flow issue not reproducing during field investigation and several suction alarms observed with low flows based on log, clinical review.
32941 | 1220648-2023-04813 | 18075543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18075543&pc=OZD | 10/07/2023 | 02/15/2024 |  | 11/06/2023 |  |  | 10/07/2023 | 11/06/2023 |  |  | 05/31/2025 |  | 11/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024354613 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Material Puncture/Hole | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/07/2023 |  |  |  |  | The user facility reported a 61-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after 30 minutes of support there was bleeding from the red plug. gauze was taped around it but it continued to leak. as a result, the physician replaced the impella cp to address the issue. | [5 empty] | The investigation into the product damage (hole in catheter) and the access site bleeding ¿ minor has been completed since the original report was filed. the device was returned for investigation. visually inspected and found a significant hole in the catheter. the root cause of the access site bleeding issue was patient movement related per clinical review. the cause of the product damage (hole in the catheter) issue was use related due to puncture hole observed close to 80 cm marking per field investigation. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
32942 | 1220648-2023-04814 | 18075720 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18075720&pc=OZD | 10/10/2023 | 04/22/2024 |  | 11/06/2023 |  |  | 10/10/2023 | 11/06/2023 |  |  | 08/31/2025 |  | 11/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024391667 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported a 47-year-old male patient was on impella 5.5 support ongoing for a week with persistent suction alarms. the team therefore made choice to explant the 5.5 and place an impella cp instead. the probable cause was determined to be perhaps clot formation per the medical team. there was no reported patient harm. | [5 empty] | The investigation for the reported low pump flow has been completed. the impella device was not returned for evaluation. data logs show several suction alarms and low pump flow. there was no motor current spike or placement signal trend noted during the case. as per clinical practice, the doctor explanted the pump after unresolved suction alarms for 24 hours. the cause of the low pump flow issue was not determined since product was not returned and sufficient clinical information was not provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿
32943 | 1220648-2023-04815 | 18076507 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18076507&pc=OZD | 10/10/2023 | 09/21/2024 |  | 11/06/2023 |  |  | 10/10/2023 | 11/06/2023 |  |  | 07/31/2025 |  | 11/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024392402 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Improper Flow or Infusion | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | Other | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/10/2023 |  |  |  |  | The user facility reported a 70-year-old female in acute myocardial infarction/cardiogenic shock was scheduled for percutaneous coronary intervention (pci). the patient coded for several rounds and the decision was made to place an impella cp after return of spontaneous circulation before pci. during support, the pump had suction (low flow) issues the entire time. the patient had ventricular fibrillation arrests during support. the family made the choice to withdraw care. the pump had been on support for less than five hours. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation for the reported low pump flow and the ectopy. the pump was not returned for investigation. data logs shows low flow since the beginning of the case, followed by several 'impella position in aorta' events. downward placement signal trend was observed during the case run. according to the clinical notes, the doctor observed that the impella was positioned in the aorta, and the patient experienced a ventricular fibrillation arrest during the use of the pump, resulting in their passing. when the device was placed in a deeper position, the flows leveled. the cause of the positioning issues is most likely due to use. the root cause of the vt/vf was not determined.'' the instructions for use states ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ added- h6 impact code 4641. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6.
32944 | 1220648-2023-04816 | 18077111 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18077111&pc=PYX | 10/10/2023 | 01/12/2024 |  | 11/06/2023 |  |  | 10/10/2023 | 11/06/2023 |  |  | 08/31/2025 |  | 11/06/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024397299 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Increase in Pressure | Improper Flow or Infusion | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/10/2023 |  |  |  |  | The user facility reported a 38-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support during surgery. during the permanent pacer lead implant procedure, the impella rp had purge pressure/purge flow alarms. the patient was also being supported by an impella 5.5. the purge alarms were troubleshot by intervention and adding tissue plasminogen activator to the purge. the team also noted the patient displayed signs of hemolysis. the urine was tinged and plasma free hemoglobin was elevated. the team explanted the impella rp and observed a thrombus burden at the motor and cannula. | [5 empty] | The investigation into the high purge pressure and the hemolysis has been completed. the device was not returned for benchtop evaluation and investigation. per the clinical information provided, the root cause of the hemolysis and high purge pressure is most likely due to ingested biomaterial. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follow. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions: including catheter position, pre-existing medical conditions, and small left ventricular volumes may also play a role in patient susceptibility to hemolysis.¿
32945 | 1220648-2023-04817 | 18078319 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18078319&pc=OZD | 10/26/2023 | 10/24/2025 |  | 11/06/2023 |  |  | 10/26/2023 | 11/06/2023 |  |  |  |  | 11/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1273696 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 10/26/2023 |  |  |  |  | The user facility reported that during a case preparation, the automated impella controller, (aic) port was damaged. a replacement loaner was sent to the facility. no patient is involved. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | A.1 revised as none was inadvertently omitted from the initial manufacturer device report number 1220648-2023-04817. a.2 age should have been left blank on the initial manufacturer device report number 1220648-2023-04817 as it is unknown. e.2 revised as it was incorrectly entered on the initial manufacturer device report number 1220648-2023-04817. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2023-04817 and should not have been. g.1 revised reporting contact email in accordance with updated procedures since the initial manufacturer device report 1220648-2024-10844 was submitted. added reporting contact fax number as it was inadvertently omitted from the initial manufacturer device report 1220648-2024-10844. h.6 code 2199 was entered incorrectly on the initial manufacturer device report 1220648-2024-10844.
32946 | 1220648-2023-04818 | 18081208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18081208&pc=OZD | 10/07/2023 | 02/21/2024 |  | 11/06/2023 |  |  | 10/07/2023 | 11/06/2023 |  |  | 02/08/2025 |  | 11/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - APAC |  |  | ABIOMED, INC. |  |  | US |  | 2024345187 |  | 0048-0047 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | Distributor |  | 46 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/07/2023 |  |  |  |  | The distributor reported a 46-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support in medicover hospital in india. the patient had access site bleeding and a hematoma. the impella cp was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported access site bleeding has been completed since the original report was filed. pump was not returned for investigation. the cause of the access site bleeding issue was not determined since product was not returned and sufficient clinical information was not provided.
32947 | 1220648-2023-04819 | 18081210 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18081210&pc=OZD | 10/07/2023 | 08/20/2024 |  | 11/06/2023 |  |  | 10/07/2023 | 11/06/2023 |  |  | 07/31/2025 |  | 11/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024390793 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | Death | Other | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/07/2023 |  |  |  |  | The patient coded overnight, and the family decided to withdraw care, and the patient expired. per the medical safety officer review there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | The user facility reported a 69-year-old female with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the patient had no pulses present, in the rt. foot to which the md believed to be caused by clots and as a result had to perform an external bypass. according to the md, the pump¿s sheath was not flushed prior to insertion. then the patient had popliteal pulses however the vascular surgeon stated that there was nothing they could do at that point due to the patient being too sick. the patient was placed on palliative care and after coding overnight the family decided to withdraw care. |  |  |  |  | The investigation of limb ischemia and thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b2 updated to death and date of death provided. b5 updated with additional information. h1 type of reportable event updated to death. h6 health effect - impact code 1802 added. instructions for use as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings and cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-04819: in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8-should have been left blank. g1 reporting contact fax number missing. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32948 | 1220648-2023-04820 | 18081233 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18081233&pc=OZD | 10/07/2023 | 02/05/2024 |  | 11/06/2023 |  |  | 10/07/2023 | 11/06/2023 |  |  | 04/30/2025 |  | 11/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024353441 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | Thrombosis/Thrombus | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/07/2023 |  |  |  |  | The user facility reported a 67-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support during cardiac surgery. three days after implant, the team noted concern for a left ventricle thrombus. the patient remained on the impella cp support until (b)(6) 2023. there is insufficient information in the file at this time regarding the thrombus whether it was confirmed and how it was managed. | [5 empty] | The investigation into the issue of thrombus has been completed. the device was not returned for investigation. during the time of reported suspected thrombus, there were multiple suction alarms, motor current spikes and a placement signal spike. however, these conditions resolve towards the end of the case, indicating that the biomaterial most likely passed through. additionally, there was suspected sepsis reported. the cause of the infection was most likely patient condition related since the device is sterilized under validated conditions and there was no evidence to indicate a break in the sterile barrier or pump contamination. unknown relation to impella. the root cause of the thrombus is most likely due to the patient¿s condition. added h6 clinical code: 2067. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32949 | 1220648-2023-04821 | 18081242 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18081242&pc=PYX | 10/07/2023 | 03/19/2024 |  | 11/06/2023 |  |  | 10/07/2023 | 11/06/2023 |  |  | 05/31/2025 |  | 11/06/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024355141 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Pumping Stopped | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/08/2023 |  |  |  |  | The user facility reported a 58-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. the impella had high purge pressure. staff denied kinks, attempted cassette change/de-air to no avail. while waiting for tpa, the pump stopped. the impella rp was replaced without any issues. | [5 empty] | The investigation into the high purge pressure and the pump stop issues has been completed since the original report was filed. the device was not returned for investigation. the cause of the high purge pressure leading to a pump stop was most likely biomaterial, based on data log review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
32950 | 1220648-2023-04824 | 18081243 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18081243&pc=OZD | 10/08/2023 | 01/21/2025 |  | 11/06/2023 |  |  | 10/08/2023 | 11/06/2023 |  |  |  |  | 11/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1364511 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant in japan reported the automated impella controller (aic) would not recognize the purge set during priming. the aic was therefore replaced successfully. there was no reported patient harm. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - purge issue ¿ other has been completed. according to the log analysis and device analysis the root cause for the console not recognizing purge disc during priming was due to a damaged purge pressure flag. the following have been reported incorrectly or missing from manufacturer device report: 1220648-2023-04824. d2: common device name. g1: reporting contact email revised in accordance with updated procedures. g1: reporting contact fax number missing.
32951 | 1220648-2023-04825 | 18081250 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18081250&pc=OZD | 10/08/2023 | 02/09/2024 |  | 11/06/2023 |  |  | 10/08/2023 | 11/06/2023 |  |  | 07/31/2025 |  | 11/06/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024380067 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/08/2023 |  |  |  |  | The user facility reported a 81-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. patient on impella support was found to have a hematoma at the access site. imaging showed bleeding from what appeared to be from side of repo sheath. the md decided to explant the pump and have a new one implanted. additionally, patient also received blood transfusion to address the issue. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was filed. the product was returned for investigation. upon investigation of the pump, no abnormalities or damages were observed. the introducer was not returned. the root cause of the access site bleeding - major was most likely a lack of vessel recoil onto the sheath. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32952 | 1220648-2023-04800 | 18070727 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18070727&pc=OZD | 02/07/2023 | 11/05/2023 |  | 11/05/2023 |  |  | 02/07/2023 | 11/05/2023 |  |  |  |  | 11/05/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1551790 |  | 0042-0010 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported had a console failure alarm. the ifu was followed and a backup controller was utilized. there was no noted harm or injury and the support with the 5.5 pump continues. | [5 empty] | The investigation into the reported 'aic - battery failure (red alarm)" has been completed. product was returned for investigation. the device was inspected and found the failure mode was reproduced on a field service lab bench. using batttest analysis, the battery one was confirmed to have cell imbalance which caused the battery lockout internally and prevent charging which resulted eventual depletion and communication failure. the battery failure alarm was due to a defective battery.
32953 | 1220648-2023-04801 | 18070729 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18070729&pc=OZD | 01/30/2023 | 11/05/2023 |  | 11/05/2023 |  |  | 01/30/2023 | 11/05/2023 |  |  |  |  | 11/05/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1587787 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant was attempting to start an impella rp support. the automated impella controller (aic) failed and would not identify the purge cassette at the disc. the ifu could not be followed. there was no backup controller to use for the patient support. the team instead placed a competitor's device for support. | [5 empty] | The investigation into the reported "aic - purge disc/flag issues" has been completed. product was returned for investigation. the device was inspected and found that no issues with purge cassette and disc detection could be reproduced in the lab. no visible damage to the retainer or purge flag. during data analysis, there were no apparent issues detecting the purge cassette or the disc. aic console logs don't show evidence of any purge issue and were not consistent with the problem description.the root cause of the purge disc detection issues were not confirmed nor determined due to inability to reproduce.
32954 | 1220648-2023-04802 | 18070730 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18070730&pc=OZD | 03/03/2023 | 11/05/2023 |  | 11/05/2023 |  |  | 03/03/2023 | 11/05/2023 |  |  |  |  | 11/05/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1629292 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Run on Battery | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility had an automated impella controller (aic) in a box for a year. the aic was removed from the box and found to have an out of the box battery failure. no patient use and no harm or injury. the ifu can be followed and a backup console used for patient service. | [5 empty] | The investigation into the reported "aic - battery failure (red alarm)" has been completed since the original report was filed. product was returned for investigation. after console was received, the issue was reproduced. inspection of battery packs revealed both batteries fully depleted (led indicators were blank), and battery test diagnostics showed that batteries could no longer be changed due to internal lock-up caused by cell voltages below thresholds. during testing, batteries were temporarily replaced with known-good ones, and the pbm was able to charge them. per the ifu, the console should be stored with its power cord plugged into wall ac outlet, which was not the case for automated impella controller (aic). during data analysis, the console logs from the reported day of event are consistent with complaint and show battery failure alarm upon boot-up. prior boot up was eleven months before that, and there were no battery alarms.
32955 | 1220648-2023-04797 | 18070283 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18070283&pc=OZD | 10/06/2023 | 04/11/2024 |  | 11/04/2023 |  |  | 10/06/2023 | 11/04/2023 |  |  | 07/31/2025 |  | 11/04/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024390790 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella cp support experienced a tear on the insertion site that occurred during the sheath exchange process. a large hematoma and drop in hemoglobin resulted which required the patient to receive several units of blood. patient expired two days after following family decision for comfort measures. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. it was reported that vessel perforation/dissection was noted. the root cause of the vascular damage - peripheral vessels is most likely due to use as a tear was created during the sheath exchange. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32956 | 1220648-2023-04758 | 18065524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18065524&pc=OZD | 09/01/2022 | 11/03/2023 |  | 11/03/2023 |  |  | 10/05/2023 | 11/03/2023 |  |  |  |  | 11/03/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1126092 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 10/05/2023 |  |  |  |  | The user facility reported a 45-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support the team had an problem with the automated impella controller (aic) left ventricle signal. once the aic was swapped for another aic, the left ventricle signal was resolved. | [5 empty] | The investigation into the automated impella controller (aic) system error has been completed. data logs confirm the transfer of the aic equipment. there were persistent cyclic redundancy check errors. this points to errors while writing eeprom. the root cause is not determined since the product was not returned.
32957 | 1220648-2023-04759 | 18065608 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18065608&pc=OZD | 10/04/2023 | 03/04/2024 |  | 11/03/2023 |  |  | 10/04/2023 | 11/03/2023 |  |  | 06/30/2025 |  | 11/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024364826 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 62-year-old female patient had been implanted with an impella 5.5 for mechanical circulatory support. two days post impella removal, the patient had a right side cerebral vascular accident (cva). | [5 empty] | The investigation into the stroke/cva event has been completed. the device was not returned for investigation. as the clinical information was not provided, the true root cause of the stroke/cva could not be determined. | The impella device was not received from the customer as it was discarded and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32958 | 1220648-2023-04760 | 18067590 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18067590&pc=OZD | 10/04/2023 | 10/24/2025 |  | 11/03/2023 |  |  | 10/04/2023 | 11/03/2023 |  |  |  |  | 11/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1741225 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 10/04/2023 |  |  |  |  | The user facility reported that during a case preparation, the automated impella controller, (aic) had a battery communication failure alarm. troubleshooting did not resolve the problem. a replacement loaner was sent to the facility. there was no patient harm resulting from the noted issue. | [5 empty] | The investigation of automated impella controller (aic) - battery comm. failure (yellow alarm) has been completed. data analysis: during the data analysis the reported failure mode was confirmed. device analysis: the root cause for the battery 2 communication failure was not determined due to the inability to reproduce the issue. e2 was erroneously selected on the initial manufacturer device report number 1220648-2023-04760. g1 reporting contact email revised on manufacturer device report 1220648-2023-04760 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2023-04760. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | Corrections to the initial mfr report: d2 device has been updated. f6 and f8 dates have been removed. updated h3 to device evaluation anticipated.
32959 | 1220648-2023-04761 | 18067604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18067604&pc=OZD | 10/05/2023 | 06/05/2024 |  | 11/03/2023 |  |  | 10/05/2023 | 11/03/2023 |  |  | 05/31/2025 |  | 11/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024358555 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the 51-year-old male patient sustained ventricular fibrillation after the impella was advanced over guide wire into the left ventricle. return of spontaneous circulation after 2 defibrillations of 200 joules. the patient was reported to be stable following the event. | [5 empty] | The investigation into the vf/vt/af/at (ventricular fibrillation) issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary information was not provided, the root cause of the vt/vf was not determined. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32960 | 1220648-2023-04762 | 18067664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18067664&pc=OZD | 10/05/2023 | 01/22/2024 |  | 11/03/2023 |  |  | 10/05/2023 | 11/03/2023 |  |  |  |  | 11/03/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1557936 |  | 0042-0010 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Charge | Battery Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the batteries for the automated impella controller (aic) do not charge. when the power cable is plugged in, the led stays off and battery icon shows 0%. the plug icon is crossed out even when the controlled is connected to the main supply. there was no known harm or injury as a result of the noted issue. | [5 empty] | The device was returned and the investigation is ongoing. upon investigation closure, a supplemental report will be submitted. | The investigation into the aic - battery charging/other issues has been completed since the original report was filed. product was returned for evaluation. the cause of the issue was a defective power supply that was not providing power to the system or charging the batteries.
32961 | 1220648-2023-04763 | 18067690 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18067690&pc=OZD | 10/05/2023 | 02/14/2024 |  | 11/03/2023 |  |  | 10/05/2023 | 11/03/2023 |  |  |  |  | 11/03/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1109474 |  | 0042-0000-US |  | No |  |  | No |  | No | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella support experienced a "battery failure" alarm noted when console turned on. console was subsequently swapped with a back up controller for support. | [5 empty] | The investigation into the aic - battery failure (red alarm) has been completed. the console was not sent back to field services (fs) for investigation. according to the field representatives, the cause of the issue was use related. the battery transport switch was not enabled. the battery transport switch must be enabled for the batteries to be engaged and connected to the console's power battery manager (pbm). | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32962 | 1220648-2023-04764 | 18067737 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18067737&pc=OZD | 04/17/2023 | 01/23/2024 |  | 11/03/2023 |  |  | 10/05/2023 | 11/03/2023 |  |  | 11/30/2024 |  | 11/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024315881 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported that an 80-year-old male patient was implanted with an impella cp pump for mechanical circulatory support. during support, the patient pulled the bedsheets and tugged on the impella cp. the patient subsequently experienced blood loss and was hypotensive. physician therefore removed the impella cp. a femostop was placed to treat the bleed without issue. | [5 empty] | The investigation for the access site bleeding - major and the positioning issues has been completed. the impella device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the access site bleeding - major issue was most likely patient movement as the bleeding occurred after the patient accidentally pulled on the pump and the root cause of the positioning issues was most likely patient movement because the patient accidentally tugged on the pump. | The impella device was not received from the customer as it was discarded and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32963 | 1220648-2023-04765 | 18067743 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18067743&pc=OZD | 10/04/2023 | 02/21/2024 |  | 11/03/2023 |  |  | 10/04/2023 | 11/03/2023 |  |  | 07/31/2025 |  | 11/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024386920 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  |  |  | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient was coding in the er multiple times. the impella cp was therefore was placed in the patient for support but there was low flow and immediate suction. waveforms were unclear. the impella was removed and a second one placed. the patient expired in the procedure area with the second impella. | [5 empty] | The investigation into the low pump flow issue has been completed. the device was not returned for investigation. data confirmed the low flow when pump started and remained to the rest of the case that triggered auto suction response and suction alarms. the root cause of low flow was most likely low native cardiac output patient condition related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32964 | 1220648-2023-04766 | 18067891 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18067891&pc=OZD | 10/04/2023 | 03/26/2024 |  | 11/03/2023 |  |  | 10/04/2023 | 11/03/2023 |  |  | 08/31/2025 |  | 11/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024382321 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Additional information was reviewed pertaining to the outcome of the patient. following the perforation repair, the patient was converted back to ecmo support from cardiopulmonary bypass. despite the successful repair, it was noted that flows were not optimal and the maintenance of the patient's volume was difficult. there were no noted bleeds to account for the volume loss and there was minimal native heart function in-spite of utilizing multiple stimulation techniques. due to the ongoing poor condition of the patient, the decision was made to transfer the patient comfort care. the patient passed away within three hours of admission to the unit. a definitive relationship between the perforation and patient's passing was not established. | The complainant reported that impella 5.5 was implanted and it appeared to be in a good position under trans-esophogeal echocardiogram (tee). however, imaging also showed that the impella perforated the left ventricle (lv) during distal graft anastomosis and manual manipulation of the heart. the surgeon had to intervene and successfully repaired the ventricle. |  |  |  |  | The investigation for the ventricle perforation has been completed. the device was not returned for evaluation. however, clinical information provided was sufficient to determine the root cause. per the clinical account, the perforation occurred during graft anastomosis and manual manipulation of the heart. the root cause of the ventricle perforation issue was use related to improper technique. | Additional information added to b2, b5, h1, and h6 to coincide with the report of the patient's passing. should additional relevant information become available, a supplemental report will be submitted. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
32965 | 1220648-2023-04767 | 18067892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18067892&pc=OZD | 10/05/2023 | 01/12/2024 |  | 11/03/2023 |  |  | 10/05/2023 | 11/03/2023 |  |  | 07/31/2025 |  | 11/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024387778 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 71-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant had an impella cp support ongoing when on day two of the support the patient developed a hematoma. the patient was rolled and hematoma developed and urine color change occurred/pink tinge. the pump was repositioned and heparin were held off. systemic heparin paused and bicarb discussed. the patient remains on support despite the access site bleed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported access site bleeding ¿ major has been completed since the original report was filed. no product was returned. as per clinical large hematoma at impella site started after patient was rolled to change sheets and systemic heparin was paused with site redressed, angled better stabilizing bleeding. the cause of the access site bleeding issue was patient movement related with patient rolling over the bed leading to hematoma at insertion site per clinical review.
32966 | 1220648-2023-04768 | 18067958 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18067958&pc=OZD | 09/18/2023 | 09/20/2024 |  | 11/03/2023 |  |  | 10/05/2023 | 11/03/2023 |  |  | 05/31/2025 |  | 11/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2024350550 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/05/2023 |  |  |  |  | The user facility reported a 62-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was also being managed with ecmo. after withdrawal from ecmo, hemodynamic breakdown occurred and ecmo was reinserted. although the circuit was replaced, the hemolysis did not improve, so impella was determined to be the cause of the hemolysis. thereafter, the hemolysis was treated with haptoglobin, etc., but cardiac function did not improve and the patient expired. the doctor confirmed there was no relationship between the cause of death and hemolysis. additionally, there was a history of suction alarm occurrence. there was no problem with the position, but the volume was unknown. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the into the hemolysis and the low pump flow issue has been completed since the original report was filed. product was returned and investigated. per the provided clinical details and product investigation results, the root cause of the hemolysis issue was biomaterial wrapped around impeller. per the provided clinical details, data logs and product investigation results, the root cause of the low flow issue was biomaterial found on the impeller blades and purge gap. | The impella device was returned by the customer and an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6.
32967 | 1220648-2023-04769 | 18068091 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18068091&pc=OZD | 10/05/2023 | 04/23/2024 |  | 11/03/2023 |  |  | 10/05/2023 | 11/03/2023 |  |  | 05/31/2025 |  | 11/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024345526 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Improper Flow or Infusion | Air/Gas in Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/05/2023 |  |  |  |  | The user facility reported a 50-year-old male post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. shortly after insertion, the pump appeared to have large number of bubbles and no flow and the pump was immediately turned off. the introducer sheath had been aspirated and flushed with approximately 60cc of blood. this step repeated, however bubbles/air appeared again and not able to get any flow. process repeated a total of three times, all with same result. at this point, the md decided to switch sites and place cp instead. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the low pump flow has been completed. the impella device was not returned for evaluation. data logs were reviewed which showed low pump flow with several suction alarms during the case run. there were no reported events related to biomaterial interaction or improper positioning. the cause of the low pump flow issue was not determined since product was not returned and sufficient clinical information was not provided. d4-updated model number.
32968 | 1220648-2023-04770 | 18068238 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18068238&pc=OZD | 09/30/2023 | 05/29/2024 |  | 11/03/2023 |  |  | 10/05/2023 | 11/03/2023 |  |  | 11/30/2024 |  | 11/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2024289716 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 43 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/05/2023 |  |  |  |  | The user facility reported a 43-year-old male post cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the patient had gradual bleeding from the impella access site. the impella was removed and ecmo was inserted. pressure was applied and it is unknown if heparin was paused. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use referenced in the initial report is the impella cp with smart assist system, sections general patient care considerations, entering cardiac output, and potential adverse events (united states), which now includes a statement "cardiac or vascular injury (including ventricular perforation).¿ : in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32969 | 1220648-2023-04771 | 18068396 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18068396&pc=OZD | 10/05/2023 | 09/20/2024 |  | 11/03/2023 |  |  | 10/05/2023 | 11/03/2023 |  |  | 04/30/2025 |  | 11/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024349039 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Death | Yes | Yes | Pumping Stopped | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/05/2023 |  |  |  |  | The user facility reported male (unknown age and race) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. three hours after implantation, the pump stopped. the team attempted to restart the pump but they were not successful. the pump was explanted. | Additional information noted care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation for the reported pump stop has been completed. the impella cp was returned for evaluation. the pump was visually inspected, and biomaterial was ingested at outflow, the purge gap, and wrapped around the impeller. electrical resistance test was done, and all the motor phase values were within normal range. no other abnormalities were found. data logs were reviewed and found a high motor current pump stop with flow saturation at the end. as per clinical information, the pump stop was seen after 9 hours implant and restart was unsuccessful. the cause of the pump stop issue was due to biomaterial wrapped around the impeller and outflow with saturated flow observed before high motor current pump stop based on field investigation and log review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b1, b2, b5, h1, and h6.
32970 | 1220648-2023-04772 | 18068441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18068441&pc=OZD | 10/05/2023 | 01/18/2024 |  | 11/03/2023 |  |  | 10/05/2023 | 11/03/2023 |  |  | 06/30/2025 |  | 11/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2024366434 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/05/2023 |  |  |  |  | The user facility reported a patient (unknown age and gender) was implanted with an impella cp device for mechanical circulatory support. the patient had a thrombus that resulted in ischemia. heparin was turned off for six hours and this is thought to be the cause of the thrombus. the thrombus was removed via fogarty catheter and the patient's condition improved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported limb ischemia and thrombus has been completed since the original report was filed. no product was returned. as per clinical there was ischemia of the lower limb due to a blood clot at the time of impella removal. lower limb ischemia was removed by fogarty and then flow improved. the cause of the limb ischemia issue was patient condition related with thrombus flowing through lower limb and blocking the flow and unknown relation to impella with both legs being ischemic per clinical review.
32971 | 1220648-2023-04773 | 18068738 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18068738&pc=OZD | 10/07/2023 | 04/02/2024 |  | 11/03/2023 |  |  | 10/07/2023 | 11/03/2023 |  |  | 03/31/2025 |  | 11/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024326616 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/07/2023 |  |  |  |  | The user facility reported a 72-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. as support continued, a high purge pressure alarm was triggered. troubleshooting was attempted but the issue remained. the pump was explanted the following day. there was no report of patient harm. | [5 empty] | The investigation into the high purge pressure and the access site bleeding ¿ minor issues has been completed since the initial report was submitted. the product was not returned for investigation. according to clinical details, bleeding was reported from multiple access sites. high purge pressure related alarm was observed during impella use. the cause of the high purge pressure issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the access site bleeding issue was most likely patient condition related since the bleeding was reported form multiple access site. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32972 | 1220648-2023-04774 | 18069140 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18069140&pc=OZD | 10/06/2023 | 02/02/2024 |  | 11/03/2023 |  |  | 10/06/2023 | 11/03/2023 |  |  | 06/30/2025 |  | 11/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024374894 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported a 74 year old male patient was implanted with an impella cp pump for mechanical circulatory support. the patient had a femstop applied due to impella cp access site bleeding. hemolysis was noted in foley bag with no suction events on automated impella controller (aic). auto mode continued w/ ps map 120 mmhg. while in the icu, the hematoma was increasing. femstop was removed and manual pressure was applied, however the hematoma remained. the impella was therefore removed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major and the hemolysis issues has been completed. the product was not returned for review. the data logs did not record any events related to hemolysis. according to clinical details, the access site bleeding was observed due to a poor angle match and a loose tuohy-borst. the device was removed because of the bleeding after applying 45 minutes of manual pressure and using a femstop. hemolysis was detected in the foley bag. the cause of the access site bleeding issue was use related since the good practice was not followed by the doctor related to angle match and tuohy-borst, based on provided clinical information. the cause of the hemolysis issue was not determined since the product was not returned and its relation to impella remained unknown since the hemolysis was diagnosed with foley bag and no abnormalities were observed in data logs.
32973 | 1220648-2023-04775 | 18069163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18069163&pc=OZD | 10/05/2023 | 01/31/2024 |  | 11/03/2023 |  |  | 10/05/2023 | 11/03/2023 |  |  | 12/31/2024 |  | 11/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024297957 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 58-year-old patient in the eu received hemodynamic support from an impella cp pump due to acute myocardial infarction/cardiogenic shock. femoral access site bleeding has been reported. heparin anticoagulation was paused to treat the bleeding. in addition, blood units were transfused. the bleeding could be stopped. four days after the bleeding occurred, the patient was successfully weaned off the impella. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32974 | 1220648-2023-04776 | 18069208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18069208&pc=OZD | 10/10/2023 | 03/04/2024 |  | 11/03/2023 |  |  | 10/18/2023 | 11/03/2023 |  |  | 01/31/2025 |  | 11/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024312417 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/19/2023 |  |  |  |  | The user facility in japan reported a 73-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient had been on ecmo and impella for 1 week, but it was difficult to wean ecmo, and the plan was to switch to impella 5.5. before switching to impella 5.5 a ct scan showed evidence of cerebral infarction in several locations. it is unknown if this is related to the device. patient is stable. | [5 empty] | The investigation into the reported stroke/ cva has been completed since the original report was filed. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. | The impella device was not received from the customer in japan and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32975 | 1220648-2023-04777 | 18069343 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18069343&pc=OZD | 10/04/2023 | 11/03/2023 |  | 11/03/2023 |  |  | 10/04/2023 | 11/03/2023 |  |  | 06/30/2025 |  | 11/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024376100 |  | 0048-0014 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complaint in the eu reported that patient has been treated for acute myocardial infarction / cardiogenic shock and received hemodynamic support from an impella cp pump. continuous suction alarms have been reported. the user checked the impella position using fluoroscopy, and volume status of the patient has also been found acceptable. the impella device removed and replaced with a new impella cp pump. patient support continued without further issues. | [5 empty] | The investigation into the reported low pump flow has been completed. the pump was returned for investigation. data logs were not provided. during investigation, biomaterial was observed on the impellor. the pump ran as expected during functional test after removal of biomaterial, at the bench level. the root cause of the low flow issue was determined to be due to the biomaterial found on impellor. a risk assessment was performed, and no escalation is required at this time.
32976 | 1220648-2023-04778 | 18069403 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18069403&pc=OZD | 10/05/2023 | 05/29/2024 |  | 11/03/2023 |  |  | 10/05/2023 | 11/03/2023 |  |  | 03/31/2025 |  | 11/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024341855 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/05/2023 |  |  |  |  | The patient remained critically ill and the family made the decision to withdraw care. the patient expired. upon review by abiomed's medical safety officer, there is not enough information to associated the use of the device with the patient's outcome. | The user facility reported a 43-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was further reported that upon the patient arriving in icu and waking up from sedation, the patient started to move the leg and the impella became malpositioned in the ventricle. the impella was repositioned resolving the issue. |  |  |  |  | The investigation has been completed. according to the clinical information, on the first day of impella cp support, the patient awoke from sedation in the intensive care unit and moved their leg which caused the pump to become improperly positioned. shortly afterward, the pump was repositioned successfully. support continued until the patient was withdrawn from support. no product was returned for a visual inspection. data log analysis was not necessary for this complaint. the most likely cause of the positioning issues was patient movement. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 codes 2199 and 4629 were reported incorrectly on the previously submitted report. new code has been added to health effect - impact code. | The impella device has not been received from by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32977 | 1220648-2023-04779 | 18069444 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18069444&pc=OZD | 10/05/2023 | 04/02/2024 |  | 11/03/2023 |  |  | 10/05/2023 | 11/03/2023 |  |  | 07/31/2025 |  | 11/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024386895 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 75-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. after repositioning procedure, high purge pressure alarms were noted. upon inspection of the purge tubing and catheter, a twist and kink were noted on the impella catheter (cm marker 99-93). the decision to remove the impella was made and the patient was stable. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation into the high purge pressure issue has been completed since the initial report was submitted. the product was not returned for investigation. according to clinical details, bleeding was reported from multiple access sites. high purge pressure related alarm was observed during impella use. the cause of the high purge pressure issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the access site bleeding issue was most likely patient condition related since the bleeding was reported form multiple access site.
32978 | 1220648-2023-04780 | 18069493 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18069493&pc=OZD | 10/06/2023 | 11/03/2023 |  | 11/03/2023 |  |  | 10/06/2023 | 11/03/2023 |  |  | 07/31/2025 |  | 11/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024392360 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The us complainant reported a 51-year-old male patient was implanted with an impella cp pump for mechanical circulatory support. during support, the patient developed limb ischemia and loss of pulses on lower left extremities (lle). the impella cp was therefore removed and a new impella was inserted via axillary access for ongoing support. | [5 empty] | The investigation into the reported limb ischemia has been completed. the medical center did not return any impella products for benchtop analysis as it was discarded. the root cause of the reported event was determined to be most likely due to the patient¿s condition. patient had small and diseased vessels and the limb ischemia was resolved with impella removal. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32979 | 1220648-2023-04781 | 18069495 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18069495&pc=OZD | 10/07/2023 | 03/22/2024 |  | 11/03/2023 |  |  | 10/07/2023 | 11/03/2023 |  |  | 07/31/2025 |  | 11/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024380963 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/07/2023 |  |  |  |  | The user facility reported a 75-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella had suction alarms and the patient had low pulsatility. the impella was explanted and arterial duplex confirmed a focal thrombus. a thrombectomy was performed to address the issue. | [5 empty] | The investigation for the low pump flow and thrombus has been completed. the impella device was not returned for evaluation. data logs were reviewed indicating multiple suction alarms and low pump flow. clinical details provided that an arterial duplex confirmed a focal thrombus in the common femoral artery (cfa). therefore, the root cause of the thrombus was likely due to patient condition. the root cause of the low pump flow could not be determined with the device and sufficient clinical details. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/pos, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32980 | 1220648-2023-04785 | 18069520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18069520&pc=OZD | 06/14/2023 | 11/03/2023 |  | 11/03/2023 |  |  | 06/14/2023 | 11/03/2023 |  |  | 12/31/2024 |  | 11/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024325616 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/14/2023 |  |  |  |  | The user facility reported a 78-year-old female (race unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support during cardiac surgery. patient tolerated pci/impella placement well. due to hematoma, md made decision to remove impella and send patient to icu. doctor successfully managed complication with device removal and fem-stop. all drips were stopped. | [5 empty] | The investigation into a patient's access site bleeding has been completed. the cause of the access site bleeding issue was not determined since product was not returned and sufficient clinical information was not provided. instructions for use for this event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32981 | 1220648-2023-04784 | 18069521 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18069521&pc=OZD | 10/06/2023 | 02/20/2024 |  | 11/03/2023 |  |  | 10/06/2023 | 11/03/2023 |  |  | 04/30/2025 |  | 11/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024349392 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Activation, Positioning or Separation Problem | Improper Flow or Infusion | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported the 72-year-old male patient with impella cp for cp for mechanical circulatory support, experienced suction events and placement alarms. the device was found to be in a shallow position, the device was advanced and suction flow increased from 2.6 to 3.0 at p6. when attempts were made to increase the level of support to p8, there was a substantial suction event with stark drops in the diastolic numbers beyond -40mm/hg. patient responded well to the dobutamine, and so impella cp was removed and a clot was observed on the outlet of the impella. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported positioning issue, low pump flow and thrombus has been completed since the original report was filed. the product was not returned for investigation, and data logs were not provided for the last day of support. according to clinical details, the cause of the positioning issue was most likely use related. the cause of the low pump flow issue was most likely biomaterial, based on given clinical information. correction : the medical device problem code in section h6, was inadvertently left out in the initial report which has been updated in this report.
32982 | 1220648-2023-04786 | 18069526 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18069526&pc=OZD | 10/20/2023 | 02/06/2025 |  | 11/03/2023 |  |  | 10/20/2023 | 11/03/2023 |  |  | 08/31/2025 |  | 11/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024397839 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/20/2023 |  |  |  |  | The user facility reported a 52-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient was experiencing limb ischemia requiring fem fem bypass. decision was made to explant impella. the patient is stable. | [5 empty] | The investigation into the limb ischemia ¿ reversible has been completed. the device and data were not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The following have been reported incorrectly or missing from manufacturer device report 1220648-2023-04786. b1 product problem was inadvertently selected incorrectly, should have been adverse event. b2 outcomes attributed to adverse event was left blank. d4 model number. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact email revised in accordance with updated procedures. g1 reporting contact fax number missing.
32983 | 1220648-2023-04788 | 18069609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18069609&pc=OZD | 10/06/2023 | 07/08/2024 |  | 11/03/2023 |  |  | 10/06/2023 | 11/03/2023 |  |  | 07/31/2025 |  | 11/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024392251 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a 69 year old female patient undergoing impella cp pump implant, had the pump inserted through peel-away sheath without aspiration or flush. the pump attempted to flow 1.6 l then stopped. the pump was removed and there was a clot in the outlet. when the pump ran in saline at p4, multiple clots were noted. therefore, a different device was used to complete the procedure. there was no reported patient harm. | [5 empty] | The investigation for the reported thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombus could not be determined. the instructions for use referenced in the initial report is for the impella cp with smartassist for use during cardiogenic shock and high risk cpi in the following sections: section: using the automated impella controller with the impella catheter. section: using the automated impella controller with the impella rp with smartassist system catheter. added- h6 clinical code 4440 b1-corrected to adverse event. b2-selected life threatening. h1 changed to serious injury. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
32984 | 1220648-2023-04789 | 18069616 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18069616&pc=OZD | 10/05/2023 | 02/28/2024 |  | 11/03/2023 |  |  | 10/06/2023 | 11/03/2023 |  |  | 06/30/2025 |  | 11/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024359755 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 67-year-old male patient had been implanted with an impella 5.5 for mechanical circulatory support. when the 14x13 peel away sheath was removed a few days later from the right groin, the patient was grimacing in pain when the leg was handled, and pulses were lost on the following day. the patient therefore went in for emergent fasciotomy. the patient was also taken to the cath lab in the morning for intervention of the leg. however, the md was unable to open any blockages and now plans are being made to transfer patient to barnes for vascular surgical intervention. | [5 empty] | The investigation into the reported limb ischemia -reversible has been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
32985 | 1220648-2023-04791 | 18069701 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18069701&pc=OZD | 10/06/2023 | 08/19/2024 |  | 11/03/2023 |  |  | 10/06/2023 | 11/03/2023 |  |  |  |  | 11/03/2023 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1401075 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 35-year-old male patient with known cardiomyopathy was implanted with an impella cp for mechanical circulatory support. however, during support, a ¿battery temperature high¿ alarm was triggered. therefore, the automated impella controller (aic) was replaced for a new one and the issue resolved. there was no reported patient harm. | [5 empty] | The investigation for the console (aic) battery red alarm has been completed. data logs were reviewed which confirmed the battery temperature high alarm (event set #49). imc logs also showed that the battery temperature high alarm coincided with the ac power simultaneously being disconnected (event set #109), and the alarm resolved itself approximately 10 seconds later. the console was returned and evaluated which confirmed that the batteries were operational and non-defective. the reported complaint could not be reproduced. data and device analysis indicates that the alarm was most likely due to a miscommunication between the pbm and the 4-in-1, causing a false alarm and a misreporting of battery temperature. d2-corrected common device name. d4-corrected model number. | The impella device was returned by the customer and an investigation is ongoing. upon investigation closure, a supplemental mdr will be filed.
32986 | 1220648-2023-04792 | 18069740 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18069740&pc=OZD | 10/06/2023 | 02/20/2024 |  | 11/03/2023 |  |  | 10/06/2023 | 11/03/2023 |  |  | 06/30/2025 |  | 11/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024354388 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Migration | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 64-year-old male patient had been implanted with an impella cp for mechanical circulatory support. however, during support, the impella cp came out of position and was completely across the aortic valve. attempts were made to regain access to the left ventricle, but it was unsuccessful. therefore, the impella cp was explanted. | [5 empty] | The investigation into the reported positioning issue has been completed since the original report was filed. pump was not returned for investigation. the cause of the positioning issue resulting in the pump pulling out was use related due to improper technique. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32987 | 1220648-2023-04793 | 18069749 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18069749&pc=OZD | 10/06/2023 | 01/15/2024 |  | 11/03/2023 |  |  | 10/06/2023 | 11/03/2023 |  |  | 05/31/2025 |  | 11/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024358617 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/06/2023 |  |  |  |  | The user facility reported a 63-year-old male undergoing coronary artery bypass grafting x 2 was implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 was attempted to be placed via direct insertion to aorta. during insertion, the left ventricle was perforated with the impella and there was damage to the aortic valve (av). cardiopulmonary bypass (cpb) was reinitiated to repair the perforation and perform an aortotomy and av repair. the same impella was implanted via direct observation through the new av and support was initiated. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation for the reported heart valve damage and the ventricle perforation has been completed. the impella device was not returned for evaluation. clinical information provided was sufficient to determine the root cause. the same impella was implanted via direct observation through the new atrioventricular. support was initiated and cardiopulmonary bypass (cpb) was weaned successfully. cardiopulmonary bypass (cpb) was reinitiated to repair the perforation and perform an aortatomy and aortic valve replacement. the cause of both issues (left ventricle perforation and damage to the aortic valve issues) was use related due to improper technique.
32988 | 1220648-2023-04794 | 18069777 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18069777&pc=OZD | 10/06/2023 | 03/19/2024 |  | 11/03/2023 |  |  | 10/06/2023 | 11/03/2023 |  |  | 08/31/2024 |  | 11/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023271730 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fracture | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/06/2023 |  |  |  |  | The user facility reported a 63-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was being transferred into the airplane when the ambulance crew hit the white connector cable plug with the airplane door, breaking half of the white plug and leaving half of the plug stuck in aic with wires. the white cable was completely separated from the console stopping the impella cp. the impella was replaced with a new one and a different aic was used. | [5 empty] | The investigation for the reported pump issue has been completed. the pump was not returned for investigation. data logs show "impella stopped" alarm at the end of the case. according to clinical observations, the white patient cable connector was found broken in half after hitting the airplane door during transport. this resulted in half of the plug being stuck in the aic with wires, and the white cable completely separated from the console. the cause of the pump stop issue was an open electrical line based on provided clinical and data log review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
32989 | 1220648-2023-04796 | 18070000 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18070000&pc=OZD | 10/06/2023 | 05/07/2024 |  | 11/03/2023 |  |  | 10/06/2023 | 11/03/2023 |  |  | 11/30/2024 |  | 11/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024295439 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/06/2023 |  |  |  |  | The user facility reported a 81-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. there was an "impella stopped, retrograde flow" alarm present as the impella was being withdrawn from the ventricle. the impella cp was able to be explanted without any issues and there was no harm to the patient. | [5 empty] | The investigation into the removal issues has been completed since the original report was submitted. the device was not returned by the user facility for investigation. data log analysis confirmed that the device was weaned down to p-0 during removal which resulted in retrograde flow opposed to weaning down to p-2 which prevents retrograde flow. there was no problem found during the case. pump stop retrograde flow alarm triggers after the pump is set to p-0 by the user by design. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
32990 | 1220648-2023-04799 | 18070005 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18070005&pc=OZD | 10/06/2023 | 02/02/2024 |  | 11/03/2023 |  |  | 10/06/2023 | 11/03/2023 |  |  | 08/31/2025 |  | 11/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024392925 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Material Deformation | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/06/2023 |  |  |  |  | The user facility reported a 74-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella cp was originally placed without complication via 14x25 sheath. hematoma was noted at access site, however there is no known intervention for the hematoma. sheath was inspected under flouro and it was found that it was kinked in two sections. decision was made to remove the impella and sheath and place a 16f cook sheath then reinsert the impella. cook sheath placed without complication. upon insertion of the impella, pump would not advance past a certain point. flouro showed two kinks again in the cook sheath making it impossible to pass the impella through it. the physician decided to remove the impella cp and replace it with an impella 5.5. | [5 empty] | The investigation into the reported, low pump flow has been completed since the original report was filed. product was not returned for review. the root cause of the delivery issue was use related to misuse of incorrect cook 16f sheath. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32991 | 1220648-2023-04798 | 18070006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18070006&pc=OZD | 10/06/2023 | 02/01/2024 |  | 11/03/2023 |  |  | 10/06/2023 | 11/03/2023 |  |  | 07/31/2025 |  | 11/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024371232 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 27 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/06/2023 |  |  |  |  | The user facility reported a 27-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the patient on impella support experienced bleeding from the access site which lead the patient to have the site repaired surgically and receive several units of blood. | [5 empty] | The impella device was not returned> by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the access site bleeding could not be determined as there was not enough information provided in the clinical notes and no product was returned.
32992 | 1220648-2023-04787 | 18070149 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18070149&pc=OZD | 10/02/2023 | 10/30/2025 |  | 11/03/2023 |  |  | 10/06/2023 | 11/03/2023 |  |  |  |  | 11/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1383039 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | No Display/Image | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 10/06/2023 |  |  |  |  | The user facility reported a 67-year-old male in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. during the priming step, the automated impella controller (aic) blacked out. priming appears to have continued during the blackout. a reboot was attempted but a controller error alarm appeared on the screen. the aic was replaced. | The patient developed multiple organ failure and expired while on support. it is unknown if the black out of the automated impella controller (aic) contributed to the patient death, thus there is not enough evidence to exclude aic as an associated factor in the patient's outcome. |  |  |  |  | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-04787. b1 adverse event and product problem. b2 death. date of death is unknown. b5 patient outcome and death rationale statement. d2 common device name. d4 model number. f6/f8-should have been left blank. h1 type of reportable event: death. h6 health effect - impact code: death. | The investigation of automated impella controller (aic) - power on issues/blank screen has been completed. data analysis: the logs from the console confirm that the failure mode of unexpected controller shutdown occurred. device analysis: this failure was unable to be replicated during investigation across multiple pump connection attempts with power cycles in between. there were no observable anomalies with console behavior during testing. root cause: the root cause for the unexpected controller shutdown could not be determined due to the failure not being reproduced during investigation. there was no issue related power on issues/blank screen found during the case. d2 common device name revised on manufacturer device report 1220648-2023-04787 and 1220648-2023-04787-1 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2023-04787 in accordance with updated procedures.
32993 | 1220648-2023-04790 | 18070151 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18070151&pc=OZD | 10/04/2023 | 01/15/2024 |  | 11/03/2023 |  |  | 10/04/2023 | 11/03/2023 |  |  | 06/30/2025 |  | 11/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024374104 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/04/2023 |  |  |  |  | The user facility reported a 66-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support post hip surgical procedure. the impella cp pump had a high purge pressure problem. no harm to the patient was reported. it is unknown at this time if the device was explanted or replaced. | [5 empty] | The investigation for the reported high purge pressure has been completed. the impella was not returned for evaluation. however, clinical details states that purge pressures rose directly after a purge bag change. hihg pressure alarm was then triggered at the time a motor current was spiking. the root cause of the high purge pressure is most likely due to ingested biomaterial. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
32994 | 1220648-2023-04795 | 18070185 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18070185&pc=OZD | 10/06/2023 | 10/10/2024 |  | 11/03/2023 |  |  | 10/06/2023 | 11/03/2023 |  |  | 08/08/2025 |  | 11/03/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024392353 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/06/2023 |  |  |  |  | The user facility reported a 50-year-old male noted for electrophysiology ablation procedure was implanted with an impella cp device for mechanical circulatory support. during the procedure patient had signs of hemolysis. in addition, the patient required surgical repair of the femoral artery due to extravasation. the patient was given five units of packed red blood cells. issue was resolved successfully with surgical repair. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the hemolysis and the vascular damage has been completed. pump was returned in damaged condition (catheter was cut off). visual inspection found no issues on the pump. unable to conduct hemolysis test as the pump was damaged. data logs showed pump was started and run without issue for entire case. based on clinical description, the root cause of vascular damage - peripheral vessel was most likely due to patient condition. the root cause of hemolysis issue was not determined since the pump was returned in damaged state and could not be investigated for hemolysis test and insufficient clinical details were provided. the instructions for use for the impella cp with smartassist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ added-h6 component code 925, d9 device return date corrections to the initial mfr report: added- d6a/d6b. f6/f8 should have been left blank.
32995 | 1220648-2023-04731 | 18055781 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18055781&pc=OZD | 10/03/2023 | 11/02/2023 |  | 11/02/2023 |  |  | 10/03/2023 | 11/02/2023 |  |  | 02/28/2023 |  | 11/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022054991 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a 64-year old patient on impella 5.5 support experienced a fluid leak at the purge filter but the hemodynamic values were stable. there was no harm to the patient as a result of this incident. | [5 empty] | The investigation into the reported purge leak has been completed. the medical center did not return any impella products for benchtop analysis. the cause of the purge leak was unable to be determined since the product was not returned and hence, cannot be evaluated. also, there was not enough clinical information provided. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir."
32996 | 1220648-2023-04732 | 18055902 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18055902&pc=OZD | 10/14/2023 | 09/04/2024 |  | 11/02/2023 |  |  | 10/14/2023 | 11/02/2023 |  |  |  |  | 11/02/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1426097 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 10/14/2023 |  |  |  |  | The user facility reported a 55-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. on the 4th day of impella support the impella pump stopped and restarted three times in p-4 condition and the impella was removed. staff in japan additionally questioned the automated impella controller (aic) causing or contributing to the pump stoppage, and so the ifu was followed and aic was removed and replaced. there was no patient harm due to the issue. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | Engineers have determined that manufacturer device report 1220648-2023-04732 is not related to the automated impella controller (aic), it is pump related only. a regulatory report is no longer necessary. please see manufacturer device report 1220648-2023-04907 for more information regarding the pump.
32997 | 1220648-2023-04733 | 18056037 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18056037&pc=OZD | 10/07/2023 | 01/18/2024 |  | 11/02/2023 |  |  | 10/07/2023 | 11/02/2023 |  |  | 08/31/2025 |  | 11/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024386772 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/07/2023 |  |  |  |  | The user facility reported a 63-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, the patient's heparin drip was turned off due to a golf ball sized hematoma at the impella access site. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the cause of the access site bleeding issue was not determined since product was not returned and due to insufficient clinical information.
32998 | 1220648-2023-04734 | 18056054 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18056054&pc=OZD | 10/07/2023 | 06/04/2024 |  | 11/02/2023 |  |  | 10/07/2023 | 11/02/2023 |  |  | 07/31/2025 |  | 11/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024389332 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/07/2023 |  |  |  |  | The user facility reported a 66-year-old female was implanted with an impella cp as a short-term bridge to impella 5.5 and left ventricular assist device transplant (lvad). during support, the patient had a hematoma at the impella access site. the hematoma was managed with manual pressure and heparin was paused and vasoconstrictor medication drip was turned off. the patient remained on cp support for six days and was escalated to an impella 5.5 and lvad. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use referenced in the initial report is for the impella cp with smart assist system in the following sections: general patient care considerations, entering cardiac output, and potential adverse events (united states), which now includes the statement "cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
32999 | 1220648-2023-04735 | 18057725 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18057725&pc=OZD | 10/03/2023 | 12/05/2025 |  | 11/02/2023 |  |  | 10/03/2023 | 11/02/2023 |  |  |  |  | 11/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1297872 |  | 0042-0000-UK |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Female |  |  | White | Malfunction | No | Yes | Failure to Power Up | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The user facility reported that the impella cp was connected to the automated impella controller (aic) but the screen would not advance to start. therefore, a second aic was used, which worked. there was no known harm or injury to the patient. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | Corrections to the initial mfr report: d2 common device name updated. g1 mdr reporting contact email.
33000 | 1220648-2023-04736 | 18057794 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18057794&pc=OZD | 10/15/2023 | 02/21/2024 |  | 11/02/2023 |  |  | 10/15/2023 | 11/02/2023 |  |  | 08/31/2025 |  | 11/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024385405 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/15/2023 |  |  |  |  | The user facility reported a 65-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was suspected to have experienced hemolysis that began overnight. and the urine color was a dark red. the medical staff decided to wean off the patient form impella and ultimately the pump had to be explanted. | [5 empty] | The investigation into the reported hemolysis has been completed since the original report was filed. no product was returned for investigation. data logs do not show evidence of positioning issues. the root cause of the hemolysis was not determined since product was not returned and there is a lack of clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33001 | 1220648-2023-04737 | 18057873 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18057873&pc=OZD | 10/03/2023 | 08/20/2024 |  | 11/02/2023 |  |  | 10/03/2023 | 11/02/2023 |  |  |  |  | 11/02/2023 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1102534 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 was plugged into the automated impella controller (aic) and immediately received an ¿impella defective¿ alarm. the impella and the aic were replaced. there was no reported patient harm. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation for the console (aic) yellow alarm has been completed. the data logs were reviewed which confirmed the reported event. there were multiple instances of impella defective alarms (event set #3), preceded by errors when the console attempted to read the calibration data (first section) from the eeprom of the pump. specifically, the issues with reading data from the pump¿s eeprom were caused by crc errors when comparing what was read versus what was expected. the console was returned and upon evaluation, reported complaint could not be reproduced despite testing the functionality of the epm reading the pump eeprom during initial connection of a known-good pump (multiple times). the console was inspected internally, and there were no misaligned cables. the root cause was isolated to the impellatronic, because the impellatronic is responsible for reading the eeprom data and sending it to the 4-in-1 (epc). d2-corrected common device name d4-corrected model number.
33002 | 1220648-2023-04738 | 18057914 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18057914&pc=OZD | 10/03/2023 | 08/20/2025 |  | 11/02/2023 |  |  | 10/03/2023 | 11/02/2023 |  |  | 08/31/2025 |  | 11/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024390869 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 was plugged into the automated impella controller (aic) and immediately received an ¿impella defective¿ alarm. the impella and the aic were replaced. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | Corrections to the initial mfr report: g1 mdr reporting contact email. h6 component code updated to 925.
33003 | 1220648-2023-04739 | 18058102 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18058102&pc=OZD | 10/03/2023 | 09/20/2024 |  | 11/02/2023 |  |  | 10/03/2023 | 11/02/2023 |  |  | 04/30/2025 |  | 11/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP JP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024366067 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that the 60-year-old male patient was on ecpella support for 77.53 hours before the patient expired. patient was on support due to acute exacerbation of dilated cardiomyopathy. a subsequent ct scan revealed cerebral hemorrhage. time of the cerebral hemorrhage is unknown and the relationship to the impella is unknown. patient expired from cerebral hemorrhage. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation for the stroke/cva has been completed. the impella device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the stroke/cva is most likely due to the patient¿s condition. heparin was used in the purge, but it is unknown if it is the root cause. the failure mode will be monitored and trended. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6.
33004 | 1220648-2023-04740 | 18058166 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18058166&pc=OZD | 10/03/2023 | 08/20/2025 |  | 11/02/2023 |  |  | 10/03/2023 | 11/02/2023 |  |  |  |  | 11/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1131455 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the automated impella controller (aic) did not prime successfully. a different aic was used and case completed successfully. since that time, the field team had been unable to replicate the issue and have primed and ran a demo pump on this controller successfully. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | Corrections to the initial mfr report: d4 common device name updated. g1 mdr reporting contact email.
33005 | 1220648-2023-04741 | 18058248 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18058248&pc=OZD | 10/03/2023 | 09/26/2025 |  | 11/02/2023 |  |  | 10/03/2023 | 11/02/2023 |  |  | 07/31/2025 |  | 11/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024375030 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported the 53-year-old male patient had bleeding at the impella 5.5 access site and heparin was halted. the pump remained on for support despite the bleed, already for 71days. | Additional information indicated that, during patient support, the pump experienced placement signal not reliable alarm. an echocardiogram was performed to verify correct position. support was continued, and the patient was successfully weaned off the device. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding - major issue was most likely patient condition related since the bleeding was extracorporeal membrane oxygenation (ecmo) site. h6 component code added. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ g1 reporting contact fax number missing on manufacturer device report 1220648-2023-04741. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation for the optical signal issue has been completed. the investigation for the optical signal issue has been completed. no product or data logs were returned for investigation. the root cause of the optical signal issue was not determined since product and data logs were not returned for investigation. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 updated with product problem. b.5. updated with additional information. g.1. updated with correct mdr reporting contact email. h.6. medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2023-04741.
33006 | 1220648-2023-04742 | 18058609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18058609&pc=PYX | 10/09/2023 | 02/08/2024 |  | 11/02/2023 |  |  | 10/09/2023 | 11/02/2023 |  |  | 07/24/2023 |  | 11/02/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024366057 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/09/2023 |  |  |  |  | The user facility reported a 55-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. during support, the pump experienced low flow issues from what appeared to be a clot ingestion in the cannula. there was no harm to the patient because of this incident. the device was explanted and another non-impella device implanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow issue has been completed. the device was not returned for investigation. based on the clinical information provided, the case images show a clot in the inflow and outflow cages upon explant. data logs show purge pressure spike and suction alarms followed by low pump flow at p-9. the root cause of the low pump flow issue was determined to be biomaterial.
33007 | 1220648-2023-04745 | 18061164 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18061164&pc=OZD | 10/05/2023 | 04/16/2024 |  | 11/02/2023 |  |  | 10/05/2023 | 11/02/2023 |  |  | 06/30/2025 |  | 11/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024359572 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 61-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that a patient on impella support experienced bleeding at the access site and was required to have a femstop to control it. five days after that event, the patient was found to have a hematoma at the access site and patient had no flow in the distal leg. the medical team had to surgically remove the hematoma and administer blood products. the sheath was replaced in attempt and surgery was performed however they were unable to restore flow in the distal leg. decision was made to explant the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed and the limb ischemia has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was most likely due to anticoagulation management since the act was 291 at the time of the bleed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
33008 | 1220648-2023-04747 | 18061506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18061506&pc=OZD | 02/24/2023 | 11/02/2023 |  | 11/02/2023 |  |  | 02/24/2023 | 11/02/2023 |  |  |  |  | 11/02/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1402368 |  | 0042-0010-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 02/24/2023 |  |  |  |  | The user facility reported an 80-year-old female with severe multi-vessel disease was implanted with an impella cp device for mechanical circulatory support. approximately 30 minutes after securing impella in place and placing a leg immobilizer, a blockage purge alarm occurred on the automated impella controller (aic). the blockage purge alarm self-resolved and the aic displayed a high-pressure alarm. no kinks were found in purge tubing or purge side arm. troubleshooting was performed, the cassette was changed but alarm remained. automated impella controller (aic) to aic transfer was performed and the purge pressure normalized. there was no patient harm related to the aic to aic transfer. | [5 empty] | The investigation into the purge pressure high alarm on the automated impella controller has been completed. console logs from the reported day of event are consistent with complaint by the user. the root cause of high purge pressure could not be not determined. the issue could not be reproduced.
33009 | 1220648-2023-04744 | 18061508 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18061508&pc=OZD | 11/01/2023 | 09/09/2024 |  | 11/02/2023 |  |  | 11/01/2023 | 11/02/2023 |  |  |  |  | 11/02/2023 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1322914 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 11/01/2023 |  |  |  |  | The us complainant had an impella console (aic) sent to the medical facility. the console was a loaner. the loaner was removed from the packaging and found to have a crack/damage. the console was not used on a patient. no harm or injury. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the console (aic) damage has been completed. data logs were reviewed and no relevant information was found related to the reported complaint. the aic was sent back for evaluation. physical inspection confirmed that the console¿s front housing was damaged near the location of the purge door release button. the root cause of the housing damage was not determined. added- d9 device return date d2-corrected common device name. d4-corrected model number. e2-changed to no. f6/f8 should have been left blank in the initial report. g3-corrected pma number.
33010 | 1220648-2023-04746 | 18061585 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18061585&pc=OZD | 10/05/2023 | 05/02/2024 |  | 11/02/2023 |  |  | 10/05/2023 | 11/02/2023 |  |  | 04/30/2025 |  | 11/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP JP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024339522 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility in japan reported a 82-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient¿s blood tests and lab values showed positive for hemolysis. the impella cp was removed. | [5 empty] | The investigation into the hemolysis issue has been completed since the original report was submitted. the device was not returned by the user facility for investigation. the cause of the hemolysis could not be determined because no product or logs were returned and limited clinical information was given. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ to conform to updated procedures, sections d6 & h10 were revised with updated information. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions including catheter position, pre-existing medical conditions, and small left ventricular volumes may also play a role in patient susceptibility to hemolysis.¿
33011 | 1220648-2023-04748 | 18062206 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18062206&pc=OZD | 10/03/2023 | 01/31/2024 |  | 11/02/2023 |  |  | 10/03/2023 | 11/02/2023 |  |  | 04/30/2024 |  | 11/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218537 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/03/2023 |  |  |  |  | The user facility reported a 68-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. during support, the impella cp pump stopped. attempts were made to restart the pump unsuccessfully. the impella cp was explanted and a new cp implanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop issue has been completed. the product and data logs were not returned for review. the root cause of pump stop was unable to be determined as limited clinical details were provided and no product was returned for review. the pump stop occurred after 17 days of run time, which is well beyond the design life of eight days.
33012 | 1220648-2023-04749 | 18062208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18062208&pc=OZD | 07/20/2023 | 02/09/2024 |  | 11/02/2023 |  |  |  | 11/02/2023 |  |  | 12/31/2024 |  | 11/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024309181 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/03/2023 |  |  |  |  | The user facility reported a 46-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the impella metrics were reading suction and ¿impella in aorta¿. the impella was on p2 with zero flow reading despite being in good position. the impella was explanted. | [5 empty] | The investigation into the reported, low pump flow has been completed since the original report was filed. data logs showed, pump was stopped manually & disconnected from the console. visual inspection found biomaterial inside the pump housing & wrapped around the impeller. no other issues were found on the rest of the pump. the root cause of low flow was fiber and biomaterial ingestion. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
33013 | 1220648-2023-04750 | 18062305 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18062305&pc=OZD | 09/26/2023 | 01/26/2024 |  | 11/02/2023 |  |  |  | 11/02/2023 |  |  |  |  | 11/02/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1420348 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/04/2023 |  |  |  |  | The user facility reported their automated impella controller (aic) was displaying battery failure and battery communication failure error messages. troubleshooting was attempted by abiomed field service representative but was unsuccessful. the aic will be returned for evaluation. | [5 empty] | The investigation into the reported battery failure alarm has been completed. the console and the data logs were returned for evaluation. the logs were reviewed and confirmed a battery failure alarm (transport connection blown/ missing due to low pack voltage. they also revealed that there was an approximately one year period during which the console was powered off, which matches the reported complaint. this one year period is most likely when the console¿s batteries were allowed to fully deplete. the console was evaluated using the batttest utility which revealed that both batteries were fully depleted and were therefore unable to be charged or communicate with the pbm¿s charging circuitry. replacing the original depleted batteries with known working batteries confirmed that the pbm was working correctly, as it was able to communicate with the new batteries. the console remained powered on by the new batteries, with no battery failure alarm triggered, which confirmed that the original issue was solely due to the depleted batteries. once an aic¿s batteries are fully depleted, they become unable to communicate with the charging circuitry and therefore are unable to be recharged, triggering the battery failure alarm. therefore the root cause of the reported console battery alarm was due to depleted batteries. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
33014 | 1220648-2023-04751 | 18062322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18062322&pc=OZD | 10/03/2023 | 05/01/2024 |  | 11/02/2023 |  |  |  | 11/02/2023 |  |  | 06/30/2025 |  | 11/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024360927 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a 62-year old male patient on impella cp support had the pump caught on chordae tendineae and the position could not be changed. when the assistance level was raised, a suction alarm sounded and it was not possible to manage it well, so the impella cp was removed and a new pump was inserted. there was no harm to the patient as a result of this incident. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ | The investigation into the positioning issues has been completed since the original report was submitted. the device was not returned by the user facility for investigation. according to the clinical, during impella cp insertion the pump became caught in the chordae tendineae. multiple attempts were made to reposition the pump without success. the most likely cause of the positioning issues was use related. to conform with updated procedures, section d6 has been revised with updated information.
33015 | 1220648-2023-04752 | 18062581 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18062581&pc=OZD | 10/03/2023 | 11/18/2024 |  | 11/02/2023 |  |  | 10/03/2023 | 11/02/2023 |  |  | 07/31/2025 |  | 11/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024379137 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the medical team was having difficulty delivering an impella pump due to the patient's anatomy. during the process the canula kinked and a second pump had to be used, which was delivered successfully. there was no harm to the patient as a result of this incident. | [5 empty] | The impella device was returned by the customer and an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. | The investigation for the product damage has been completed. the pump was returned for investigation. cannula kink and catheter kink were noted near the motor housing of the impella. according to the clinical report, the patient had extremely tortuous vessels due to which the pump delivery was extremely difficult. the cannula and the catheter of the pump kinked in that process. the root cause of the delivery issue that led to the product damage-cannula kink issue was a kinked cannula due to patients tortuous vessels. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock states the following: section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: cautions ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ corrections to the initial mfr report: added d6a/d6b
33016 | 1220648-2023-04753 | 18062588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18062588&pc=OZD | 10/03/2023 | 02/02/2024 |  | 11/02/2023 |  |  | 10/03/2023 | 11/02/2023 |  |  | 04/30/2025 |  | 11/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024351288 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 59-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient¿s labs were hemolyzing. patient given 500ml of albumin to treat the condition and the hemolysis was reported to be resolving. support was maintained with the implanted pump. | [5 empty] | The investigation for the reported hemolysis issue has been completed. data logs were reviewed which revealed the pump was in improper position then back to proper position corresponding with pump reposition per clinical description. product was not returned for review. based on clinical description and data analysis, the root cause of hemolysis was most likely due to pump was not in the proper position. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33017 | 1220648-2023-04754 | 18062597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18062597&pc=OZD | 10/03/2023 | 04/09/2024 |  | 11/02/2023 |  |  | 10/03/2023 | 11/02/2023 |  |  | 07/31/2025 |  | 11/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024389496 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 72-year-old male patient was implanted with an impella cp pump for mechanical circulatory support. following insertion of the impella cp for patient support, purge pressure alarms occurred and flows dropped after the initial uploading. the impella was subsequently repositioned, however, due to the noted issue, the decision was made to replace the impella pump. there was no patient harm resulting from the reported event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the high purge pressure issue has been completed since the initial report was submitted. the device was not returned for investigation. the data logs indicate that the pump flow was consistently reported as low from the beginning of the case, with no noted motor current spikes or abnormalities in the placement signal. the doctor encountered the same low flow issue with the second pump. the cause of the low pump flow issue was most likely patient condition related since the second placed pump also experienced the same issue.
33018 | 1220648-2023-04755 | 18062617 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18062617&pc=OZD | 10/03/2023 | 04/12/2024 |  | 11/02/2023 |  |  | 10/03/2023 | 11/02/2023 |  |  | 05/31/2025 |  | 11/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024355458 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Air/Gas in Device | [6 empty] | Embolism/Embolus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported that 36-year old male patient was implanted with an impella 5.5 for mechanical circulatory support. during impella support, it was noted that "bubbles" were observed in the aorta on echo. additionally, there was a purge system open alarm. all connections were correct, removing and re-inserting purge cassette didn't resolve alarm, and purge wasn't exiting the purge line at the yellow connection during de-air. new purge cassette and bag was placed and all values are now within correct parameters. | [5 empty] | The investigation into the purge leak ¿ pump, the embolism, and the priming issue issues has been completed since the original report was submitted. the device was not returned for investigation. per clinical details provided, the cause of the purge leak and the priming (air in purge) issues was not determined since product was not returned and limited clinical information was provided. the purge leak condition was resolved after successfully completing the cassette change procedure. the cause of the embolism issue was not determined since the product was not returned, and limited clinical information was provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
33019 | 1220648-2023-04756 | 18062622 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18062622&pc=PYX | 10/03/2023 | 05/28/2024 |  | 11/02/2023 |  |  | 10/03/2023 | 11/02/2023 |  |  | 04/30/2025 |  | 11/02/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024337987 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Crack | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/03/2023 |  |  |  |  | There is no association between the impella use and the patient outcome. patient's peri-procedural condition was poor. | The user facility reported a 64-year-old female in acute myocardial infarction/cardiogenic shock post cardiac surgery was implanted with an impella rp device for mechanical circulatory support. the patient was also being supported by an impella cp. upon removing the sheath introducer from the impella rp, the sheath cracked resulting in a constant ooze from the sheath during the insertion process. blood products were given to the patient. the sheath was removed and peeled away immediately after the impella rp was placed into the peripheral artery successfully. |  |  |  |  | The investigation into the access site bleeding ¿ major and the introducer damage ¿ mechanical issues has been completed since the original report was submitted. the device was not returned for investigation. according to the clinical, during the impella rp insertion process the introducer sheath was cracked while being removed resulting in constant oozing from the access site throughout the insertion process. a balloon tipped catheter and blood products were used to resolve the issue. the cause of the mechanical introducer damage was not determined because no product was returned, and not enough clinical information was given. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33020 | 1220648-2023-04757 | 18062628 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18062628&pc=OZD | 10/03/2023 | 01/25/2024 |  | 11/02/2023 |  |  | 10/03/2023 | 11/02/2023 |  |  | 07/31/2025 |  | 11/02/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024377888 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/03/2023 |  |  |  |  | The user facility reported a 62-year-old male in cardiogenic shock was to be emergently implanted with an impella cp device for mechanical circulatory support in addition to va ecmo. upon removal of 14f peel away sheath, impella migrated into aorta. under fluoroscopy, the pigtail was noted not to be across the aortic valve. use of an .035 wire was advanced through the re-access port to assist in advancement of device back across the valve. after several failed attempts, decision was to replace impella catheter with new device. | [5 empty] | The investigation into the positioning issues has been completed since the original report was filed. the product was reviewed, and the pigtail was slightly deformed and there was a kink in the cannula. the data logs were reviewed, and it was confirmed that the pump was out of position. there were no anatomical abnormalities with the patient that could have caused the pump to migrate into the aorta. the root cause of the positioning issues is most likely due to use. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
33021 | 1220648-2023-04704 | 18047039 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18047039&pc=OZD | 10/10/2023 | 04/22/2024 |  | 11/01/2023 |  |  | 10/10/2023 | 11/01/2023 |  |  | 07/31/2025 |  | 11/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024371455 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/10/2023 |  |  |  |  | The user facility in the eu reported a male with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 showed flow issues (purge pressure rose to 1050mmhg and purge flow rate dropped to 1.5ml/hr), there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow issue has been completed since the original report was submitted. the device was not returned for investigation. based on the log review, the cause of the high purge pressure issue was most likely biomaterial.
33022 | 1220648-2023-04705 | 18047074 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18047074&pc=OZD | 06/17/2023 | 11/01/2023 |  | 11/01/2023 |  |  | 06/17/2023 | 11/01/2023 |  |  |  |  | 11/01/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1185326 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The us complainant had an automated impella controller (aic) with a red battery alarm. the aic was therefore replaced and support proceeded without harm or injury from delay. | [5 empty] | The investigation into the reported battery cell failure has been completed. the automated impella controller (aic) logs confirmed the reported failure. the aic was received for evaluation. testing revealed battery 1 had no communication and lockout as opposed to battery 2 that had information indicating well charged and balanced cells. upon visual inspection of the interior of the aic, the lcd screen on battery 1 was noted to be blank indicating cell lockout. the root cause of the aic issue was determined to be a defective battery. this report is being filed as part of a retrospective review of historical records.
33023 | 1220648-2023-04706 | 18047143 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18047143&pc=OZD | 06/20/2023 | 11/01/2023 |  | 11/01/2023 |  |  | 06/20/2023 | 11/01/2023 |  |  |  |  | 11/01/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1656816 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The us complainant had a automated impella controller (aic) returned after annual service and console was found to have damage. the aic had been placed on the cart and the back handle broke right off. the damage prompted the team to want the console looked to again. there was no reported patient harm. | [5 empty] | The investigation into the reported automated impella controller (aic) damage has been completed. the aic was received for evaluation. the reported event was confirmed. visual inspection revealed missing/broken gray handle in the posterior of the console. the cause of the aic issue was determined to be defective housing. this report is being filed as part of a retrospective review of historical records.
33024 | 1220648-2023-04707 | 18047184 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18047184&pc=OZD | 06/30/2021 | 11/01/2023 |  | 11/01/2023 |  |  | 06/30/2021 | 11/01/2023 |  |  | 10/20/2020 |  | 11/01/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 2022058029 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  | White | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 06/30/2021 |  |  |  |  | The user facility reported a 64-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the us complainant had an automated impella controller (aic) fail. the ifu was followed and a backup controller was used. no harm or injury. | [5 empty] | The console was returned for analysis. root cause is due to a malfunctioning pud board. the root cause of the malfunctioning pud board is unable to be determined. before returning to customer, field service remedied the issue.
33025 | 1220648-2023-04708 | 18047212 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18047212&pc=OZD | 10/02/2023 | 02/20/2024 |  | 11/01/2023 |  |  | 10/02/2023 | 11/01/2023 |  |  | 08/31/2025 |  | 11/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024386811 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Sensing Intermittently | Optical Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient was to be implanted with an impella cp prior to high risk percutaneous coronary intervention (hrpci) for support . during intervention, the left main (lm) required shockwave. after pulses of the shockwave, the optical sensor was knocked out. the physician finished the intervention, attempted to wean the patient from the impella but was unable to. therefore, the cp was replaced with a new cp in the same access sheath prior to sending the patient to the icu. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the optical signal issue has been completed. the device was not returned for investigation. per the clinical information provided, the cause of the optical signal issue was a damaged sensor due to interaction with shockwave device.
33026 | 1220648-2023-04709 | 18047393 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18047393&pc=OZD | 10/05/2023 | 01/25/2024 |  | 11/01/2023 |  |  | 10/05/2023 | 11/01/2023 |  |  | 07/31/2025 |  | 11/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024377551 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/05/2023 |  |  |  |  | The user facility reported a 68-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. during support, the patient had impella access site bleeding. the patient received fibrinogen and the nurse applied a new pressure dressing. heparin was put on hold due to bleeding. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as foll: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for benchtop investigation. according to the clinical information provided, patient was bleeding at access site and pressure dressing, fibrinogen administration resolved bleeding. the cause of the access site bleeding issue was not determined due to insufficient clinical information and since product was not returned.
33027 | 1220648-2023-04710 | 18047517 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18047517&pc=OZD | 10/11/2023 | 01/11/2024 |  | 11/01/2023 |  |  | 10/11/2023 | 11/01/2023 |  |  |  |  | 11/01/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1401275 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 10/11/2023 |  |  |  |  | The user facility reported that during the preparation the impella console (aic) had a broken purge disc. a replacement was sent to address the issue. no patient was involved. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the console was returned for investigation. upon examination of the console, it was confirmed that the purge disk retainer on this console was broken. part where retainer was fastened to the purge pressure sensor assembly was no longer intact. the root cause of this issue is a broken purge retainer. g.3 date device was returned to the manufacturer was incorrect on the initial manufacture device report. the correct date (16-oct-2023) was added in this report.
33028 | 1220648-2023-04711 | 18047863 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18047863&pc=OZD | 10/02/2023 | 09/26/2025 |  | 11/01/2023 |  |  | 10/02/2023 | 11/01/2023 |  |  | 08/31/2025 |  | 11/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024380597 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | Improper Flow or Infusion |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/02/2023 |  |  |  |  | The user facility reported a 44-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support, as the patient had admitted with stemi. the pump had lost optical signal and had purge system blocked alarms. the pump support was explanted on day 8 of the support and replaced by a new 5.5 pump. no harm or injury noted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation of high-pressure purge has been completed. the impella device was not returned for evaluation. the cause of the high purge pressure issue was most likely kinked purge line, based on data log review. d6b added the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-04711: d4 model number, f6/f8-should have been left blank. g1 reporting contact fax number missing. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation for the optical signal issue has been completed. data logs showed multiple placement signal spiked to 3000 mmhg that triggered alarms placement signal not reliable during first 23 hours into support then back to normal range. about 148 hours later, placement signal spiked again to 3000 mmhg with the same alarms. this event remained to the rest of the case. product was not returned for review. the cause of the optical signal issue was most likely damaged fiber line. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 updated with product problem. d3 country updated. g1 updated with correct mdr reporting contact email. h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2023-04711.
33029 | 1220648-2023-04696 | 18049877 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18049877&pc=OZD | 10/01/2023 | 02/20/2024 |  | 11/01/2023 |  |  | 10/01/2023 | 11/01/2023 |  |  | 07/31/2025 |  | 11/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024385765 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a 62-year-old male was implanted with an impella cp for mechanical circulatory support. after the impella was placed, flows decreased and impella was noted to be back into the aorta. the impella appeared kinked near the outlet, therefore the pump was removed and a new one was successfully placed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation into the product damage (cannula kink) issue has been completed. the device was not returned for investigation. per the clinical information provided, the root cause of the low pump flow was a cannula kink likely due to interaction with tavr.
33030 | 1220648-2023-04694 | 18049879 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18049879&pc=OZD | 10/03/2023 | 02/05/2024 |  | 11/01/2023 |  |  | 10/03/2023 | 11/01/2023 |  |  | 07/31/2025 |  | 11/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024388111 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/03/2023 |  |  |  |  | The user facility reported a 68-year-old male with heart transplant rejection was implanted with an impella cp device for mechanical circulatory support. the patient was also on ecmo support. during support, there were multiple suction alarms and multiple instances of p-level decreasing to p1. a couple of hours after these occurrences, high purge pressure alarms and a ¿purge system blocked¿ alarm occurred. the impella was exchanged for a new one. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported high purge pressure issue has been completed. the impella device and the data logs were returned for evaluation. the data logs were reviewed which showed the purge pressure gradually increased from 600 mmhg to 1000 mmhg within 30 minutes and remained to the rest of the case. the pump was visually inspected and found biomaterial inside the whole pump housing, purge gap & wrapped around the impeller. no other issues were found on the rest of the pump. the root cause of high purge pressure was due to biomaterial ingestion.
33031 | 1220648-2023-04695 | 18049882 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18049882&pc=OZD | 09/25/2023 | 09/26/2024 |  | 11/01/2023 |  |  | 10/01/2023 | 11/01/2023 |  |  | 04/30/2025 |  | 11/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024339797 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/01/2023 |  |  |  |  | The user facility in japan reported a 70-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient had a cerebral infarction and cerebral hemorrhage and expired four days later. the patient had been on impella cp support for 328 hours before expiring. the doctor stated that it is impossible to say with certainty that there is no connection to the relationship of the impella and cause of death. there were no product malfunctions that occurred with the impella. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the stroke/cva has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. based on the clinical information provided, the root cause of the stroke/cva event was most likely patient condition related and the relationship to impella is unknown. the failure mode will be monitored and trended. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-04695 was provided in sections b2, g3, and h1, and h6.
33032 | 1220648-2023-04697 | 18049887 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18049887&pc=OZD | 10/04/2023 | 02/29/2024 |  | 11/01/2023 |  |  | 10/05/2023 | 11/01/2023 |  |  | 03/31/2025 |  | 11/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024327079 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/05/2023 |  |  |  |  | The user facility reported a 42-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. patient had a 5.5 support that was placed via the axillary to support a patient in need of care escalation from the cp. the access site was bleeding and the team gave 2 units of blood. the site was also washed out. the 5.5 pump supported till removal and placement of the durable vad and ecmo. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported access site bleeding -major has been completed since the original report was filed. product was not returned for investigation. the cause of the access site bleeding issue was not determined since product was not returned and sufficient clinical information was not provided.
33033 | 1220648-2023-04712 | 18050276 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18050276&pc=OZD | 10/12/2023 | 05/07/2024 |  | 11/01/2023 |  |  | 10/12/2023 | 11/01/2023 |  |  | 06/30/2025 |  | 11/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024372862 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/12/2023 |  |  |  |  | The user facility reported a 57-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. upon initial insertion of the impella cp while attempting to position, the device jumped forward into the ventricle resulting in low impella position and low flow alarm. the physician pulled the device back and the alarm resolved. following the patient transfer to the icu, a suction event occurred. the device was once again repositioned, and the alarm resolved. there was no known harm or injury to patient. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the positioning issue was use related. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
33034 | 1220648-2023-04713 | 18051773 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18051773&pc=OZD | 10/31/2023 | 09/02/2024 |  | 11/01/2023 |  |  | 10/31/2023 | 11/01/2023 |  |  | 05/31/2025 |  | 11/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024369464 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/31/2023 |  |  |  |  | The user facility reported a 77-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp support ongoing for 4 days when there were signs of hemolysis. lab values prompted the team to believe hemolysis to be ongoing. the pump remained on for support without any intervention. conversations are being made regarding escalation to a 5.5 pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the reported hemolysis has been completed. the pump was not returned for investigation. data logs were investigated and no positioning issue, motor current high or suction alarms were observed at the time of the reported failure. clinical information provided was insufficient and merely suspected hemolysis based off of drop in platelets. therefore, the root cause of the hemolysis issue was not determined since the product was not returned and insufficient clinical information was provided. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ b.2 revised as previous selection was entered incorrectly on the initial manufacturer device report number 1220648-2023-04713. b.7 revised as information was inadvertently omitted from the initial manufacturer device report number 1220648-2023-04713. d.4 revised model number from the initial submission of manufacturer device report number 1220648-2023-04713 in accordance with updated procedures. d.6a and d.6b revised from the initial submission of manufacturer device report number 1220648-2023-04713 in accordance with updated procedures. e.1 phone number was revised as it was entered incorrectly on the initial manufacturer device report number 1220648-2023-04713. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2023-04713 and should not have been. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2023-04713 was submitted. h.6 code 4641 was entered incorrectly on the initial manufacturer device report number 1220648-2023-04713. a new code has been added. h.6 code 925 was added since the initial manufacturer device report number 1220648-2023-04713 was submitted in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures.
33035 | 1220648-2023-04714 | 18051814 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18051814&pc=OZD | 10/02/2023 | 08/30/2024 |  | 11/01/2023 |  |  | 10/02/2023 | 11/01/2023 |  |  | 04/30/2025 |  | 11/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024391379 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 68-year-old female patient was scheduled for escalation to an impella 5.5 pump for medical support. the patient's blood vessels were noted to be extremely severe and tortuous in nature making the 5.5 insertion unsuccessful. it was further reported that the catheter kinked. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for delivery issue has been completed. the cause of delivery issue was patient condition related due to strong tortuosity vessels. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use: section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ a.1 revised as was previously entered incorrectly on the initial manufacturer device report number 1220648-2023-04714. d.4 revised serial number and unique identifier (udi) # as they were previously entered incorrectly on the initial manufacturer device report number 1220648-2023-04714. d.6a and d.6b revised from the initial submission of manufacturer device report number 1220648-2023-04714 in accordance with updated procedures. e.1 revised as information was inadvertently omitted on the initial manufacturer device report number 1220648-2023-04714. g.1 revised reporting contact fax number in accordance with updated procedures since the initial manufacturer device report 1220648-2023-04714 was submitted. h.6 codes 2199, 2981 and 4114 were previously entered incorrectly on the initial manufacturer device report number 1220648-2023-04714. a new code has been added to type of investigation codes. h.10 additional manufacturer narrative instructions for use were inadvertently omitted from the previously submitted manufacturer device report number 1220648-2023-04714. additional manufacturer narrative on the previously submitted report stated that the device was not returned but the device was discarded. type of investigation codes has been updated to reflect discarded.
33036 | 1220648-2023-04715 | 18051871 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18051871&pc=OZD | 10/02/2023 | 11/01/2023 |  | 11/01/2023 |  |  | 10/02/2023 | 11/01/2023 |  |  |  |  | 11/01/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1089014 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Unexpected Shutdown | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The us complainant reported the automated impella controller (aic) had a device shutdown so the aic was switched out. there was no reported patient harm. | [5 empty] | The investigation into the reported shutdown or restart issue has been completed. data analysis shows that the automated impella controller (aic) ran the pump until (b)(6) 2023 when the io-graphics process crashed. the console rebooted in response followed by an io-graphics core dump and an unexpected shutdown alarm confirming what was reported clinically. therapy was resumed for another hour before the pump was unplugged. the failure mode was unable to be reproduced naturally while running a pump during the investigation at the bench level. the io-graphics was forced to shutdown using a telnet command and the console responded the same as in the field. the root cause of the aic issue was determined to be the third party hardware/software that runs io-graphics.
33037 | 1220648-2023-04716 | 18052001 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18052001&pc=OZD | 10/02/2023 | 03/01/2024 |  | 11/01/2023 |  |  | 10/02/2023 | 11/01/2023 |  |  | 07/31/2025 |  | 11/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024387450 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/02/2023 |  |  |  |  | The user facility reported a 61-year-old male patient was implanted with an impella device for mechanical circulatory support. the complainant reported that on day four of impella support sedation holiday was performed. the patient was not moving the right upper extremity or tracking to the right side. a computed tomography¿(ct) scan of the head was performed and there was a small, stable wedge-shaped area of hypodensity near the frontoparietal junction, representing a small acute to subacute infarction. the cardiologist believes that it could possibly be related to either impella support or previous balloon pump support. support continued with the implanted pump. | [5 empty] | The investigation into the stroke/cva has been completed. the device was not returned for investigation. as the clinical information was not provided, the true root cause of the stroke/cva could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33038 | 1220648-2023-04717 | 18052023 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18052023&pc=OZD | 10/02/2023 | 02/02/2024 |  | 11/01/2023 |  |  |  | 11/01/2023 |  |  | 06/30/2025 |  | 11/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024364840 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  |  | Malfunction | No | Yes | Migration | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported that 67-year old male patient was implanted with an impella 5.5 as bridge to transplant. the impella 5.5 migrated into the aorta while the patient was transferred 'out of bed' to commode. metrics consistent with "impella in aorta" alarm, and radiographic evidence confirmed that cannula was in innominate takeoff. the surgeon suspects the lock was "greasy." the pump was explanted as a result of the noted issue. | [5 empty] | The investigation into the positioning issues has been completed. the product was returned for review. visual inspection found no issues on the pump. the cath lock was intact and locked well on the catheter. the cath lock passed the pull test as the lock was unable to move on the catheter when it locked. the root cause of positioning issue was most likely a cath anchor use related. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿
33039 | 1220648-2023-04718 | 18052097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18052097&pc=OZD | 10/02/2023 | 02/20/2024 |  | 11/01/2023 |  |  |  | 11/01/2023 |  |  | 08/31/2025 |  | 11/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024385349 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a 53-year-old male patient was implanted with an impella cp pump for mechanical circulatory support. upon removal of the 14f peel away sheath and advancement of re-positioning sheath, the impella cp catheter advanced into the left ventricle and was caught in the papillary muscle. interventional cardiologist attempted to reposition the device and met resistance. the body of the cannula kinked under fluoroscopy. the decision was therefore made to remove the impella and replace it with a new impella cp. there was no reported patient harm. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed since the original report was filed. the device was returned and tested after arriving in fs. visual inspection found a kink on can about 15 mm from of cage & can bonding site. the root cause of positioning issue was use related due to improper technique removing 14f peel away sheath and advancement of re-positioning sheath.
33040 | 1220648-2023-04719 | 18052147 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18052147&pc=OZD | 10/02/2023 | 01/23/2024 |  | 11/01/2023 |  |  |  | 11/01/2023 |  |  | 11/30/2024 |  | 11/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023279074 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 85-year-old female was implanted with an impella 5.5 for mechanical circulatory support post coronary artery bypass graft (cabg) procedure. upon ramping up p levels, it was noticed that the patient had a dissection in their aorta. it was unclear to the local team and medical staff if it was caused by the impella or if it was an anatomical issue with the patient. the impella 5.5 had to therefore be explanted. | [5 empty] | The investigation into patient's vessel damage has been completed. no product was returned. per the clinical investigation, the root cause of the vascular damage-great vessel is most likely due to the patients condition. the relationship between the impella and the vascular damage is unknown. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
33041 | 1220648-2023-04720 | 18052356 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18052356&pc=OZD | 10/02/2023 | 11/01/2023 |  | 11/01/2023 |  |  |  | 11/01/2023 |  |  |  |  | 11/01/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1690036 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Power Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant in japan reported that the automated impella controller (aic) controller error alarm occurred during impella cp priming. the light from the optical sensor was not coming out from the connector connection. there was no reported patient harm. | [5 empty] | The investigation into the reported automated impella controller (aic) - controller error (yellow alarm)issue has been completed. data analysis confirmed the reported failure. log shows that on the complaint date, the console received the yellow controller error alarm before connecting the pump on the initial startup, then the aic restarted twice, and every time upon startup received the controller error (defective optical sensor lamp). during functional testing at the bench level, the failure mode was able to be reproduced. the root cause of the optical signal issue was determined to be due to the defective lamp. the failure modes will be monitored and trended.
33042 | 1220648-2023-04721 | 18052358 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18052358&pc=OZD | 10/02/2023 | 08/13/2025 |  | 11/01/2023 |  |  |  | 11/01/2023 |  |  |  |  | 11/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1613395 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/02/2023 |  |  |  |  | The user facility reported an unspecified patient was implanted with an impella device for mechanical circulatory support. upon staff powering on the aic, it would not maintain a charge. a red battery failure and yellow alert battery communication failure noted. a new console was used, as a result. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
33043 | 1220648-2023-04722 | 18052361 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18052361&pc=OZD | 10/02/2023 | 11/17/2025 |  | 11/01/2023 |  |  |  | 11/01/2023 |  |  |  |  | 11/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1145133 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | While preparing an impella heart pump, the automated impella controller (aic) skipped the final preparation step. therefore, the backup controller was used. this meant that the impella heart pump could be started without any further problems. there was no reported patient harm. | [5 empty] | A.2 age should have been left blank on the initial manufacturer device report as it is unknown. a.3 unknown should have been selected on the initial manufacturer device report as it is unknown. d.1 and d.2 brand name and common device name were revised in accordance with updated procedures since the initial manufacturer device report was submitted. d.4 model number and serial number were revised in accordance with updated procedures since the initial manufacturer device report was submitted. e.1 name prefix/title, last name and address line 1 were reported incorrectly on the initial manufacturer device report and were revised. e.3 revised as it was incorrectly reported on initial manufacturer device report. g.1 revised reporting contact email and manufacturing site fax number in accordance with updated procedures since the initial manufacturer device report was submitted. h.6 code 2199 was reported incorrectly on the initial manufacturer device report. | The investigation has been completed. console logs show that after the pump is connected, the console exists case wizard. device analysis: the device was returned for investigation. inspection of the internal circuitry of the console didn't reveal any anomaly. a known good pump and purge line was run with console with no observable anomaly. conclusion: the root cause for the complaint was likely use related. the case start procedure is skipped when an already initialized pump has already gone through case start, which includes pump/cassette connection and priming. the console operated as designed. h3: added information regarding the status of the investigation. h10: added the results of the investigation. | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
33044 | 1220648-2023-04723 | 18052364 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18052364&pc=OZD | 09/14/2023 | 03/04/2024 |  | 11/01/2023 |  |  |  | 11/01/2023 |  |  | 04/30/2025 |  | 11/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024339951 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/02/2023 |  |  |  |  | The user facility reported a 72-year-old male post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. the patient expired after 69.5 hours of impella support. the cause of death was determined to be cerebral hemorrhage due to cerebral infarction. the relationship between the impella and the cerebral hemorrhage is unknown, therefor the relationship between the cause of death and the impella is unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported stroke/cva has been completed since the original report was filed. since enough clinical information was not provided, the true root cause of the stroke/cva could not be determined.
33045 | 1220648-2023-04724 | 18052386 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18052386&pc=OZD | 10/02/2023 | 08/07/2024 |  | 11/01/2023 |  |  |  | 11/01/2023 |  |  | 06/30/2025 |  | 11/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024373488 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 60-year-old female was implanted with an impella cp device for mechanical circulatory support. while in the cardiac catheterization lab, the pump was noted to have been left slightly deep. an echo was done and the pump was found to have moved back towards the aortic valve. even though hemodynamics were stable the pump was ultimately removed at bedside. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the positioning issue was patient movement. the instructions for use referenced in the initial report is for the impella cp with smartassist system in the section: use of echocardiography for positioning of the impella catheter. added- h6 component code 925 added-d6a/d6b.
33046 | 1220648-2023-04725 | 18052388 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18052388&pc=OZD | 10/02/2023 | 02/02/2024 |  | 11/01/2023 |  |  |  | 11/01/2023 |  |  | 08/31/2025 |  | 11/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024387894 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Improper Flow or Infusion | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/02/2023 |  |  |  |  | The user facility reported a 67-year-old female with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the patient was scheduled for possible va ecmo decannulation. after the impella placement there were flow saturations with inverted aortic/left ventricle signals and flat motor current. impella was removed due to clot ingestion. the patient was sent back to unit on va ecmo. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the saturated flow and thrombus issues has been completed since the original report was filed. the device was returned for investigation. data logs showed several motor current spikes followed by saturated flow and high purge pressure. catheter kinks and cannula kinks were observed near the motor housing. upon further investigation, biomaterial was observed in the purge gap. during testing, the pump ran as expected after removing the biomaterial from the purge gap. the root cause of the saturated flow issue was biomaterial, based on the investigation of the returned product and the log analysis.
33047 | 1220648-2023-04726 | 18052395 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18052395&pc=OZD | 10/02/2023 | 01/26/2024 |  | 11/01/2023 |  |  |  | 11/01/2023 |  |  |  |  | 11/01/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | ABIOMED, INC. |  |  | US |  | 1143274 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 10/02/2023 |  |  |  |  | The user facility reported that during set-up for an impella cp implant, the automated impella controller (aic) purge disc was not recognized. the aic was swapped out without issue. the aic will be returned for investigation. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the console was returned for investigation. this complaint was investigated in collaboration with fs and aic was tested. upon examining the aic for any visible defects, it was evident that the purge disc flag was broken which would result in ¿purge disc not recognized¿ alarms. the root cause of this issue is a broken purge disc flag.
33048 | 1220648-2023-04727 | 18052432 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18052432&pc=OZD | 10/02/2023 | 11/01/2023 |  | 11/01/2023 |  |  | 10/02/2023 | 11/01/2023 |  |  |  |  | 11/01/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1427009 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Activation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an automated impella controller (aic), failed to recognize an installed purge disc. the device was removed from service as a result of the noted issue. there was no reported patient harm. | [5 empty] | The investigation into the reported purge disc issues has been completed. the device was returned for evaluation. during functional testing at the bench level, the console was repeatedly unable to detect the purge. upon further analysis, the purge disc retainer and flag were confirmed to be broken. the root cause of the reported failure was determined to be a missing purge flag. should additional relevant information become available, a supplemental report will be submitted.
33049 | 1220648-2023-04728 | 18052464 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18052464&pc=OZD | 10/02/2023 | 01/19/2024 |  | 11/01/2023 |  |  | 10/02/2023 | 11/01/2023 |  |  | 08/31/2025 |  | 11/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024381526 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that 58-year-old female was scheduled for impella 5.5 implant as bridge to decision. after two hours of attempting insertion, and over an hour of fluoro time, decision made to stop and discuss mini-sternotomy with family given complexity of patient with limited bridge options. when the impella 5.5 was removed, there was visible kinking of white catheter just below the motor. there was no known harm or injury to the patient. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation into delivery issues and cannula kink has been completed. impella 5.5 pump was returned for evaluation. visually inspected the pump and found a significant catheter kink close to 35 cm marking. no other abnormalities were found. as per clinical, the impella 5.5 advanced and had resistance at subclavian/aorta junction with patient having calcified vessel conditions, decision made to abort insertion. a visible kinking of white catheter below the motor was observed after removal. the cause of the delivery issue was patient condition related with calcified vessel condition due to resistance at subclavian/aorta junction and catheter kink observed during insertion.
33050 | 1220648-2023-04729 | 18052489 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18052489&pc=OZD | 10/02/2023 | 03/01/2024 |  | 11/01/2023 |  |  | 10/02/2023 | 11/01/2023 |  |  | 06/30/2025 |  | 11/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024373889 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Improper Flow or Infusion | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/02/2023 |  |  |  |  | The user facility reported a 79-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient was on impella support for 10 minutes when there were persistent suctions alarms during transition to repositioning sheath. the physician noted there was a link in the cannula. despite repositioning, the impella cp was removed and replaced with a new cp without issue. | [5 empty] | The impella device was received from the customer and an evaluation is ongoing. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow and the product damage (cannula kink) events has been completed since the original report was filed. data confirmed the low flow when pump started and remained to the rest of the case that triggered auto suction response and suction alarms. kink was also found. the root cause of low flow was kinked cannula.
33051 | 1220648-2023-04730 | 18052494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18052494&pc=OZD | 10/02/2023 | 04/02/2024 |  | 11/01/2023 |  |  | 10/02/2023 | 11/01/2023 |  |  | 07/31/2025 |  | 11/01/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024385931 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/02/2023 |  |  |  |  | The user facility reported a 85-year-old male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the medical team had difficulty delivering the impella pump and the pump showed immediate low flows/continuous suction despite repositioning. the md attempted to re-wire and re-advance with the same issues. it was noted that impella cannula was kinking near outlet. impella ultimately replaced with new impella cp without problem. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the product damage (cannula kink) issue has been completed since the initial report was submitted. the product was not returned for investigation. based on the clinical description and data analysis, the cause of the low pump flow was cannula kink that occurred during insertion due to improper technique.
33052 | 1220648-2023-04691 | 18038212 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18038212&pc=OZD | 10/02/2023 | 01/23/2024 |  | 10/31/2023 |  |  | 10/02/2023 | 10/31/2023 |  |  | 05/31/2025 |  | 10/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024359492 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/02/2023 |  |  |  |  | The user facility reported a 73-year-old female with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support, the patient had bleeding from impella right groin access site. one unit packed red blood cells transfused at that time. systemic heparin was placed on hold, purge solution was transitioned to bicarb, and a fem-stop was applied. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the issue of access site bleeding ¿ major has been completed. the device was not returned for investigation. as per clinical patient had bleeding at access site around the sheath, blood was given and femstop was applied. clinical advised the nurse for angle matching and placing the sheath up during the next dressing change, and bleeding was stopped with angle being matched the following day. the cause of the access site bleeding issue was most likely use related based on the clinical review.
33053 | 1220648-2023-04692 | 18038326 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18038326&pc=OZD | 10/07/2023 | 02/20/2024 |  | 10/31/2023 |  |  | 10/07/2023 | 10/31/2023 |  |  | 07/31/2025 |  | 10/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024380552 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/07/2023 |  |  |  |  | The user facility reported a 50-year-old male was implanted with an impella 5.5 as a bridge to transplant. the patient had a hematoma at the access site. heparin was being held due to hematoma. issue was resolved and patient was on support for 21 days post the issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported access site bleeding major has been completed since the original report was filed. the pump was not returned for investigation. according to clinical records, heparin was stopped after the hematoma was discovered. the cause of the access site bleeding issue was most likely high patient act values, based on provided clinical details. removing heparin resolved bleeding.
33054 | 1220648-2023-04693 | 18038442 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18038442&pc=OZD | 10/02/2023 | 06/06/2024 |  | 10/31/2023 |  |  | 10/02/2023 | 10/31/2023 |  |  | 02/28/2025 |  | 10/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024337803 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/02/2023 |  |  |  |  | The user facility reported a 65-year-old female (race unknown) who underwent cardiac surgery was implanted with an impella cp device for mechanical circulatory support. during support, the patient was bleeding from the impella access site. angle matching was performed, and pressure dressing applied. partial thromboplastin time (ptt) are therapeutic. the plan is for the next dressing change to attempt a new angle and give the sheath forward tension to see if the bleeding will slow or subside. systemic heparin was stopped. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. impella cp with smart assist system: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
33055 | 1220648-2023-04698 | 18040319 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18040319&pc=OZD | 10/13/2023 | 06/20/2024 |  | 10/31/2023 |  |  | 10/13/2023 | 10/31/2023 |  |  | 07/31/2025 |  | 10/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024386944 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/13/2023 |  |  |  |  | The user facility reported a 66-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had placed a cp via the femoral artery for a patient admitted with multiple underlying cardiac comorbidities and systemic issues, such as gastro intestinal bleed and deep vein thrombosis (dvt). the pump was placed, and the limb showed signs of limb ischemia. the team lacked the equipment for a bypass placement and so the leg was monitored. the limb had calcification and had to have a vascular repair done at the time of explant. the support was for under 11 hours. after explant a 5.5 was placed via axillary insertion. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported limb ischemia and vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia and vascular damage could not be determined. the instructions for use referenced in the initial report is the impella cp with smartassist system in section: potential adverse events (united states) which now includes statement "cardiac or vascular injury (including ventricular perforation.¿ d4-corrected model number in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
33056 | 1220648-2023-04699 | 18040484 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18040484&pc=OZD | 10/02/2023 | 01/25/2024 |  | 10/31/2023 |  |  | 10/02/2023 | 10/31/2023 |  |  | 07/31/2025 |  | 10/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024380574 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Material Deformation | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/02/2023 |  |  |  |  | The user facility reported a 58-year-old female with cardiomyopathy was to be implanted with an impella 5.5 for mechanical circulatory support. the medical staff was having difficulty delivering the impella pump and, after multiple failed attempts and techniques used, they aborted the procedure and pulled the pump out and noticed that a kink had been formed in the cannula just below the motor. decision made to remove device and attempt second device but the outcome was the same. | [5 empty] | The impella device was received from the customer and an evaluation is underway. upon completion, a supplemental mdr will be filed. | The investigation into the into the delivery issues- use has been completed since the original report was filed. returned complaint 5.5 pump visually inspected. catheter kink near motor housing observed. the root cause of the delivery issue was patient condition related due to calcified vessels condition that kinked the catheter while attempted insertion process.
33057 | 1220648-2023-04700 | 18041345 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18041345&pc=OZD | 10/04/2023 | 02/02/2024 |  | 10/31/2023 |  |  | 10/04/2023 | 10/31/2023 |  |  | 02/28/2025 |  | 10/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024309705 |  | 0048-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation | Distributor |  | 0 YR |  | Female |  |  |  | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/04/2023 |  |  |  |  | A distributor reported an israeli female patient (unknown age and race) was being supported by an impella cp when the impella had low pump flows. "no pump flow in all p-level; motor current in range; purge fluid/purge pressure in range; well positioned; filled left ventricle (lv); va-ecmo supports with 2.5l; suspicion of thrombus in impella inlet cage/cannula; therefore discuss internal to change pump; the decision is still pending." there was no report of patient harm related to this event. no further details have been shared. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow issue has been completed. the product and data logs were not provided for investigation. the cause of the low pump flow issue was not determined since product was not returned and sufficient clinical information was not provided.
33058 | 1220648-2023-04701 | 18041566 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18041566&pc=OZD | 10/04/2023 | 01/12/2024 |  | 10/31/2023 |  |  | 10/04/2023 | 10/31/2023 |  |  | 10/30/2024 |  | 10/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024292376 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | Material Deformation | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/04/2023 |  |  |  |  | The user facility reported an 83-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the physician was unhappy with the impella pump flows, and there was a kink on cannula near outlet. decision made to remove the pump. pump removed and new pump utilized. there was no patient harm related to this event. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow and the product damage (cannula kink) has been completed since the original report was filed. the pump was returned for evaluation. visually inspected the pump and the cannula was found to be kinked. no other physical abnormalities were found. it was reported that the pump had low flows with suction alarms at the end of case. shortly after implant, patient flows dropped, suction event followed with noted bend at cannula leading to limited support for patient. the cause of the low pump flow issue was due to a kinked cannula per field investigation, clinical and data log review.
33059 | 1220648-2023-04702 | 18041803 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18041803&pc=OZD | 10/05/2023 | 10/31/2023 |  | 10/31/2023 |  |  | 10/05/2023 | 10/31/2023 |  |  |  |  | 10/31/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1133094 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 10/05/2023 |  |  |  |  | The user facilites biomed department reported a malfunction with the impella console (aic). it was found to have a purge clip breakage. the malfunction was observed outside of patient use. | [5 empty] | The console was returned for a benchtop analysis and was visually inspected under a digital microscope. cracks on the lateral side of the purge retainer as well as a missing purge flag were observed. the root cause was a defective purge pressure sensor .
33060 | 1220648-2023-04703 | 18042885 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18042885&pc=OZD | 10/01/2023 | 01/31/2024 |  | 10/31/2023 |  |  |  | 10/31/2023 |  |  | 07/31/2025 |  | 10/31/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024390903 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient was implanted with an impella cp for hemodynamic support. the patient continued with belt pressure on femostop and received 1 unit of packed red blood cells (prbc) overnight due to continued groin oozing. the impella cp was successfully weaned and explanted four days post implantation. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the product was not returned for review. based on clinical description, the root cause of the access site bleed was most likely due to patient movement.
33061 | 1220648-2023-04677 | 18029199 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18029199&pc=OZD | 09/30/2023 | 05/29/2024 |  | 10/30/2023 |  |  | 09/30/2023 | 10/30/2023 |  |  | 07/31/2025 |  | 10/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024378435 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 40-year-old male was implanted with an impella 5.5 for mechanical circulatory support. the patient developed a hematoma and needed washout at the surgical site and systemic heparin was on hold. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use noted in the initial report is from the impella 5.5 with smart assist for use during cardiogenic shock section: : potential adverse events (united states) , which now includes the statement "cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
33062 | 1220648-2023-04678 | 18029270 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18029270&pc=OZD | 09/30/2023 | 02/02/2024 |  | 10/30/2023 |  |  | 09/30/2023 | 10/30/2023 |  |  | 12/31/2024 |  | 10/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024309168 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient was experiencing oozing from the access site of the impella cp and there was a significant drop in hemoglobin. the physician had already tightened the preclose but oozing continued. a decision was therefore made to remove the impella cp. during the removal, the patient started to require neo and epi boluses. intra-aortic balloon pump (iabp) was placed. code blue called and patient expired. | It is unknown if the device was related to the patient's death. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. clinical details states that that the patient's activated coagulation time (act) was prolonged. therefore, the root cause of the access site bleed was most likely due to improper anticoagulation management. section b5: information was inadvertently left out in the inital report which has been added to this report.
33063 | 1220648-2023-04679 | 18029323 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18029323&pc=OZD | 10/04/2023 | 09/04/2024 |  | 10/30/2023 |  |  | 10/04/2023 | 10/30/2023 |  |  | 08/31/2025 |  | 10/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024385513 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/04/2023 |  |  |  |  | The user facility in japan reported a 58-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. during impella insertion, the impella was inadvertently pulled out of the left ventricle. the hemodynamics were stabilized by increasing the dose of catecholamine, and fluoroscopy and echocardiography confirmed that impella had been successfully reinserted into the left ventricle. there was no harm to the patient as a result of this incident, | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | Upon review of this file it was determined that there was no adverse event associated with impella. the impella was pulled out from inside the left ventricle (lv) for an off-pump coronary artery bypass surgery that was performed while impella was inserted. a regulatory report is no longer necessary.
33064 | 1220648-2023-04680 | 18031143 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18031143&pc=OZD | 10/01/2023 | 02/02/2024 |  | 10/30/2023 |  |  | 10/01/2023 | 10/30/2023 |  |  | 03/31/2025 |  | 10/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024343524 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/01/2023 |  |  |  |  | The user facility reported a 42-year-old male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support, the patient had an echo which showed that the pump appeared to be pointing at mitral valve. patient had cola colored urine and md wanted to reposition impella device. the physician was repositioning the pump with echo guidance and the patient was transferred to cath lab for reposition with fluoro guidance. the pump was then explanted and care was escalated to the imeplla 5.5 pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the positioning issues has been completed. the product was not returned for review. based on the clinical details, the root cause of positioning issue was due to use related. b.1 adverse event was selected on the original manufacturer device report (mdr) and product problem should have been selected. b.2 required intervention to prevent permanent impairment/damage was selected on the previously submitted mdr and should not have been. h.1 type of reportable event was incorrect on the previously submitted mdr and malfunction should have been selected. h.10 the following information was included on the additional manufacturer narrative previously submitted and should not have been: instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33065 | 1220648-2023-04681 | 18031210 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18031210&pc=OZD | 10/01/2023 | 02/21/2024 |  | 10/30/2023 |  |  | 10/01/2023 | 10/30/2023 |  |  | 07/31/2025 |  | 10/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024388643 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/01/2023 |  |  |  |  | The user facility reported a 73-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support post cardiac surgery. during support, the pump was inadvertently pulled back into the patient's aorta. an echo was performed and it showed that the patient had ejection fraction of about 45%. decision was made by the medical team to remove impella one day earlier than planned. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed. the device was not returned for investigation. per the clinical information provided, the cause of the positioning issue was use related due to improper technique because impella was accidentally pulled back into the aorta.
33066 | 1220648-2023-04682 | 18031287 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18031287&pc=OZD | 10/26/2023 | 01/25/2024 |  | 10/30/2023 |  |  | 10/26/2023 | 10/30/2023 |  |  | 08/31/2025 |  | 10/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024392558 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/26/2023 |  |  |  |  | The user facility reported a 51-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. after the pump access site was closed, the team observed a blood leak at the sleeve and suture hub. the team made choice to explant and replace the pump. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation into the product damage (hole in catheter) has been completed since the original report was filed. the product was returned and tested. upon investigation of the pump, no damages were noted. bleeding at access site was also reported where additional intervention was required, and pump was explanted. the root cause of the access site bleeding - major was determined to be use related due to the repositioning unit being too close to the access site; blood was able to enter through the catheter anchor.
33067 | 1220648-2023-04683 | 18031292 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18031292&pc=OZD | 10/01/2023 | 08/13/2024 |  | 10/30/2023 |  |  | 10/01/2023 | 10/30/2023 |  |  | 07/31/2025 |  | 10/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024388086 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/01/2023 |  |  |  |  | The user facility reported a 63-year-old male post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support when the patient was decompensating after coronary artery bypass surgery. during support, reported the patient lost distal pulses and the impella had to be removed due to the ischemia. their leg became mottled and cold. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. the instructions for use for the impella cp with smartassist system states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
33068 | 1220648-2023-04684 | 18031327 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18031327&pc=OZD | 10/27/2023 | 12/04/2023 |  | 10/30/2023 |  |  | 10/27/2023 | 10/30/2023 |  |  | 09/30/2025 |  | 10/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024398411 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/27/2023 |  |  |  |  | The user facility reported a 57-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient had been admitted with multiple cardiac comorbidities and a planned tavr. the pump stopped after just under 5 hours. when the pump was explanted, they observed calcium in the pump. a new 5.5 pump was placed. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop has been completed. according to the clinical review, the patient had preexisting calcium around the aortic valve (av). the returned pump was visually inspected and calcium was observed in the outflow cage. the field rtlog showed several pump stops following a spike in motor current. the cause of the pump stop was ingested calcium pulled in from the diseased aortic valve.
33069 | 1220648-2023-04685 | 18031359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18031359&pc=OZD | 09/26/2023 | 05/01/2024 |  | 10/30/2023 |  |  | 09/26/2023 | 10/30/2023 |  |  | 03/31/2025 |  | 10/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024344386 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 10/02/2023 |  |  |  |  | The user facility reported a 53-year-old male in cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was on support for just under 13 hours. when the pump support was ongoing , the patient developed limb ischemia. the leg with the impella placed lost pulses distally. there was no pedal pulse. the team additionally noted the patient had signs of hemolysis. the urine color had changed and the team made choice to explant the pump. the pump site also had an access site blood ooze. the ooze was not treated aggressively, but rather by just nursing care to the site. the patient was escalated to an impella 5.5 via the axillary site. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the limb ischemia, access site bleed, and hemolysis has been completed. according to the clinical, shortly after beginning impella cp support the patient was taken off heparin because of oozing at the groin site. no pulse could be found in the right leg and the patient's urine was a dark amber color. the decision was then made to explant the pump. no product was returned for a visual inspection. data log analysis confirmed no motor current (mc) spikes, positioning issues, or other abnormalities. the cause of the access site bleeding-major was not determined because no product was returned, and not enough clinical information was given. the cause of the hemolysis was not determined because no product was returned, and not enough clinical information was given. the cause of the limb ischemia could not be determined without additional clinical details. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ to conform with updated procedures, sections d6 & h10 have been revised with updated information.
33070 | 1220648-2023-04686 | 18031682 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18031682&pc=OZD | 10/27/2023 | 03/25/2024 |  | 10/30/2023 |  |  | 10/27/2023 | 10/30/2023 |  |  | 08/31/2025 |  | 10/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED INC. |  |  | US |  | 2024395866 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/27/2023 |  |  |  |  | The user facility reported a 72-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported that a hematoma formed at the access site when the patient was moving. pressure was held and heparin was paused until the next morning. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the initial report was submitted. the product was not returned for investigation. the root cause of the access site bleeding was most likely due to patient movement since patient moved during foley placement.
33071 | 1220648-2023-04688 | 18032405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18032405&pc=OZD | 10/01/2023 | 09/16/2025 |  | 10/30/2023 |  |  | 10/01/2023 | 10/30/2023 |  |  | 08/31/2025 |  | 10/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024383002 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was one isolated episode of an impella console (aic) failure during the support with an impella 5.5 pump. the nurse was able to restart the console and resume support at previous p-level with no further issues, but the team has a backup console ready in case of further malfunction. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation for the reported aic - controller failure (red alarm) issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the data logs from the period of date of occurrence could not be obtained. all available logs were not applicable. a known good pump and purge line was run with the console with no observable anomaly. there was no confirmation of the reported failure mode from any of the available logs. the root cause for the failure could not be determined due to insufficient information.
33072 | 1220648-2023-04687 | 18032485 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18032485&pc=OZD | 10/02/2023 | 01/31/2024 |  | 10/30/2023 |  |  | 10/02/2023 | 10/30/2023 |  |  | 01/21/2025 |  | 10/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024326939 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/02/2023 |  |  |  |  | The user facility reported a 66-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, there was concern that the impella left ventricle (lv) wave forms looked strange. lv waveform was higher than aorta (ao) waveform. motor current spiked, and pump set itself to p0. they then set it to p2, and it was reading a flow of 3 liters at p2. patient was taken to cath lab for replacement of the impella pump with low flows. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported low pump flow has been completed. the impella device was not returned for evaluation. however, data logs were reviewed which confirmed a high motor current pump stop followed by saturated flow. clinical details states that after restarting the pump saturated flow was also recorded. without the actual device, the root cause of the high motor current pump stop was unable to be determined because no product was returned.
33073 | 1220648-2023-04689 | 18032621 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18032621&pc=OZD | 10/30/2023 | 08/19/2024 |  | 10/30/2023 |  |  | 10/30/2023 | 10/30/2023 |  |  | 12/31/2024 |  | 10/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2024303741 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/30/2023 |  |  |  |  | The user facility in japan reported a 67-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. complainant reported the patient had a right frontal lobe subcortical hemorrhage. the causal relationship between bleeding and impella is unknown. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d4-corrected model number added- d6a f6/f8 should have been left blank in the initial report.
33074 | 1220648-2023-04690 | 18032742 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18032742&pc=OZD | 10/01/2023 | 08/08/2024 |  | 10/30/2023 |  |  | 10/01/2023 | 10/30/2023 |  |  | 01/31/2025 |  | 10/30/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024371793 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user in the eu reports on a 63-year-old patient in cardiogenic shock after an acute st-elevation myocardial infarction who was treated with an impella cp pump for hemodynamic support. the patient was observed to have hemolysis and developed compartment syndrome following access site bleeding. | It was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the access site bleed and hemolysis has been completed. according to the clinical, the patient began to bleed from the access site in their leg. shortly after this was discovered the patient also developed hemolysis and began to experience compartment syndrome. imaging confirmed the pump was making contact with the mitral valve. although a recommendation was made to reposition the pump, the patient expired on support before any troubleshooting could be performed. no product was returned for an investigation. data logs were not returned for analysis. the most likely cause of the hemolysis was positioning issues. the cause of the major access site bleeding was not determined because no product was returned and not enough clinical details were given. impella cp with smart assist system instructions for use states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added-b2-selected death, b5 mso review, h6 impact code 1802, h6 component code 925 h1 type of reportable event changed to death.
33075 | 1220648-2023-04676 | 18026652 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18026652&pc=OZD | 10/01/2023 | 01/25/2024 |  | 10/28/2023 |  |  | 10/01/2023 | 10/28/2023 |  |  | 06/30/2025 |  | 10/28/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024379155 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/01/2023 |  |  |  |  | The user facility reported a 62-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after impella was placed, impella flows decreased and the impella was back in the aorta. upon further observation, the impella appeared kinked near the outlet so impella was removed and new one placed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow has been completed. the device was not returned for benchtop investigation. according to the clinical information provided, the root cause of the low pump flow was cannula kink due to improper technique.
33076 | 1220648-2023-04611 | 18021428 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18021428&pc=OZD | 09/25/2023 | 10/27/2023 |  | 10/27/2023 |  |  | 09/27/2023 | 10/27/2023 |  |  | 02/29/2024 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023191004 |  | 005066-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/27/2023 |  |  |  |  | The user facility reported a 48-year-old female with cardiomyopathy was implanted with an impella 5.0 for mechanical circulatory support. during support, the device was intermittently triggering an outflow blocked alarm and the catheter position was adjusted each time. after the catheter was pulled out a little to adjust the position on september 25th, the impella outflow blocked alarm kept alarming. there was no problem in the depth of the catheter when confirmed using echocardiogram. at present, there are no findings that would positively suggest hemolysis. the impella was replaced with no complications. no reported harm to patient was noted. | [5 empty] | The impella device was returned, and an evaluation was completed. a review of the data logs noted a high motor current pump stop during support coinciding with impella outflow blocked alarms and impella position unknown alarms. a visual inspection of the returned pump found a large clot on the outflow cage. upon conclusion of the investigation, the cause of the reported issue was the presence of biomaterial in the device. per the ifu: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
33077 | 1220648-2023-04612 | 18021523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18021523&pc=OZD | 09/27/2023 | 11/14/2025 |  | 10/27/2023 |  |  | 09/27/2023 | 10/27/2023 |  |  |  |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1162029 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 09/27/2023 |  |  |  |  | The user facility's biomedical engineer reported that the automated impella controller (aic) was powered on to check for software status and a "controller error-switch to back up" alarm populated. the console was taken out of the lab and will be returned for investigation. there was no patient involvement. | [5 empty] | A.1 added information as it was inadvertently omitted from the initial manufacturer device report number 1220648-2023-04612. a.2 age should have been left blank in accordance with updated procedures on the initial manufacturer device report number 1220648-2023-04612 as no patient was involved. d.2 common device name was revised in accordance with updated procedures since the initial manufacturer device report number 1220648-2023-04612 was submitted. d.4 model number was revised in accordance with updated procedures since the initial manufacturer device report number 1220648-2023-04612 was submitted. e.2 revised selection as it was reported incorrectly from the initial manufacturer device report number 1220648-2023-04612. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2023-04612 and should not have been. g.1 revised mdr reporting contact email since the initial submission of manufacturer device report number 1220648-2023-04612 in accordance with updated procedures. h.6 code 3221 was reported incorrectly on the initial manufacturer device report number 1220648-2023-04612 in accordance with updated procedures. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation has been completed. data analysis: console imc logs confirm that the controller error (loss of comm with keypad or between the kbd and i2c for kbd hw revision 1) - alarm was issued. device analysis: inspection of the internal circuitry of the console revealed no anomaly. a known good pump and purge line was run with the console with observable anomaly. conclusion: the root cause of the kbd communication error was unable to determine due to the intermittent nature of the failure and the inability to reproduce.
33078 | 1220648-2023-04613 | 18021607 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18021607&pc=OZD | 09/27/2023 | 02/19/2024 |  | 10/27/2023 |  |  | 09/27/2023 | 10/27/2023 |  |  | 04/30/2025 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024351981 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  |  |  | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/27/2023 |  |  |  |  | The user facility reported a 59-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. initially, the impella was placed without complications; however, upon transfer to hospital bed, impella pulled back across the valve due to touhy borst not being tightened. decision made by ic to urgently replace device. new impella cp placed without complications. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed. the device and data were not returned for investigation. based on the clinical information provided, the cause of the positioning issue was touhy- borst related due to improper technique because impella pulled back across the valve due to touhy-borst not being tightened.
33079 | 1220648-2023-04614 | 18021642 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18021642&pc=OZD | 09/27/2023 | 02/02/2024 |  | 10/27/2023 |  |  | 09/27/2023 | 10/27/2023 |  |  | 07/31/2025 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024376896 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 36 YR |  | Female |  |  |  | Malfunction | No | Yes | Device Dislodged or Dislocated | Material Twisted/Bent | Positioning Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/27/2023 |  |  |  |  | The user facility reported a 36-year-old female (unknown race) in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support prior to percutaneous coronary intervention. the patient is enrolled in the dtu study. the impella jumped back and kinked in the patient's aortic arch. the physician attempted to place the impella back in the left ventricle but was unsuccessful due to the kink. the impella was removed. | [5 empty] | The investigation for the reported delivery issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the delivery issue was use related, due to improper technique of removing the sheath and placing the repo unit. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
33080 | 1220648-2023-04615 | 18021718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18021718&pc=OZD | 09/27/2023 | 01/25/2024 |  | 10/27/2023 |  |  | 09/27/2023 | 10/27/2023 |  |  | 05/31/2025 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024356532 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Mechanical Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/27/2023 |  |  |  |  | The user facility reported a 78-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. during implant of impella cp, while advancing into descending aorta, the 0.018" came out of the left ventricle. wire and catheter were pulled out of 14fr sheath. attempt was made to backload .018" wire back onto impella catheter unsuccessfully. per staff, complaint wire was unable to exit window of outlet after multiple attempts. md made decision to open new pump as patient was decompensating and was hemodynamically unstable. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation into the into the delivery issues- use has been completed since the original report was filed. the device was returned and evaluated. no issues were found on the pump. gw & red lumen were not returned for review. based on clinical description & pump analysis, the root cause of delivery issue was most likely due to use issue.
33081 | 1220648-2023-04616 | 18022016 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18022016&pc=OZD | 09/28/2023 | 01/22/2024 |  | 10/27/2023 |  |  | 09/28/2023 | 10/27/2023 |  |  | 05/31/2025 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2024353105 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Fluid/Blood Leak | Increase in Pressure |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/28/2023 |  |  |  |  | The user facility reported a 65 year-old male post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. due to a hairline crack between the pressure relief valve and the bacterial filter, purge fluid leaked out. after several attempts, the crack was finally glued and the decision was made to continue using the implanted pump, despite a notable purge flow of 1.8 and purge pressure of 1088mmhg. | [5 empty] | The investigation into the of the purge leak ¿ pump and the high purge pressure issues has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. the root cause of the purge leak-pump is most likely due to damage to the joint to reservoir filter, but it could not be confirmed due to lack of product returned. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
33082 | 1220648-2023-04617 | 18022053 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18022053&pc=OZD | 09/28/2023 | 03/10/2025 |  | 10/27/2023 |  |  | 09/28/2023 | 10/27/2023 |  |  | 02/29/2024 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023193098 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient experienced recurring hemolysis during impella cp support. although the hemolysis seemed to initially clear up on its own, it eventually returned and the patient was placed on va ecmo for additional support. plans were made to explant the cp, for solitary ecmo support, however the patient decompensated and required the pump and an lv vent. as a result, the patient was escalated to an impella 5.5 pump for ongoing support. following explant, the patient experienced persistent bleeding from the cutdown site prompting utilization of a jp drain. support continued with the subsequently implanted impella 5.5 pump. | [5 empty] | The root cause of the access site bleeding issue was not determined since product was not returned and limited clinical details were provided. the root causes of hemolysis could not be determined as insufficient clinical details were provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33083 | 1220648-2023-04618 | 18022152 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18022152&pc=OZD | 09/28/2023 | 05/02/2024 |  | 10/27/2023 |  |  | 09/28/2023 | 10/27/2023 |  |  | 06/30/2025 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024379336 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Infusion or Flow Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The u.s. complainant had impella 5.5 support ongoing for 16 days for a patient transferred to the tertiary center for advanced care. the pump support was ongoing with a hematoma treated and flow saturation. the flow saturation concerned the team that the pump may stop. the pump was explanted after the 16th day with underlying sepsis. | [5 empty] | The investigation into the saturated flow and the access site bleeding ¿ major issues has been completed since the original report was submitted. the device was not returned by the user facility for investigation. according to the clinical, flow saturation was recorded and followed by a pump stop. the next day low platelet counts were also recorded. the cause of the pump stop, the major access site bleeding, and the thrombocytopenia could not be determined because no product was returned, and limited clinical information was given. to conform to updated procedures, section d6 was revised with updated information. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33084 | 1220648-2023-04619 | 18022170 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18022170&pc=PYX | 09/28/2023 | 05/15/2024 |  | 10/27/2023 |  |  | 09/28/2023 | 10/27/2023 |  |  | 06/30/2025 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024370614 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/28/2023 |  |  |  |  | The user facility reported a 64-year-old male post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella rp for mechanical circulatory support. the impella rp failed delivery due to the loss steering ability of the wire/pump combination. the pump was removed and a kink was observed in the catheter. a new impella rp was prepped for implant. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation for the delivery issues has been completed. the product was returned and there was a visible kink in the catheter. there were no other damages found. the root cause of the delivery issue is most likely due to use as the catheter was kinked due to excessive force. in accordance with updated procedures, section d6 has been revised from the initial submission.
33085 | 1220648-2023-04620 | 18022684 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18022684&pc=OZD | 09/28/2023 | 01/18/2024 |  | 10/27/2023 |  |  | 09/28/2023 | 10/27/2023 |  |  | 07/31/2025 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024388694 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Unknown |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Advance | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/28/2023 |  |  |  |  | The user facility reported a 61-year-old patient with cardiomyopathy was to be implanted with an impella cp for mechanical circulatory support. the physician noted the impella catheter kept "buckling" in the ventricle during insertion. the physician stated the catheter felt weird. the physician made multiple attempts to straighten and reposition the impella. a second impella was implanted successfully. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation into delivery issues- use and product damage (cannula kink) has been completed since the original report was filed. the pump was returned for investigation. during visual inspection, a kink in the cannula was noted near the motor housing. no other physical abnormalities were observed. clinically, it was reported that the doctor was unable to place the impella in the ventricle despite multiple attempts with the first pump. the second impella was placed without any issues. the cause of the delivery issue was use related, as the kink in the cannula was noted during insertion of first pump, and the second device was placed without any problem.
33086 | 1220648-2023-04621 | 18022743 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18022743&pc=OZD | 09/29/2023 | 10/27/2023 |  | 10/27/2023 |  |  | 09/29/2023 | 10/27/2023 |  |  |  |  | 10/27/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1483134 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | Activation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that upon regular preventative maintenance inspection, when the power plug was unplugged and the automated impella controller (aic) was left running on battery for about 10 to 20 minutes, a short, small "pip" sound was heard. there should be no sound when the aic battery switch is on. it was noted the sound was not a long beep, but a short "pip" sound. there was no patient involvement. | [5 empty] | The investigation into the reported "aic - battery charging/other issues" has been completed. product was returned for investigation. during device analysis, the console was tested after arriving for analysis. the alarm sound was reproduced as reported through the following procedure, (1) unplugged the ac power. (2) disengaged the battery switch (discharging the supercapacitor). (3) plugged the ac cord for approx. 30 sec and unplugged (partially charging the supercapacitor). (4) observed the alarm sound that degraded in a few seconds (alarm was audible for approx. 33 seconds); this alarm sound was the same as in the video evidence uploaded in the system. during data analysis, no log was found related to the periodic inspection. after the last usage of the console with the pump at the customer site on (b)(6) 2023, the console was turned on only for two minutes on the complaint date (b)(6) 2023. no log was found to confirm the statement that ¿the aic was left running on battery for about 10 to 20 minutes¿. upon reviewing the data of the aic¿s last usage from the customer site, no issue was found with the battery cell¿s voltage. the reproduced alarm sound is a known safety mechanism to alert not to have the battery switch disengaged unless the aic is under shipping. the root cause for the alarm sound was due to the use issue.
33087 | 1220648-2023-04622 | 18022869 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18022869&pc=OZD | 09/29/2023 | 01/18/2024 |  | 10/27/2023 |  |  | 09/29/2023 | 10/27/2023 |  |  | 06/30/2025 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2024382489 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A patient with acute st-elevation myocardial infarction required hemodynamic support from an impella cp with smartassist heart pump. the user reported that the heart pump could not be started after implant. the pump was subsequently replaced with another impella cp pump for patient support. | [5 empty] | The investigation for the reported "pump did not start" issue has been completed. since the medical center did not return the impella device, no benchtop analysis was possible. the root cause of the pump did not start issue was not determined since product and data logs were not returned. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
33088 | 1220648-2023-04623 | 18022963 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18022963&pc=OZD | 09/29/2023 | 11/14/2025 |  | 10/27/2023 |  |  | 09/29/2023 | 10/27/2023 |  |  |  |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1467028 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Activation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 | [18 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a purge system failure alarm. the automated impella controller (aic) was exchanged. no further alarms or issues were noted. there was no patient harm resulting from the failure. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | Corrections to the initial mfr report: g1 mdr reporting contact email. d2 device name has been updated. h3 has been updated to device eval completed. | The investigation has been completed. data analysis: analysis of device logs (imc and rtlog data) indicated that during the case with pump 464162, purge cassette had been replaced 9 times (10 purge cassettes were used). reported purge issue started while purge cassette sn (b)(6) was in use and resolved - on the same console - after this purge cassette was replaced with sn (b)(6) (which was also used on the backup console). device analysis: console was tested in the lab, and no anomalies of purge system were observed. purge pressure sensor calibration was verified to be within specification, and no issues with purge motor or the rest of purge system occurred, while using a known-good engineering purge cassette. root cause: console testing and inspection, as well as data analysis, indicated that reported purge issues were unrelated to the console, and most likely caused by defective purge cassette 54480 or kinked catheter of pump 464162 (subject to separate investigation 23-18839-1). h3 updated. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated.
33089 | 1220648-2023-04624 | 18023035 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18023035&pc=OZD | 09/29/2023 | 11/26/2025 |  | 10/27/2023 |  |  | 09/29/2023 | 10/27/2023 |  |  |  |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1126306 |  | 0042-0000 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during transport to the icu, the white cable was tugged when the bed stopped moving and the plug pulled out of the console just slightly. there was no interruption in flow, but there were spikes in the aorta placement signal. the nurse recognized this and pushed it back in and the waveform returned to normal aorta placement signal. there was no reported patient harm. | [5 empty] | The automated impella controller (aic) device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
33090 | 1220648-2023-04625 | 18023048 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18023048&pc=OZD | 09/28/2023 | 10/31/2023 |  | 10/27/2023 |  |  | 09/28/2023 | 10/27/2023 |  |  | 08/31/2025 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024384892 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 50-year-old was implanted with an impella 5.5 device for mechanical circulatory support. the patient was on the impella 5.5 pump for two weeks, when the pump stopped. the pump could not be restarted and was pulled back and explanted. post explant the team upped the medical/pharmaceutical support and then transitioned the patient to palliative care. the patient lacked other options due to underlying condition. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the low pump flow has been completed since the original report was filed. product was returned for evaluation. upon investigation of the pump, a catheter kink was found right at the red handle on the proximal side. the sterile sleeve was detached on both sides. the pump failed electrical testing. opening the red handle revealed broken lines. the root cause of the pump stop issue was determined to be an open electrical line due to excessive force.
33091 | 1220648-2023-04626 | 18023061 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18023061&pc=OZD | 09/29/2023 | 02/21/2024 |  | 10/27/2023 |  |  | 09/29/2023 | 10/27/2023 |  |  | 08/31/2025 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024385573 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a 71-year-old male patient just had a successful insertion of impella cp through the right femoral artery (rfa). however, during sheath exchange and pump start, impella was pulled back across the atrioventricular (av). insertion of new impella cp was performed due to concern of clot formation in motor. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed. the device was not returned for investigation. per the clinical information provided, the cause of the positioning issue was use related due to improper technique because during sheath exchange and pump start impella was pulled back across the av.
33092 | 1220648-2023-04627 | 18023075 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18023075&pc=OZD | 09/30/2023 | 01/15/2024 |  | 10/27/2023 |  |  | 09/30/2023 | 10/27/2023 |  |  | 07/31/2025 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024380949 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/30/2023 |  |  |  |  | The user facility reported a 58-year-old male in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the placement signal alarm was active after the patient was coughing. no aortic waveform or numbers displayed. cable connection was checked. alarm audio was then disabled. additionally, the ecmo cannulas were flushed and clamped due to extremely low flow. the cp was increased to p8 and pump remained on for support for 19 days despite the saturation of the flows, was then weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the optical signal issue and the low pump flow has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the root cause of the saturated flow cannot be determined. a "placement signal not reliable" occurred on 9/30 and the 5s parameters are indicative of the patient cable being torqued. the root cause of the optical signal issue is most likely due to a damaged fiber line.
33093 | 1220648-2023-04628 | 18023083 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18023083&pc=OZD | 09/28/2023 | 01/24/2024 |  | 10/27/2023 |  |  | 09/28/2023 | 10/27/2023 |  |  | 03/31/2025 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024357295 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 33-year-old male was implanted with an impella cp device for mechanical circulatory support. the patient developed a hematoma at the groin. the staff made attempts to manage the hematoma, but ultimately the decision was made to remove the impella. manual pressure was held and hemostasis was achieved. | [5 empty] | The investigation into the issue of infection/sepsis has been completed. the device was not returned for investigation. review of the clinical notes revealed the location of the hematoma to be the groin. based on the clinical information, the root cause of the access site bleeding was most likely lack of vessel recoil onto the repo sheath. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33094 | 1220648-2023-04629 | 18023085 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18023085&pc=OZD | 09/27/2023 | 09/03/2024 |  | 10/27/2023 |  |  | 09/27/2023 | 10/27/2023 |  |  |  |  | 10/27/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 2023252322 |  | IC8168 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Defective Component | Activation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the purge disc not detected alarm occurred on the automated impella controller (aic) controller, after the cassette and bag were changed. the issue was resolved with new tubing. | [5 empty] | Upon review of the provided information there is currently no allegation of a malfunction or serious injury associated with the automated impella controller (aic). a regulatory report is no longer necessary. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
33095 | 1220648-2023-04630 | 18023115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18023115&pc=OZD | 09/29/2023 | 02/28/2024 |  | 10/27/2023 |  |  | 09/29/2023 | 10/27/2023 |  |  | 06/30/2025 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024368526 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/29/2023 |  |  |  |  | The user facility in the eu reported patient of unknown age, gender and race noted as high risk percutaneous coronary intervention (hrpci) was implanted with an impella cp device for mechanical circulatory support. during support, the patient developed right leg ischemia. a compartment syndrome was diagnosed. due to the necessity of further therapy (fasciotomy /compartment release) a detailed discussion with the patient's relatives took place. in view of a clearly formulated patient's will with rejection of a prolonged intensive medical maximum therapy in case of unclear outcome or in case of a reduction of the quality of life to be expected in the perspective even in a favorable course, change to palliative therapy, discontinuation of therapy. | [5 empty] | The investigation into the limb ischemia ¿ reversible has been completed. the device and data were not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33096 | 1220648-2023-04632 | 18023137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18023137&pc=OZD | 09/27/2023 | 01/24/2024 |  | 10/27/2023 |  |  | 09/27/2023 | 10/27/2023 |  |  | 08/31/2025 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024384961 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Improper Flow or Infusion | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/27/2023 |  |  |  |  | The user facility reported a 59-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support after surgery was complete. prior to leaving cath lab, the physician stated they wanted to leave the peel away sheath in. angio of the groin showed the peel away was occlusive so an antegrade reperfusion catheter was placed. the micro wire buckled part way down the superficial femoral artery (sfa). angio to mid sfa showed 100% occlusion. this was balloon opened and flow restored to the left leg. antegrade reperfusion sheath placed and pulses dopplered in bilateral feet. at the end of the procedure there were suction alarms present. the device was not replaced and the patient remained on support for six days. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reversible limb ischemia and the low pump flow has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. data logs show suction, position unknown alarms, and low pump flow. placement signal is trending downwards at the time of low flow. the root cause of the reversible limb ischemia issue was most likely use related due to leaving the peel-away sheath in place. the root cause of the low pump flow was most likely patient condition related. the failure modes will be monitored and trended.
33097 | 1220648-2023-04631 | 18023138 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18023138&pc=OZD | 09/28/2023 | 01/25/2024 |  | 10/27/2023 |  |  | 09/28/2023 | 10/27/2023 |  |  | 07/31/2025 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024380541 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Deformation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The u.s. complainant had impella 5.5 support ongoing for 12-days when there was an observation made of sterile sleeve damage. the team troubleshot the damage by the placement of tape. the pump support continued for 8 more days until the pump was weaned and explanted. | [5 empty] | The investigation into the product damage (sterile sleeve damage) has been completed. the device was not returned for benchtop investigation. according to the clinical information provided, the root cause of the sterile sleeve damage issue was not determined since product was not returned and sufficient clinical information was not provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
33098 | 1220648-2023-04634 | 18023221 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18023221&pc=OZD | 09/29/2023 | 02/07/2024 |  | 10/27/2023 |  |  | 09/29/2023 | 10/27/2023 |  |  | 07/31/2025 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024372378 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an ongoing purge flow increase/decrease advisory while on impella 5.5 pump support. the purge system was reviewed via impella connect and was found to have a consistent drop in purge pressure and increase in purge flow at the same time. the clinician discovered that the purge flow increase/decrease alerts was due to a kink in the catheter proximal to the red plug. | [5 empty] | The investigation has been completed. according to the clinical information, a purge pressure dipping every 5th cycle began to occur. no product was returned for a visual inspection. log analysis confirmed that when the purge cassette was connected an issue with the purge pressure dipping every drive cycle began. after switching this purge cassette with another purge cassette, the purge pressure dipping was resolved prior to an automated impella controller (aic) change. the cause of the priming issue was not determined since the product was not returned. the catheter kink is unrelated since the priming issue was resolved by replacing the purge cassette. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
33099 | 1220648-2023-04633 | 18023252 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18023252&pc=OZD | 09/29/2023 | 01/24/2024 |  | 10/27/2023 |  |  | 09/29/2023 | 10/27/2023 |  |  | 04/30/2025 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024330678 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Improper Flow or Infusion | [6 empty] | Hematoma | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported there was a hematoma at the access site and heparin was halted. additionally, the doctor mentioned they had to remove the impella 5.5 because of a clot ingestion and ordered heparin after washout of axillary pocket. the motor currents where high and the pump was explanted prior to any possible pump stop. the patient was then supported by a balloon pump (iabp). | [5 empty] | The investigation for the reported low pump flow, access site bleeding, and thrombus has been completed. the device and data logs were returned for evaluation. the root cause of the access site bleeding issue was not determined since insufficient clinical information was provided. the root cause of the saturated flow issue was biomaterial found in the outflow track and on the impellers. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33100 | 1220648-2023-04635 | 18023265 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18023265&pc=OZD | 09/29/2023 | 08/07/2024 |  | 10/27/2023 |  |  | 09/29/2023 | 10/27/2023 |  |  | 06/30/2025 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024361265 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 was inserted in the right axillary without difficulty but the 5.5 was angled towards the sub-valvular apparatus and was unable to be repositioned. after rotating and trying to transmit the torque, it was pulled back into the aorta and was unable to be re-advanced. a new impella 5.5 was therefore used, advanced, and placed successfully. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation for the positioning issue has been completed. according to the clinical, the pump was placed in the right axillary without difficulty. the pump was then angled towards the mitral valve and was unable to be repositioned. after rotating and trying to transmit the torque it was pulled back into the aorta and was unable to be re-advanced. the team didn¿t feel comfortable re-advancing the same device and decided to place a new 5.5. the new 5.5 was advanced and placed without issue. no product was returned for an investigation. data log analysis was not necessary for this complaint. the most likely cause of the positioning issue was cath anchor use related. the instructions for use referenced in the initial report is for the impella 5.5 with smartassist for use during cardiogenic shock from the following section: section: warnings & cautions: cautions: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ added- d6a/d6b.
33101 | 1220648-2023-04636 | 18023435 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18023435&pc=OZD | 03/27/2023 | 10/27/2023 |  | 10/27/2023 |  |  | 03/27/2023 | 10/27/2023 |  |  | 02/28/2025 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024308466 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that an impella 5.5 pump developed a purge leak during a 76-year-old female patient support. a purge bypass was performed and support continued with the implanted pump. there was no patient harm. | [5 empty] | (h3) the investigation into the reported ¿purge leak ¿ pump¿ has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the nurse called to let the team know they experienced a purge leak during the night. when they performed the purge sidearm bypass, the issue was resolved. the root cause of the purge leak is most likely due to a damaged sidearm, since the bypass resolved the issue. however, the location of the leak is unknown. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
33102 | 1220648-2023-04637 | 18023791 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18023791&pc=OZD | 10/10/2023 | 05/07/2024 |  | 10/27/2023 |  |  | 10/10/2023 | 10/27/2023 |  |  | 03/31/2025 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024346429 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Malposition of Device | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/10/2023 |  |  |  |  | The user facility reported a 62-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the impella pump was placed and started, then the device moved back across the aortic valve. the doctor attempted to reposition the device, but the catheter kinked and they were unable to reposition. the pump was explanted to address the issue and a second pump was placed successfully. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported product damage has been completed. the impella device was not returned for evaluation. the data logs were not relevant for this issue. however, clinical details state that shortly after beginning impella cp support, the physician accidentally pulled the pump back across the aortic valve. in an attempt to reposition the pump the cannula was also kinked. the decision was then made to explant the pump and replace it with another impella device. the most likely cause of the product issue was due to a positioning issue, which was use related.
33103 | 1220648-2023-04639 | 18024242 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18024242&pc=OZD | 09/05/2022 | 10/27/2023 |  | 10/27/2023 |  |  | 09/05/2022 | 10/27/2023 |  |  | 05/31/2024 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023220074 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/05/2022 |  |  |  |  | The user facility reported a 68-year-old male was implanted with an impella 5.5 device for mechanical circulatory support. on day 16 of support, the placement signal from the 5.5 pump was lost. the loss of the optical signal prompted the team to explant and replace the pump. it had been used beyond indication. no lasting harm or injury. | [5 empty] | The field lt log was reviewed and a placement signal not reliable alarm was confirmed to have triggered with a sudden ps spike to 3000 mmhg on 2022-sept-05 at 2:12pm. the characteristic of the placement signal metrics indicate a sensor wear. the device was not returned for evaluation. based on the provided information, the cause of the placement signal issue was most likely sensor silicone wear. this report is being filed as part of a retrospective review of historical records.
33104 | 1220648-2023-04638 | 18024283 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18024283&pc=OZD | 04/21/2022 | 10/27/2023 |  | 10/27/2023 |  |  | 04/21/2022 | 10/27/2023 |  |  | 01/31/2024 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023184036 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/21/2022 |  |  |  |  | The user facility reported a 66-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the us complainant had a 5.5 replaced due to lost position signal. there was no harm or injury from the interruption in the support. | [5 empty] | The data logs were returned and showed suction alarms. on the 2nd day, the p-level was lowered and the pump was disconnected while running. the pump was plugged back in and when plugged in, it produced placement signal not reliable alarms. the placement signal did not return for the remainder of the case. the product was returned and the placement signal not reliable alarm was reproduced. the functionality of the sensor was checked and found to be functioning properly. biomaterial was found in the outflow and wrapped around the impeller. the biomaterial was removed but the alarm persisted. the sensor was removed and examined and biomaterial was found on the surface of the sensor. the root cause of the placement signal issue was most likely biomaterial on the sensor face. this report is being filed as part of a retrospective review of historical records.
33105 | 1220648-2023-04640 | 18024418 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18024418&pc=OZD | 02/07/2023 | 10/27/2023 |  | 10/27/2023 |  |  | 02/07/2023 | 10/27/2023 |  |  | 08/31/2024 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 SMARTASSIST SET, JP |  |  | ABIOMED, INC. |  |  | US |  | 2023262826 |  | 1000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 27 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/07/2023 |  |  |  |  | The user facility reported a 27-year-old male patient was implanted with an impella 5.5 device for mechanical circulatory support. roughly 57 days after implant, the impella pump stopped functioning. the pump was explanted as a result of the noted failure. there was no patient harm resulting from the failure. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
33106 | 1220648-2023-04641 | 18024526 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18024526&pc=OZD | 06/02/2021 | 10/27/2023 |  | 10/27/2023 |  |  | 06/02/2021 | 10/27/2023 |  |  | 02/28/2023 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022054952 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Dislodged or Dislocated | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/02/2021 |  |  |  |  | The user facility reported a 65-year-old male patient was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that a patient on impella support, despite excellent placement and best practices in securing the device, had the pump migrate to the aorta allegedly due to patient's anatomy. as a consequence, the pump had to be explanted. there was no reported patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. based on the provided information, it was determined that the pump could not be re-advanced due to improper technique. it has been noted that the device should not be repositioned without a guidewire per ifu number 0550-9002. this report is being filed as part of a retrospective review of historical records.
33107 | 1220648-2023-04642 | 18024705 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18024705&pc=OZD | 02/24/2022 | 02/12/2024 |  | 10/27/2023 |  |  | 02/24/2022 | 10/27/2023 |  |  | 10/31/2023 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023159346 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 60-year-old male noted for high risk percutaneous cardiac insertion was implanted with an impella 5.5 device for mechanical circulatory support. during support, the a purge system blocked alarm was triggered. it was noted that the impella 5.5 was not secured to the patient with 3 point fixation. surgeons involved agreed that an exchange would be the best plan of action for the patient¿s care. the patient was taken to the or for insertion of a new impella 5.5. | [5 empty] | The investigation into the access site bleeding has been completed. the device was not returned for investigation. based on the clinical information, the cause of the high purge pressure was most likely due to a kinked purge line since kink in the catheter was observed upon removal of the device. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
33108 | 1220648-2023-04643 | 18024720 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18024720&pc=OZD | 10/13/2021 | 09/22/2024 |  | 10/27/2023 |  |  | 10/13/2021 | 10/27/2023 |  |  | 06/30/2023 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMART ASSIST SET, EU |  |  | ABIOMED INC. |  |  | US |  | 2022117232 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | No | Other | Foreign | Health Professional | User facility | Company representation |  |  | 33 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [5 empty] | ECMO | [18 empty] | A |  |  |  | Initial submission | Followup | Followup |  |  | 10/13/2021 |  |  |  |  | The complaiant in the eu had a impella 5.5 support ongoing for a patient admitted in cardiogenic shock. on day 6 of the support there were suction alarms and pump was malpositioned. the team did not explant the pump however as they were bridging to possible tranplant or lvad. | Survival outcome at explant noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The medical center in the eu did not return for analysis the impella pump product. upon completion of the investigation, a final report will be forthcoming. this complaint is being reopened as a part of a retrospective review. | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the positioning issue was use related, several attempts were made to turn the pump around in lv but no success as per the clinical description. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-04643 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
33109 | 1220648-2023-04644 | 18024813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18024813&pc=OZD | 05/01/2022 | 01/25/2024 |  | 10/27/2023 |  |  | 05/01/2022 | 10/27/2023 |  |  | 09/30/2023 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022151931 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 55-year-old male in st elevated myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. during support, the nurse noted bleeding at the access site and 100 mls were drained via jp drain. the patient was given 1 unit prbc and 2 units plt to treat the bleed. systemic heparin was stopped and purge heparin was restarted the following morning once the bleeding had resolved. support with the implanted pump was maintained. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleeding event has been completed. the device was not returned for evaluation. without sufficient clinical details, the root cause of the access site bleeding could not be determined.
33110 | 1220648-2023-04645 | 18024820 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18024820&pc=OZD | 12/09/2021 | 03/05/2025 |  | 10/27/2023 |  |  | 12/09/2021 | 10/27/2023 |  |  | 09/30/2023 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2022140623 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 18 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/09/2021 |  |  |  |  | The medical center had placed an impella 5.5 to support a younger patient admitted with cardiomyopathy. the pump support was ongoing for over a week when the team noted the patient was afebrile and acting septic and vasoplegic. the team infused antibiotics. the pump remained on for support and was explanted after 2 weeks. | [5 empty] | The medical center has not returned any pump product for benchtop analysis. root cause has not been determined. at conclusion of investigation a final report will be filed. this complaint is being reopened as part of a retrospective review. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
33111 | 1220648-2023-04646 | 18024863 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18024863&pc=OZD | 05/17/2022 | 01/15/2024 |  | 10/27/2023 |  |  | 05/17/2022 | 10/27/2023 |  |  | 08/31/2023 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022128761 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/17/2022 |  |  |  |  | The user facility reported a 49-year-old male patient was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that following implant of the impella 5.5 pump, the patient experienced persistent bleeding from the axillary site. a jackson-pratt drain was placed and heparin was removed from the purge to control the bleed. support continued with the implanted pump. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the reported access site bleeding has been completed since the original report was filed. no product was returned. the cause of the access site bleeding could not be determined since the product was not returned and enough clinical information was not provided.
33112 | 1220648-2023-04647 | 18024984 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18024984&pc=OZD | 05/23/2022 | 09/24/2025 |  | 10/27/2023 |  |  | 05/23/2022 | 10/27/2023 |  |  | 02/29/2024 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023195761 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/23/2022 |  |  |  |  | The user facility reported a 48-year-old male patient was implanted with an impella 5.5 device for mechanical circulatory support. the complainant noted that an impella pump experienced high purge pressure followed by saturated flow during patient support. the patient was tolerating lower support levels at the time of the failure and the decision was made to explant the pump. there was no patient harm. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. the root cause of saturated flow could not be determined since the product was not returned and insufficient clinical details were provided. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2023-04647. d6a and d6b revised from the initial manufacturer device report number 1220648-2023-04647 in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2023-04647. g1 reporting contact email revised on manufacturer device report 1220648-2023-04647 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2023-04647.
33113 | 1220648-2023-03703 | 18025027 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18025027&pc=OZD | 11/15/2022 | 01/26/2024 |  | 10/27/2023 |  |  | 11/15/2022 | 10/27/2023 |  |  | 09/30/2024 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023258452 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/15/2022 |  |  |  |  | The user facility reported a 46-year-old male was presented for impella 5.5 placement for mechanical circulatory support. the complainant reported that during support, a patient developed a large hematoma at their shoulder. the patient was taken to the operating room the following day to have the impella pocket evacuated. support continued following the intervention. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for investigation. per the clinical information provided, the root cause of the access site bleeding was unable to be determined as limited clinical details and no product were returned for investigation.
33114 | 1220648-2023-04648 | 18025031 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18025031&pc=OZD | 10/14/2022 | 10/06/2025 |  | 10/27/2023 |  |  | 10/14/2022 | 10/27/2023 |  |  |  |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Advance | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/14/2022 |  |  |  |  | The user facility reported a 55-year-old male patient was scheduled for impella 5.5 implant for mechanical circulatory support. the complainant reported failure to deliver the impella device due to patient's calcification. upon final failed attempt, the pump was removed and the anastomosis pulled off the site. arterial repair was completed to control the bleeding. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. the investigation of delivery issues and vascular damage - peripheral vessel has been completed. delivery issues: the root cause of the delivery issue was most likely patient condition related owing to the calcification and tortuosity of the patient's blood vessels. vascular damage - peripheral vessel: the root cause of the vascular damage - peripheral vessel in the left axillary artery was most likely use issue relating to the multiple attempts of trying to insert the impella device from the access site. b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2023-04648. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2023-04648. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2023-04648 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2023-04648 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2023-04648.
33115 | 1220648-2023-04649 | 18025042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18025042&pc=OZD | 05/17/2023 | 05/29/2025 |  | 10/27/2023 |  |  | 05/17/2023 | 10/27/2023 |  |  | 12/31/2023 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023177933 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pericardial Effusion | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/17/2023 |  |  |  |  | The user facility reported a 63-year-old male patient was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient developed a pericardial effusion after being escalated from an impella cp to a 5.5, but prior to the pci procedure. intervention was administered to extract the effusion and perform a pericardial window. it's unclear how the effusion originated, however, impella could not be ruled out. support continued with the implanted pump. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
33116 | 1220648-2023-04650 | 18025197 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18025197&pc=OZD | 10/01/2023 | 02/02/2024 |  | 10/27/2023 |  |  | 10/01/2023 | 10/27/2023 |  |  | 08/31/2025 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024387781 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Problem | Unintended Movement | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/01/2023 |  |  |  |  | The user facility reported a 36-year-old female with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, the patient coughed and an "impella in aorta" alarm occurred. the cardiologist attempted to reposition impella despite not being able to visualize the impella. the rn stated that the impella was pushed in 20cm. after which, the patient complained of ¿feeling funny.¿ a head and neck ct was performed confirming the impella was in the right carotid artery and ischemia to middle cerebral artery on perfusion study. the patient was brought to the or and the vascular surgeon snared the impella from lt groin access and impella was explanted. a cerebral angiogram was performed and showed no dissection or abnormal flows. | [5 empty] | The impella device was received from the customer and an evaluation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issues has been completed. the pump was returned for investigation. a catheter kink was observed, and the pump stop-motor current high alarm was appeared during the test. data logs indicate the only abnormality related to pump position in the aorta at the end of the case run. as per the provided clinical information, the impella position in the aorta alarm appeared while the patient was coughing. during repositioning, the doctor operated blindly and found the pump too deep into the ventricle. the cause of the positioning issue was use related since the pump was pulled out due to reported patient movement while coughing, based on the given clinical details and data log review.
33117 | 1220648-2023-04651 | 18025225 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18025225&pc=OZD | 10/11/2023 | 04/22/2024 |  | 10/27/2023 |  |  | 10/11/2023 | 10/27/2023 |  |  | 08/31/2025 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024385598 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Infusion or Flow Problem | Positioning Problem | [5 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/11/2023 |  |  |  |  | The user facility reported a 69-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. oozing was noted from the impella access site. an attempt at manual pressure was made without any improvement. another suture around the repo sheath was done to see if that would help the oozing but it was unsuccessful. an order was placed for a unit of fresh frozen plasma (ffp). the patient was already in the process of receiving a unit of paced red blood cells (prbc) for volume resuscitation. a fem-stop was placed over the right femoral site until the ffp could be transfused. it was further reported that the pump appeared to be too deep. the physician slowly pulled back on the impella catheter. after the catheter was pulled back, the aic displayed the ¿impella in aorta¿ alarm. without any adjustment, the alarm resolved on its own. shortly after arrival to the intensive care unit (icu) from cath lab, and before formal echo assessment of position, the patient went into ventricular fibrillation arrest. received one shock at 150 joules and returned to sinus rhythm. albumin and a unit of blood given for fluid resuscitation. in icu the clinician was unable to detect pedal pulses bilaterally. patient had impella cp placed in right femoral artery and va-ecmo in left femoral artery. it is unknow if the impella caused/contributed to the limb ischemia. the patient remained on support for seven more days. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported issues has been completed since the original report was submitted. the device was not returned for investigation. the cause of the access site bleeding issue was most likely patient condition related since bleeding was resolved after placing femstop. as all the necessary information and return product were not provided, the root cause of the limb ischemia and the vt/vf was not determined. the cause of the low pump flow issue was most likely patient condition related, based on data log review and provided clinical details. the cause of the positioning issue was most likely use related since ee revealed that the impella was found deep, and repositioning resolved the issue.
33118 | 1220648-2023-04652 | 18025239 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18025239&pc=OZD | 10/11/2023 | 05/07/2024 |  | 10/27/2023 |  |  | 10/11/2023 | 10/27/2023 |  |  | 08/31/2025 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024392318 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/11/2023 |  |  |  |  | The user facility reported a 56-year-old female with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the patient had a cp support ongoing for patient suffering from pneumonia and other comorbidities like covid and st elevation myocardial infarction( stemi). the pump was exhibiting "in aorta alarms" for mispositioning. the pump was adjusted, and support does continue to date. no patient harm was reported due to the issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed since the original report was submitted. the device was not returned by the user facility for investigation. the most likely cause of the positioning issues was use related.
33119 | 1220648-2023-04653 | 18025688 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18025688&pc=OZD | 04/04/2022 | 05/03/2024 |  | 10/27/2023 |  |  | 04/04/2022 | 10/27/2023 |  |  | 11/30/2023 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2023166331 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/04/2022 |  |  |  |  | The user facility reported a 65-year-old male patient was implanted with an impella 5.5 pump for mechanical circulatory support. following impella implant, it was noted that the lv signal was trending upwards and that signs of pump saturation were present. shortly after, high purge pressure alarms were triggered, and the decision was ultimately made to explant the pump. there was no patient harm. | [5 empty] | The impella device was not returned by the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported saturated flow issue has been completed. the impella was not returned for evaluation. the data logs were reviewed which revealed high motor currents (mc). the root cause of the saturated flow was most likely biomaterial since high mc is observed in the logs several times. added- d6a/d6b, d4 updated model number.
33120 | 1220648-2023-04654 | 18026031 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18026031&pc=OZD | 03/26/2022 | 03/10/2025 |  | 10/27/2023 |  |  | 03/26/2022 | 10/27/2023 |  |  | 12/31/2023 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC |  |  | US |  | 2023177831 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/26/2022 |  |  |  |  | The user facility reported a 48-year-old male patient was implanted with an impella 5.5 pump for mechanical circulatory support. roughly one day following implant, a left atrial thrombus was discovered. the patient was taken to the operating room where the thrombus was surgically removed. support with the implanted pump continued following the intervention. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | As the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined".
33121 | 1220648-2023-04655 | 18026132 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18026132&pc=OZD | 12/13/2021 | 05/29/2024 |  | 10/27/2023 |  |  | 12/13/2021 | 10/27/2023 |  |  | 07/31/2023 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED INC. |  |  | US |  | 2022112652 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/13/2021 |  |  |  |  | The medical team placed an impella 5.5 for support of a patient admitted with angina and found to have multivessel coronary artery disease. a bypass (cabg) procedure was planned with the impella providing hemodynamic support post-operatively. the team found the patient to have bleeding that prompted intervention in the form of 7 units of red blood cells, 7 units of ffp, and 3 platelets. the location of the blood loss was unknown, and as such relationship to the use of impella causing or contributing could not be determined. the pump remained on support for 2 days. | There is not enough information to associate use of device with outcome. |  |  |  |  | The medical center has not returned any product for investigation. when the investigation is completed, a final report will be filed. of note the ifu states: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury." | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding issue was patient condition related as there was continuous bleeding at the access site and blood products were also given as per the clinical description provided.
33122 | 1220648-2023-04656 | 18026188 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18026188&pc=OZD | 02/22/2022 | 01/25/2024 |  | 10/27/2023 |  |  | 02/22/2022 | 10/27/2023 |  |  | 11/30/2023 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC |  |  | US |  | 2023168265 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/22/2022 |  |  |  |  | The user facility reported a 73-year-old male patient was implanted with an impella 5.5 pump for mechanical circulatory support. during and following implant, the patient experienced oozing from their access site. multiple blood products were given to counteract the oozing. support was able to continue with the implanted pump; however, heparin was not introduced into the purge due to the bleed. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
33123 | 1220648-2023-04657 | 18026389 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18026389&pc=PYX | 08/26/2023 | 01/19/2024 |  | 10/27/2023 |  |  | 09/28/2023 | 10/27/2023 |  |  | 02/28/2025 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | RP FLEX W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2024348408 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 62 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/28/2023 |  |  |  |  | The complainant reported a 62 year old female patient presenting for impella support post-op lvad. during support, hemolysis developed and was confirmed by plasma-free hemoglobin testing. in response, multiple blood products were delivered. | [5 empty] | The impella device was not received by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the hemolysis issue has been completed. the root cause of the hemolysis cannot be determined as there was insufficient clinical information and no product returned. the data logs do not show an indicative motor current spike or positioning issues. the relationship between the hemolysis and the impella is unknown as it cannot be confirmed that it occurred during support, resolved with removal or a reduction in p-level. the instructions for use states the following: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿
33124 | 1220648-2023-04659 | 18026448 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18026448&pc=OZD | 09/27/2023 | 02/02/2024 |  | 10/27/2023 |  |  | 09/27/2023 | 10/27/2023 |  |  | 07/31/2025 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024385814 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported patient came in for impella cp assisted chronic total occlusion (cto). thrombosis and vessel damage were when the 14fr sheath was removed. the team therefore performed balloon tamponade and stenting to treat the vessel dissection. the team made decisions after stenting of what to do for the vessel thrombosis. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported vascular damage and thrombus has been completed since the original report was filed. product was not returned for review. based on clinical description, the root cause of vascular damage was most likely due to use related. the root cause of thrombus was most likely was due to patient condition.
33125 | 1220648-2023-04658 | 18026450 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18026450&pc=OZD | 09/27/2023 | 02/05/2024 |  | 10/27/2023 |  |  | 09/27/2023 | 10/27/2023 |  |  |  |  | 10/27/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1167410 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Nurse | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had an automated impella controller (aic) with purge disc failure/breakage observed. the purge disc was not reading. there was no patient use or harm. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation into the console "purge disc/flag" issues has been completed since the original report was filed. the console was returned for investigation. the console was evaluated finding that the purge disc retainer was confirmed to be broken. the cause of the issue was a defective component.
33126 | 1220648-2023-04660 | 18026451 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18026451&pc=OZD | 09/27/2023 | 02/08/2024 |  | 10/27/2023 |  |  | 09/27/2023 | 10/27/2023 |  |  | 04/30/2025 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024355671 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | Company representation |  |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/27/2023 |  |  |  |  | The user facility reported a 62-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced oozing from the impella site. manual pressure was held, and hemostasis achieved. two units ofpacked red blood cells were given for hemoglobin (hgb) of 5. prior to the procedure, the patient¿s hgb was 7. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
33127 | 1220648-2023-04661 | 18026452 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18026452&pc=OZD | 09/27/2023 | 01/15/2024 |  | 10/27/2023 |  |  | 09/27/2023 | 10/27/2023 |  |  | 05/31/2025 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024371338 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had an impella cp placed for hemodynamic support. the support was for under 24 hours when the pump was explanted. the pump was explanted after there were access site bleed/hematoma at the right internal jugular (ij) and impella access sites. the team treated by removal of heparin and pump explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the root cause of the access site bleeding issue was not determined since the product was not returned and insufficient clinical details were provided.
33128 | 1220648-2023-04663 | 18026453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18026453&pc=OZD | 09/28/2023 | 10/19/2024 |  | 10/27/2023 |  |  | 09/28/2023 | 10/27/2023 |  |  |  |  | 10/27/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1131452 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Alarm Not Visible | Optical Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that their automated impella controller (aic) failed to display a placement signal and began to alarm with a 'placement signal not reliable' alarm. patient support continued with no know patient harm or injury. upon evaluation of the aic, it was determined that the optical signal issue was due to an eeprom corruption issue that is unable to be correlated with the pump or aic. | [5 empty] | The investigation into the reported optical issues has been completed. console logs from the reported day of event are consistent with complaint and show that after pump was connected to console, the pump eeprom usage data area 1 and 2 could not be correctly read (due to crc errors). this resulted in the console¿s software not recognizing the pump as previously used, which caused the missing prompt to resume pump at previous speed (with the same p-level as the last p-level on prior console). because entire usage data was not read correctly, the erroneous value of 0 was used for calibrated g0, which resulted in placement signal out of range, due to incorrect scaling. the pump was unplugged and re-connected in an attempt to resolve the issue, but prior to that, the usage data area 1 and 2 sections were overwritten by aic software with erroneous value of g0 but correct (matching) values of crc, so that after the pump was reconnected, there were no issues with reading its eeeprom. nevertheless, the placement signal was still out of range (as indicated by alarm #1048 ¿placement signal not reliable¿ that followed), due to pump being recognized as ¿previously initialized¿ and having calibrated value of g0=0. the case continued, with uninterrupted hemodynamic support, although without placement signal, until its completion. log analysis indicated that during eeprom reading issues, consistently repeating mismatch between read and calculated values of crc indicates that incorrect values were not the result of communication issues during eeprom reading but were recorded as such in the eeprom. therefore, this issue might be related to the prior console, on which the pump¿s eeprom had been written to. the console was returned and extensively tested in the lab: entire communication path between pump¿s red handle and console¿s epc/4-in-1 carrier was monitored, and data was recorded and decoded (i2c between pump and impellatrnic and rs232 between impellatronic and 4-in-1). no issues with eeprom writing resulting in crc mismatch were observed, however some irregularities were observed. by externally corrupting (mis-programming) a test pump, so that crc values would not match the usage data area values, we were able to reproduce scenario of console being unable to recognize pump as previously used and assuming incorrect (0) value of g0. this scenario was not observed occurring on its own, despite multiple repetitions of the test. the exact cause of the issue (and whether it is due to pump or console), although known to be related to the mechanism of writing to pump¿s eeprom (inside the red handle), could not be determined because the issue could not be reproduced, and the pump used was not returned for analysis. the failure modes will be monitored and trended. | Corrected data: h6.
33129 | 1220648-2023-04662 | 18026455 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18026455&pc=OZD | 09/27/2023 | 02/16/2024 |  | 10/27/2023 |  |  | 09/27/2023 | 10/27/2023 |  |  | 05/31/2025 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024356626 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that patient developed a hematoma and impella cp access site was still oozing around the sheath. forward tension sutured deployed, redressed by maintaining angle of insertion site, and manual pressure was applied for 30 minutes. bleeding stopped and hematoma was managed. two days post implant, the patient was successfully weaned off the impella cp. the device was explanted and the patient is noted to be doing well. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed. the device was not returned for investigation. based on clinical description, the root cause of access site bleeding was most likely due to lack of vessel recoil as the bleeding occurred after switching to the gwru.
33130 | 1220648-2023-04664 | 18026456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18026456&pc=OZD | 09/28/2023 | 10/27/2023 |  | 10/27/2023 |  |  | 09/28/2023 | 10/27/2023 |  |  |  |  | 10/27/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | ABIOMED, INC. |  |  | US |  | 1147285 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Communication or Transmission Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 09/28/2023 |  |  |  |  | The user facility returned an automated impella controller for investigation of a connectivity/wi/fi issue. during the investigation of the device, a controller error (red alarm) displayed. there was no patient involvement. | [5 empty] | The investigation into the connectivity issue/wi-fi issue has been completed. data analysis: ic2823 received frequent errors on 7/14 related to the purge assembly for "controller failure (purge pressure sensor defective)" due to a pressure sensing voltage out of range. device analysis: the failure mode was able to be reproduced during investigation. there was a severely kinked ribbon cable within the purge assembly. replacing the purge pressure sensor resolved the controller failure alarm. root cause: the cause of the controller failure was due to a defective purge pressure sensor with a kinked ribbon cable. the failure modes will be monitored and trended.
33131 | 1220648-2023-04667 | 18026457 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18026457&pc=OZD | 09/28/2023 | 01/25/2024 |  | 10/27/2023 |  |  | 09/28/2023 | 10/27/2023 |  |  |  |  | 10/27/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1193365 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Alarm System | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/28/2023 |  |  |  |  | The user facility reported a 69-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. during support, the automated impella controller (aic) showed a controller error message. the aic would not display the placement signal or left ventricle (lv) waveform. the motor current showed normal function and flows were appropriate for p level. after switching aics, the placement signal and lv waveform showed normal function. there was no patient harm related to this event. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - controller error (yellow alarm) has been completed since the original report was filed. the console was returned for evaluation and was tested after arriving in fs. the root cause of the controller error was due to an aic software issue. the failure mode will be monitored and trended.
33132 | 1220648-2023-04665 | 18026458 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18026458&pc=OZD | 09/28/2023 | 10/30/2025 |  | 10/27/2023 |  |  | 09/28/2023 | 10/27/2023 |  |  | 07/31/2025 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024389423 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 77-year-old male was implanted with an impella cp for mechanical circulatory support to vent. the impella was placed through two edwards tavr valves. the original valve was snared and pulled into the annulus upside down. this meant to valve skirt was facing the wrong direction. the new valve was deployed through the upside down valve. the impella pump stopped and there was a moto current high alarm that occurred. the impella was attempted to re-started but was unsuccessful. therefore, decision was made to place new pump and console. there was no reported patient harm. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
33133 | 1220648-2023-04666 | 18026459 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18026459&pc=OZD | 09/28/2023 | 10/27/2023 |  | 10/27/2023 |  |  | 09/28/2023 | 10/27/2023 |  |  | 05/31/2025 |  |  |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024352069 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | [45 empty] | Invalid/Incomplete |  |  |  | Initial submission
33134 | 1220648-2023-04671 | 18026460 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18026460&pc=OZD | 09/29/2023 | 01/12/2024 |  | 10/27/2023 |  |  | 09/29/2023 | 10/27/2023 |  |  | 02/28/2025 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024321937 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Improper Flow or Infusion | Positioning Problem | [5 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/29/2023 |  |  |  |  | The user facility reported a 45-year-old female who underwent cardiac surgery and had an intra-aortic balloon pump implanted was also implanted with an impella cp device for mechanical circulatory support. after the peel-away was removed and repo sheath was being advanced, cannula moved forward deep into the patient's left ventricle (lv). despite repositioning and volume, there was continuous suction at all p-levels attempted. due to very delayed acute myocardial infarction presentation and numerous resuscitations, they suspected a thrombus ingestion. additionally, pulses in the left leg did not feel adequate. 5fr sheath was placed in left femoral artery and side port to side port hook up with 11fr iabp sheath. pulses were present via doppler. the impella cp was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported low pump flow, limb ischemia and positioning issue has been completed since the original report was filed. no product was returned. as per clinical pump cannula moved deep into the lv during repositioning sheath exchange and patient experienced vt. there were continuous suction alarms and ventricle position alarms at all p-levels attempted. pump flows were low at 0.7 l/min on p-2 and labs were hemolyzed. also patient did not have sufficient pulses in left leg and side port to side port bypass done. data logs were reviewed and there were suction alarms followed by low flows during entire case with placement signal trending down. improper positioning of pump was seen with several ventricle position alarms. the cause of the low pump flow issue was most likely patient condition related with suction alarm observed with low flows and down trending placement signal per log and clinical review. the cause of the limb ischemia issue was most likely patient condition related per clinical review.
33135 | 1220648-2023-04668 | 18026461 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18026461&pc=OZD | 09/28/2023 | 01/24/2024 |  | 10/27/2023 |  |  | 09/28/2023 | 10/27/2023 |  |  | 05/31/2025 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024366830 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/28/2023 |  |  |  |  | The user facility reported a 71-year-old female in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient was noted as high risk percutaneous intervention. while on support, the nurse changed the bed sheets and a hematoma developed in the access site. impella sheath appeared more up and down. initial attempts looked to have resolved the hematoma. however, shortly after hematoma resurfaced and the patient vagal down and became hypotensive. fluids started and pressure rebounded. the physician decided that the impella needed to be removed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access site bleeding was most likely due to patient movement during bed sheet change.
33136 | 1220648-2023-04670 | 18026462 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18026462&pc=PYX | 09/28/2023 | 10/27/2023 |  | 10/27/2023 |  |  | 09/28/2023 | 10/27/2023 |  |  | 08/31/2024 |  |  |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024289354 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | [45 empty] | Invalid/Incomplete |  |  |  | Initial submission
33137 | 1220648-2023-04669 | 18026463 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18026463&pc=OZD | 09/28/2023 | 10/27/2023 |  | 10/27/2023 |  |  | 09/28/2023 | 10/27/2023 |  |  | 01/31/2025 |  |  |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024310018 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Injury | Yes | No | [45 empty] | Invalid/Incomplete |  |  |  | Initial submission
33138 | 1220648-2023-04672 | 18026464 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18026464&pc=OZD | 09/29/2023 | 02/02/2024 |  | 10/27/2023 |  |  | 09/29/2023 | 10/27/2023 |  |  | 05/31/2025 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024367603 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 61-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the purge tubing was changed with minimal improvement in purge flows. inability to palpate distal pulses in impella leg was noted. the patient¿s foot was cool and cap refill was sluggish. the patient was therefore taken to the ccl for fem to fem bypass. there were lots of clot/thrombus noted in right leg below the impella site. fem to fem bypass was successful and the impella was successfully weaned and explanted a day post implant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported limb ischemia and thrombus has been completed since the original report was filed. product was not returned for review. based on clinical description, the root cause of limb ischemia and thrombus was most likely due to patient condition.
33139 | 1220648-2023-04673 | 18026465 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18026465&pc=OZD | 09/29/2023 | 02/06/2024 |  | 10/27/2023 |  |  | 09/29/2023 | 10/27/2023 |  |  | 06/30/2025 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024369055 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/29/2023 |  |  |  |  | The user facility reported a 50-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the pump was being weaned for explant when the impella cp motor stopped. the clinician attempted to restart the impella at p2 but was unsuccessful. the impella was explanted. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the reported pump stop has been completed since the original report was filed. according to the clinical, a pump stop occurred on the fourth day of treatment. multiple restart attempts and an aic change did not resolve the issue. product evaluation electrical testing confirmed one open motor cable line. upon further inspection the white motor cable line appeared to be kinked and showed evidence of necking. the motor cable damage was located within the red handle on the catheter side. log analysis confirmed alarm 115 and an immediate pump stop with no prior motor current spikes. the cause of the pump stop is an open electrical line. most likely due to excessive force as there were signs of necking.
33140 | 1220648-2023-04674 | 18026466 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18026466&pc=OZD | 09/29/2023 | 02/02/2024 |  | 10/27/2023 |  |  | 09/29/2023 | 10/27/2023 |  |  | 05/31/2025 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024368909 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 64-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that a patient on impella support was bleeding from the access site and had to have the pump explanted as a consequence. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. therefore, the root cause of the reported access site bleed was due to patient movement as the patient had been moving spontaneously.
33141 | 1220648-2023-04675 | 18026468 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18026468&pc=OZD | 09/29/2023 | 05/02/2024 |  | 10/27/2023 |  |  | 09/29/2023 | 10/27/2023 |  |  | 06/30/2025 |  | 10/27/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024371835 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Ischemia | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 73-year-old male was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient had periods of non-pulsatility. additionally, there was a lot of bleeding from the impella access site. a femostop placed and one unit of packed red blood cells (prbc's) was transfused. furthermore, waveforms and imaging suggested the need for repositioning. impella repositioned with fluro, and slack was removed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported limb ischemia, positioning issue, and the access site bleed has been completed. according to the clinical, shortly after beginning impella cp support the patient had periods of non-pulsatility. additionally, bleeding was observed coming from the impella access site after being transferred to the icu. after resolving the bleeding issue with femstop and blood products, suction was recorded and imaging confirmed the pump was shallow. following a successful repositioning procedure, the pump remained in the patient until they were weaned off of support. no product was returned for a visual inspection. data log analysis was not necessary for this complaint. the cause of the major access site bleeding was not determined because no product was returned, and not enough clinical information was given. the most likely cause of the positioning issues was use related. the cause of the limb ischemia could not be determined since no additional clinical details were provided. potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added-b1-selected product problem. to conform to updated procedures, sections d6 and h10 were revised with updated information.
33142 | 1220648-2023-04587 | 18010720 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18010720&pc=OZD | 09/26/2023 | 10/26/2023 |  | 10/26/2023 |  |  | 09/26/2023 | 10/26/2023 |  |  | 06/30/2025 |  | 10/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024377349 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Insufficient Flow or Under Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The us complainant had an impella cp support ongoing with a patient with known corevalve. the md was aware and did discuss the placement of the impella was off label through the corevalve. the pump was positioned and was supporting when there was a coronary percutaneous coronary intervention (pci) issue that prompted the patient to be sent urgently to the or. the pump lost appropriate flows and upon explant and the team noted pump damage at the impeller motor/blades. there were no biomaterial, lv thrombus, or cannula kinks observed. the patient's act was reported to be between 160-180 seconds and no blood volume was given. | [5 empty] | The investigation into the reported low pump flow issues has been completed. upon investigation of the pump, impeller damage was observed. there was no biomaterial, or any other abnormalities found. the root cause of the low pump flow issue was determined to be fractured impeller blades. the fracture was characteristic of an interaction with another device, such as the medtronic tavr/corevalve used in this case.
33143 | 1220648-2023-04588 | 18010739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18010739&pc=OZD | 09/26/2023 | 02/21/2024 |  | 10/26/2023 |  |  | 09/26/2023 | 10/26/2023 |  |  | 05/31/2025 |  | 10/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  |  |  | 2024360958 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had an impella cp support for 3 weeks when the pump stopped. the team was aware the pump had been used beyond indication. they attempted to troubleshoot by replacement of the automated impella controller (aic). this did not allow for a successful re-start and so the pump was removed and replaced in the or. | [5 empty] | The investigation into the pump stop issue has been completed. the device was not returned for investigation. the root cause of the pump stop was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
33144 | 1220648-2023-04589 | 18010798 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18010798&pc=OZD | 09/26/2023 | 02/19/2024 |  | 10/26/2023 |  |  | 09/26/2023 | 10/26/2023 |  |  | 08/31/2025 |  | 10/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024385518 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant in japan reported that during a ct imaging study, the patient moved, the impella 5.5 malpositioned out of the left ventricle. the patient was taken to the catheterization lab for repositioning. impella repositioning was successful and no patient harm was reported. | [5 empty] | The investigation into the positioning issues has been completed. the device and data were not returned for investigation. based on clinical description, the root cause of positioning issue was due to patient movement. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿use fluoroscopy for placement impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), can help confirm the position of the impella catheter after placement, tee does not allow visualization of the entire catheter assembly and is inadequate for reliably placing the impella catheter across the aortic valve¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire¿.
33145 | 1220648-2023-04590 | 18010817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18010817&pc=OZD | 09/26/2023 | 11/14/2025 |  | 10/26/2023 |  |  | 09/26/2023 | 10/26/2023 |  |  |  |  | 10/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1112884 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The us based hospital biomed department noted a red battery alarm on a console. there was no patient harm or injury as there was no noted patient use at time of observation. | [5 empty] | The impella console device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: imc logs confirm multiple instances of battery failure alarm; [26/09/23 08:09:11 imcgui: event set: #360 from 17 (0)]. it is plausible that the battery transport was engaged and reengaged within a short period leading to the aforementioned alarm. lt logs reveal no anomaly with the battery of the console. in the lab, batttest and logs revealed console operating within spec. device analysis: console was received at fs danvers for further investigation. running the console on a known good pump for 5 minutes didn't recreate the alarm described. the root cause was most likely disengaged/reengaged battery transport switch. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d2 common device name updated.
33146 | 18010836 | 18010836 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18010836&pc=OZD | 10/07/2023 | 10/17/2023 | 10/17/2023 | 10/26/2023 | 10/17/2023 | 10/26/2023 |  | 10/26/2023 |  |  |  |  | 10/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | Yes |  |  | Not Available |  |  | No | Health Professional |  |  | User Facility report | No | Risk Manager | [6 empty] | 22265 DA |  | Male | 90 KG |  | Black Or African American | Malfunction | No | Yes | [7 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Patient was scheduled for a percutaneous left ventricular assistive device (lvad) insertion. equipment malfunction in the motor of the impella device (lvad device) 2 hours after insertion. the patient returned to the cath lab and a new device was inserted successfully.
33147 | 1220648-2023-04591 | 18011533 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18011533&pc=OZD | 09/28/2023 | 10/26/2023 |  | 10/26/2023 |  |  | 09/28/2023 | 10/26/2023 |  |  |  |  | 10/26/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1148833 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 09/28/2023 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) had a red alarm indicating battery failure. there was no patient involvement. | [5 empty] | The investigation into the aic battery failure-red alarm has been completed. there was a battery failure alarm due to low pack voltage. additionally, the lt logs showed that battery 1 had a cell imbalance on cell 4. the failure mode was reproduced on an engineering bench. the battery 1 pack voltage was measured to be 0 v rather than the expected 16 v, and the battery lcd indicator was blank (fully discharged). battest confirmed battery 1 cell 4 was outputting 200mv which triggered the cell imbalance condition. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure.
33148 | 1220648-2023-04592 | 18012063 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18012063&pc=OZD | 10/09/2023 | 02/05/2024 |  | 10/26/2023 |  |  | 10/09/2023 | 10/26/2023 |  |  | 07/31/2025 |  | 10/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024380095 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/09/2023 |  |  |  |  | The user facility reported a 59-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient had a cp support ongoing in which despite all troubleshooting the pump would not achieve appropriate flows. the pump had suction and upon pump explantation no thrombus was observed. | [5 empty] | The investigation into the low pump flow issue has been completed. the device was returned for investigation. the returned complaint cp pump was visually inspected. no cannula kink or broken blades observed. no biomaterial observed. complaint pump run for 30 minutes in wat in the investigation lab. flows were in specification range at any p-level. no flow issues. no alarm triggered. no other abnormality observed. data log reviewed. suction alarm occurred often but no indication for issue cause. the cause of the suction alarm most likely was patient condition related. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a [supplemental/final] report will be filed.
33149 | 1220648-2023-04593 | 18012097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18012097&pc=OZD | 10/09/2023 | 02/02/2024 |  | 10/26/2023 |  |  | 10/09/2023 | 10/26/2023 |  |  | 05/31/2025 |  | 10/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024355605 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/09/2023 |  |  |  |  | The user facility reported a 61-year-old male was implanted with an impella 5.5 for mechanical circulatory support prior to a high risk surgery. the support was ongoing for 16 days when the team observed suction alarms. the alarms were present when they attempted to adjust p levels. the pump remained on for support despite the malfunction without any harm or injury noted to patient. | [5 empty] | The investigation for the reported low flow issue has been completed. the data logs only were returned for evaluation and they confirmed low flow and multiple suction alarms when running on p-9 during support. however, without the device return, the root cause of the low flow issue could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. when the investigation concludes a supplemental mdr will be filed.
33150 | 1220648-2023-04594 | 18012271 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18012271&pc=OZD | 06/27/2023 | 10/21/2025 |  | 10/26/2023 |  |  | 09/26/2023 | 10/26/2023 |  |  |  |  | 10/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1245285 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 09/26/2023 |  |  |  |  | The user facility reported a 73-year-old male post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. during support, the automated impella controller (aic) displayed an error message stating "pump not recognized by the aic." a second impella 5.5 was subsequently utilized for patient support. upon further investigation, a review of the data logs discovered that the failure was caused by a known issue with the aic console. there was no patient harm related to this event. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. corrected information for the initial mfr 1220648-2023-04594 was provided in sections: a1 patient identifier added. d1/d2 common device name corrected. d4 serial number, lot number, and model number corrected. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
33151 | 1220648-2023-04595 | 18013423 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18013423&pc=OZD | 09/25/2023 | 02/07/2024 |  | 10/26/2023 |  |  | 09/26/2023 | 10/26/2023 |  |  | 11/30/2024 |  | 10/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023282037 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [11 empty] | Death | [5 empty] | ECMO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/26/2023 |  |  |  |  | The user facility in japan reported a male of unknown age, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was on ecpella support for 161 hours before the patient expired. patient had a left ventricle thrombus and an extensive cerebral infraction occurred. it is unknown when the thrombus occurred. the relationship between the impella and cause of death is unknown. there were no impella malfunctions that occurred during support. | [5 empty] | The investigation has been completed since the original report was filed. the medical center did not return the impella device. data logs were not provided. based on clinical findings, the doctor found a ventricular thrombus during treatment. the patient died from a cerebral infection. the cause of the stroke issue was most likely biomaterial but its relation to impella remained unknown due to insufficient clinical information. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33152 | 1220648-2023-04596 | 18013500 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18013500&pc=OZD | 09/26/2023 | 08/15/2025 |  | 10/26/2023 |  |  | 09/26/2023 | 10/26/2023 |  |  |  |  | 10/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1103720 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  |  | Black Or African American | Malfunction | No | Yes | Application Program Freezes, Becomes Nonfunctional | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) froze. none of the soft keys worked. the aic was replaced. there was no reported patient harm. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | D.1 and d.2 brand name and common device name were revised in accordance with updated procedures since the initial manufacturer device report number 1220648-2023-04596 was submitted. d.4 model number and serial number were revised in accordance with updated procedures since the initial manufacturer device report number 1220648-2023-04596 was submitted. g.1 revised mdr reporting contact email since the initial submission of manufacturer device report number 1220648-2023-04596 in accordance with updated procedures. added reporting contact fax number in accordance with updated procedures since the initial manufacturer device report number 1220648-2023-04596 was submitted. h.3 revised to device evaluation anticipated as it was reported incorrectly on initial manufacturer device report number 1220648-2023-04596. h.6 codes 2199 and 3221 were reported incorrectly on the initial manufacturer device report number 1220648-2023-04596 in accordance with updated procedures.
33153 | 1220648-2023-04597 | 18013573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18013573&pc=OZD | 09/26/2023 | 08/20/2024 |  | 10/26/2023 |  |  | 09/26/2023 | 10/26/2023 |  |  | 07/31/2024 |  | 10/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023251878 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Perforation of Vessels | Thrombosis/Thrombus | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/26/2023 |  |  |  |  | The updated outcome of the patient was death. per medical safety officer review use of the device cannot conclusively be associated with the outcome. | The user facility reported a 63-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that a patient developed a mottled left leg and had oozing from their impella access site during support. a femostop was placed to treat the oozing and pump escalation was decided upon due to the ischemia. upon removal of the cp pump, the percloses did not take and a cutdown revealed four moderate sized clots. the clots were removed and a perforation was then noted at the site. a covered stent was placed to treat the perforation and the site was closed. flow was restored to the leg following the intervention. |  |  |  |  | The investigation of limb ischemia and vascular damage has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ b5 updated with additional information h6 component code added the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-04597. d4 model number in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8-should have been left blank. g1 reporting contact fax number missing. h.6 code 1888 and 4440 were reported incorrectly. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
33154 | 1220648-2023-04598 | 18013760 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18013760&pc=OZD | 09/26/2023 | 02/26/2024 |  | 10/26/2023 |  |  | 09/26/2023 | 10/26/2023 |  |  | 07/31/2025 |  | 10/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024377856 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 79 year old patient was implanted with an impella cp pump for mechanical circulatory support. following implant, it was documented that the patient had no flow distal of the access site. a balloon catheter was swapped with a glidecath for better imaging and the flow was said to have been restored. support continued with the implanted pump. | [5 empty] | The investigation into the limb ischemia ¿ reversible issue has been completed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33155 | 1220648-2023-04600 | 18013786 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18013786&pc=OZD | 09/28/2023 | 02/02/2024 |  | 10/26/2023 |  |  | 09/28/2023 | 10/26/2023 |  |  | 01/31/2025 |  | 10/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024326542 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Frayed | Detachment of Device or Device Component | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/28/2023 |  |  |  |  | The user facility reported a 78-year-old male with heart failure and on an intra-aortic balloon pump was implanted with an impella cp for mechanical circulatory support. when nurse was assessing the groin, the impella was unlocked and the sterile sleeve was ripped. the lock was locked and a tegederm was placed and the team was made aware. decision was made to escalate to impella 5.5, however, the medical team wanted to wait a little longer due to pending labs. the patient continued to be supported with the impella cp for three additional days and was successfully weaned and transferred to cleveland clinic. | [5 empty] | The investigation into the product damage (sterile sleeve damage) has been completed since the original report was filed. the device was not returned for investigation. the root cause of the sterile sleeve damage was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
33156 | 1220648-2023-04602 | 18015229 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18015229&pc=OZD | 09/29/2023 | 03/20/2024 |  | 10/26/2023 |  |  | 09/29/2023 | 10/26/2023 |  |  | 02/28/2025 |  | 10/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2024325834 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/29/2023 |  |  |  |  | The user facility in japan reported a 66-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. on the 16th day of support, the team planned to explant the impella cp due to the long duration of use. the pump stopped, prompting the team to explant the pump. the patient's vitals were normal thereafter. | [5 empty] | The investigation into the pump stop issue has been completed since the original report was filed. the device was not returned for investigation. the cause of the pump stop issue was not determined since product was not returned and sufficient clinical information was not provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
33157 | 1220648-2023-04599 | 18015535 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18015535&pc=OZD | 09/26/2023 | 02/28/2024 |  | 10/26/2023 |  |  | 09/26/2023 | 10/26/2023 |  |  | 06/30/2025 |  | 10/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024375550 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/26/2023 |  |  |  |  | The user facility reported a 64-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support post coronary artery bypass grafting and mitral valve repair. the pump was explanted after 2 days and care was escalated to a 5.5 pump. the cp limb had signs of limb ischemia and was treated by explant and fasciotomy due to compartment syndrome. | [5 empty] | The investigation into the limb limb ischemia ¿ irreversible has been completed. the device and data were not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33158 | 1220648-2023-04601 | 18015950 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18015950&pc=OZD | 09/28/2023 | 01/24/2024 |  | 10/26/2023 |  |  |  | 10/26/2023 |  |  |  |  | 10/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1219727 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Failure of Device to Self-Test | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 09/28/2023 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) failed the system self-check. the aic will be returned for investigation. there was no patient involvement. | [5 empty] | The investigation for the reported console issue has been completed. the console and the data logs were returned for evaluation. the log file and system configuration file contained data that was directly indicative of a software upgrade failure, specifically indicating a failure of communication between the 4-in-1 and the epm of the impellatronic board. the system self-check failure was confirmed upon bootup. firmware check on the console¿s subsystems revealed that the msc1210 of the impellatronic¿s eep circuitry was unable to communicate with the 4-in-1 and the eep msc1210 were monitored revealing that there was no communication on the signal one from the 4-in-1 to the eep msc1210. the root cause of this issue was identified as a faulty esd suppression diode on the impellatronic board that was causing a short between the signal line and ground, causing electrical activity on the signal line to be suppressed and preventing communication between the 4-in-1 and the eep msc 1210. r215.1 was removed and resolved the issue. the cause of the aic issue was a defective impellatronic board, which prevented communication between the 4-in-1 and the epm on the impellatronic board during the software upgrade process, causing the software upgrade to fail. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
33159 | 1220648-2023-04603 | 18016036 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18016036&pc=OZD | 09/27/2023 | 02/06/2024 |  | 10/26/2023 |  |  | 09/27/2023 | 10/26/2023 |  |  | 06/30/2025 |  | 10/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024379661 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported there were access site bleeding complications with the impella cp. a femstop was applied and 4 units of packed red blood cells (prbcs) were given for hemoglobin of 7. the patient was successfully weaned off the impella cp a day post implant. | [5 empty] | The investigation for the reported access site bleeding complications has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine a root cause. the most likely cause of the reported access site bleeding was most likely patient condition related since bleeding was noted from multiple access sites. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33160 | 1220648-2023-04605 | 18016072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18016072&pc=OZD | 09/27/2023 | 02/02/2024 |  | 10/26/2023 |  |  | 09/27/2023 | 10/26/2023 |  |  | 06/30/2025 |  | 10/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024377969 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had placed an impella cp to help with left ventricle (lv) offload. the patient was therapeutically anticoagulated but, the pump had issues with flow. the pump was noted to have migrated further into the ventricle. the flow dropped and could not be remedied by repositioning. therefore, the pump was explanted and replaced. | [5 empty] | The investigation has been completed since the original report was filed. the device was returned for investigation. data logs showed about 10 minutes into support, the pump position was in the aorta that triggered impella position in aorta alarms. about one minute later, the pump was stopped manually. no low flow was found on the data. visual inspection found a kink on the cannula about 15mm from the of cage and the cannula bonding sire. no other issues were found on the rest of the pump. the root cause of positioning issue was due to attempt to re-cross aortic valve wirelessly. the cannula kink most likely occurred during re-crossing the valve. h.10 the following was included on the original manufacturer device report and does not belong in the report: instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it¿.
33161 | 1220648-2023-04606 | 18016145 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18016145&pc=OZD | 10/09/2023 | 05/29/2024 |  | 10/26/2023 |  |  | 10/09/2023 | 10/26/2023 |  |  | 07/31/2025 |  | 10/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024390843 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Death | No | Yes | Deformation Due to Compressive Stress | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/09/2023 |  |  |  |  | The patient expired. upon review by abiomed's medical safety officer, there is not enough information to exclude the patient's outcome from the use of the device. | The user facility reported a 73-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient who had just had an impella cp implant had the flow reduced. images were taken and a kink was noticed in the cannula just distal to the motor. the medical team attempted to manipulate without success and decision was made to explant the device and implant a new one. no patient harm was reported due to the issue. |  |  |  |  | The investigation has been completed since the original report was submitted. according to the clinical information, shortly after beginning impella cp support the flows began to decrease. imaging confirmed the presence of a cannula kink that could not be resolved after multiple attempts. the decision was made to explant the pump and replace it with another impella device. no product was returned for a visual inspection. data log analysis confirmed suction alarms with gradual decrease in flow. the most likely cause of the low flow was a cannula kink. b.7 revised as this information was inadvertently omitted from the previously submitted report. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.1 was revised due to medical safety officer review. h.6 codes 2199 and 2964 1248 were reported incorrectly on the previously submitted report. new code has been added to health effect - impact code. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
33162 | 1220648-2023-04607 | 18016215 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18016215&pc=OZD | 09/25/2023 | 09/12/2024 |  | 10/26/2023 |  |  | 09/25/2023 | 10/26/2023 |  |  |  |  | 10/26/2023 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1663002 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Communication or Transmission Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported the automated impella controller (aic) had a ¿controller error¿ yellow alarm. the staff reported the issue resolved itself and the alarm was cleared. the representative logged into the impella connect, but it did not capture the alarm. therefore, the aic was replaced. there was no reported patient harm. | [5 empty] | The investigation of automated impella controller (aic) - battery comm. failure has been completed. the logs confirmed the reported failure mode given there was a controller error alarm. the controller error and loss of placement signal is due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received. the following have been reported incorrectly or missing from manufacturer device report. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
33163 | 1220648-2023-04604 | 18016259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18016259&pc=OZD | 09/29/2023 | 02/21/2024 |  | 10/26/2023 |  |  |  | 10/26/2023 |  |  | 07/31/2025 |  | 10/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024385834 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/29/2023 |  |  |  |  | The user facility reported a 56-year-old male in anterior st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient endured an access site injury. there was a retroperitoneal bleed treated by pump explant and surgical repair in the or. the plan was to transfer the patient out to the (b)(6). | [5 empty] | The investigation into the reported access site bleeding -major has been completed since the original report was filed. product was not returned for review. the root cause of the access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33164 | 1220648-2023-04555 | 18016571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18016571&pc=OZD | 06/27/2023 | 01/18/2024 |  | 10/26/2023 |  |  |  | 10/26/2023 |  |  | 05/31/2025 |  | 10/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 |  |  | ABIOMED, INC. |  |  | US |  | 2024352604 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  |  |  | Malfunction | No | Yes | Device Sensing Problem | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/27/2023 |  |  |  |  | The user facility reported that an 82-year-old female patient implanted with the impella 5.5 for mechanical circulatory support experienced low pump flows and optical signal issues. the flows were at 0.5 l/min on insertion and motor current was flat at p-6. the patient's activated clotting time (act) was greater than 250 on implant, central venous pressure (cvp) was 7, and her ventricle measured 5cm. 500 ml of albumin was administered. pump flows remained between 0.5-1.5 l/min at p-6 with non-pulsatile motor current and the device continued to false position in aorta alarms despite being in good position. the medical staff was unable to achieve adequate flows. of note, the patient was noted to have a small left ventricle and calcified aortic valve and arch, and a "stenotic valve with low volume". no further information was provided. there was no reported harm to the patient. | [5 empty] | The investigation into the reported low pump flow has been completed since the original report was filed. the device was not returned by the medical facility. data logs were reviewed, several suction alarms were observed with low flow issue. retrograde flow was detected with shifting pump speed (rpm) followed by suction and then rpm goes back to set value resolving suction and again detecting retrograde flow. this cycle of retrograde flow and suction is seen during entire case. also, several false alarm of 'impella position in aorta' were found with good position verified. the cause of the low pump flow issue was most likely patient condition related with suction alarms observed during entire case and patient having small left ventricle, stenotic heart valve. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
33165 | 1220648-2023-04608 | 18017093 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18017093&pc=OZD | 12/31/2021 | 10/26/2023 |  | 10/26/2023 |  |  | 12/31/2021 | 10/26/2023 |  |  | 04/30/2023 |  | 10/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022074336 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 64-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella 5.5 device for mechanical circulatory support. the pump stopped and it was unable to restart. there was no mention of harm to the patient. | [5 empty] | (h3): the investigation into the reported "pump stop" has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. it was determined that the cause of the pump stop was most likely blood ingress as a result of high purge pressure. the cause of the purge leak was most likely sidearm yellow luer damage due to the use of sodium bicarbonate in the purge solution. a risk assessment has been initiated to escalate this issue. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
33166 | 1220648-2023-04609 | 18017097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18017097&pc=OZD | 09/27/2023 | 02/19/2024 |  | 10/26/2023 |  |  | 09/27/2023 | 10/26/2023 |  |  | 07/31/2025 |  | 10/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024379663 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 45 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Material Integrity Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant in japan had a impella 5.5 support ongoing for cardiomyopathy patient when after 20 days the sleeve was damaged. the sterilization sleeve detached from the anchoring ring, which exposed the shaft. there was no reported patient harm. | [5 empty] | The investigation into the product damage (sterile sleeve damage) issue has been completed. the device was not returned for investigation. the root cause of product damage was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
33167 | 1220648-2023-04610 | 18017142 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18017142&pc=OZD | 02/06/2023 | 10/26/2023 |  | 10/26/2023 |  |  | 02/06/2023 | 10/26/2023 |  |  | 11/30/2024 |  | 10/26/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023283403 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/06/2023 |  |  |  |  | The complainant reported a 42 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, saturated flows were exhibited that led to device explant. the patient was deemed stable. although a clot was found wrapped around the motor housing, there was no injury sustained. | [5 empty] | The product and corresponding data logs were returned and analyzed. data logs confirmed two instances of high purge pressure and flow saturation. the product was evaluated and a large clot was found wrapped around the impeller and around the purge gap. the clinical also notes this patient had clinical factors making them susceptible to clotting. based on this evidence, the root cause of the high purge pressure was likely biomaterial. this report is being filed as part of a retrospective review of historical records.
33168 | 1220648-2023-04550 | 18002137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18002137&pc=OZD | 03/18/2022 | 10/25/2023 |  | 10/25/2023 |  |  | 03/18/2022 | 10/25/2023 |  |  | 12/31/2023 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023173221 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/18/2022 |  |  |  |  | The user facility reported that a 61-year-old female patient was implanted with an impella 5.5 for mechanical circulatory support. the site suspected a clot had been ingested during insertion of the device due to a non-pulsatile motor current and abnormal purge flow/pressure. tissue plasminogen activator (tpa)) was introduced into the purge as a troubleshooting technique. an echo was performed, and the pump was noted to be positioned toward the mitral valve resulting in hemolysis. a possible clot was also visualized near the inflow. the device was subsequently explanted as a result. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported thrombus, hemolysis, and low purge flows was completed. the device was not returned for evaluation. based on the available information and no device return, the root cause of the reported high purge pressure/thrombus is undetermined. the cause for the hemolysis was believed to be most likely related to improper positioning of the device by the user. this report is being filed as part of a retrospective review of historical records.
33169 | 1220648-2023-04551 | 18002257 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18002257&pc=OZD | 07/19/2022 | 10/25/2023 |  | 10/25/2023 |  |  | 07/19/2022 | 10/25/2023 |  |  | 02/29/2024 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2023196637 |  | 0550-0006 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/19/2022 |  |  |  |  | The user facility reported a 53-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. during support, a high purge flow was observed coinciding with a fluid leak from a crack in the yellow luer connection. a purge bypass was completed to resolve the leak; however, a high purge pressure alarm was noted the following day. the pump was subsequently explanted as a result of the noted issue. | [5 empty] | The device was returned and an investigation was completed. the data logs confirmed low purge pressure. the rt log revealed that purge pressure rose above 1000mmhg and then suddenly dropped close to 0mmhg prior to the purge pressure low alarm. the purge flow was close to 0ml/hr prior to the alarm and then spiked to 30ml/hr after the alarm. a visual inspection of the pump revealed a crack in the yellow luer. therefore, the root cause of the purge leak was a damaged sidearm yellow luer due to the use of bicarbonate in the purge solution. the data logs also confirmed high purge pressure later in the case. a visual inspection of the pump revealed a kink in the catheter. upon conclusion of the investigation, the root cause of the high purge pressure was a kink in the purge line. per the ifu: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
33170 | 1220648-2023-04552 | 18002477 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18002477&pc=OZD | 09/25/2023 | 06/25/2024 |  | 10/25/2023 |  |  | 09/25/2023 | 10/25/2023 |  |  | 08/31/2025 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024384914 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | There is not enough information to associate use of the device with the outcome. | The us complainant reported there was bleeding from the right groin, which was the site of the impella cp access. 1 unit of packed red blood cells (prbc's) were given to the patient. the plan was to wean the patient and to remove the impella and surgical repair was to be performed in the or. however, when the patient was moved to the or table, the patient had pulseless electrical activity arrest. the patient had a dnr and the patient expired after 15.28 hours of support. no allegations were made towards the impella cp for cause of death. |  |  |  |  | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the vascular damage issue was not determined since product was not returned and sufficient information is not provided in the clinical description. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33171 | 1220648-2023-04553 | 18002497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18002497&pc=OZD | 05/29/2023 | 10/25/2023 |  | 10/25/2023 |  |  | 05/29/2023 | 10/25/2023 |  |  | 01/31/2025 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP- JP |  |  | OZD |  |  | IMPELLA CP- JP |  |  | ABIOMED, INC. |  |  | US |  | 2024315938 |  | 0048-0024 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 71-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant had a cp-ecmo support ongoing for 2.5 hours when due to a purge line leak the pump was replaced. the replacement was at the suggestion of the physician in conversation with the medical team. no harm from interruption in support. the second impella cp supports the pump to date. | [5 empty] | The investigation into the reported purge leak has been completed since the original report was filed. product was returned for investigation. the leak test was performed and leak was found through the luer neck of filter. the cause of the leak was due to a crack at the filter joint. instructions for use for the related event are as follows: "do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
33172 | 1220648-2023-04554 | 18002608 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18002608&pc=OZD | 05/02/2023 | 10/25/2023 |  | 10/25/2023 |  |  | 05/02/2023 | 10/25/2023 |  |  | 07/31/2024 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023253847 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/02/2023 |  |  |  |  | The user facility reported a 29-year-old female was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was high purge pressure on the pump. the pump was moved to a new console and the issue resolved. | [5 empty] | The investigation for the reported high purge pressure has been completed. the data logs were reviewed and confirm a rise in the purge pressure and the issue was brief and resolved. however, the actual cause of the high purge pressure could not be determine without the device for evaluation. the instructions for use provides troubleshooting for high pressure issues. it states unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ this report is being filed as part of a retrospective review of historical records.
33173 | 1220648-2023-04556 | 18003454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18003454&pc=OZD | 08/02/2022 | 10/25/2023 |  | 10/25/2023 |  |  | 08/02/2022 | 10/25/2023 |  |  | 11/30/2023 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023169248 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/02/2022 |  |  |  |  | The user facility reported that a 41-year-old female patient implanted with the impella 5.5 for mechanical circulatory support experienced positioning issues on the repositioning sheath, where the blue suture hub, and the white portion (where the t-lock is) swivels. the physician believed this made the pump torque out of position. the physician then put a stitch in to hold it in place on day 3 of the support. no excessive force was reported. three days later it was later reported that the issue was persisting and so the issue was minimized when a zip tie was placed. no further information was provided. there were no reports of harm to the patient. the pump support continued for another 6 days til pump was explanted. the patient remained on ecmo support after the 5.5 explanted. | [5 empty] | The investigation into the reported positioning issues was completed. the pump was returned for evaluation. visual inspection of the returned pump did not reveal anything abnormal. based on available information, the root cause of the positioning issue was use related. this report is being filed as part of a retrospective review of historical records.
33174 | 1220648-2023-04557 | 18003472 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18003472&pc=OZD | 09/22/2023 | 09/26/2024 |  | 10/25/2023 |  |  | 09/25/2023 | 10/25/2023 |  |  | 01/31/2025 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024315350 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant in japan reported the impella cp access site was bleeding. patient received a blood transfusion. the patient had circulatory system failure due to bleeding at the insertion site. the patient was on impella support for 48.65 hours before expiring. cause of death was determined to be circulatory system failure and multiple organ failure. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was determined to be anticoagulation management because the activated clotting time values (185) are higher than the therapeutic range. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ a.3 revised from the initial submission in accordance with updated procedures. b.3 date of event revised as previously entered incorrectly. d.6a and d.6b revised from the initial submission in accordance with updated procedures. e.1 first name revised as previously entered incorrectly. address line 2 was revised from the initial submission in accordance with updated procedures. g.1 revised reporting contact fax number from the initial submission in accordance with updated procedures. h.1 revised as previously entered incorrectly. h.6 code 4114 was reported incorrectly on the previously submitted report. new code has been added to type of investigation codes. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. additional manufacturer narrative on the previously submitted report stated that the device was not returned but the device was discarded. type of investigation codes has been updated to reflect discarded. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-04557 was provided in sections b2, g3, and h1, and h6. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33175 | 1220648-2023-04558 | 18003622 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18003622&pc=OZD | 09/26/2023 | 06/19/2024 |  | 10/25/2023 |  |  | 09/26/2023 | 10/25/2023 |  |  | 12/31/2024 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 SET AU |  |  | ABIOMED, INC. |  |  | US |  | 2024286281 |  | 1000361 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 33 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/26/2023 |  |  |  |  | The user facility in australia reported a 33-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. on day 11 of the support, the team observed the sterile sleeve to be damaged. the damaged sleeve was troubleshot by the team with sticky tape. the pump remains on for support for now over 17 days. there was no noted harm or infection. | [5 empty] | The investigation into the product damage (sterile sleeve damage) has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the of the product damage is most likely due to use based on the provided clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
33176 | 1220648-2023-04559 | 18003724 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18003724&pc=OZD | 10/06/2023 | 07/16/2024 |  | 10/25/2023 |  |  | 10/06/2023 | 10/25/2023 |  |  |  |  | 10/25/2023 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER, |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1363939 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Difficult to Setup or Prepare | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 10/06/2023 |  |  |  |  | The user facility reported a 81-year-old male noted as high risk percutaneous coronary intervention (hrpci) was implanted with an impella cp device for mechanical circulatory support. there was an "impella purge disk not detected" alarm present as the staff was setting up the device. purge disk was quickly advanced into the proper position within the automated impella controller (aic ) and the alarm was resolved. | [5 empty] | The investigation for the aic purge flag/disc issues has been completed. the impella device was not returned for evaluation. data logs were reviewed which showed that the console recorded an rfid error for six seconds, which indicates that the cassette was not completely inserted into the aic. upon the third insertion of the cassette, the console displayed ¿purge disc not detected ¿ alarm¿ event #402, most likely due to the delay in the disc insertion. this aligns with the reported problem description. after 9 seconds, event #409 was cleared, and the disc insertion was recorded as ¿imcgui: purger stat: sensor=1¿. this aligns with the claim that the ¿purge disk was quickly advanced into the proper position within the aic, and the alarm was resolved by staff. the root cause of the reported failure mode was most likely a use issue. the ¿purge disc not detected¿alarm¿ was due to the delay in the disc insertion. this is an intended alarm without a disc after the cassette was inserted. there was no evidence of product malfunction. d2-corrected common device name. f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
33177 | 1220648-2023-04560 | 18004029 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18004029&pc=OZD | 10/07/2023 | 09/19/2025 |  | 10/25/2023 |  |  | 10/07/2023 | 10/25/2023 |  |  | 07/31/2025 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024374298 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  | American Indian Or Alaskan Native | Malfunction | No | Yes | Device Sensing Problem | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/07/2023 |  |  |  |  | The user facility reported a 66-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. after the patient had a coughing episode, suction alarms and false impella in aorta alarms occurred. echo showed impella to be in proper position. no intervention was needed to address the issue. | [5 empty] | The investigation into the reported, low pump flow has been completed since the original report was filed. product was not returned for review. the root cause of the low flow was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed with root cause. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data confirmed the low flow when pump started & remained to the rest of the case that triggered auto suction response & suction alarms. logs also showed impella position in aorta alarms were triggered during support. product was not returned for review. based on clinical description & data analysis, the root cause of optical signal issue was most likely patient condition. all pertinent information available to abiomed has been submitted at this time. h.6. updated to include optical signal issue coding. d6a / d6b updated. d.10. concomitant medical products and therapy dates updated. f.6. and f.8. should not have been filled out in the initial report. g.1. reporting contact email and reporting contact fax number updated.
33178 | 1220648-2023-04561 | 18004176 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18004176&pc=OZD | 04/30/2022 | 10/25/2023 |  | 10/25/2023 |  |  | 04/30/2022 | 10/25/2023 |  |  | 07/31/2023 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022115118 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | Improper Flow or Infusion | Pressure Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/30/2022 |  |  |  |  | The user facility reported a 55 year-old male was implanted with an impella 5.5 device mechanical circulatory support. the complainant reported a high purge pressure that was ultimately resolved with adding heparin in the purge solution. prior to this issue, a leak was noted at the yellow luer of the device's side arm. there was no harm to the patient in regards to the reported issues. | [5 empty] | The investigation for the purge pressure high issues and the leaking purge sidearm has been completed. the site did not return the device for evaluation. however, the clinical report was evaluated. based on the clinical account that the hpp was resolved by adding heparin, the root cause of the high purge pressure was most likely due to biomaterial. the cause of the purge leak was not determined as product was not returned. this report is being filed as part of a retrospective review of historical records.
33179 | 1220648-2023-04562 | 18004207 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18004207&pc=OZD | 04/24/2023 | 10/25/2023 |  | 10/25/2023 |  |  | 04/24/2023 | 10/25/2023 |  |  | 12/31/2024 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 202430044 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/24/2023 |  |  |  |  | The user facility reported 77yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that on day 2 of the support, there were high purge pressure alarms. the team troubleshot by the use of tissue plasminogen activator (tpa). this did not resolve the issue and so the team explanted and replaced the 5.5 with another pump. there was no patient harm reported. | [5 empty] | The investigation for the high purge pressure has been completed. the data log shows that there are suction alarms throughout the case as well as a motor current spike in the beginning of the case. the high purge pressure does not resolve during the case. the returned pump was investigated and biomaterial was observed in the purge gap. the pump was purged and the high purge pressure was reproduced after purging for several minutes. the pump was then run and saturated flow was noted. the cause of the high purge pressure was biomaterial. the instructions for use provides troubleshooting suggestions on how to manage high purge pressures. it states unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ this report is being filed as part of a retrospective review of historical records.
33180 | 1220648-2023-04563 | 18004218 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18004218&pc=OZD | 06/11/2022 | 09/24/2024 |  | 10/25/2023 |  |  | 06/11/2022 | 10/25/2023 |  |  | 03/31/2024 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023201276 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  |  | Injury | Yes | Yes | Fluid/Blood Leak | Adverse Event Without Identified Device or Use Problem | [5 empty] | Tachycardia | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/11/2022 |  |  |  |  | It was reported that the patient's care was withdrawn, and the patient expired while on impella support. patient experienced supraventricular tachycardia and then went in to ventricular tachycardia. . medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | The user facility reported a 58 year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the pump had a purge leak. the team effectively troubleshot the leak by using the purge bypass technique. there was no patient harm reported. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-04563 was provided in sections b1, b2, b5, d6, h1, and h6. | The investigation into the side arm purge leak has been completed. the device was not returned for evaluation. however, the clinical report was evaluated. the cause of the sidearm purge leak was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records.
33181 | 1220648-2023-04564 | 18004317 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18004317&pc=OZD | 04/01/2023 | 10/25/2023 |  | 10/25/2023 |  |  | 04/01/2023 | 10/25/2023 |  |  | 10/31/2024 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023283640 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 80-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant had an imepella cp support ongoing for approximately two days with signs of hemolysis. the hemolysis was not treated nor was the pump explanted. the team made choice to transition the patient to care/comfort measures only. | [5 empty] | The investigation into the reported "hemolysis" has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the logs does not indicate any suction or unsuitable pump position events. the cause of the issue was not determined since the product was not returned. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33182 | 1220648-2023-04565 | 18004330 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18004330&pc=OZD | 07/03/2022 | 10/25/2023 |  | 10/25/2023 |  |  | 07/03/2022 | 10/25/2023 |  |  | 08/31/2022 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 1496975 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/03/2022 |  |  |  |  | The user facility reported a 55 year old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was pump saturation issues. ther pump was replaced. there was no patient harm reported. | [5 empty] | The investigation for the flow saturation issues has been completed. the device was not returned for evaluation. however, data logs were reviewed and confirmed the saturated flow as well as multiple high motor current pump pauses. the root cause of the saturated flow was unable to be determined as no product was returned for evaluation. the root cause of the pump stop was unable to be determined as no product was returned for evaluation. this report is being filed as part of a retrospective review of historical records.
33183 | 1220648-2023-04566 | 18005599 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18005599&pc=OZD | 05/30/2023 | 09/19/2024 |  | 10/25/2023 |  |  | 05/30/2023 | 10/25/2023 |  |  | 09/30/2024 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023277382 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Pumping Stopped | Infusion or Flow Problem | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The complainant had an impella cp support for a 68-year-old female patient with severe mitral regurgitation (mr). the impella cp was placed but, had saturated flows. the saturated flows did not prompt the pump to be replaced or explanted. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-045 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding. | The investigation into the reported "saturated flow" has been completed. product was returned for investigation. the device was inspected and found that biomaterial was near the impeller of the pump. there were no other abnormalities found. while reviewing impella connect logs of the pump at p4 showed possible flow saturation. patient's act remained high once in the unit so heparin was not stared in the purge. the pump experienced pump saturation and pump stoppage. the pump was able to restarted at p3 and maintained at p3 overnight. the investigation determined that the root cause of the issue was biomaterial near the impeller. instructions for use for the related event are as follows: "maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿
33184 | 1220648-2023-04431 | 18005702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18005702&pc=OZD | 09/01/2022 | 10/24/2023 |  | 10/25/2023 |  |  | 09/01/2022 | 10/25/2023 |  |  | 04/30/2024 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023213158 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/01/2022 |  |  |  |  | The user facility reported a 45yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that twice after a patient stretched his arms up, the motor current spiked and the pump stopped. the pump was able to be restarted. the console was changed out without any further issue. the patient was eventually explanted and transplanted with a heart. there were no reports of patient harm. | [5 empty] | The investigation into the pump stop has been completed. neither the pump nor the data logs were returned for evaluation. however, the clinical information was sufficient in determining the root cause. the cause of the pump stop was most likely patient movement since the issue occurred during movement and resolved after restart. instructions for use for the related event are as follows: ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33185 | 1220648-2023-04414 | 18005704 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18005704&pc=OZD | 10/05/2023 | 10/24/2025 |  | 10/25/2023 |  |  | 10/05/2023 | 10/25/2023 |  |  |  |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1299516 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 10/05/2023 |  |  |  |  | The user facility reported that during a case preparation, the automated impella controller, (aic) had a battery failure . a replacement loaner was sent to the facility. no patient is involved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Corrections to the initial mfr report: d2 updated common device name d9 added device return date f6 and f8 should have been left blank updated h3 to device evaluation anticipated. | The investigation of automated impella controller (aic) - battery failure (red alarm) has been completed. data analysis: the aic logs confirmed the reported failure. there was a battery failure alarm (transport connection blown/missing ¿ event set #360) and a battery 2 communication failure alarm (smbus communication loss or sda short ¿ event set #353) due to low pack voltage device analysis: this console was tested after arriving at fs. visually confirmed that the battery 2 lcd display was off and the batttest result also confirmed that battery 2 was not communicating, most likely due to cell imbalance causing the battery pack¿s internal electronics to shut down. root cause: the root cause of the battery failure (red alarm)/battery cell failure was most likely a single cell failure in battery 2. e2 was erroneously selected on the initial manufacturer device report number 1220648-2023-04414. g1 reporting contact email revised on manufacturer device report 1220648-2023-04414 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2023-04414.
33186 | 1220648-2023-04419 | 18005705 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18005705&pc=OZD | 05/11/2023 | 09/11/2024 |  | 10/25/2023 |  |  | 05/11/2023 | 10/25/2023 |  |  | 12/31/2024 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024315849 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  |  |  | Injury | Yes | Yes | Increase in Pressure | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding the patient that endured impella partial malfunction with a "definite" relationship to the impella device used: "impella 5.5, malfunctioned and had to be replaced with new one. impella was partially malfunctioning on giving off high pressure and was replaced. the patient recovered with no sequelae. | The user facility reported that observed unusual amount of bleeding at the end of the graft where the graft is hubbed to the sheath. this bleeding/oozing occurred despite the graft being secured using the graft locks. upon examination of the sheath, it was appeared as if there were lined "channels" on the surface of the sheath. the sheath of the surface is normally smooth. the issue was resolved by securing the graft using suture tightened around the graft. |  |  |  |  | The investigation for high purge pressure has been completed. the impella device was not returned for evaluation. the data logs do not show any purge pressure high or purge system blocked alarms. the purge pressure and purge flows were within normal ranges. there was no problem found as the root cause of high purge pressure. b1 updated to reflect product problem in addition to the previously reported ae. b5 updated with failure mode information. h6 health effect - impact code, medical device problem code, and component code updated to reflect new failure mode reported. as stated in the instructions for use: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ this complaint was reopened as part of a retrospective review of historical records to assess coding, severity, and reportability. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-04419. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. g1 reporting contact fax number missing. h3 was the device evaluated by the manufacture missing | The investigation into the reported access site bleeding has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the returned pictures show score lines on the sheath body. the cause of the access site bleeding was an introducer sheath split along the scorelines. this caused the introducer to not seal when the graft locks were deployed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33187 | 1220648-2023-04427 | 18005710 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18005710&pc=OZD | 04/21/2023 | 10/24/2023 |  | 10/25/2023 |  |  | 04/21/2023 | 10/25/2023 |  |  | 12/31/2024 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED INC. |  |  | US |  | 2024314483 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  |  | Malfunction | No | Yes | Use of Device Problem | Material Deformation | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/21/2023 |  |  |  |  | The user facility in the eu reported a 66-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. after a difficult implantation due to the graft, the placement signal was lost. by checking the catheter of the 5.5, a kinking was spotted. replaced the pump with a second 5.5. patient stable. | [5 empty] | No data logs were returned. multiple kinks were observed on the catheter when examined upon the benchtop. the catheter was then stripped to examine the optical fiber line. a kink was observed in the fiber line close to the motor housing. the cause of the optical signal issue was a damaged optical fiber line. the damage occurred during/with improper technique at delivery. this report is being filed as part of a retrospective review of historical records.
33188 | 1220648-2023-04435 | 18005711 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18005711&pc=OZD | 09/24/2022 | 10/24/2023 |  | 10/25/2023 |  |  | 09/24/2022 | 10/25/2023 |  |  | 03/31/2024 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023206013 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/24/2022 |  |  |  |  | The user facility reported a 79 year old male with st elevated myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that an impella position in aorta alarm occurred. when reduced to p2, the flows remained high. the pump was replaced due to flow saturation. there was no patient harm reported. | [5 empty] | The investigation into the flow saturation issue has been completed. neither the product nor the data logs were returned for investigation. the cause of the issue was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records.
33189 | 1220648-2023-04428 | 18005714 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18005714&pc=OZD | 08/04/2022 | 09/19/2024 |  | 10/25/2023 |  |  | 08/04/2022 | 10/25/2023 |  |  | 04/30/2024 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023214722 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/04/2022 |  |  |  |  | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 57 year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was a sudden pump stop on p8. a new console was added and a restart was attempted unsuccessfully. the pump was pulled into aorta. the patient was quickly transitioned to end of life care and died within an hour of pump stop. there was no patient harm reported. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b1, b2, b5, h1, and h6. | The investigation into the pump stop has been completed. data logs showed the pump stopped immediately upon the sudden occurrence of "impella stopped (rpm deviations" alarm (alarm 115). when the impella stopped, the motor current spiked above 30,000 and the placement signal spiked to 3000. all metrics were within specification prior to the pump stop. however, since the product was not returned, the cause of the pump stop was not determined. instructions for use for the related event are as follows: ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
33190 | 1220648-2023-04432 | 18005718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18005718&pc=OZD | 12/04/2022 | 10/24/2023 |  | 10/25/2023 |  |  | 12/04/2022 | 10/25/2023 |  |  | 09/30/2024 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023262494 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Neuropathy | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 55-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported administering gabapentin to alleviate pain symptoms suspected from impella insertion thought to be brachial plexus related pain. | [5 empty] | The investigation into the reported nerve pain, has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that due to limited clinical details and the pump not being returned the root cause of the injury is not determined. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33191 | 1220648-2023-04429 | 18005719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18005719&pc=OZD | 02/22/2023 | 10/24/2023 |  | 10/25/2023 |  |  | 02/22/2023 | 10/25/2023 |  |  | 11/30/2024 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023262474 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 56-year-old male in st elevated myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. the complainant had the impella 5.5 pump placed for escalation from the impella cp. the impella 5.5 pump had been supporting for over three days when there was an observation made of hemorrhagic bleed/cva. the pump remained on for support and the patient was successfully weaned the following day. | [5 empty] | The investigation into the reported "stroke/cva" has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the cause of the issue was not determined since product was not returned and sufficient clinical information was not provided. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33192 | 1220648-2023-04422 | 18005720 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18005720&pc=OZD | 01/27/2022 | 10/24/2023 |  | 10/25/2023 |  |  | 01/27/2022 | 10/25/2023 |  |  | 10/31/2023 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022157665 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Embolism/Embolus | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/27/2022 |  |  |  |  | The user facility reported a 63-year-old male was implanted with an impella device for mechanical circulatory support. the complainant was prepping a patient for 5.5 support when the team observed air in the (left ventricle) lv. the pump was not yet on in the lv but, was positioned within it. the air resolved within about 30 seconds without any intervention or harm. | [5 empty] | The returned pump and purge cassette were visually inspected, and no abnormalities were observed. the purge cassette was plugged into a console and de-aired after which the pump was connected and purged normally. the field data logs were reviewed, and no abnormalities were noted. a relationship between impella and the reported ventricular bubbles could not be confirmed. this report is being filed as part of a retrospective review of historical records.
33193 | 1220648-2023-04430 | 18005721 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18005721&pc=OZD | 04/21/2022 | 10/24/2023 |  | 10/25/2023 |  |  | 04/21/2022 | 10/25/2023 |  |  | 09/30/2023 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022143644 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Break | [6 empty] | Hematoma | Pain | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/21/2022 |  |  |  |  | The user facility reported a 51 year-old male was implanted with an impella 5.5 device mechanical circulatory support. it was reported that a hematoma developed overnight during support. the patient was taken for swan placement and debridement of the axillary site. the patient complained of ¿a lot of pain¿ at the insertion site after having hematoma evacuation. additionally, it was reported that the red plug by the purge system broke. a bypass was performed. there was no patient harm reported. | [5 empty] | The investigation into the reported access site bleeding/hematoma and the broken purge sidearm has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the root cause of the access site bleeding was unable to be determined as no product was returned for analysis. the root cause of the purge leak was unable to be determined as no product was returned for analysis so the location of the leak could not be identified. the purge leak occurred beyond the 5.5 design life of 60 days. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
33194 | 1220648-2023-04410 | 18005726 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18005726&pc=OZD | 07/07/2022 | 10/24/2023 |  | 10/25/2023 |  |  | 07/07/2022 | 10/25/2023 |  |  | 01/31/2024 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023182460 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/07/2022 |  |  |  |  | The user facility reported a 66 year-old male was implanted with an impella 5.5 device for mechanical circulatory support as a bridge to transplant. it was reported that there were intermittent purge flow/pressure alarms with a noticeable leak in the purge sidearm. the pump was replaced. there was no patient harm reported. | [5 empty] | The investigation into the purge sidearm leak has been completed. the device was not returned for evaluation. the data logs were returned and were reviewed along with clinical report. data logs showed there multiple "purge pressure low" alarms with erratic purge pressure metrics. the root cause of the purge leak was unable to be determined because the product was not returned so the location of the leak could not be identified. this report is being filed as part of a retrospective review of historical records.
33195 | 1220648-2023-04420 | 18005727 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18005727&pc=OZD | 06/29/2022 | 02/21/2025 |  | 10/25/2023 |  |  | 06/29/2022 | 10/25/2023 |  |  | 03/31/2024 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023206917 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/29/2022 |  |  |  |  | The user facility reported a 63 year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that a purge sidearm leak occurred after sodium bicarbonate was added to the purge. the pump had been supporting for 13 days when this occurred. there was no patient harm reported. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | The impella was returned with the entire catheter removed, including the pump handle, purge sidearm, and patient cable. after reviewing the clinical information, it was determined that the cause of the purge leak was most likely sidearm yellow luer damage due to the use of sodium bicarbonate in the purge solution.
33196 | 1220648-2023-04423 | 18005729 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18005729&pc=OZD | 02/27/2023 | 10/24/2023 |  | 10/25/2023 |  |  | 02/27/2023 | 10/25/2023 |  |  | 12/31/2024 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024293967 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a 57-year-old male patient with cardiomyopathy was implanted with an impella 5.5 for support. towards the end of impella support, and prior to explant, a blood leak was noted at red impella plug. the patient received a heart transplant and impella was explanted at this time. the patient was considered to be stable following the events. | [5 empty] | The investigation into the reported access site bleeding has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that root cause of the bleeding cannot be determined since the complaint device was not returned for investigation. the instructions for use are as follows: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
33197 | 1220648-2023-04433 | 18005731 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18005731&pc=OZD | 03/29/2023 | 10/24/2023 |  | 10/25/2023 |  |  | 03/29/2023 | 10/25/2023 |  |  | 12/31/2024 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024286883 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a 67-year-old patient was scheduled for impella 5.5 implant for mechanical circulatory support. the surgeon felt that the vessel was small but without disease and at appropriate depth to attempt an impella 5.5 implant for patient support. surgeon attempted and cannula kept kinking on itself and also bent 0.018 wire in half. the impella took trajectory towards head despite multiple manipulations and removing from sheath kit and reinserting. therefore, the surgeon felt that continuing would dissect the artery and pump implant of impella 5.5 aborted. there was no reported patient harm. | [5 empty] | The investigation into the reported 'delivery issues" has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. based on the provided information, the root cause of the delivery issue was patient condition related, specifically the patient's small vessels. this report is being filed as part of a retrospective review of historical records.
33198 | 1220648-2023-04418 | 18005740 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18005740&pc=OZD | 07/23/2022 | 10/24/2023 |  | 10/25/2023 |  |  | 07/23/2022 | 10/25/2023 |  |  | 03/31/2024 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023201446 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/23/2022 |  |  |  |  | The user facility reported a 69 year-old male in st elevated myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the pump had a purge leak noted within the purge sidearm retainer. a purge bypass procedure was complete without further issues. the purge was running sodium bicarbonate within the concentration. there was no patient harm reported. | [5 empty] | The investigation into the purge sidearm leak has been completed. the pump and the data logs were returned for evaluation. review of the data logs showed the purge pressure immediately dropped to almost 0 upon the sudden occurrence of the "purge pressure low" alarm. leak reproduction testing was performed on the returned sidearm, and leaking was observed at the filter neck. therefore, it was determined that the cause of the purge leak was damage to the filter to tubing joint. this report is being filed as part of a retrospective review of historical records.
33199 | 1220648-2023-04409 | 18005741 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18005741&pc=OZD | 06/27/2022 | 10/24/2023 |  | 10/25/2023 |  |  | 06/27/2022 | 10/25/2023 |  |  | 03/07/2022 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023180579 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/27/2022 |  |  |  |  | The user facility reported a 66 year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was a purge system open alarm. after troubleshooting, a purge sidearm leak was discovered. a purge bypass was performed. there was no harm or injury reported. | [5 empty] | The investigation into the purge sidearm leak has been completed. the device was not returned for evaluation. however, the clinical report has been reviewed. the root cause of the purge leak was unable to be determined because the product was not returned so the location of the leak could not be identified. this report is being filed as part of a retrospective review of historical records.
33200 | 1220648-2023-04415 | 18005742 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18005742&pc=OZD | 07/05/2022 | 10/24/2023 |  | 10/25/2023 |  |  | 07/05/2022 | 10/25/2023 |  |  | 03/31/2024 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023209732 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Improper Flow or Infusion | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/05/2022 |  |  |  |  | The user facility reported a 58yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that an impella had low flow and had to be removed after 35 minutes. on removal, the 5.5 was severely kinked and had a clot formation on the inlet and outlet. the doctor performed a mini sternotomy and a new pump was placed. there was no patient harm reported. | [5 empty] | The investigation for the reported low flow issue and clot has been completed. no product or data was returned for evaluation. however, the clinical report provides sufficient information to determine the root cause. the cause of the low pump flow was most likely due to the ingestion of a clot. the instructions for use contains details on how to prevent thrombus formation. it states ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿. this report is being filed as part of a retrospective review of historical records.
33201 | 1220648-2023-04434 | 18005743 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18005743&pc=OZD | 10/04/2022 | 10/24/2023 |  | 10/25/2023 |  |  | 10/04/2022 | 10/25/2023 |  |  | 06/30/2024 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023217168 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Electrical /Electronic Property Problem | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/04/2022 |  |  |  |  | The user facility reported a 67 yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the pump abruptly stopped with associated power spike. bedside staff attempted restart procedures without success. the patient was taken to the operating room for pump removal. upon removal, the connection to the red impella plug was obviously separated with exposed wires. there was no patient harm reported. | [5 empty] | The investigation into the pump stop has been completed. an image received from the field revealed the catheter had come out of the kink protector bond, exposing the motor cable wires. data logs showed that the pump stopped immediately upon the sudden occurrence of the "impella stopped - motor current high" alarm. when the pump stopped, the motor current spiked to over 30,000 and the placement signal went to 3000 immediately. after the pump stop, there were further motor spikes, accompanied frequent impella stopped alarms (alarms 13 and 115). no product was returned. in conclusion, it was determined that the cause of the pump stop was most likely an open electrical line stemming from damage to the motor cables when the catheter was separated from the red handle. the cause of this separation was not determined since the product was not returned. instructions for use for the related event are as follows: ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
33202 | 1220648-2023-04413 | 18005744 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18005744&pc=OZD | 07/12/2022 | 10/24/2023 |  | 10/25/2023 |  |  | 07/12/2022 | 10/25/2023 |  |  | 02/29/2024 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023190571 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 19 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/12/2022 |  |  |  |  | The user facility reported a 19 year-old female was implanted with an impella 5.5 device for mechanical circulatory support as a bridge to transplant. it was reported that while the team was replacing the purge cassette, there was a purge leak and the purge pressure dropped. the team attempted an unsuccessful bypass and eventually the pump was replaced. no lasting harm or injury from the interruption and pump replacement. | [5 empty] | The investigation into the purge sidearm leak after a purge cassette change has been completed. the device was not returned for evaluation. however, the clinical report was evaluated. the evaluation found damage was noted to be at the yellow luer and the sidearm retainer was in use. it is unknown if excessive force was used. a sidearm bypass was performed as a result. the root cause of the purge leak was unable to be determined because the product was not returned so the location of the leak could not be identified. this report is being filed as part of a retrospective review of historical records.
33203 | 1220648-2023-04425 | 18005745 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18005745&pc=OZD | 02/14/2023 | 10/24/2023 |  | 10/25/2023 |  |  | 02/14/2023 | 10/25/2023 |  |  | 12/31/2024 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024289544 |  | 1000211 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/14/2023 |  |  |  |  | The user facility reported a 52-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant in japan had a delivery of the 5.5 that was not possible. the team attempted all troubleshooting meausres of wire exchange and imaging of the anatomy, but when delivery failed the team instead placed a cp. there was no harm or injury to patient from delay or attempted insertions. the 5.5 did kink in the process. | [5 empty] | The returned catheter was inspected and both catheter and cannula kinks were observed. no data logs were returned. the cause of the delivery issue was patient condition related and related to the anatomy of the patient's vessels causing the catheter to kink. this report is being filed as part of a retrospective review of historical records.
33204 | 1220648-2023-04416 | 18005750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18005750&pc=OZD | 07/17/2022 | 10/24/2023 |  | 10/25/2023 |  |  | 07/17/2022 | 10/25/2023 |  |  | 11/30/2023 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023166224 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/17/2022 |  |  |  |  | The user facility reported a 78 year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was a purge sidearm leak at the yellow luer. a purge bypass was completed. there was sodium bicarbonate being used in the purge. there was no patient harm reported. | [5 empty] | The investigation into the purge sidearm leak has been completed. the device was not returned for evaluation. however, the clinical report has been reviewed. the clinical reported that the leak was coming from the yellow luer, however no product was returned for analysis. therefore, the root cause of the purge leak was most likely a damaged sidearm yellow luer. this report is being filed as part of a retrospective review of historical records.
33205 | 1220648-2023-04426 | 18005753 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18005753&pc=OZD | 01/24/2023 | 10/24/2023 |  | 10/25/2023 |  |  | 01/24/2023 | 10/25/2023 |  |  | 09/30/2023 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022151070 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a placement signal issue during impella support. after crossing the aortic valve, both the ao and left ventricle (lv) waveforms became flat. the pump was subsequently removed and replaced to resolve the noted issue. the pump was returned and a significant cannula kink was observed. this was determined to be a positioning issue rather than a placement signal issue as the cannula was kinked during initial positioning causing abnormal pressures in the area of the optical sensor. there was no patient harm as a result of the reported issue. | [5 empty] | The investigation into the reported issue has been completed. the pump was returned for investigation. during the evaluation, the 5s parameters were found to be in normal range. a visual inspection noted a significant cannula kink. upon conclusion of the investigation, the root cause of the reported issues was most likely use related as a cannula kink occurred during initial positioning which prevented the optical sensor from getting an accurate reading. this report is being filed as part of a retrospective review of historical records.
33206 | 1220648-2023-04567 | 18005754 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18005754&pc=OZD | 06/05/2023 | 10/25/2023 |  | 10/25/2023 |  |  | 06/05/2023 | 10/25/2023 |  |  |  |  | 10/25/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1630372 |  | 0042-0010 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Run on Battery | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant in the uae reported a battery failure (red alarm) on the automated impella controller (aic) unit. | [5 empty] | The investigation into the reported 'aic - battery failure (red alarm)" has been completed. product was returned for investigation. the device was inspected and found the batteries were confirmed to be unable to power the console and could not charge. testing was performed which showed that both batteries were sufficiently depleted such that they were locked out. replacing the batteries, with batteries that were not locked restored charging capabilities and provided full voltage and communication capabilities to the system. the investigation determined that the root cause of the automated impella controller (aic) issue was a depleted battery.
33207 | 1220648-2023-04411 | 18005760 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18005760&pc=OZD | 09/24/2023 | 10/31/2023 |  | 10/25/2023 |  |  | 09/24/2023 | 10/25/2023 |  |  | 07/31/2025 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024375081 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 red plug was leaking post-implant. the leaking stopped and the physician decided to hold off changing the pump. it was further reported that impella 5.5 had a defective alarm and a pump stop. the patient was then being transferred to a new facility so the automated impella controller (aic) needed to be exchanged. after exchanging the aic's and plugging in the impella, the "impella defective" alarm occurred and the impella had a pump stop at the same time. it was plugged back into the old aic but the defective alarm still occurred. therefore, the pump was replaced. | [5 empty] | The investigation into the pump not detected issue and the low pump flow issue has been completed since the original report was filed. product was returned for evaluation. the root cause of the pump not detected is due to blood ingress in the red handle from a hole in the catheter around the 30cm mark. the failure mode will be monitored and trended. | The investigation into the reported pump stop/purge leak has been completed. the root cause of the pump not detected is due to blood ingress in the red handle from a hole in the catheter around the 30cm mark. the failure modes will be monitored and trended. instructions for use for the related event are as follows: ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
33208 | 1220648-2023-04424 | 18005761 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18005761&pc=OZD | 10/27/2021 | 01/23/2024 |  | 10/25/2023 |  |  | 10/27/2021 | 10/25/2023 |  |  | 08/31/2023 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022126642 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Female |  |  |  | Malfunction | No | Yes | Activation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/27/2021 |  |  |  |  | The user facility reported a 34-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported a delay in initiation of the impella pump. the team troubleshot and got the pump to start on 3rd attempt. the pump was ultimately started and used without issue. there was no harm to the patient from the delay. | [5 empty] | The investigation into the pump did not start issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. the root cause of the pump being unable to start initially was not determined as no product or data logs were provided for evaluation. after multiple attempts, the pump was able to start and provided support from (b)(6) 2021. | The root cause of the pump being unable to start initially was not determined as no product or data logs were provided for evaluation. after multiple attempts, the pump was able to start and provided support from (b)(6) 2021. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
33209 | 1220648-2023-04412 | 18005767 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18005767&pc=OZD | 10/21/2022 | 10/24/2023 |  | 10/25/2023 |  |  | 10/21/2022 | 10/25/2023 |  |  | 01/31/2024 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023180551 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/21/2022 |  |  |  |  | The user facility reported an 82-year-old female was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the physician inadvertently pulled the impella out of the left ventricle while attempting to reposition the pump after the patient was transferred to the operating room. multiple attempts were made to re-cross the aortic valve, however these efforts were unsuccessful. the decision was then made to explant the impella and replace it with an iabp (intra-aortic balloon pump). there was no harm to the patient. | [5 empty] | The investigation for the positioning issue has been completed. the device was not returned for evaluation. however, the clinical report provided sufficient details to determine the root cause. it was ascertained that the cause of the positioning issue was wireless re-access. the impella device is not indicated for wireless re-access. this report is being filed as part of a retrospective review of historical records.
33210 | 1220648-2023-04421 | 18005768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18005768&pc=OZD | 01/07/2023 | 10/24/2023 |  | 10/25/2023 |  |  | 01/07/2023 | 10/25/2023 |  |  | 11/30/2024 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023278490 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a 65-year-old male was implanted with an impella 5.5 pump for long term care as the patient waited for transplant. following implant, the patient developed a hematoma at the access site. heparin was subsequently reduced and act/ptt's were sub therapeutic. coinciding with the reported hematoma there was arm swelling and a large amount of drainage from the jp drain noted. the hematoma subsequently resolved on its own, however, the patient was administered one unit of blood as a result of the condition. there was no lasting harm to the patient. | [5 empty] | The investigation into the reported access site bleeding has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the root cause of the of the issue was patient condition related. instructions for use are as follows: ¿assess access site for bleeding and hematoma.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
33211 | 1220648-2023-04417 | 18005769 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18005769&pc=OZD | 04/14/2023 | 02/20/2024 |  | 10/25/2023 |  |  | 04/14/2023 | 10/25/2023 |  |  | 09/30/2024 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023264103 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the physician attempted to pull the boston v-18 guide wire out with considerable force, but it could not be removed from the impella 5.5. the physician pulled the impella catheter back out of the graft and was able to remove the wire. it was discovered that the wire had twisted and bent and this preventing the wire from being able to be removed while device was inserted in the patient. impella appeared to be damaged at the base of the motor and 9f catheter. a second impella device was placed without difficulty. | [5 empty] | The investigation into the delivery issues has been completed. the device was not returned for investigation. based on the clinical information provided, the doctor found a damaged guidewire during impella removal after having difficulty crossing the heart valve. second impella was placed without any issue. the cause of the delivery issue was most likely use related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿.
33212 | 1220648-2023-04568 | 18005891 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18005891&pc=OZD | 09/20/2023 | 03/08/2024 |  | 10/25/2023 |  |  | 09/25/2023 | 10/25/2023 |  |  | 05/31/2025 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2024356128 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic | Asian | Malfunction | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/25/2023 |  |  |  |  | The user facility reported a 70 year-old male was implanted with an impella cp. the complainant in japan reported that red urine was observed in the patient. the pump position was adjusted because the indwelling position was deep based on ultrasound. afterwards, the urine changed from red to brown after position adjustment, but renal function deteriorated and urine output decreased, so continuous hemodiafiltration (chdf) was performed. | [5 empty] | The investigation into the hemolysis and the purge leak ¿ cassette issues has been completed since the original report was filed. the pump was not returned for review, but the pc was returned for leaking analysis. no leak was found under leak test. the cause of the high purge flow was not determined since the pump was not returned and a leak was unable to be reproduced on the returned purge cassette. however, based on clinical description and data analysis, the root cause of hemolysis was most likely due to pump was not in the proper position. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
33213 | 1220648-2023-04569 | 18006092 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18006092&pc=OZD | 09/25/2023 | 09/20/2024 |  | 10/25/2023 |  |  | 09/25/2023 | 10/25/2023 |  |  | 05/31/2025 |  | 10/25/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1420740 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 85 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Fracture | Connection Problem | Adverse Event Without Identified Device or Use Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant in japan reported that a "purge pressure transmitter connection failure" alarm appeared. the problem was solved by replacing the automated impella controller (aic). there was no reported patient harm. | Additional information was provided that noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-04569 was provided in sections b1, b2, b5, d6a, d6b, h1, and h6.
33214 | 1220648-2023-04570 | 18006422 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18006422&pc=OZD | 09/25/2023 | 01/22/2024 |  | 10/25/2023 |  |  | 09/25/2023 | 10/25/2023 |  |  | 06/30/2025 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024365438 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/25/2023 |  |  |  |  | A patient of unknown age and gender has been treated for myocardia infraction and cardiogenic shock. for hemodynamic support, an impella cp smart assist pump was to be inserted. during the preparation, an issue with the purge system occurred and the decision was made to utilize second pump. another impella cp device was successfully used. no patient harm reported. | [5 empty] | The investigation into the purge leak issue has been completed since the original report was filed. impella pump and the purge cassette was returned, no kinks or damages were observed on the purge cassette or the pump. the purge cassette was able to be successfully primed in the lab. the purge cassette functioned properly, and no leaks were observed. the purge pressure and flow metrics were within normal ranges. there was no priming issue found during the case. the root cause of the reported leak issue could not be determined. the device functioned/operated to specification. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
33215 | 1220648-2023-04572 | 18006803 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18006803&pc=OZD | 09/26/2023 | 02/26/2024 |  | 10/25/2023 |  |  | 09/26/2023 | 10/25/2023 |  |  | 06/30/2025 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024368864 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 68-year-old male was implanted with an impella 5.5 for mechanical circulatory support. during closure of left axillary artery and impella 5.5 removal, the left arterial line showed loss of pressure. an ultrasound of the left upper arm showed loss of flow. cut down on the brachial artery was performed and clot retrieval was successfully performed. | [5 empty] | The investigation into the reported limb ischemia has been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33216 | 1220648-2023-04573 | 18006877 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18006877&pc=OZD | 06/13/2023 | 10/25/2023 |  | 10/25/2023 |  |  | 06/13/2023 | 10/25/2023 |  |  | 01/31/2025 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024330946 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a 47-year-old male patient with a stemi was admitted to the hospital and required an impella implant. the impella cp was explanted due to hemolysis and the patient had an iabp placed on the opposite site/limb for support. | [5 empty] | The investigation into the reported "hemolysis" has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the cause of the hemolysis issue was not determined since product was not returned, and sufficient clinical information and data logs were not provided. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33217 | 1220648-2023-04574 | 18006918 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18006918&pc=OZD | 09/25/2023 | 02/26/2024 |  | 10/25/2023 |  |  | 09/25/2023 | 10/25/2023 |  |  | 08/31/2025 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024386650 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 63 year old patient was implanted with an impella cp pump for mechanical circulatory support. the patient was supported by both ecmo and the cp when there was noted limb ischemia. a doppler pulse was notably difficult to find in the impella leg. an ultrasound revealed bilateral slow flow. the team therefore placed the distal perfusion catheter and support was continued. | [5 empty] | The investigation into the reported limb ischemia has been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33218 | 1220648-2023-04575 | 18006986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18006986&pc=OZD | 09/26/2023 | 09/22/2024 |  | 10/25/2023 |  |  | 09/26/2023 | 10/25/2023 |  |  | 07/31/2025 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024380178 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | Black Or African American | Death | No | Yes | Pumping Stopped | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The us complainant had placed an impella cp for support of a glasgow coma scale (gcs) patient in heart block with aortic valve disease. the patient was taken to the or for surgical intervention and aortic valve replacement. the pump stopped in the or and was troubleshot to re-start. the patient was allowed to come off bypass and have care withdrawn. the valve was not viable. the death is not attributed to the use of the cp. the device was discarded and will not be returned for evaluation. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. as noted in the impella ifu: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrected information for the initial mfr 1220648-2023-04575 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding. | The investigation into the reported pump stop has been completed since the original report was filed. the case details indicated necrotic aortic valve tissue present at the time of use. as well as the presence of high motor current alarms at the beginning of the case within the logs (after only 30 mins of use) it could be concluded that probable biomaterial ingestion was the root cause of this pump stop. instructions for use for the related event are as follows: maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The medical center has not returned any product or logs to date. upon closure of the investigation, a final report will be filed.
33219 | 1220648-2023-04576 | 18006999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18006999&pc=OZD | 09/25/2023 | 01/22/2024 |  | 10/25/2023 |  |  | 09/25/2023 | 10/25/2023 |  |  | 03/31/2025 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024346534 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had implanted an impella cp via the right femoral artery for support of a patient admitted for a valve replacement. the day of insertion the leg lost pulses. a perfusion catheter was placed to treat the condition. the pump remained on for 3 days and was weaned off. | [5 empty] | The investigation into the patient's limb ischemia has been completed. product was not returned for investigation. per the clinical review, the root cause of the limb ischemia issue was patient condition related to small vessel since limb ischemia was resolved after placing distal perfusion cannula. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33220 | 1220648-2023-04578 | 18007155 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18007155&pc=OZD | 09/26/2023 | 01/11/2024 |  | 10/25/2023 |  |  | 09/26/2023 | 10/25/2023 |  |  | 08/31/2025 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024385542 |  | 1000211 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant in japan reported that using surgical mode, a left ventricle perforation was detected before weaning from cardiopulmonary bypass after mitral valve repair (mvr) surgery, and left ventricular repair surgery was performed after impella 5.5 was removed. it was speculated that interference between the fragile myocardial tissue and the tip of the impella 5.5 inflow cage caused the perforation. | [5 empty] | The investigation into the ventricle perforation has been completed. the product was not returned. per the clinical information provided, the root cause of the ventricle perforation issue was patient condition related to the narrowing of ventricular lumen leading to interference between the fragile myocardial tissue and the tip of the impella inflow cage. the failure mode will be monitored and trended. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33221 | 1220648-2023-04577 | 18007161 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18007161&pc=OZD | 09/25/2023 | 06/06/2024 |  | 10/25/2023 |  |  | 09/25/2023 | 10/25/2023 |  |  | 06/30/2025 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024356972 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/25/2023 |  |  |  |  | The user facility reported a 57-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, the patient had a hematoma at the access site and heparin was discontinued. patient had a hematoma evacuation and support with the implanted pump was maintained. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the pump was not returned for investigation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 code 1888 was reported incorrectly on the previously submitted report. new code has been added to health effect - clinical code. added code 925 as it was inadvertently left off the previously submitted manufacture device report. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted report incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33222 | 1220648-2023-04579 | 18007198 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18007198&pc=OZD | 01/23/2022 | 10/25/2023 |  | 10/25/2023 |  |  | 01/23/2022 | 10/25/2023 |  |  | 10/31/2023 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022153718 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/23/2022 |  |  |  |  | The user facility reported a 56 year-old male was implanted with an impella 5.5 device for bridge to transplant. it was reported that there was a purge side arm leak. the patient was taken to the operating room for replacement. there was no patient injury reported. | [5 empty] | The investigation into the reported sidearm leak has been completed. the medical center did not return any impella products for benchtop analysis or the data logs for review; only the clinical report was evaluated. of note, sodium bicarbonate was used as a purge solution. the root cause of the reported purge sidearm leak was most likely a damaged yellow luer due to sodium bicarbonate usage. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33223 | 1220648-2023-04580 | 18008425 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18008425&pc=OZD | 09/07/2023 | 10/25/2023 |  | 10/25/2023 |  |  | 09/25/2023 | 10/25/2023 |  |  | 07/31/2025 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024376345 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/25/2023 |  |  |  |  | Long-term outcome and quality indicator (loqi) impella registry notification received 2023-09-25 for ventricular arrhythmias occurring during support. required multiple cardioversions, lidocaine boluses, dopamine bolus, amio bolus, mag, procainamide etc. the patient presented for impella support for cardiomyopathy. | [5 empty] | The impella device was not received by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
33224 | 1220648-2023-04581 | 18008429 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18008429&pc=OZD | 09/23/2023 | 10/25/2023 |  | 10/25/2023 |  |  | 09/25/2023 | 10/25/2023 |  |  | 06/30/2025 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024360925 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  | Asian | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant in japan reported the impella cp was inserted according to the instructions, but the pump did not start. the cable was reconnected and the automated impella controller (aic) was replaced, but the pump still did not work. although they thought they had followed the procedure exactly, it was found out that about 3 cm of the easyguide lumen remained in the pump. the impella pump was therefore replaced. | [5 empty] | The investigation into the reported pump not starting issues has been completed. the pump and guidewire were returned for investigation. partial guidewire was found to be wrapped around the impeller. there were no other physical abnormalities observed. the root cause of the pump not starting was determined to be a red lumen removal issue since partial red lumen was found wrapped around impellor. the failure modes will be monitored and trended.
33225 | 1220648-2023-04582 | 18008430 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18008430&pc=OZD | 09/21/2023 | 02/07/2025 |  | 10/25/2023 |  |  | 09/25/2023 | 10/25/2023 |  |  |  |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1084502 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported the automated impella controller (aic) had a broken purge clip. there was no reported patient harm. | [5 empty] | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. according to the device analysis the cause of the issue was a defective component. the following have been reported incorrectly or missing from manufacturer device report. b3 date of event d2 common device name revised on manufacturer device report in accordance with updated procedures e1 name prefix/title e2 should have been selected as no. e3 occupation g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. g1 reporting contact fax number missing. | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
33226 | 1220648-2023-04583 | 18008431 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18008431&pc=OZD | 09/25/2023 | 02/20/2024 |  | 10/25/2023 |  |  | 09/25/2023 | 10/25/2023 |  |  | 06/30/2025 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024377959 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant was implanted with an impella cp. it was reported that while the patient was being cleaned, the console started alarming pump in aorta. both the placement signal and lv waveform matched and looked like an aortic waveform, and motor current were flat. flows were reduced to p2. the patient was taken to the ccl to remove the pump. fluoroscopy demonstrated that the pump was in the aorta. therefore, the pump was explanted. | [5 empty] | The investigation into the positioning issues has been completed. the device was not returned for investigation. based on the clinical information provided, the cause of the positioning issue (impella pulled out) was patient movement. staff was turning and moving patient to clean. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿
33227 | 1220648-2023-04584 | 18008434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18008434&pc=OZD | 09/25/2023 | 04/19/2024 |  | 10/25/2023 |  |  | 09/25/2023 | 10/25/2023 |  |  | 04/30/2025 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024384532 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Female |  | Hispanic | Asian | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that when the impella cp was being inserted over the aortic arch the cannula kinked. after crossing the aortic valve, the impella cp started with suction alarms and flows of less than 1. there was a kink in the cannula by the motor housing. the physician tried to reposition the impella multiple times to get the kink out of the cannula but was unsuccessful. the impella was therefore removed and a new impella was inserted successfully. | [5 empty] | The investigation into the low pump flow issue has been completed since the original report was submitted. the device was not returned for investigation. according to the clinical notes, the doctor reported cannula kink during pump use. the cause of the low pump flow issue was kinked cannula, based on given clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿
33228 | 1220648-2023-04585 | 18008435 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18008435&pc=OZD | 09/25/2023 | 01/12/2024 |  | 10/25/2023 |  |  | 09/25/2023 | 10/25/2023 |  |  | 06/30/2025 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024376426 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Material Deformation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/25/2023 |  |  |  |  | The user facility reported a 40-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. upon impella implant there were suction alarms with a flow of 0.8. staff noticed the cannula was kinked up by the outlet cage. doctor attempted to reposition the pump, but was unsuccessful and the kink could not come out. the pump was replaced successfully. no patient harm reported. | [5 empty] | The investigation into the reported low pump flow and product damage has been completed since the original report was filed. no product was returned. as per clinical, the pump immediately got low flows of 0.8 l/min with constant suction alarm and the doctor confirmed the cannula was kinked up by the outlet cage. data logs were reviewed and multiple suction alarms with flow less than 1 l/min was found. the cause of the low pump flow issue was most likely due to kinked cannula occurred during delivery of the pump and low flow with suction alarms observed per log and clinical review. | The impella device was not received from the customer, therefore, investigation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
33229 | 1220648-2023-04586 | 18008436 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18008436&pc=OZD | 09/25/2023 | 01/16/2024 |  | 10/25/2023 |  |  | 09/25/2023 | 10/25/2023 |  |  | 04/30/2025 |  | 10/25/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024338775 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  |  | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/25/2023 |  |  |  |  | The user facility reported a 64-year-old male (race unknown) in non-st elevated myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. during support, the labs were hemolyzed and the pump was positioned incorrectly against the mitral valve. the md was unable to reposition the pump and after multiple attempts, the decision was made to remove the impella 5.5. | [5 empty] | The investigation into the positioning issues and the hemolysis has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the root cause of hemolysis was most likely due to pump was not in the proper position and the root cause of positioning issue was most likely due to use related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
33230 | 1220648-2023-04392 | 17994506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17994506&pc=OZD | 06/07/2022 | 10/20/2025 |  | 10/24/2023 |  |  | 06/07/2022 | 10/24/2023 |  |  | 03/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023198279 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  |  | White | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/07/2022 |  |  |  |  | The user facility reported a 66-year-old female patient was implanted with an impella device for mechanical circulatory support. the complainant reported an increase in purge pressure that was resolved by adding tissue plasminogen activator (tpa) in the purge fluid. purge flow returned to normal and the patient was not harmed. | [5 empty] | The device was returned and the investigation is ongoing. upon conclusion of the investigation, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records.
33231 | 1220648-2023-04393 | 17994573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17994573&pc=OZD | 07/12/2023 | 05/23/2025 |  | 10/24/2023 |  |  | 07/12/2023 | 10/24/2023 |  |  | 03/31/2025 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024324184 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/12/2023 |  |  |  |  | The user facility reported that a 59-year-old male patient implanted with the impella 5.5 device for mechanical circulatory support experienced a purge leak and pump stop. the purge fluid leak was observed from the area of catheter and grey junction on distal end of the red impella plug. the following day the motor current was increased, and pump failure occurred. the pump was subsequently exchanged without issue. there was no patient harm reported. | [5 empty] | The device was returned for evaluation and an investigation is ongoing. upon investigation closure, a supplemental mdr will be filed. per the ifu: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records. | The cause of the pump stop was blood ingress into the motor due to a purge leak. the purge leak stems from the failure of the purge lumen bond at the stainless steel connector within the red handle. the root cause of the purge leak - pump could not be determined as insufficient clinical details were provided. this report is being filed as part of a retrospective review of historical records.
33232 | 1220648-2023-04394 | 17994706 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17994706&pc=OZD | 05/18/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 05/18/2022 | 10/24/2023 |  |  | 02/29/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023190583 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/18/2022 |  |  |  |  | The user facility reported a 61 year old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there were saturated flows on the pump. the pump remains on for support without harm or injury despite the flow alarms. | [5 empty] | The investigation into the saturation issues was completed. neither the product or data log were returned for evaluation. therefore, the cause of the saturated flow could not be determined. this report is being filed as part of a retrospective review of historical records.
33233 | 1220648-2023-04395 | 17994725 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17994725&pc=OZD | 05/19/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 05/19/2022 | 10/24/2023 |  |  | 01/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023182462 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | Pressure Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/19/2022 |  |  |  |  | The user facility reported a 67 year-old female was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the 5.5 was removed and replaced at day 8 due to a purge sidearm leakage. in addition to the purge leakage, there was continuous purge pressure low alarms. there was no harm or injury from the interruption in the support. | [5 empty] | The investigation into the purge sidearm leak at the filter has been completed. the device was returned without any data logs for evaluation. visual inspection was performed finding no damage or abnormalities on the pump. leak testing was performed and did not observe any leaking from the purge sidearm. during the testing, the purge metrics were within specifications and there were no alarms. the cause of the purge leak could not be determined as the device operated as intended. this report is being filed as part of a retrospective review of historical records.
33234 | 1220648-2023-04396 | 17994762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17994762&pc=OZD | 05/16/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 05/19/2022 | 10/24/2023 |  |  | 01/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 SMARTASSIST SET, JP |  |  | ABIOMED, INC. |  |  | US |  | 2023179434 |  | 1000211 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/19/2022 |  |  |  |  | The user facility reported a male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that after 3 days, support was ended due to low flows. the pump was explanted without harm or injury. thrombotic material was on the pump at explant. the adhered thrombus did not cause harm. | [5 empty] | The investigation for the reported low flow issue has been completed. the returned pump was visually inspected and biomaterial was observed against the purge gap. the pump was connected to a purge cassette was was run through de-air and high purge pressure was initially reproduced but eventually cleared with a purge flow of 3.3 ml/hr. the cause of the high purge pressure was no heparin in the purge likely leading to the biomaterial ingestion. this report is being filed as part of a retrospective review of historical records.
33235 | 1220648-2023-04397 | 17994768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17994768&pc=OZD | 01/19/2022 | 09/16/2024 |  | 10/24/2023 |  |  | 01/19/2022 | 10/24/2023 |  |  | 09/30/2023 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022149886 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/19/2022 |  |  |  |  | Additional information noted patient care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 74 year-old male was implanted with an impella 5.5 device mechanical circulatory support. it was reported that after starting support and obtaining optimal positioning, the pump had purge flow (pf) low and purge pressure (pp) high issues. the surgeon tried troubleshooting by de-airing purge, changing the 5% dextrose fluids, purge tubing, removing the sidearm protector, straightening the red plug, and changing the console. however, none of these steps resolved the issue. the surgeon used tpa (tissue plasminogen activator) to resolve the purge issue. there was no patient injury reported. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-04397 was provided. | The investigation into the reported purge flow low and purge pressure high issues has been completed. the medical center did not return any impella products for benchtop analysis. however, the clinical report was evaluated. the root cause of the high purge pressure and purge flow low issues was most likely ingested biomaterial since tpa resolved the issue. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ this report is being filed as part of a retrospective review of historical records.
33236 | 1220648-2023-04398 | 17994800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17994800&pc=OZD | 05/22/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 05/23/2022 | 10/24/2023 |  |  | 01/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023183971 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/23/2022 |  |  |  |  | The user facility reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the procedure was aborted due to a kink in the catheter. there was no patient harm reported. | [5 empty] | The investigation for the reported kinked catheter has been completed. the device was not returned for evaluation. however, after reviewing the clinical information, it was determined that the cause of the delivery issue was patient condition related. the patient's vascular anatomy caused the catheter to kink. this report is being filed as part of a retrospective review of historical records.
33237 | 1220648-2023-04399 | 17994820 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17994820&pc=OZD | 06/06/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 06/06/2022 | 10/24/2023 |  |  | 02/29/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023189549 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Break | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/06/2022 |  |  |  |  | The user facility reported a 65 year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there were abnormal flows during an electrophysiology (ep) study. when the ep study was completed, the flows returned to normal. additionally, the pump had a purge line break that prompted the team to perform a bypass. the team had just replaced the purge. there was no patient harm reported. | [5 empty] | The investigation into the saturation flow issues and the sidearm damage has been completed. the device was not returned for evaluation. however, the clinical report was evaluated. the cause of the purge leak was most likely sidearm yellow luer damage from excessive force. no bicarbonate in the purge. the cause of the temporary saturated flow upon the electrophysiology procedure was not determined since product was not returned. the issue was related to the vt ablation procedure since it resolved following the procedure. this report is being filed as part of a retrospective review of historical records.
33238 | 1220648-2023-04400 | 17994880 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17994880&pc=OZD | 06/06/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 06/06/2022 | 10/24/2023 |  |  | 05/31/2023 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022086832 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/06/2022 |  |  |  |  | The user facility reported a 67 year old female was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was possible oversaturation. the impella flows were not corresponding with the p-level. there were multiple troubleshooting attempts but were unsuccessful. the catheter was exchanged for new device. there was no patient harm reported. | [5 empty] | The investigation for the oversaturation issue has been completed. the field data logs were reviewed and found that pump flow was low from the start of the run until a motor current spike which was followed by the saturated flow. the biomaterial was likely present at the start of the run causing low pump flow and then was ingested into the impeller causing the spike in motor current and then the resistance to the impeller resulting in the saturated flow for the remainder of the run. the cause of the saturation issues was due to biomaterial. this report is being filed as part of a retrospective review of historical records.
33239 | 1220648-2023-04401 | 17994948 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17994948&pc=OZD | 06/09/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 06/09/2022 | 10/24/2023 |  |  | 10/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023190521 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Other | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/09/2022 |  |  |  |  | The user facility reported a 70yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that during placement of the pump, the catheter kinked on itself due to angle of entry. the device was swapped out. there was no patient harm reported. | [5 empty] | The investigation for the reported kinked cannula has been completed. the device was returned for evaluation. a cannula kink was found near the outflow cage. no other damage abnormalities were found. after reviewing the clinical information, it was determined that the cause of the delivery issue was a steep insertion angle. this report is being filed as part of a retrospective review of historical records.
33240 | 1220648-2023-04402 | 17994969 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17994969&pc=OZD | 06/10/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 06/10/2022 | 10/24/2023 |  |  | 02/29/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023188524 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/10/2022 |  |  |  |  | The user facility reported a 59 year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was a purge system leak at the yellow luer causing the system to not be able to maintain adequate pressure. the pump has been used beyond indication. there was no harm or injury. | [5 empty] | The investigation into the purge system leak has been completed. the device was not returned for evaluation. however, the clinical report has been reviewed. it was found that prior to reported leak issue, the customer stated that the yellow luer had to be soaked in warm water to loosen the connection and bicarbonate was being used in the purge. since the device was not returned, the root cause of the purge leak nor the exact location of the leak could not be determined. this report is being filed as part of a retrospective review of historical records.
33241 | 1220648-2023-04403 | 17994986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17994986&pc=OZD | 06/14/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 06/14/2022 | 10/24/2023 |  |  | 03/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023203053 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/14/2022 |  |  |  |  | The user facility reported a 54yo male was implanted with an impella 5.5 device for mechanical circulatory support as a bridge to heart transplant. it was reported that there was a leak at the sidearm micron filter. the pump remained on for support. there was no patient harm reported. | [5 empty] | The investigation for the reported purge leak has been completed. neither the device nor data logs were returned for evaluation. therefore, the root cause of the reported side leak could not be determined. the instructions for use provides details on the maintenance of the impella sidearm. it states ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g.,infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿. this report is being filed as part of a retrospective review of historical records.
33242 | 1220648-2023-04404 | 17995011 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17995011&pc=OZD | 06/16/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 06/16/2022 | 10/24/2023 |  |  | 01/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023183969 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/16/2022 |  |  |  |  | The user facility reported a 59yo female was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the pump stopped during transport of the patient. the patient remained on support. there was no patient harm reported. | [5 empty] | The investigation into the pump stop has been completed. the device nor the data logs were returned for evaluation. however, the clinical information was sufficient in determining the root cause, given the results of prior failure investigations. the movement of the patient most likely contributed to the pump stop. instructions for use for the related event are as follows: ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
33243 | 1220648-2023-04405 | 17995034 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17995034&pc=OZD | 06/23/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 06/24/2022 | 10/24/2023 |  |  |  |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/24/2022 |  |  |  |  | The user facility reported a 65yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was a side arm leak prompting a purge bypass configuration. the issue was resolved with no harm to the patient. | [5 empty] | The investigation for the reported purge sidearm leak has been completed. no product or data was returned to the bhl. the usage of bicarbonate most likely contributed to the purge leak. with no product returned, the location of the purge leak is undetermined. this report is being filed as part of a retrospective review of historical records.
33244 | 1220648-2023-04406 | 17995075 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17995075&pc=OZD | 06/04/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 06/04/2022 | 10/24/2023 |  |  | 04/30/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023211220 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/04/2022 |  |  |  |  | The user facility reported a 48 year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that blood was leaking from the red plug on the pump. bone-wax was applied. the pump was supporting for 2 days when this occurred. the pump was replaced. there was no patient harm reported. | [5 empty] | The investigation into the blood leakage from the red handle has been completed. the device was not returned for evaluation. however, the clinical report was evaluated. the root cause of the blood in the red handle/product damage was most likely a result of a puncture in the catheter as the blood would have had to enter through the catheter. this report is being filed as part of a retrospective review of historical records.
33245 | 1220648-2023-04407 | 17995126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17995126&pc=OZD | 06/30/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 06/30/2022 | 10/24/2023 |  |  | 02/29/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023188883 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/30/2022 |  |  |  |  | The user facility reported a 76yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that while attempting to reposition the impella with tee, the placement signal was noted to show that the device had been pulled across the aortic valve and into the aorta. several attempts were made to reposition the impella. it was then decided to abort repositioning and not to implant the device due to the risk aortic dissection as well as bleeding complications that the patient was exhibiting. | [5 empty] | The investigation for the reported positioning issue has been completed. data logs showed recurrent "suction" alarms before the placement signal waveform rapidly became aortic. soon after the "impella position in aorta" alarm occurs and all pump performance metrics suddenly drop to 0. no damage or abnormalities were found on the returned pump. after reviewing the clinical information, it was determined that the cause of the positioning issue was wireless re-access. the impella device is not indicated for wireless re-access. this report is being filed as part of a retrospective review of historical records.
33246 | 1220648-2023-04408 | 17995157 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17995157&pc=OZD | 07/03/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 07/03/2022 | 10/24/2023 |  |  | 03/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023206007 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  |  | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/03/2022 |  |  |  |  | The user facility reported a 50 year-old male was implanted with an impella 5.5 device for mechanical circulatory support as a bridge to transplant. it was reported that there was a purge sidearm leak for which a bypass was performed . there was sodium bicarbonate running through the purge. the team explanted and replaced the pump. there was no harm reported. | [5 empty] | The investigation into the purge leak has been completed. the pump and data logs were returned for evaluation. data logs were reviewed and showed recurrent "purge pressure low" alarms upon a sudden drop in purge pressure to almost zero, followed by erratic rises and drops in pressure. after about 45 minutes, the purge metrics stabilized. the pump was inspected and found no notable damage or abnormalities on the returned sidearm. however, the purge tubing between the sidearm and the pump plug was cut off. no leaking was observed on any part of the sidearm during reproduction testing. after reviewing the clinical information, it was determined that the cause of the purge leak was most likely damage to the sidearm filter to tubing joint due to excessive force from the patient. this report is being filed as part of a retrospective review of historical records.
33247 | 1220648-2023-04436 | 17995994 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17995994&pc=OZD | 09/25/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 09/25/2022 | 10/24/2023 |  |  | 06/30/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023232941 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/25/2022 |  |  |  |  | The user facility reported a 64 year-old male was implanted with an impella 5.5 device for mechanical circulatory support for bridge to transplant. it was reported that there was a flow saturation noted. a purge bypass performed but eventually the pump stopped. the patient was taken to the operating room for a pump replacement. there was no patient harm reported. | [5 empty] | The investigation into the purge leak and the pump stop has been completed. the data logs only were returned for evaluation. the data logs were analyzed and confirm a low purge pressure/ purge leak event. the logs show the purge pressure recovering after the bypass was performed. a day after the purge pressure event, a sharp spike in purge pressure is seen, this is indicative of a kink in the purge line. purge system blocked alarms are triggered and the motor current begins to rise without a change in p-level. the flow also rises and becomes saturated as a result of the increased motor current. a high motor current pump stop is triggered when the increasing motor current reaches the threshold. the slow increase in motor current following the spike in purge pressure is characteristic of blood ingress due to kinked purge line. the most likely root cause of the pump stop was determined to be a kinked purge line. the root cause of the purge sidearm leak was unable to be determined as no product or additional clinical details were provided so the location of the leak could not be identified. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33248 | 1220648-2023-04437 | 17996017 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17996017&pc=OZD | 11/12/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 11/12/2022 | 10/24/2023 |  |  | 06/30/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023229213 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  | White | Malfunction | No | Yes | Fluid/Blood Leak | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/12/2022 |  |  |  |  | The user facility reported a 65yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was a leak at the purge sidearm. a purge bypass was performed but was unsuccessful. saturated flows followed by a pump stop occurred. there was no patient harm reported. | [5 empty] | The investigation into the pump stop has been completed. no product was returned for investigation. the cause of the pump stop was a leak from the sidearm but the exact location of the leak was unknown. it is also unknown how the damage to the sidearm occurred. instructions for use for the related event are as follows: ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
33249 | 1220648-2023-04438 | 17996065 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17996065&pc=OZD | 10/05/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 10/05/2022 | 10/24/2023 |  |  | 05/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023223644 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/05/2022 |  |  |  |  | The user facility reported a 52 year-old female was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the purge sidearm cracked after 5 days of use. a purge sidearm bypass initiated which lasted for a little over 12 hours. then leaking started occurring in red side arm plug which was not able to be remedied. the pump was removed earlier than anticipated at bedside by surgeon but did not require pump replacement or escalation of current therapy. there was no patient harm reported. | [5 empty] | The investigation into the purge sidearm leak has been completed. the device was not returned for evaluation. the clinical report was reviewed stating that the leak was at the yellow luer but it was unknown if the leak was from the luer or luer bond. however, no bicarbonate was being used. without the device, the root cause of the purge leak could not be determined. this report is being filed as part of a retrospective review of historical records.
33250 | 1220648-2023-04439 | 17996095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17996095&pc=OZD | 10/12/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 10/12/2022 | 10/24/2023 |  |  | 06/30/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023237485 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/12/2022 |  |  |  |  | The user facility reported a 54yo male was implanted with an impella 5.5 device for mechanical circulatory support. multiple suction alarms were observed. an echocardiogram verified the positioning of the pump and the device's performance level was decreased. the issue remained, therefore, patient support continued under advised monitoring. | [5 empty] | The investigation for the reported low flow has been completed. the data logs were reviewed and confirm suction alarms and low motor current pulsatility. no spikes in motor current indicating biomaterial ingestion were observed. the product was returned, the catheter was cut but no biomaterial of kinks were found. the root cause of the low flow was most likely patient condition. this report is being filed as part of a retrospective review of historical records.
33251 | 1220648-2023-04440 | 17996123 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17996123&pc=OZD | 09/17/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 09/17/2022 | 10/24/2023 |  |  | 04/30/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023214217 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/17/2022 |  |  |  |  | The user facility reported a 65 year old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that an extensive clot was found both in the inflow and outflow cages after explant. there were no alarms or patient issues during use. | [5 empty] | The investigation into the clot formation in the pump has been completed. no product was returned. images received from the field confirmed the presence of considerable masses of biomaterial ingested into the inflow cage and the outflow cage. data logs showed that all relevant pump performance metrics were within specification throughout support, with no relevant alarms occurring. it was noted that the pump flow waveform became more diastolic as weaning began about 48 hours prior to explant; however, no pump performance metrics were affected at this time. no heparin was used in the purge solution. after reviewing the clinical information, it was determined that the cause of the thrombus ingestion was most likely patient condition related. the instructions for use provides suggestions for prevention of thrombus entering the catheter as it states ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ this report is being filed as part of a retrospective review of historical records.
33252 | 1220648-2023-04441 | 17996144 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17996144&pc=OZD | 10/17/2022 | 09/22/2024 |  | 10/24/2023 |  |  | 10/17/2022 | 10/24/2023 |  |  | 04/30/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023254989 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Improper Flow or Infusion | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/17/2022 |  |  |  |  | The procedural outcome noted that the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | The user facility reported a 71yo female was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that after implant, the readings showed a low flow through the pump coinciding with suction. a transesophageal echo (tee) probe showed that the pump had flipped positioning towards the mitral. however, upon attempting to reposition the pump, a placement signal unreliable alarm was displayed and the ao and lv waveforms were lost. good positioning was obtained and patient support continued with the implanted pump. there was no patient harm reported. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-04441 was provided in sections b1, b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding. | The investigation into the reported low flow issue has been completed. the device nor data logs were returned for evaluation. however, based upon clinical information, the root cause of the low flow was determined to be improper positioning based on the clinical details. this report is being filed as part of a retrospective review of historical records.
33253 | 1220648-2023-04442 | 17996166 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17996166&pc=OZD | 10/16/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 10/16/2022 | 10/24/2023 |  |  | 07/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023252166 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Defective Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/16/2022 |  |  |  |  | The user facility reported a 51yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the customer was having issues with torque on the catheter. when troubleshooting the catheter, a "controller error" alarm occurred. the controller was swapped out without success. the decision was made to replace the catheter and the issue resolved using the initial console. there was no patient harm reported. | [5 empty] | The investigation for the "controller error" has been completed. the device and the data logs were returned for evaluation. the logs were reviewed and show the placement signal cut out before the patient cable is unplugged. the log pattern suggests a fiber break in the red handle. when the fiber is broken and the pump is plugged into a new console it can trigger a console failure because the console is unable to read the optical parameters. upon evaluation of the device, the patient cable bond was found not intact which resulted in broken optical fiber in the red handle. the root cause of the optical signal issue was determined to be a broken fiber line due to an insufficient patient cable bond. this report is being filed as part of a retrospective review of historical records.
33254 | 1220648-2023-04443 | 17996187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17996187&pc=OZD | 10/17/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 10/17/2022 | 10/24/2023 |  |  | 03/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023209653 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/17/2022 |  |  |  |  | The user facility reported a 44yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that after a scheduled explantation of a pump, a clot was noted. despite the administration of systemic anticoagulants, the patient's partial thromboplastin time was not within a therapeutic range during support. there was no medical intervention was required. | [5 empty] | The investigation for the thrombus has been completed. the pump was returned for evaluation. upon visual inspection, biomaterial was confirmed in the outflow. therefore the cause of the thrombus was related to biomaterial in the pump as a result of subtherapeutic lab results. the instructions for use discussed prevention of thrombus to subsequently avoid pump stop. it states: "maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ this report is being filed as part of a retrospective review of historical records.
33255 | 1220648-2023-04444 | 17996249 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17996249&pc=OZD | 04/24/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 04/24/2023 | 10/24/2023 |  |  | 12/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024292407 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 62-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella 5.5 device for mechanical circulatory support. the patient had significant bleeding from his right axilla requiring several units of blood (prbc) and emergent trip to the or for exploration of bleed. the graft/axillary artery anastamosis had separated and required surgical repair. there was no problem with the impella pump, just the surgical attachement of the graft to the artery. | [5 empty] | The investigation into the reported bleeding/vascular damage has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the cause of the access site bleeding was most likely use related and related to the initial attachment of the graft to the artery. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33256 | 1220648-2023-04445 | 17996258 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17996258&pc=OZD | 10/27/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 10/27/2022 | 10/24/2023 |  |  | 07/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023254793 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/27/2022 |  |  |  |  | The user facility reported a 59yo female was implanted with an impella 5.5 device for mechanical circulatory support. it was that the patient experienced hemolysis during impella support. the pump was repositioned in an attempt to troubleshoot the issue but the condition initially remained. impella support continued, however, the pump was ultimately explanted. | [5 empty] | The investigation for the hemolysis event was completed. the pump was returned for evaluation. the inside of the pump was visually inspected with a borescope. no sharp edges, broken blades, or biomaterial were observed. the catheter was cut in the field so no hemolysis test could be performed. "impella position unknown" alarms were observed in the data logs through out the case. these alarms lasted less than a minute. no other evidence of the pump out of the position was observed with the logs. flows were within specification and no suction alarms were triggered. the clinical report specified repositioning did not resolve the issue. no evidence of biomaterial was found. the root cause of the hemolysis was most likely the patient condition because no abnormalities found in the pump. the instructions for use as it relates to the occurrence of hemolysis states the following: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. it also states that ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
33257 | 1220648-2023-04446 | 17996291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17996291&pc=OZD | 10/30/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 10/30/2022 | 10/24/2023 |  |  | 06/30/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023232099 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/30/2022 |  |  |  |  | The user facility reported a 51 year old male with heart failure/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was a spike in motor current, followed by motor current flattening, saturated lv signals and saturated pump flow issues. the impella 5.5 was removed. there was no patient harm reported. | [5 empty] | The investigation into the motor current issues has been completed. the pump and data logs were returned for evaluation. the data logs were reviewed and confirm a motor current spike followed by a lack of pulsatility in motor current which is indicative of ingested biomaterial. the pump was evaluated and biomaterial was found in the outlet cage of the pump. the root cause of the saturated flow was determined to be biomaterial ingestion adding additional resistance to the impeller, increasing motor current and flow. instructions for use for the related event are as follows: ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33258 | 1220648-2023-04447 | 17996297 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17996297&pc=OZD | 09/26/2023 | 01/22/2024 |  | 10/24/2023 |  |  | 09/27/2023 | 10/24/2023 |  |  | 06/30/2025 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024375555 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/26/2023 |  |  |  |  | The user facility reported a 58-year-old male in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient had red urine on impella support. echo performed bedside showed impella in deep and directed towards the mitral valve. md repositioned under echo guidance. staff plans to give volume for suspected hemolysis. the patient received fresh frozen plasma. | [5 empty] | The investigation into the positioning issues and the hemolysis has been completed. the device was not returned for investigation. based on clinical description, the data logs were reviewed and there were no positioning alarms or significant motor current spikes. the placement signal was trending up towards the end of the case and there were suction conditions present. the clinical stated imaging was used and confirmed that the pump was too deep and directed towards the mv. it was also stated that blood volume was given to resolve the hemolysis. the root cause of the positioning issues is most likely due to use while the root cause of the hemolysis was most likely due to the patient¿s condition. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
33259 | 1220648-2023-04448 | 17996407 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17996407&pc=OZD | 11/02/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 11/02/2022 | 10/24/2023 |  |  | 04/30/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023214236 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/02/2022 |  |  |  |  | The user facility reported a 39 year-old female was implanted with an impella 5.5 device for mechanical circulatory support. it was reported after a purge cassette change, a purge system open alarm appeared. multiple troubleshooting attempts were made including purge cassette changes, de-airing purge system, purge bypass with extension tubing and securing sidearm with tongue depressor. a leak was observed on the yellow luer lock portion that sits inside the purge sidearm plastic encasing. the pump was replaced. there was no patient harm reported. | [5 empty] | The investigation into the purge leak has been completed. the device was not returned for evaluation. however, the clinical report has been reviewed. the root cause of the purge leak was unable to be determined because the product was not returned so the location of the leak could not be identified. this report is being filed as part of a retrospective review of historical records.
33260 | 1220648-2023-04449 | 17996426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17996426&pc=OZD | 04/01/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 04/01/2023 | 10/24/2023 |  |  | 11/30/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024295016 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 52-year-old male noted as high risk percutaneous cardiac intervention was implanted with an impella cp device for mechanical circulatory support. the complainant had an impella cp placed for support but after twenty hours the team found the pump to have come out of left ventricle (lv) positioning and so the team explanted. | [5 empty] | The investigation into the reported "positioning issues" has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation of the data logs show several "position in aorta" alarm. the cause of the issue was not determined since product was not returned and sufficient clinical information was not provided. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
33261 | 1220648-2023-04450 | 17996443 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17996443&pc=OZD | 11/11/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 11/11/2022 | 10/24/2023 |  |  | 05/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023226007 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/11/2022 |  |  |  |  | The user facility reported a 76 year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the catheter built up excessive torque while manipulating into the ventricle causing an ¿impella defective¿. the pump had to be replaced. no patient harm was reported. | [5 empty] | The investigation into the pump stop has been completed. the pump and the data logs were returned for evaluation. data logs showed the "impella defective" alarm triggered immediately when attempting to start the pump. the pump failed to activate as all pump performance metrics remained at 0. the pump failed to be recognized when connected to an aic. inspection of the pump revealed the bond between the patient cable and the kink protector had failed, as the patient cable was able to freely rotate and translate relative to the kink protector. within the red handle, all wires between the patient cable and pcb were tightly wound around each other. electrical resistance testing revealed an open line in the sda connection (yellow wire). further inspection confirmed there was considerable damage to the sda wire due to the twisting. in conclusion, it was determined that the cause of the pump's failure to start was an open electrical line due to the failure of the bond between the patient cable and kink protector. instructions for use for the related event are as follows: ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
33262 | 1220648-2023-04451 | 17996487 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17996487&pc=OZD | 11/15/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 11/15/2022 | 10/24/2023 |  |  | 06/30/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023246712 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/15/2022 |  |  |  |  | The user facility reported a 50 year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that during a purge cassette change, the threading of the side arm yellow luer connection broke. a purge bypass was performed, after cutting the purge sidearm off. there was no additional information provided. no patient harm was reported. | [5 empty] | The investigation into the yellow luer damage was completed. the device was not returned for evaluation. however, the clinical information was sufficient in determining the root cause, given previous investigations for the same failure. it was concluded that the cause of the purge sidearm damage was as a result of excessive force, as the staff used clamps to disconnect. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33263 | 1220648-2023-04452 | 17996498 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17996498&pc=OZD | 04/24/2023 | 09/26/2025 |  | 10/24/2023 |  |  | 04/24/2023 | 10/24/2023 |  |  | 12/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024307339 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | Difficult to Advance | Material Twisted/Bent |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was escalated from a intra-aortic ballon pump to an impella 5.5 via standard axillary surgical approach. device passage was difficult and then aborted after a kink formed in the catheter preventing device from being advanced further. a new wire and device were inserted and successful. there was no patient harm. | The complainant reported that a patient was escalated from a intra-aortic ballon pump to an impella 5.5 via standard axillary surgical approach. device passage was difficult and then aborted after a kink formed in the catheter preventing device from being advanced further. a new wire and device were inserted and successful. no patient harm. additionally, placement signal lost. |  |  |  |  | The investigation into the reported 'delivery issues" has been completed. product was returned for investigation. the device was visually inspected and found the catheter was kinked below mh and the guidewire was kinked and damaged. there were no other issues found on the rest of the pump. per clinical description, the patient had six year old bio-prosthetic atrioventricular (av) which was very challenging to cross. therefore, the root cause of delivery issue was most likely due to the patient condition. this report is being filed as part of a retrospective review of historical records. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. visual inspection found the catheter was kinked below mh & the gw was kinked & damaged. no other issues were found on the rest of the pump. per clinical description, the patient had 6 year old bioprosthetic av which was very challenging to cross. therefore, the root cause of delivery issue was most likely due to the patient condition. the rt logs are not long enough to show any 5s parameters. the cause of the optical signal issue was a damaged optical fiber line which was kinked at the location of the catheter kink next to the kink protector. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issue "2917" as part of a retrospective review of optical signal issues in accordance with fda recommendation.
33264 | 1220648-2023-04453 | 17996503 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17996503&pc=OZD | 10/16/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 10/16/2022 | 10/24/2023 |  |  | 03/31/2023 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022069633 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/16/2022 |  |  |  |  | The user facility reported a 82 year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that a patient rolled over and the sidearm of the impella broke. a purge bypass was done. there was no patient harm reported. | [5 empty] | The investigation into the purge sidearm leak has been completed. the device was not returned for evaluation. however, the clinical report has been reviewed. the root cause of the purge leak was most likely patient movement. this report is being filed as part of a retrospective review of historical records.
33265 | 1220648-2023-04454 | 17996516 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17996516&pc=OZD | 05/10/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 05/10/2023 | 10/24/2023 |  |  | 12/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023284596 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Optical Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a placement signal unreliable after adjusting position during impella 5.5 insertion, but after the access incision was closed. of note the pump was torqued/moved after access was closed. there was no harm done to the patient. the pump remained on for support without signal being reliable. the 5.5 was placed as care was escalated from the cp in the bridge to lvad device. | [5 empty] | The investigation into the optical signal loss is completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the data logs show that at the time of repositioning the pump was unplugged from the console and replugged in. upon replug the imc logs show a series of errors. the console is unable to read the usage data off of the pump. this causes a placement signal unreliable alarm. these alarms are triggered due to eeprom corruption. the root cause of the system-double eeprom corruption is unable to be determined as no product was returned for evaluation.
33266 | 1220648-2023-04455 | 17996527 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17996527&pc=OZD | 11/10/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 11/11/2022 | 10/24/2023 |  |  | 07/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023255191 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/11/2022 |  |  |  |  | The user facility reported a 68 year-old male in cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the purge sidearm was broken. a purge system bypass was configured and support proceeded. there was no patient harm reported. | [5 empty] | The investigation for the damaged side arm has been completed. the device was not returned for investigation. however, the clinical information was sufficient in determining the root cause, given the the nurse reportedly broke the sidearm with excessive force. it was concluded that the cause of the broken purge sidearm was due damage from excessive force. this report is being filed as part of a retrospective review of historical records.
33267 | 1220648-2023-04456 | 17996550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17996550&pc=OZD | 11/10/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 11/10/2022 | 10/24/2023 |  |  | 07/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023252325 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/10/2022 |  |  |  |  | The user facility reported a 67 year-old female was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that during support, the purge flow increased and the purge pressures dropped. during evaluation, the nurse noticed a leak in the extension tubing coming from the yellow luer on the sidearm. a purge bypass was performed and the flows and pressures returned to normal. there was no patient harm reported. | [5 empty] | The investigation into the purge leak was completed. neither the product nor the data logs were returned for investigation. however, the clinical information was sufficient in determining the root cause, given the results of prior failure investigations. it was concluded that the cause of the purge leak was sidearm yellow luer damage due to sodium bicarbonate use. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33268 | 1220648-2023-04459 | 17996604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17996604&pc=OZD | 11/23/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 11/23/2022 | 10/24/2023 |  |  | 09/30/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023259440 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/23/2022 |  |  |  |  | The user facility reported a 59 year-old male with an impella 5.5 device for mechanical circulatory support as a bridge to transplant. it was reported that leaking was noted at the red plug. after troubleshooting, the leak seemed to occur towards the end of the red plug or purge side arm. no additional information was provided. there was no patient harm reported. | [5 empty] | The investigation into the purge leak has been completed. the device was not returned for evaluation. however, the clinical report has been reviewed. without additional information, the root cause of the purge leak was unable to be determined because the product was not returned. this report is being filed as part of a retrospective review of historical records.
33269 | 1220648-2023-04458 | 17996605 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17996605&pc=OZD | 02/07/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 02/07/2023 | 10/24/2023 |  |  | 11/30/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023277989 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Detachment of Device or Device Component | [6 empty] | Transient Ischemic Attack | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 67-year-old male noted as high risk percutaneous cardiac intervention (hrpci) was implanted with an impella 5.5 device for mechanical circulatory support. the patient had a transient ischemic attack (tia) that altered mental status, facial droop, slurring of speech lasting approximately 2-3 minutes, and then recovered. patient has a history of tias and stroke. patient stable and recovering further during the next days. additionally, the device's cannula became separated from the remainder of the device during difficult removal from graft after the 28 day support. | [5 empty] | The investigation into the reported stroke/tia and cannula detachment has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the root cause of the neurological event is due to the patient's condition because of their history of strokes and transient ischemic attack's (tia). the root cause of the cannula detachment is due to a use issue of using excessive force. instructions for use for the related event are as follows: ¿avoid manual compression of the inlet and outlet areas of the cannula assembly.¿ ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33270 | 1220648-2023-04457 | 17996607 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17996607&pc=OZD | 04/27/2023 | 09/21/2024 |  | 10/24/2023 |  |  | 04/27/2023 | 10/24/2023 |  |  | 01/31/2025 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024296616 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Failure to Advance | Material Twisted/Bent | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported that the surgeon was unable to push pump across stenosis within left axillary artery. many attempts were made, with no success. a cannula kink was observed. multiple techniques were subsequently used thereafter in an attempt to advance the pump, but each was unsuccessful. there was no patient harm resulting from the complication. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b1, b2, b5, h1, and h6. | The investigation into the reported "delivery issue" has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the cause of the issue was patient condition related to stenosis in the axillary artery. this report is being filed as part of a retrospective review of historical records.
33271 | 1220648-2023-04461 | 17996632 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17996632&pc=OZD | 05/08/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 05/08/2023 | 10/24/2023 |  |  | 04/30/2025 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024329634 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a failure to deliver the device due to 77-year-old female patient¿s anatomy. it was noted that the device would kink during attempts to advance the device. the pump insertion was aborted and an impella cp was used without further issue. there was no harm to the patient. | [5 empty] | The investigation into the reported "delivery issues" has been completed. the product was returned for investigation. the device was visually inspected and found there were kinks observed on the catheter near the motor housing and on the cannula. as per clinical investigation, the physician was trying to insert a 5.5 via right axillary artery. it was noted that the patient had tortuous vessels. the investigation determined that the root cause of the issue was related to the patient's tortuous vessels. this report is being filed as part of a retrospective review of historical records.
33272 | 1220648-2023-04460 | 17996646 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17996646&pc=OZD | 01/15/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 01/15/2023 | 10/24/2023 |  |  | 09/30/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023263525 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 58-year-old male noted as high risk percutaneous cardiac intervention (hrpci) was implanted with an impella 5.5 device for mechanical circulatory support. the complainant had an impella 5.5 support ongoing when after over six days the pump was removed and replaced due to flow saturation. | [5 empty] | The investigation into the reported "saturated flow" has been completed. product was returned for investigation. the device was visually inspected and found biomaterial was observed in the purge gap. the pump was connected to an automated impella controller (aic) and placement signal was -25/-25 mmhg. there was no loose patient cable or catheter. the investigation determined that the root cause of the saturated flow was biomaterial in the motor. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
33273 | 1220648-2023-04463 | 17996692 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17996692&pc=OZD | 12/01/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 12/01/2022 | 10/24/2023 |  |  | 05/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023218337 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/01/2022 |  |  |  |  | The user facility reported 83yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there were high purge pressure alarms with low purge flows on the pump. there were no kinks in purge tubing observed. patient was started on tissue plasminogen activator (tpa) protocol and the purge flows returned to normal. there was no patient harm reported. | [5 empty] | The investigation for the high purge pressure has been completed. the data log confirms high purge pressure and purge system blocked alarms. the purge pressure begins to rise on the first day of the case and rises over the course of about 25 minutes until there is the first high purge pressure alarm. the logs then show the high purge pressure resolve the next day. a visual inspection of the pump did not reveal any abnormalities. the pump was attached to a purge cassette and a console and ran in the lab and the high purge pressure was not reproduced. the cause of the high purge pressure was most likely biomaterial. the high purge pressure was resolved with the infusion of tpa. this report is being filed as part of a retrospective review of historical records.
33274 | 1220648-2023-04462 | 17996696 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17996696&pc=OZD | 06/08/2023 | 03/07/2025 |  | 10/24/2023 |  |  | 06/08/2023 | 10/24/2023 |  |  | 03/31/2025 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024326598 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | Material Puncture/Hole | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 67-year-old male in st elevated myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that blood was oozing out of the red plug during impella support. despite troubleshooting with the red plug, the bleeding continued. as a result of the noted issue, the pump was removed and replaced with another impella 5.5 pump. during visual inspection of the returned pump a hole in the catheter was found at the 30cm marking. there was no lasting patient harm as a result of the issue. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported product damage issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that a hole in the catheter was found at the 30cm marking. the cause of the product damage was blood ingress due to a hole in the catheter resulting from improper use. this report is being filed as part of a retrospective review of historical records.
33275 | 1220648-2023-04464 | 17996707 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17996707&pc=OZD | 07/07/2022 | 09/20/2024 |  | 10/24/2023 |  |  | 07/07/2022 | 10/24/2023 |  |  | 01/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 PUMP SET JP |  |  | ABIOMED, INC. |  |  | US |  | 2023178299 |  | 1000211 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Restricted Flow rate | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/07/2022 |  |  |  |  | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility in japan reported a 73-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported a suspected positioning issue observed by motor current waveform abnormality during support. the patient ultimately expired, however, there was no allegation the impella was a causal/contributory factor in the patient's out come. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-04464 was provided in sections b2, b5, d6a, d6b, h1, and h6. | The filed lt log was reviewed and a motor current spike which is indicative of ingested biomaterial occurred 17 hours into the run and then pump flow became saturated with no pulsatility. the returned pump was visually inspected and no abnormalities were observed. the pump was connected to an aic and placement signal was normal. the pump was run on p9 and pump flow was within specification at 4.9 l/min with suction. marks in the cannula surface and delaminated nitinol were observed. the pump was torn down and biomaterial was observed in the purge gap. the biomaterial was likely pulled in when the pump was in contact with the mitral valve. the cause of the saturated flow was biomaterial in the purge gap. this report is being filed as part of a retrospective review of historical records.
33276 | 1220648-2023-04466 | 17996769 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17996769&pc=OZD | 01/17/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 01/17/2022 | 10/24/2023 |  |  | 07/31/2023 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022121433 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Difficult to Remove | Adverse Event Without Identified Device or Use Problem | [5 empty] | Adhesion(s) | Device Embedded In Tissue or Plaque | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/17/2022 |  |  |  |  | The user facility reported a 61-year-old male undergoing cardiac bypass surgery was implanted with an impella device for mechanical circulatory support. the 5.5 pump adhered to the patient's sternum, the physician had to open the patient's chest to take the pump out of body, the pump was used for 645 hours which is around 26 days and this is beyond the instructions for use. | [5 empty] | The cause of the impella removal issue was not determined since neither the product nor sufficient clinical information was provided. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. this report is being filed as part of a retrospective review of historical records.
33277 | 1220648-2023-04468 | 17996787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17996787&pc=OZD | 11/21/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 11/21/2022 | 10/24/2023 |  |  | 03/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023207840 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  |  | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/21/2022 |  |  |  |  | The user facility reported a 82 year-old male with an impella 5.5 device for mechanical circulatory support. it was reported that there was leaking noted at yellow-to-yellow connection site. the problem was resolved with extension tubing at the connection point and 2 iv tourniquets secured around connection. no further leaking or purge alarms. bicarbonate was used in the purge. there was no patient harm reported. | [5 empty] | The investigation into the purge sidearm leak has been completed. no product was returned for investigation. the lt log confirmed the low purge pressure alarm. based on the provided clinical information, the cause of the purge leak was most likely sidearm yellow luer damage as a result of bicarbonate use. this report is being filed as part of a retrospective review of historical records.
33278 | 1220648-2023-04465 | 17996796 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17996796&pc=OZD | 01/24/2023 | 09/19/2025 |  | 10/24/2023 |  |  | 01/24/2023 | 10/24/2023 |  |  | 06/30/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023237048 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Break | Fluid/Blood Leak | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/09/2022 |  |  |  |  | The user facility reported 65 yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that a pump, that had been used beyond indication, had developed a leak from the sidearm filter. support continued. however, it was also noted that there was damage to the sterile sleeve. the team troubleshot by placement of a sterile dressing and the pump remains on for support. there was no patient harm reported. | Additional information indicated that, during patient support, the pump experienced placement signal issues. |  |  |  |  | The investigation for the side arm leak and the damaged sterile sleeve has been completed. the device was not returned for evaluation. however, clinical information has been provided is sufficient to determine the root cause of the purge leak. the cause of the purge leak was most likely cracked sidearm filter damage. device used beyond indication. the cause of the sterile sleeve tear was not determined since product was not returned. device used beyond indication. this report is being filed as part of a retrospective review of historical records. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the cause of the optical signal issue was not determined since product and data was not returned. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated with additional information. d6a/d6b updated with additional information. f6 and f8 should have been left blank. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures from the initial manufacturer device report number 1220648-2023-04465 h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2023-04465.
33279 | 1220648-2023-04467 | 17996800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17996800&pc=OZD | 06/13/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 06/13/2023 | 10/24/2023 |  |  | 04/30/2025 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024339387 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Failure to Advance | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 73-year-old female patient was scheduled for impella 5.5 implant. it was documented that resistance was consistently met during attempted delivery despite the use of multiple techniques. the pump was eventually removed after particulate in the cannula could not be cleared and the cannula bent. another impella 5.5 was successfully used. | [5 empty] | The investigation into the reported 'delivery issues" has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the cause of the delivery issue was due to calcified vessels that led to cannula kink. this report is being filed as part of a retrospective review of historical records.
33280 | 1220648-2023-04469 | 17996821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17996821&pc=OZD | 12/15/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 12/15/2022 | 10/24/2023 |  |  | 09/30/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023259499 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/15/2022 |  |  |  |  | The user facility reported a 62yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported the during support, the pump failed to re-cross valve after the pump fell out of the ventricle. a new impella 5.5 was used with no harm to the patient. | [5 empty] | The investigation for the reported positioning issues has been completed. the device was not returned for evaluation. however, the clinical report provided details to determine the root cause. the root cause of the positioning issue was due to patient movement as the patient pulled out the pump while turning. this report is being filed as part of a retrospective review of historical records.
33281 | 1220648-2023-04470 | 17996833 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17996833&pc=OZD | 11/07/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 11/07/2022 | 10/24/2023 |  |  | 06/30/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023233376 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/07/2022 |  |  |  |  | The user facility reported a 52 year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that a purge side arm leak was noticed at the sidearm yellow luer lock resulting in low purge pressures and high purge flows. a purge bypass was performed to resolve the issue. sodium bicarbonate was used in the purge. there was no patient harm reported. | [5 empty] | The investigation for the purge leak has been completed. neither the product nor the data logs were returned for investigation. however, the clinical information was sufficient in determining the root cause, given the results of prior failure investigations. it was concluded that the cause of the purge leak was sidearm yellow luer damage due to sodium bicarbonate use. this report is being filed as part of a retrospective review of historical records.
33282 | 1220648-2023-04471 | 17996850 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17996850&pc=OZD | 06/14/2023 | 09/30/2025 |  | 10/24/2023 |  |  | 06/14/2023 | 10/24/2023 |  |  | 11/30/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED INC. |  |  | US |  | 2023283748 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  |  | Death | Yes | Yes | Disconnection | Activation, Positioning or Separation Problem | Device Sensing Problem |  |  |  |  | Unspecified Vascular Problem | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 06/14/2023 |  |  |  |  | The user facility in the eu reported a 63-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. placement signal and left ventricle signal failed during the support after approximately 1 month. no harm done to the patient. additionally the sterile sleeve became compromised. the damage occurred during repositioning, but the pump was not explanted. the pump remained on despite the signal loss and sleeve damage for 60 days until family made choice to withdraw care. | [5 empty] | The investigation into the optical signal issue and the product damage (sterile sleeve damage) has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. per the clinical details, the cause of the optical signal issue was not determined since product was not returned. the root cause of the sterile sleeve damage issue was use related as there was excessive force. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. per the clinical details, the cause of the issue was not determined since product was not returned. the root cause of the sterile sleeve damage issue was use related to excessive force. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added death "1802" as well as optical signal issue "2917" as part of a retrospective review of optical signal issues in accordance with fda recommendation.
33283 | 1220648-2023-04472 | 17996864 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17996864&pc=OZD | 06/06/2023 | 01/23/2024 |  | 10/24/2023 |  |  | 06/06/2023 | 10/24/2023 |  |  | 12/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024315891 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 39-year-old female noted as high risk percutaneous cardiac insertion was implanted with an impella 5.5 device for mechanical circulatory support. the complainant had a 5.5 delivered via the axillary site. during support, the patient developed some bleeding from their access site and a jp drain was placed. three units of packed red blood cells were administered to counteract the bleed. support was maintained with the implanted pump following the intervention. | [5 empty] | The investigation into the access site bleeding ¿ major and the positioning issues has been completed since the original report was filed. the device was returned for investigation and was tested upon return. the root cause of the access site bleeding issue was most likely patient condition related since bleeding was reported from arteriotomy. the patient was morbidly obese, and the vessel was reported to be deep down. the root cause of the positioning issue was most likely use related due to improper technique. low pump was also detected. the root cause of the low pump flow issue was biomaterial. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
33284 | 1220648-2023-04473 | 17996873 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17996873&pc=OZD | 12/29/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 12/29/2022 | 10/24/2023 |  |  | 05/31/2023 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022085247 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | Pressure Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/29/2022 |  |  |  |  | The user facility reported 78yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there were low flow and high purge pressure alarms on the device. there were no kinks in the line and the cassette was changed. they were only running sodium bicarbonate and no systemic heparin. finally, tissue plasminogen activator (tpa) was added to the purge which then resolved the high purge pressure issue. there was no patient harm reported. | [5 empty] | The investigation for the reported low flows and the purge pressure issue has been completed. no product was returned for investigation. the logs shows the purge pressure at approximately 400mmhg at the beginning of the case. the purge pressure then rises to around 700mmhg over the next two days. after a purge bag change the purge pressure then increases over 1000mmhg until the high purge pressure alarm is triggered. the bag change procedure is then completed several times in a row and the alarm does not resolve. the purge pressure then begins to decline likely when tissue plasminogen activator (tpa) is given and the issue resolves. the cause of the high purge pressure was most likely biomaterial since the issue resolved with the tpa infusion. this report is being filed as part of a retrospective review of historical records.
33285 | 1220648-2023-04475 | 17997009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17997009&pc=OZD | 12/31/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 12/31/2022 | 10/24/2023 |  |  | 06/30/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023247648 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/31/2022 |  |  |  |  | The user facility reported a 61 year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the yellow luer began to leak where the purge tubing is connected. there were no alarms reported. no patient harm was reported. no additional information was provided. | [5 empty] | The investigation into the purge sidearm leak has been completed. the device was not returned for evaluation. the logs were returned and confirmed purge pressure low alarms on the day of the reported leak. however, without evaluation of the device, the cause of the purge leak was not determined. this report is being filed as part of a retrospective review of historical records.
33286 | 1220648-2023-04474 | 17997016 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17997016&pc=OZD | 09/27/2023 | 02/19/2024 |  | 10/24/2023 |  |  | 09/28/2023 | 10/24/2023 |  |  | 06/30/2025 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024358652 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 27 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/28/2023 |  |  |  |  | The user facility reported a 27-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support as a bridge in left ventricular assist device. the patient required a video-assisted thoracoscopic surgery (vats) for a lung mass. during positioning for vats, 5.5 flows went from 5lpm to 0. the patient decompensated status post intubation and cannulation of va ecmo. the impella 5.5 was removed and surgical closure performed of axillary pocket. large clot was discovered on pump outflow extending into catheter and a small clot observed in inlet. an impella cp was placed. | [5 empty] | The investigation into the low pump flow issue has been completed. the device was not returned for investigation. per clinical team, clots were found in if & of cages when pump was removed. based on clinical description and data review, the root cause of low flow was most likely biomaterial ingestion due to lack of heparin. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
33287 | 1220648-2023-04476 | 17997038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17997038&pc=OZD | 01/11/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 01/11/2023 | 10/24/2023 |  |  | 04/30/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023217101 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/11/2023 |  |  |  |  | The user facility reported a 63 year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was purge fluid leaking from the side arm retainer. a purge sidearm bypass was performed. there was no patient harm reported. | [5 empty] | The investigation into the purge sidearm leak has been completed. the device was not returned for evaluation. however, the clinical report has been reviewed. the root cause of the purge leak was unable to be determined because the product was not returned so the location of the leak could not be identified. this report is being filed as part of a retrospective review of historical records.
33288 | 1220648-2023-04477 | 17997067 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17997067&pc=OZD | 01/12/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 01/12/2023 | 10/24/2023 |  |  | 04/30/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023216162 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/12/2023 |  |  |  |  | The user facility reported 74yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the impella 5.5 was placed via a direct aortic insertion. while the physician was getting the device into position, a "purge system blocked" error was received, which could not be resolved despite multiple attempts at troubleshooting. new pump was placed with no issue. | [5 empty] | The investigation for the high purge pressure has been completed. the data log was reviewed and it was noted that the purge pressure increased over 7 minutes until the "purge system blocked" alarm triggered, which was accompanied by flow saturation and a non pulsatile motor current. it was also noted that the purge heparin concentration was 0 iu/ml. therefore the root cause of the high purge pressure was most likely ingested biomaterial. this report is being filed as part of a retrospective review of historical records.
33289 | 1220648-2023-04478 | 17997086 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17997086&pc=OZD | 10/19/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 10/19/2022 | 10/24/2023 |  |  | 05/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023223616 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Unknown |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 40-year-old patient was scheduled for impella 5.5 implant. the team attempted wireless delivery and backloading on the amplatz wire only to kink the pump. the 5.5 was replaced without any harm or injury and support was initiated. | [5 empty] | The investigation into the reported "delivery issues" has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated and it was determined that the cause of the delivery issues was most likely patient condition related, specifically the patient's tortuous vessels. this report is being filed as part of a retrospective review of historical records.
33290 | 1220648-2023-04479 | 17997121 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17997121&pc=OZD | 01/15/2023 | 09/19/2025 |  | 10/24/2023 |  |  | 01/15/2023 | 10/24/2023 |  |  | 11/30/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023278445 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Thrombosis/Thrombus | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/15/2023 |  |  |  |  | The user facility reported 51yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that an impella pump was placed and supported the patient for 10 minutes when it was removed due to pump coming out of position. at the time the pump was removed and replaced there was thrombosis noted at the site of implant. there was no lasting harm or injury from the pump replacement and observed thrombosis. | [5 empty] | Data logs showed that the ps waveform was generally within normal limits despite recurrent "impella position in aorta" alarms from the beginning of support. notably, the motor current was non-pulsatile and the pump flow waveform was moderately saturated, with some moderate fluctuation with no p-level change. in conclusion, when considering the clinical information, it was determined that the cause of the optical signal issue was most likely ingested biomaterial. the instructions for use contains information for the prevention of thrombus. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ this report is being filed as part of a retrospective review of historical records. | H6 updated to include optical signal issue coding. corrections that were made from the initial manufacturer device report number 1220648-2023-04479. b7 other relevant history, including preexisting medical conditions, updated. d4 model number updated. d6a / d6b implant and explant dates updated. d10 concomitant medical products and therapy dates updated. f6 and f8 erroneously included in initial report. g1, reporting contact email, reporting contact facility name, and reporting contact fax number updated. h6 health effect - impact code updated for thrombus failure mode.
33291 | 1220648-2023-04480 | 17997191 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17997191&pc=OZD | 12/20/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 12/20/2022 | 10/24/2023 |  |  | 09/30/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023257889 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/20/2022 |  |  |  |  | The user facility reported a 54yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that right as the graft was being fixated to the anastomoses during insertion, the pump slid back through the tunnel into the chest. an attempt was made to pull it back through with kelly clamps but without success. the device and graft was then removed. there was a kink noted in catheter shaft. a decision was to replace with new device. there was no patient harm reported. | [5 empty] | The investigation for the positioning issue has been completed. clinical review confirmed the impella slide back through the tunnel into the chest. the physician attempted to pull it back through with no success. there was a kink noted in catheter shaft after removed. root cause was improper use technique. device not returned for investigation. this report is being filed as part of a retrospective review of historical records.
33292 | 1220648-2023-04482 | 17997220 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17997220&pc=OZD | 11/17/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 11/17/2022 | 10/24/2023 |  |  | 09/30/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023263570 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | [6 empty] | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/17/2022 |  |  |  |  | The user facility reported a 57 year-old male with an impella 5.5 device for mechanical circulatory support. it was report. it was reported that there was a damaged yellow luer. bicarbonate was used in the purge system. the pump was replaced. there was no patient harm reported. | [5 empty] | The device was not returned for evaluation. however, the clinical information was sufficient in determining the root cause, given previous investigations for the same failure. it was concluded that the cause of the purge leak was sidearm yellow luer damage due to sodium bicarbonate use. this report is being filed as part of a retrospective review of historical records.
33293 | 1220648-2023-04481 | 17997225 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17997225&pc=OZD | 01/02/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 01/03/2023 | 10/24/2023 |  |  | 09/30/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023260006 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/03/2023 |  |  |  |  | The user facility reported a 33 year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that an agitated patient was restless and rolling in bed causing a crack in the purge sidearm. there were purge alarms and eventually the device was replaced. the pump was discarded. there was no patient harm reported. | [5 empty] | The investigation into the purge sidearm damage has been completed. the device was not returned for evaluation. however, the clinical information was sufficient in determining the root cause. it was concluded that the cause of the purge sidearm damage was to patient movement. this report is being filed as part of a retrospective review of historical records.
33294 | 1220648-2023-04483 | 17997234 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17997234&pc=OZD | 03/03/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 03/03/2023 | 10/24/2023 |  |  | 06/30/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023247700 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/03/2023 |  |  |  |  | The user facility reported a 62-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported a delivery issue involving an impella pump. it was documented that the pump could not be advanced as a result of the patient's tortuous anatomy. additionally the placment signal was lost in the insertion attempts. the team removed the pump and replaced with another. there was no patient harm as a result of the delivery failure. | [5 empty] | Based on the clinical data the patient had a tortuosity vessels condition that obstructed the impella to turn and deliver. root cause most likely patient condition related. complaint device not returned for investigation. data logs reviewed and the metrics are consistent with an optical fiber bend. this report is being filed as part of a retrospective review of historical records.
33295 | 1220648-2023-04484 | 17997242 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17997242&pc=OZD | 01/29/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 01/29/2023 | 10/24/2023 |  |  | 12/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023284036 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/29/2023 |  |  |  |  | The user facility reported a 69 year old male was implanted with an impella 5.5 device for mechanical circulatory support as a bridge to transplant. it was reported that there was a purge fluid leak at red impella plug site. no leak noted at purge side arm. there were no alarms. there was no patient harm reported. | [5 empty] | The investigation into the purge sidearm leak has been completed. the device was not returned for evaluation. however, the data logs were reviewed and no sustained periods of low purge pressure were observed. the root cause of the purge leak was unable to be determined because the product was not returned so the location of the leak could not be identified. this report is being filed as part of a retrospective review of historical records.
33296 | 1220648-2023-04485 | 17997271 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17997271&pc=OZD | 01/31/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 01/31/2023 | 10/24/2023 |  |  | 11/30/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023282598 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/31/2023 |  |  |  |  | The user facility reported a 81 year old male with st elevated myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that a "purge pressure low" alarm occurred on the console and the yellow luer on the purge sidearm was leaking. they attempted to change the purge cassette again to resolve the alarm but it did not solve the issue. a purge bypass was performed. there have been no alarms since the replacement. the purge was running sodium bicarbonate. there was no patient harm reported. | [5 empty] | The investigation into the purge sidearm leak has been completed. the pump was returned for evaluation. upon visual inspection, a crack in the yellow luer sidearm was confirmed. the clinical report stated that there was bicarbonate being used. the root cause of the purge leak was due to damage to the yellow luer sidearm. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33297 | 1220648-2023-04486 | 17997299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17997299&pc=OZD | 01/28/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 01/28/2023 | 10/24/2023 |  |  | 09/30/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023263381 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/28/2023 |  |  |  |  | The user facility reported a 62 year old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was a flow saturation issue with the device. the pump was repositioned and there was temporary resolution of flow saturation. the pump was ultimately removed. there was no patient harm reported. | [5 empty] | The investigation into the saturation flow issue has been completed. the pump was returned for evaluation. upon evaluation of the pump, it was noted that there were small amounts of biomaterial. the data logs show that there is saturated flow present after there¿s a spike in motor current. the flow fluctuates with no change in p-level, and temporarily resolves. the root cause of the saturated flow is most likely due to biomaterial. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ this report is being filed as part of a retrospective review of historical records.
33298 | 1220648-2023-04487 | 17997314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17997314&pc=OZD | 02/04/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 02/04/2023 | 10/24/2023 |  |  | 10/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2023272155 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/04/2023 |  |  |  |  | The user facility reported 33yo male was implanted with an impella 5.5 device for mechanical circulatory support as a bridge to transplant. it was reported that there were flow saturation issues followed by a pump stop. the device was removed and not replaced. | [5 empty] | The investigation for the flow saturation and the resulting pump stop have been completed. the returned pump was visually inspected and biomaterial was observed entraining the impeller and in outflow blocking the purge gap. the cause of the saturated pump flow and resulting pump stop was biomaterial. instructions for use for the related event are as follows: ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
33299 | 1220648-2023-04488 | 17997323 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17997323&pc=OZD | 08/20/2021 | 10/24/2023 |  | 10/24/2023 |  |  | 08/20/2021 | 10/24/2023 |  |  | 04/30/2023 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022080694 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the pump was accidentally started while outside of the patient. another device was subsequently used for patient support. there was no harm to the patient. | [5 empty] | The investigation into the reported issue has been completed. the used pump was returned for investigation. as per clinical investigation, the staff accidentally started the 5.5 pump and turned onto p4 while outside of the body. it was running underneath blue towels when the field rep entered the room. the root cause of the case start issue was a premature pump start-up. this report is being filed as part of a retrospective review of historical records.
33300 | 1220648-2023-04489 | 17997386 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17997386&pc=OZD | 02/05/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 02/05/2023 | 10/24/2023 |  |  | 01/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023181406 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 79 YR |  | Male |  |  |  | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/05/2023 |  |  |  |  | The user facility reported a 79 year old male with an st elevated myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there were signs of saturated flow during impella support. it was noted that the motor current was flat and that the lv waveform was at the top of the display. furthermore, the console alarmed with an impella in aorta alarm, despite the pump being properly placed. as a result of the noted issue, the decision was made to explant the pump. there was no patient harm reported. | [5 empty] | The investigation into the saturation flow issue has been completed. the pump was returned for evaluation. a visual inspection of the pump revealed that there was bond damage at the patient cable to kink protector bond. it also revealed that there was biomaterial in the cannula and near the purge gap. when the pump was run the biomaterial migrated to the impeller and the saturated flow was reproduced. the cause of the saturated flow was biomaterial. the cause of the product damage was bond failure at the patient cable to kink protector bond. as noted in the impella ifu: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ this report is being filed as part of a retrospective review of historical records.
33301 | 1220648-2023-04490 | 17997451 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17997451&pc=OZD | 02/01/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 02/01/2023 | 10/24/2023 |  |  | 11/30/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023283393 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/01/2023 |  |  |  |  | The user facility reported a 45 yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that during the support, there was persistent suction alarms and pump positioning issues. the pump position was sub-optimal and the decision was made to explant the pump and replace by ecmo support. there was no patient harm reported. | [5 empty] | The investigation for the reported suction/low flow and positioning issues has been completed. data logs showed that the motor current spiked to 1200 shortly before the suction alarm triggered. subsequently, the pump flow reduced and the motor current lost pulsatility with no change in motor speed. following the low pump flow, the flow then became saturated and gradually increased with no change in p-level until the pump was inadvertently removed. visual inspection of the returned pump revealed a mass of biomaterial within the cannula on the impeller. no other abnormalities were found. therefore, it was determined that the cause of the low pump flow was ingested biomaterial. the cause of the positioning issues was wireless re-access; the impella device is not indicated for wireless re-access. this report is being filed as part of a retrospective review of historical records.
33302 | 1220648-2023-04491 | 17997636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17997636&pc=OZD | 02/10/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 02/10/2023 | 10/24/2023 |  |  | 11/30/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2023278996 |  | 1000100 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/10/2023 |  |  |  |  | The user facility reported 58yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that after insertion of the 5.5 device and while increasing p-levels, the console alarmed "purge system blocked". upon troubleshooting for the issue, it was found that the a hard clamp which is not recommended had been been placed over the graft. additionally, there was a kink noted near the motor during insertion that was corrected upon discovery. the pump had to be removed and replaced with a new impella 5.5. there was no patient harm reported. | [5 empty] | The investigation for the reported high purge pressure has been completed. data logs show that the "purge pressure high" alarm occurred after a 3-minute increase in purge pressure while the pump was ramping up. the alarms persisted for the majority of support, as the purge pressure remained high with intermittent, brief drops in purge pressure. the alarms resolved once the sudden drops ended, however the purge pressure remained high for the remainder of support. given the logs and the clinical information, it was determined that the cause of the high purge pressure was most likely a kinked purge line in the catheter. this report is being filed as part of a retrospective review of historical records.
33303 | 1220648-2023-04492 | 17997659 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17997659&pc=OZD | 12/31/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 12/31/2022 | 10/24/2023 |  |  | 08/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023276894 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the surgeon was attempting to place the impella 5.5 for support for multiple hours. the team tried all guides for support and wire exchanges. the team made decision to abort the placement after the device kinked. no vessel harm or injury noted from attempts. | [5 empty] | The investigation into the reported issue has been completed. product was returned for investigation. the device was visually inspected and found the presence of a significant cannula kink near the outflow cage and a significant catheter kink near the motor housing. the catheter kink showed evidence of buckling due to excessive force. there were no other damage or abnormalities found. as per clinical investigation, the clinicians were unable to advance the impella past the patient's vasculature despite multiple attempts with various techniques. the patient's anatomy rendered it impossible to push the pump past the turn from the subclavian into the aorta. during these attempts, the catheter kinked. ultimately, the decision was made to abort delivery. after reviewing the clinical information, it was determined that the cause of the delivery issues was patient condition related, specifically the patient's vascular anatomy. this report is being filed as part of a retrospective review of historical records.
33304 | 1220648-2023-04493 | 17997786 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17997786&pc=OZD | 10/21/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 10/21/2022 | 10/24/2023 |  |  | 05/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023229029 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported surgeon was unable to insert impella 5.5 into 53-year-old male patient's left ventricle (lv). the patient has an extremely tortuous innominate and ascending aorta. the impella 5.5 implantation was aborted after about 10 attempts. the catheter was noted to be severely buckled. surgeon successfully switched to impella cp. there was no harm done to the patient. patient remained hemodynamically stable during remainder of procedure. | [5 empty] | The investigation into the reported "delivery issue" has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. upon conclusion of the investigation, the cause of the delivery issue was related to the patient's tortuous vessel. this report is being filed as part of a retrospective review of historical records.
33305 | 1220648-2023-04494 | 17997806 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17997806&pc=OZD | 02/13/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 02/13/2023 | 10/24/2023 |  |  | 11/30/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023280117 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 58 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/13/2023 |  |  |  |  | The user facility reported a 58yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that a pump stop occurred. the pump was replaced. | [5 empty] | The investigation into the pump stop has been completed. the returned pump was visually inspected and biomaterial was observed in the outflow. the cause of the pump stop was biomaterial. instructions for use related to the pump stop: ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
33306 | 1220648-2023-04495 | 17997926 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17997926&pc=OZD | 01/23/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 01/23/2023 | 10/24/2023 |  |  | 11/30/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023282636 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/23/2023 |  |  |  |  | The user facility reported a 66 year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was a fluid leak from the catheter during patient support. there was no patient harm as a result of the leak. no additional information was provided. | [5 empty] | The investigation into the purge leak has been completed. the device was not returned for evaluation. review of the clinical report did not provide any details about the reported issue. the customer discarded the device. therefore, the root cause of the purge leak could not be determined. this report is being filed as part of a retrospective review of historical records.
33307 | 1220648-2023-04497 | 17998231 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17998231&pc=OZD | 02/10/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 02/10/2023 | 10/24/2023 |  |  | 11/30/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2023276343 |  | 1000100 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/10/2023 |  |  |  |  | The user facility reported 52yo male was implanted with an impella 5.5 device for mechanical circulatory support as a bridge to transplant. it was reported that there was an infection at the implant site. the patient was treated with antibiotics and remained on support. there was no patient harm reported. | [5 empty] | The investigation for the reported insertion site infection has been completed. there was no device returned for evaluation. additionally, without additional clinical details, there is insufficient evidence to indicate that the infection was in relation to the impella. therefore, the root cause of the infection could not be determined. the instructions for use states that "the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ this report is being filed as part of a retrospective review of historical records.
33308 | 1220648-2023-04498 | 17998305 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17998305&pc=OZD | 02/10/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 02/10/2023 | 10/24/2023 |  |  | 09/30/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023259072 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/10/2023 |  |  |  |  | The user facility reported a 60yo male was implanted with an impella 5.5 device for mechanical circulatory support as a bridge to transplant. it was reported that while there was oozing at the access site, the patient was experiencing pain with intermittent numbness down their arm. as a result of the access site ooze, a pressure dressing was applied to achieve hemostasis. support continued and the patient's condition improved. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device was not returned for evaluation. however, the clinical report is sufficient to provide details to determine a root cause for the reported issue. the root cause of the limb ischemia was most likely due to improper technique and patient condition. the instructions for use lists known potential adverse events and it states:¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
33309 | 1220648-2023-04496 | 17998329 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17998329&pc=OZD | 08/25/2022 | 09/19/2024 |  | 10/24/2023 |  |  | 08/25/2022 | 10/24/2023 |  |  | 01/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023178292 |  | 1000211 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male |  |  | Asian | Injury | Yes | Yes | Restricted Flow rate | Increase in Pressure | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Survival outcome at explant noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | The user facility reported a 51-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported a high purge pressure with the impella device. the pump was removed and replaced with another impella device. there was no harm to the patient. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-04496 was provided in sections b2, b5, h1, and h6. | The investigation into the reported issue been completed. product was returned for investigation. the device was visually inspected and found the catheter had a kink and clamp marks, which match the clinical account. the pump was run in the lab and pp was around 900 mmhg and no alarms were seen. . as per clinical investigation, there was a history of biting the shaft of impella with a clamp during insertion, and although the pf value continued to be low immediately after, it recovered after a few days. however, there was a sudden drop in pf again, and the pump was replaced in consideration of the risk of stopping the pump. the root cause of the high purge pressure during the case is most likely due a kinked purge line as a result of the clamping of the catheter. this report is being filed as part of a retrospective review of historical records.
33310 | 1220648-2023-04499 | 17998418 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17998418&pc=OZD | 02/25/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 02/25/2023 | 10/24/2023 |  |  | 11/30/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023275532 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/27/2023 |  |  |  |  | The user facility reported a 49 year old female was implanted with an impella 5.5 device for mechanical circulatory support as a bridge to transplant. it was reported that there was a break in the yellow luer/sidearm. a purge bypass was performed. there was no patient harm reported. | [5 empty] | The investigation into the purge sidearm leak has been completed. the pump was returned for evaluation. visual inspection of the returned product revealed significant sidearm yellow luer damage. no other abnormalities were found. the clinical report was reviewed finding that sodium bicarbonate was being used as the purge solution. therefore, it was determined that the cause of the purge leak was sidearm yellow luer damage due to the use of sodium bicarbonate. this report is being filed as part of a retrospective review of historical records.
33311 | 1220648-2023-04500 | 17998488 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17998488&pc=OZD | 02/28/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 02/28/2023 | 10/24/2023 |  |  | 12/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024292183 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/28/2023 |  |  |  |  | The user facility reported a 66yo female was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that after starting tissue plasminogen activator (tpa) due to high purge pressure alarms, the patient began having motor current spikes. the impella pump then experienced a stop. attempts were made restart multiple times but failed. the patient refused to bridge to left ventricular assist device (lvad) so a decision was made to not replace the impella device. there was no patient harm reported. | [5 empty] | The investigation into the pump stop has been completed. the field data logs were reviewed finding that at the end of the run, a sudden spike in purge pressure with high purge pressure alarms followed by a spike in motor current and pump stop had occurred. the pump was running normal up to the sudden purge pressure spike. the the sudden spike in purge pressure with no motor current spike prior suggests a sudden kinked purge line. the returned pump was visually inspected and biomaterial was observed in the outflow. the catheter was trimmed and blood was observed in the purge line. a yellow tinge was observed with the returned purge cassette purge disc interior. the purge cassette was eventually able to purge, then connected to the pump and high purge pressure reproduced. the catheter was then cut nearest to the pump and no fluid expelled from the purge line. the pump was then cut at the 60 cm mark and fluid expelled from the purge line with a purge open alarm. the purge line was then removed from the cut portion of catheter and blood and tinged fluid was observed in the purge line. no leaks observed from the sidearm. the cause of the issue was most likely a temporary kinked purge line in the purge cassette tubing. the instructions for use for the event states the following: ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
33312 | 1220648-2023-04502 | 17998742 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17998742&pc=OZD | 11/26/2021 | 10/24/2023 |  | 10/24/2023 |  |  | 11/26/2021 | 10/24/2023 |  |  | 09/30/2023 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022141252 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 65-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported false positioning alarms despite position verification during support. the pump was subsequently removed and replaced with no harm to the patient. | [5 empty] | The investigation into the reported issue was completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that data logs showed the alternating occurrences of the "impella position in aorta" and "impella position in ventricle" alarms upon activation of the pump. placement signal metrics were generally within normal limits. there were no other relevant abnormal metrics noted. in conclusion, it was determined that the cause of the false impella position alarms was most likely patient condition. this report is being filed as part of a retrospective review of historical records.
33313 | 1220648-2023-04501 | 17998767 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17998767&pc=OZD | 03/04/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 03/04/2023 | 10/24/2023 |  |  | 12/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024301323 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 40 YR |  | Female |  |  |  | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/04/2023 |  |  |  |  | The user facility reported 40year old female was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there were flow saturation numbers and waveforms on the screen. no additional information was provided. there was no reported harm to the patient. | [5 empty] | The investigation for the flow saturation issue has been completed. the returned pump was visually inspected and biomaterial was observed on the impeller and in the purge gap. the cause of the issue was biomaterial. this report is being filed as part of a retrospective review of historical records.
33314 | 1220648-2023-04503 | 17998836 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17998836&pc=OZD | 12/25/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 12/25/2022 | 10/24/2023 |  |  | 10/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023266061 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | Material Too Soft/Flexible |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/25/2022 |  |  |  |  | The user facility reported a 67-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported a positioning issue with the impella device. it was noted that the pump was having intermittent suction issues and positioning attempts led to the pump coming across the valve and into the aorta. the medical director then had difficulty advancing the pump as the device got stuck in the sheath. the catheter subsequently bent back on itself in an accordion fashion and the decision was made to remove the pump. there was no harm to the patient. | [5 empty] | Returned compliant pump visually inspected. catheter kink found around the 40cm mark in catheter. data logs showed alarms for positioning, suction and there was low pump flow as a result of this catheter kink due to use related issues that caused the pump not to be in proper position. this report is being filed as part of a retrospective review of historical records.
33315 | 1220648-2023-04504 | 17998865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17998865&pc=OZD | 06/08/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 06/08/2023 | 10/24/2023 |  |  | 12/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024315476 |  | 1000100 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/08/2023 |  |  |  |  | The user facility reported 75yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported during placement of the pump, the physician encountered major resistance advancing the impella pump in the patient¿s heart. the procedure was aborted with plans for future surgery. there was no harm to the patient as a result of this incident. | [5 empty] | The investigation for the positioning issue during insertion has been completed. no product was returned for evaluation. as per clinical, doctor attempted to reposition and with catheter being stuck extreme force was applied, catheter kink was noted close to the locking mechanism. the cause of the positioning issue was cath anchor use related. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33316 | 1220648-2023-04505 | 17998887 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17998887&pc=OZD | 09/21/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 09/21/2022 | 10/24/2023 |  |  | 02/29/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023194882 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/21/2022 |  |  |  |  | The user facility reported a 60-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the 5.5 support initiated via direct access for a (coronary artery bypass grafting) cabg patient unable to come off pump. the patient was also on (extracorporeal membrane oxygenation) ecmo. the pump was replaced due to low flows. the 2nd pump was placed and support initiated. | [5 empty] | The medical team did return data logs for analysis of the reported low flow. flow was observed to be abnormally normally with a spike of motor current at the time a suction alarm went off. the root cause of the low flows was not determined as product has not been returned. this report is being filed as part of a retrospective review of historical records.
33317 | 1220648-2023-04506 | 17998890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17998890&pc=OZD | 12/31/2021 | 10/24/2023 |  | 10/24/2023 |  |  | 12/31/2021 | 10/24/2023 |  |  | 03/31/2023 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022062013 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 19 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/31/2021 |  |  |  |  | The user facility reported that a 19-year-old male patient was implanted with an impella 5.5 for mechanical circulatory support. on day 49 of support, multiple pump stops occurred during the patient's physical therapy. the pump restarted successfully each time. the rn noted that the 3-point fixation was no longer in place and stated it had moved down the patient's side and leg. the pump was subsequently placed back into the 3-point fixation and the issue resolved. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported pump stop was completed. the device was returned for evaluation. visual inspection of the pump revealed that the catheter was cut. the sidearm was able to purge to the cut line. there was no biomaterial or blood ingress observed. there was a kink in the catheter observed. the root cause of the pump stop was most likely use related since 3-point fixation was not in place and the patient was moving at the time of the pump stop. impella 5.5 ifu ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
33318 | 1220648-2023-04507 | 17999083 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17999083&pc=OZD | 03/05/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 03/05/2023 | 10/24/2023 |  |  | 10/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023266198 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/05/2023 |  |  |  |  | The user facility reported a 49 year old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was a purge sidearm crack at the side arm luer. a purge bypass was performed. there was no patient harm reported. | [5 empty] | The investigation into the purge sidearm damage has been completed. neither the product nor the data logs were returned for investigation. however, the clinical information was sufficient in determining the root cause, given the results of prior failure investigations. it was concluded that the cause of the purge damage was sidearm yellow luer damage due to sodium bicarbonate use. this report is being filed as part of a retrospective review of historical records.
33319 | 1220648-2023-04508 | 17999127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17999127&pc=OZD | 08/20/2021 | 10/24/2023 |  | 10/24/2023 |  |  | 08/20/2021 | 10/24/2023 |  |  | 05/31/2023 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022088350 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 27 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/20/2021 |  |  |  |  | The user facility reported that a 27-year-old patient was implanted with an impella 5.5 device for mechanical circulatory support. leading up to explant, it was noted that the placement signal was lost while the patient was on the operating room table. upon removal, a significant amount of biomaterial was observed to be ingested into the inflow cage. the patient's partial thrombin time (ptt) was noted to be within normal limits. no further information was provided. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported thrombus was completed. an image of the pump provided from the field, confirmed a significant mass of biomaterial had been ingested into the inflow cage. the pump and part of the catheter were returned for evaluation. visual inspection of the pump confirmed a moderate amount of remaining biomaterial was found on the impeller, including biomaterial covering the optical sensor. analysis of the data logs revealed numerous "impella position unknown" alarms several hours prior to explant. in conclusion, it was confirmed that biomaterial was observed on the pump at the time of pump removal. the biomaterial was most likely ingested by the pump during the removal process. this report is being filed as part of a retrospective review of historical records.
33320 | 1220648-2023-04509 | 17999257 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17999257&pc=OZD | 09/24/2023 | 02/16/2024 |  | 10/24/2023 |  |  | 09/24/2023 | 10/24/2023 |  |  | 07/31/2025 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024388323 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that 81-year old male was implanted with an impella cp to unload and alleviate vf arrest vt/storm. following impella placement, the patient became confused and pulled on the impella cp. the pump then stopped. the automated impella controller (aic) was switched and the pump was unable to restart. therefore, it was advised to exchange the pump and pump was explanted two days post implant. there was no reported patient harm. | [5 empty] | The investigation into the pump stop issue completed since the original report was filed. the console was returned for investigation. no biomaterial or other issues found under inspection visually. pump could start but stop immediately under bench test. the root cause of the pump stop was an open line due to excessive force from user. | The device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
33321 | 1220648-2023-04510 | 17999278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17999278&pc=OZD | 09/24/2023 | 01/17/2024 |  | 10/24/2023 |  |  | 09/24/2023 | 10/24/2023 |  |  | 03/31/2025 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024326621 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/24/2023 |  |  |  |  | The user facility reported a 39-year-old male in cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the pump stopped, but immediately restarted. the patient is currently stable. the impella 5.5 was explanted without issue seven days after implant. | [5 empty] | The investigation for the reported pump stop has been completed. the device nor data logs were returned for evaluation. therefore, the root cause of pump stop was unable to be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
33322 | 1220648-2023-04511 | 18000077 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18000077&pc=OZD | 03/06/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 03/06/2023 | 10/24/2023 |  |  | 12/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024292157 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/06/2023 |  |  |  |  | The user facility reported 64yo male was implanted with an impella 5.5 device for mechanical circulatory support it was reported that there were high purge pressure alarms noted. however, no kinks in the purge system were noted. heparin was to be added to the purge system and close monitoring of the patient was done. there was no patient harm reported. | [5 empty] | The investigation for the reported high purge pressure has been completed. the data logs confirm the low purge pressure even 30 minutes into support. the event lasts about 2 minutes. following this event purge pressure begins to rise and the high purge pressure persists throughout the remainder of the case. the product was returned and the high purge pressure was reproduced. the purge blockage was isolated to the impeller/motor housing. the cannula was removed and biomaterial was found under the impeller. the root cause of the high purge pressure was determined to be biomaterial buildup under the impeller. this report is being filed as part of a retrospective review of historical records.
33323 | 1220648-2023-04513 | 18000098 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18000098&pc=OZD | 03/09/2023 | 10/24/2023 |  | 10/24/2023 |  |  |  | 10/24/2023 |  |  | 12/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2023279052 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/09/2023 |  |  |  |  | The user facility reported 63yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was a high purge pressure alarm less than 5 minutes after the pump implant. there were no visible kinks or issues found during troubleshooting. however, the pump was replaced. there was no delay in care. no patient harm reported. | [5 empty] | The investigation for the reported high purge pressure has been completed. the data logs were returned and show that the pump was purging for about 25 minutes before initial insertion. the pump is then inserted and then removed. the pump is reinserted and started. the motor current pulsatility is low and pump in aorta alarms are triggered. the pump is them removed again for 8 minutes and then reinserted into the body. the pump is then ramped up and approximately 3 minutes after ramp up high purge pressure is produced. the starting purge pressure was 500mmhg and it increased to 1100mmhg 3 minutes after ramp up. the product was returned and some biomaterial was found in the purge gap. the pump was purged and high purge pressure was originally reproduced, then the purge pressure dropped and normalized back to a value of about 500mmhg. purging the pump likely removed the biomaterial causing the high purge pressure. the root cause of the high purge pressure was determined to be ingested biomaterial that was likely picked up during the multiple insertions of the device. this report is being filed as part of a retrospective review of historical records.
33324 | 1220648-2023-04514 | 18000108 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18000108&pc=OZD | 03/10/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 03/10/2023 | 10/24/2023 |  |  | 10/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2023274782 |  | 1000100 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 66 YR |  | Male |  |  |  | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/10/2023 |  |  |  |  | The user facility reported 66 yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that after running sodium bicarbonate in purge, it was reported that flow saturation was occurring. it was recommended to add heparin to the purge after thoracentesis. as soon as heparin was added, the flows returned to normal. patient was weaned and explanted the following day without issue. | [5 empty] | The investigation for the saturated flow issues has been completed. the field data log was reviewed and after an initial purge flow that was on the low side trending back to normal, pump flow began to gradually increase until saturated flow was observed. a bag change was performed and the saturated flow resolved. the cause of the saturated flow was most likely biomaterial since the issue resolved with heparin added to the purge. this report is being filed as part of a retrospective review of historical records.
33325 | 1220648-2023-04515 | 18000116 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18000116&pc=OZD | 03/14/2023 | 09/25/2024 |  | 10/24/2023 |  |  |  | 10/24/2023 |  |  | 12/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 SMARTASSIST SET, JP |  |  | ABIOMED, INC. |  |  | US |  | 2024286368 |  | 1000211 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Fluid/Blood Leak | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/15/2023 |  |  |  |  | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 37yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there were crystals in the purge reservoir outside the sidearm on the side of the blue balloon. there were no abnormalities in the pump function or purge pressure flow rate. there was no patient harm as a result of the failure. it was later documented that the patient expired during support, however there was no issue or causal relationship to the impella device alleged. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-04515 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding. | The investigation for the reported purge leak has been completed. the returned pump was visually inspected and crystalized dextrose was found on and in the retainer. the catheter was observed to be cut. the retainer was removed and crystalized dextrose was observed in and on the reservoir. the sidearm was pressurized with colored fluid and a leak was unable to be reproduced. the cause of the purge leak was most likely use related since the leak was unable to be reproduced and the purge solution was likely introduced externally since there were no reported leaks for the remaining 15 days of support. this report is being filed as part of a retrospective review of historical records.
33326 | 1220648-2023-04516 | 18000122 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18000122&pc=OZD | 03/13/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 03/13/2023 | 10/24/2023 |  |  | 12/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024292761 |  | 1000100 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/13/2023 |  |  |  |  | The user facility reported 35yo male was implanted with an impella 5.5 device for mechanical circulatory support as a bridge to heart transplant. it was reported that high purge pressure developed. tissue plasminogen activator (tpa) was added to the purge and the pressure gradually decreased and purge values returned to within specification. there was no patient harm reported. | [5 empty] | The investigation for the reported high purge pressure has been completed. the field left log was reviewed and no suction triggered prior to the gradual rise in purge pressure. tpa was reported to be added to the purge, purge pressure gradually decreased and purge values were within specification for the remainder of the run. the cause of the high purge pressure was most likely biomaterial since the issue resolved during support 18 hours after the addition of tpa to the purge. the instructions for use discusses troubleshooting for high purge pressures. it states unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ this report is being filed as part of a retrospective review of historical records.
33327 | 1220648-2023-04517 | 18000125 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18000125&pc=OZD | 03/14/2023 | 10/24/2023 |  | 10/24/2023 |  |  |  | 10/24/2023 |  |  | 01/31/2025 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024294825 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/21/2023 |  |  |  |  | The user facility reported 55yo male was implanted with an impella 5.5 device for mechanical circulatory support as a bridge to transplant. it was reported that the patient had severe hemolysis with a drop in hemoglobin. it is unknown how the patient was treated. there was no patient harm reported. no additional information was provided. | [5 empty] | The investigation for the reported hemolysis has been completed. returned complaint pump visually inspected. no (broken blades, delamination, cage) issues observed. biomaterial ingested around impella. clinical details mentioned the patient had severe hemolysis with drop in hemoglobin, therefore they sent back the pump for evaluation even they believed the hemolysis could be caused by medication. biomaterial removed prior the hemolysis testing. hemolysis testing conducted successfully for extra pump evaluation and mih=3.76 (modified index of hemolysis) which indicated no hemolysis due to device malfunction. root cause most likely biomaterial ingestion due to patient condition related. the instructions for use states the following: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ it also states ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
33328 | 1220648-2023-04519 | 18000159 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18000159&pc=OZD | 03/27/2023 | 10/24/2023 |  | 10/24/2023 |  |  |  | 10/24/2023 |  |  | 01/31/2025 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024297510 |  | 1000100 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/28/2023 |  |  |  |  | The user facility reported 62yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the 5.5 supporting a patient as escalated care when the purge pressure and flow saturation was alarming. the team placed tissue plasminogen activator (tpa) into the purge for troubleshooting. the high purge pressure and the flow saturation resolved and the pump remains on for support without harm or injury. | [5 empty] | The investigation for the saturated flow issue has been completed. data logs show immediate flow saturation upon starting the pump. the purge pressure rises for about 10 minutes following pump activation until the purge pressure high alarm triggered. the high purge pressure and saturation gradually resolved over the next 2 days. in conclusion, when considering the clinical information, it was determined that the cause of the high purge pressure was most likely ingested biomaterial since tissue plasminogen activator (tpa) resolved the issue and the patient had a pre-existing left ventricle thrombus. this report is being filed as part of a retrospective review of historical records.
33329 | 1220648-2023-04518 | 18000161 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18000161&pc=OZD | 03/24/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 03/24/2023 | 10/24/2023 |  |  | 01/31/2025 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024294821 |  | 1000100 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  |  | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/24/2023 |  |  |  |  | The user facility reported 49yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that an agitated patient accidentally pulled the impella pump as the catheter got caught on the patient's body. patient was actively being weaned and stayed stable post impella pull. repositioning was attempted and unsuccessful. swan was placed and numbers were evaluated. physician felt patient was stable enough to remove impella. | [5 empty] | The investigation for the positioning issue has been completed. the device was not returned for evaluation. clinical information was sufficient to determine a root cause. the root cause was determined that the cause of the positioning issues was patient movement. this report is being filed as part of a retrospective review of historical records.
33330 | 1220648-2023-04520 | 18000162 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18000162&pc=OZD | 04/01/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 04/01/2023 | 10/24/2023 |  |  | 01/31/2025 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024297539 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/01/2023 |  |  |  |  | The user facility reported a 73 yo female was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there were high purge alarms. they used the tissue plasminogen activator (tpa) protocol but it did not clear the line. the pump was removed. there was no patient harm reported. | [5 empty] | The investigation for the high purge pressure has been completed. he returned pump was visually inspected and biomaterial was observed around the purge gap. the pump was then connected to a purge cassette to run the purge system and high purge pressure alarm reproduced. the cannula was then cut off and biomaterial was confirmed in the purge gap. the cause of the issue was biomaterial. the instructions for use provides troubleshooting for high pressure issues. it states unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ this report is being filed as part of a retrospective review of historical records.
33331 | 1220648-2023-04521 | 18000170 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18000170&pc=OZD | 03/28/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 03/28/2023 | 10/24/2023 |  |  | 01/31/2025 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024296587 |  | 1000100 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Activation Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/28/2023 |  |  |  |  | It was reported that during a difficult placement of the pump, the device was torqued multiple times trying to cross the aortic valve. after getting it placed successfully, the pump would not start and had to be removed. there was no patient harm reported. | [5 empty] | The investigation for the failure to start of the pump has been completed. the pump nor the data logs were returned for evaluation. the cause of the issue was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records.
33332 | 1220648-2023-04522 | 18000171 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18000171&pc=OZD | 03/28/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 04/03/2023 | 10/24/2023 |  |  | 10/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2023274573 |  | 1000100 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/03/2023 |  |  |  |  | The user facility reported a 81yo female was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that during support, there were waveform changes and increase in lactate dehydrogenase (ldh). heparin was added to the purge solution and the issue resolved. after the pump was explanted, a clot was located in the inlet cage of the pump. | [5 empty] | The investigation for the reported thrombus has been completed. no product was returned for evaluation. the cause of the thrombus was most likely patient condition related. the instructions for use provides details on how prevent thrombus formation in the pump. it states the following: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ this report is being filed as part of a retrospective review of historical records.
33333 | 1220648-2023-04523 | 18000174 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18000174&pc=OZD | 04/06/2023 | 10/24/2023 |  | 10/24/2023 |  |  |  | 10/24/2023 |  |  | 01/31/2025 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024294811 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/06/2023 |  |  |  |  | The user facility reported a 53yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there were high purge pressure alarms with quick transition to purge system blocked alarms. the cassette was swapped with no improvement. the customer refused a tissue plasminogen purge (tpa) and therefore, the pump was removed to prevent failure while in use. there was no patient harm reported. | [5 empty] | The investigation for the reported high pressure issue has been completed. the data log shows a rise in purge pressure at the end of the case. the purge pressure rises from approximately 600mmhg to approximately 1050mmhg over the course of 20 minutes. a suction event is also seen right before the rise in purge pressure. a visual inspection of the pump did not reveal any kinks or biomaterial. the pump was attached to a purge cassette and run in a bucket of water and the purge pressure was between 800-1050mmhg with no high purge pressure alarms. the pump was cut near the motor housing and fluid was purged through without issue. no kinks or blood ingress was noted. the cause of the high purge pressure was most likely biomaterial based on the data log analysis. the instructions for use provides troubleshooting for high pressure issues. it states unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ this report is being filed as part of a retrospective review of historical records.
33334 | 1220648-2023-04524 | 18000176 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18000176&pc=OZD | 04/09/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 04/09/2023 | 10/24/2023 |  |  | 10/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023268994 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | Pressure Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/09/2023 |  |  |  |  | The user facility reported a 44yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that high purge pressures occurred during support. tissue plasminogen activator (tpa) was added to the purge fluid and pressure normalized. the following day, a leak was found at the pump's side arm, prompting a purge system bypass configuration. there was no patient harm reported. | [5 empty] | The investigation for the purge leak and the high purge pressure has been completed. the cause of the high purge pressure was most likely biomaterial since tpa in the purge resolved the issue. the cause of the pump purge leak was a leak from the sidearm but the exact location of the leak was unknown. this report is being filed as part of a retrospective review of historical records.
33335 | 1220648-2023-04525 | 18000183 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18000183&pc=OZD | 04/08/2023 | 10/24/2023 |  | 10/24/2023 |  |  |  | 10/24/2023 |  |  | 11/30/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2023278482 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Positioning Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/11/2023 |  |  |  |  | The user facility reported 51yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that blood was found within the anti-contamination sleeve during impella support. it was noted that there was some oozing at the access site coinciding with this observation. manual pressure was initially applied in an attempt to control the condition, however, sutures were eventually placed at the access site to slow the bleed. it was further documented that the physician had attempted to reposition the device, however, the lock failed to disengage when the blue button was pressed. there was no patient harm reported. | [5 empty] | The investigation for the reported access site bleed and the positioning issue has been completed. the returned pump was investigated and it was found that the catheter anchor assembly moves as expected when the blue button is compressed. blood was observed in the sterile sleeve. no tears in the sterile sleeve were observed. the catheter anchor had been rotated relative to blue suture hub from the original position. evidence of excessive force from attempting to move catheter anchor forcefully during repositioning was observed on the catheter. bunching of the catheter under the cath anchor was observed. the cause of the access site bleeding was use related and related to placing the blue suture hub into the arteriotomy. the cause of the positioning issue was a cath anchor use issue. the instructions for use states that the following: ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records.
33336 | 1220648-2023-04526 | 18000212 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18000212&pc=OZD | 04/16/2023 | 10/24/2023 |  | 10/24/2023 |  |  |  | 10/24/2023 |  |  | 02/28/2025 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024311658 |  | 1000100 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Improper Flow or Infusion | [6 empty] | Hemolysis | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/16/2023 |  |  |  |  | The user facility reported 37yo male was implanted with an impella 5.5 device for mechanical circulatory support it was reported that there were saturated flows and hemolyzed labs during impella support. the pump was not removed and support proceeded until the patient was successfully bridged to an left ventricular assist device. there was no patient harm reported. | [5 empty] | The investigation into the hemolysis that led to flow saturations has been completed. no product was returned for investigation. the data log shows a motor current spike at the time that the flow becomes saturated indicating that the impeller likely contacted biomaterial. the cause of the saturated flow was most likely biomaterial. the data log also shows many instances of "impella position in aorta" and "impella position unknown" alarms occurring mostly within the second half of the case. the cause of the hemolysis was not determined since product was not returned and there is insufficient clinical information regarding the hemolysis. this report is being filed as part of a retrospective review of historical records.
33337 | 1220648-2023-04527 | 18000217 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18000217&pc=OZD | 04/28/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 04/28/2023 | 10/24/2023 |  |  | 10/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2023274713 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/28/2023 |  |  |  |  | The user facility reported 57yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that during patient support, the physician had difficulty repositioning / advancing the catheter due to the lock being stuck on the catheter shaft. the blue hub was subsequently forced further into the skin tract in order to advance the pump to a more favorable position. patient support continued following the repositioning. there was no patient harm. | [5 empty] | The investigation for the positioning issue during support has been completed. visual inspection found catheter kink at 46 cm mark. most likely, catheter was kink while repositioning of the pump. blue pin lock could slide back & forth on catheter easily. no other issues were found on the blue pin lock & the rest of the pump. the root cause of the positioning issue was most likely due to cath anchor use issue. this report is being filed as part of a retrospective review of historical records.
33338 | 1220648-2023-04528 | 18000219 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18000219&pc=OZD | 05/02/2023 | 10/24/2023 |  | 10/24/2023 |  |  |  | 10/24/2023 |  |  | 03/31/2025 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024327553 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  |  | White | Injury | Yes | No | Pressure Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/02/2023 |  |  |  |  | The user facility reported 82 y o male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the pump had a high purge pressure followed by a low purge flow during impella support. following the initial high purge pressure alarm, troubleshooting was attempted, however the issue remained. the decision was made to explant the pump and during the removal procedure it was observed that the purge flow dropped. a purge system blocked alarm was triggered and, upon explant, a clot was documented within the outflow cage. a new impella 5.5 pump was subsequently implanted for ongoing support. | [5 empty] | The investigation for the reported high purge pressure and the thrombus has been completed. the pump was examined. no kinks noted. there was a small amount of biomaterial visualized near the purge gap. the pump was purged and the high purge pressure was not reproduced. the cause of the high purge pressure was biomaterial based on the clinical information. the instructions for use provides details on the prevention of thrombus formation by stating ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ this report is being filed as part of a retrospective review of historical records.
33339 | 1220648-2023-04529 | 18000221 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18000221&pc=OZD | 05/16/2023 | 09/19/2025 |  | 10/24/2023 |  |  | 05/16/2023 | 10/24/2023 |  |  | 01/31/2025 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024297506 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | Pressure Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/16/2023 |  |  |  |  | The user facility reported 52yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the complainant reported a high purge pressure alarm during impella support. the patient was provided tissue plasminogen activator (tpa) and the purge pressure resolved. the patient remained stable throughout the event and support continued with the implanted pump. there was no patient harm reported. | Additional information indicated that, during patient support, the pump experienced placement signal not reliable alarm continued. |  |  |  |  | The investigation for the reported high purge pressure has been completed. the pump was not returned for evaluation. however, the information provided from the clinical report is sufficient to determine the root cause of the reported issue. the cause of the high purge pressure was most likely biomaterial since the issue resolved with tissue plasminogen activator (tpa). this report is being filed as part of a retrospective review of historical records. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the data log was reviewed and placement signal not reliable alarm was confirmed but pump optical parameters were normal. the root cause of the optical signal issue corruption was not determined since product was not returned. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated with additional information. d4 model number updated. d6a/d6b updated with additional information. f6 and f8 should have been left blank. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures from the initial manufacturer device report number 1220648-2023-04529. h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2023-04529.
33340 | 1220648-2023-04530 | 18000225 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18000225&pc=OZD | 05/21/2023 | 09/25/2024 |  | 10/24/2023 |  |  | 05/21/2023 | 10/24/2023 |  |  | 04/30/2025 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024330364 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic |  | Injury | Yes | Yes | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/21/2023 |  |  |  |  | Survival outcome at explant noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | The user facility reported 73year-old female was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that after support was ongoing for 5 days, there was flow saturation. the team performed troubleshooting with tissue plasminogen activator (tpa) for the purge pressures. the troubleshooting did not resolve fully the issue, but the pump remained on until the patient expired on support. there was no allegation of the pump causing the death. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding. | The investigation for the reported flow saturation has been completed. the returned pump was visually inspected. biomaterial observed around impella and inlet cage. no other abnormality found. root cause of the saturated flow was biomaterial ingested due to patient condition related. this report is being filed as part of a retrospective review of historical records.
33341 | 1220648-2023-04531 | 18000246 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18000246&pc=OZD | 04/26/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 04/26/2023 | 10/24/2023 |  |  | 03/31/2025 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024320175 |  | 1000100 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 37 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/26/2023 |  |  |  |  | The user facility reported 37yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was serosanguineous fluid, possibly coming from the device's blue suture pad, dripping into the anticontamination sleeve. a drop in purge flow was also observed. purge tubing was not kinked and impella catheter appears to be kink free as well. the purge bag & cassette was changed out with no further issues reported. no other issues reported during the following days. there was no patient harm reported. | [5 empty] | The investigation for the reported flow issues/pressure issues has been completed. no purge abnormalities are seen in the logs. there are no low purge pressure or high purge pressure alarms. the cause of the product damage, specifically fluid in the anti-contamination sleeve was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records.
33342 | 1220648-2023-04533 | 18000247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18000247&pc=OZD | 09/24/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 09/24/2023 | 10/24/2023 |  |  |  |  | 10/24/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 2024374154 |  | IC5552 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Failure to Charge | Battery Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the battery on the automated impella controller (aic) being utilzed for an impella 5.5 was not charging. the battery stayed at 50% and when plugged in, the charge never increased. there was no reported patient harm. | [5 empty] | The investigation into the reported battery charging issues has been completed. data analysis of the automated impella controller (aic) logs confirmed the reported event. the reported event was also confirmed during functional testing at the bench level. the root cause was determined to be a defective battery due to cell imbalance. the failure modes will be monitored and trended.
33343 | 1220648-2023-04532 | 18000248 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18000248&pc=OZD | 06/03/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 06/03/2023 | 10/24/2023 |  |  | 04/30/2025 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024330887 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Power Up | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/03/2023 |  |  |  |  | The user facility reported 70yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that after placing the pump and trying to start it at p2, a pump failure stopped alarm appeared. attempts were made to restart but the pump never started. pump was then removed and new pump was inserted. there was no patient harm reported. | [5 empty] | The investigation for the failure to start on the pump had been completed. the pump was sent back for evaluation, and electrical testing revealed open motor current lines for all three phases. investigation showed open motor lines near the catheter side within the red handle and a significant gap between the kink protector and red handle. data logs indicated a motor current spike followed by a sudden pump stop. the root cause of the issue was an open electrical line due to excessive force applied. instructions for use for the related event are as follows: ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
33344 | 1220648-2023-04534 | 18000249 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18000249&pc=OZD | 05/15/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 05/15/2023 | 10/24/2023 |  |  | 03/31/2025 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024322182 |  | 1000100 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Female |  |  | Black Or African American | Malfunction | Yes | No | Infusion or Flow Problem | Pressure Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/15/2023 |  |  |  |  | The user facility reported 31yo female was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that high purge pressure and blocked purge flow occurred during impella support. the purge cassette was replaced to resolve the issue, however, the pump flow had already decreased considerably. when troubleshooting failed to resolve the low flow issue, the decision was made to replace the pump. upon removal, a clot was noted around the outlet of the device. there was no patient harm reported. | [5 empty] | The investigation for the reported low flows, high purge pressure, and thrombus has been completed. pump was not returned for investigation. data log shows several pump stop with the high motor current alarm. doctor reported thrombus found in left atrium. the cause of the pump stop issue was biomaterial based on given clinical details and data log review. the instructions for use states that the following regarding thrombus prevention: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ it also states ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ this report is being filed as part of a retrospective review of historical records.
33345 | 1220648-2023-04535 | 18000251 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18000251&pc=OZD | 06/05/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 06/05/2023 | 10/24/2023 |  |  | 04/30/2025 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024336987 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/05/2023 |  |  |  |  | The user facility reported a 61yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was pump saturation after a few minutes on support. troubleshooting was performed but the the device had to be removed and replaced. there was no patient harm reported. | [5 empty] | The investigation for the reported flow saturation has been completed. the product was returned and no biomaterial or other abnormalities were found. the pump was primed with no issues and run in a bucket of water. the pump was run at p-levels 4 and 5 to recreate field performance and ran within specification. the pump was then run at p-9 and was again within specification. no functional issues were found with the product. the data logs from the field were analyzed and show multiple motor current spikes suggesting something was impacting the impeller and creating additional resistance. a large motor current spike and speed deviation without a p-level change occurs right before the period of flow saturation. this is characteristic of biomaterial ingestion. the root cause of the flow saturation was most likely biomaterial as the log evidence points to interaction with impeller and sustained additional resistance.this report is being filed as part of a retrospective review of historical records.
33346 | 1220648-2023-04536 | 18000252 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18000252&pc=OZD | 05/25/2023 | 09/25/2024 |  | 10/24/2023 |  |  | 05/25/2023 | 10/24/2023 |  |  | 03/31/2025 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024321574 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Insert | Failure to Advance | Pressure Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/25/2023 |  |  |  |  | It was reported during insertion the impella, it was first attempted to be inserted over wire and it was unable to advance into the ascending aorta. the impella was removed out of body multiple times with idle time out of body. angiography showed large calcium buildup that was not allowing impella to pass. the medical team decided to use a balloon to open the vessel and a 8x 40 balloon was applied. the impella was reloaded and successfully passed the lesion. the patient's family made the decision to transition the patient to comfort care and no further interventions were given to the patient. the patient expired while on impella 5.5 support. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | The user facility reported 75 yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that when implanted with the impella 5.5, a high purge pressure alarm occurred. the tissue plasminogen activator (tpa) protocol was initiated in the purge fluid and the alarm resolved. there was no patient harm reported. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-04536 was provided in sections b1, b2, b5, d6a, d6b, h1, and h6. | The investigation for the reported high purge pressure has been completed. no data logs were provided. the root cause of the high purge pressure issue was not determined since product was not returned and insufficient information provided. this report is being filed as part of a retrospective review of historical records.
33347 | 1220648-2023-04537 | 18000256 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18000256&pc=OZD | 06/19/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 06/19/2023 | 10/24/2023 |  |  | 01/31/2025 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024296622 |  | 1000100 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/19/2023 |  |  |  |  | The user facility reported 45yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the impella 5.5 had a decreasing purge flow and rising purge pressure. tissue plasminogen activator (tpa) protocol was started. there were no further issues with pump reported. | [5 empty] | The investigation for the high purge pressure has been completed. no logs were returned for investigation. per the provided clinical information, the root cause of the high purge pressure issue was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records.
33348 | 1220648-2023-04538 | 18000257 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18000257&pc=OZD | 02/15/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 02/15/2023 | 10/24/2023 |  |  | 12/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023284588 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/15/2023 |  |  |  |  | The user facility reported 66yo male with cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that while on impella support, a high purge pressure alarm occurred. an attempt to troubleshoot using tissue plasminogen activator (tpa) protocol but was not able to resolve the issue, thus decision was made to remove impella. there was no patient harm reported. | [5 empty] | The investigation for the high purge pressure has been completed. no product was returned for evaluation. data logs were returned and high purge pressure, purge system blocked alarms were observed. the purge pressure step wise rise was observed along with stalling of purge flow ticks due to likely catheter kinking and later immediately purge pressure decreasing during unkinking with increase in purge flow. the cause of the high purge pressure issue was most likely a kinked purge line in the catheter as per data log review. this report is being filed as part of a retrospective review of historical records.
33349 | 1220648-2023-04539 | 18000260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18000260&pc=OZD | 06/23/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 06/23/2023 | 10/24/2023 |  |  | 04/30/2025 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024332348 |  | 1000100 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/23/2023 |  |  |  |  | The user facility reported a 56yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient was on support for 3 days when there were high purge pressure and subsequently saturated flows. the issue resolved spontaneously. there was no patient harm reported. | [5 empty] | The investigation for the reported saturated flows has been completed. the cause of the saturated flow cannot be determined since the complaint device not returned for investigation and data logs review didn't show any indication regarding the issue on the time stated for the complaint. this report is being filed as part of a retrospective review of historical records.
33350 | 1220648-2023-04540 | 18000267 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18000267&pc=OZD | 07/09/2023 | 10/24/2023 |  | 10/24/2023 |  |  | 07/09/2023 | 10/24/2023 |  |  | 04/30/2025 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024337420 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/09/2023 |  |  |  |  | The user facility reported male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that after 2 days of support, the pump had high purge pressure alarms. troubleshooting was performed by with tissue plasminogen activator (tpa) and stopping/starting the pump without any resolution. the pump was explanted in the operating room. there was no patient harm reported. | [5 empty] | The investigation for the reported high purge pressures has been completed. the root cause of the high purge pressure was due to biomaterial in the purge gap. the pump was not able to start and the high purge pressure reproduced, however it was soaked in warm water with tergazyme for 10 minutes then the device ran fine and the high purge pressure resolved. this report is being filed as part of a retrospective review of historical records.
33351 | 1220648-2023-04541 | 18000287 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18000287&pc=OZD | 02/25/2023 | 01/12/2024 |  | 10/24/2023 |  |  | 02/25/2023 | 10/24/2023 |  |  | 09/30/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023266023 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/25/2023 |  |  |  |  | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that a patient developed a large hematoma in the right chest wall following right axillary insertion of an impella 5.5 pump. the patient was given three units of packed red blood cells (prbc¿s) to counteract the resulting drop in hemoglobin levels. although hemolysis was noted, no intervention was required, and the condition resolved without issue. support continued with the implanted pump following the administration of blood products. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
33352 | 1220648-2023-04542 | 18000291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18000291&pc=OZD | 10/19/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 10/19/2022 | 10/24/2023 |  |  | 06/30/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023240797 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/19/2022 |  |  |  |  | The user facility reported a 45-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a 5.5 delivery fail due to native anatomy and kinking. there was no lasting harm or injury. the team removed and replaced the pump. | [5 empty] | The product was returned and the inflow cage was confirmed to be damaged. the root cause of the delivery issue was determined to be patient condition as the pump could not be delivered through the patients old bioprosthetic aortic valve replacement avr), this report is being filed as part of a retrospective review of historical records.
33353 | 1220648-2023-04543 | 18000306 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18000306&pc=OZD | 07/13/2023 | 01/18/2024 |  | 10/24/2023 |  |  | 07/13/2023 | 10/24/2023 |  |  | 04/30/2025 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024339810 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/13/2023 |  |  |  |  | The user facility reported a 67-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the medical staff was unable to advance the cannula tip beyond the tunneled graft anastomosis. the pump was removed and the catheter was shown to have kinks in several places. a new pump was inserted successfully. there was no harm to the patient as a result of this incident. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella has not been returned. however, clinical details provided are sufficient to determine root cause. as per clinical detail, the doctor was unable to pass the cannula tip beyond the graft. doctor found kinked catheter shaft. the pump was exchanged with new one and delivered successfully. the cause of the delivery issue was most likely use related, based on provided clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
33354 | 1220648-2023-04544 | 18000309 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18000309&pc=OZD | 03/23/2023 | 11/29/2024 |  | 10/24/2023 |  |  | 03/23/2023 | 10/24/2023 |  |  | 01/31/2025 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024294809 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/23/2023 |  |  |  |  | The user facility reported a 59-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had an inability to wirelessly deliver the impella 5.5 for a surgical patient undergoing (coronary artery bypass graft) cabg. the patient was unable to come off pump and so the team attempted to deliver a 5.5 but, only the axillary insertion kit and sheath were delivered. the team attempted all wire and diagnostic catheter placements, but failed and so the 5.5 placement was aborted. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issue was wireless insertion. | The investigation for the reported issue has been completed. the impella device was not received from the customer. the root cause of the delivery issue was wireless insertion. the impella device is not indicated for wireless insertion . should the device or any new information be received, a supplemental mdr will be filed. h6: investigation findings and conclusion codes were inadvertently reported incorrect in the initial report which are corrected in this report. | The root cause of the delivery issue was wireless insertion. the impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
33355 | 1220648-2023-04545 | 18000310 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18000310&pc=OZD | 08/29/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 08/29/2022 | 10/24/2023 |  |  | 06/30/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  |  |  | 2023232071 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/29/2022 |  |  |  |  | The user facility reported a 60-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported high purge pressure during use of the device. the pump was removed and replaced with no harm to the patient. | [5 empty] | Product was visually inspected. biomaterial was observed at the impeller purge gap. high purge pressure is most likely due to the biomaterial in the impeller purge gap. rt logs were reviewed and a spike in motor current was observed when the pump was initially turned on. this report is being filed as part of a retrospective review of historical records.
33356 | 1220648-2023-04546 | 18000316 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18000316&pc=OZD | 09/18/2022 | 09/25/2024 |  | 10/24/2023 |  |  | 09/18/2022 | 10/24/2023 |  |  | 05/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023220075 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [9 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/18/2022 |  |  |  |  | It was reported that the trans-esophageal echocardiogram (tee) in the cardiovascular operating room revealed thrombus throughout vasculature, including large thrombus on or near pulmonic valve and pulmonary embolism. a thrombectomy was scheduled for the following morning. it was reported during impella support the patient experienced bleeding and was transfused with platelets and blood. the patient condition began to deteriorate following impella alarms for suction events and impella in aorta alarm. the pump was turned down to p2, and tee was used to visualize the impella. the impella was found to still be in the left ventricle, but shallow. the impella was repositioned, but suction alarms continued above p5. the impella's inlet did not appear obstructed under tee. the patient continued to deteriorate with decrease in mean arterial pressures and increase in vasopressor needs. patient went into ventricular tachycardia and was defibrillated. later went into pulseless electrical activity with multiple rounds of cardiopulmonary resuscitation initiated. the code was eventually terminated and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | The user facility reported that a 66-year-old female patient implanted with the impella 5.5 for mechanical circulatory support experienced low pump flows and suction alarms. the pump flow decreased suddenly upon the occurrence of "suction" alarms about 5 days into support. the pump flow was unable to exceed 3 l/min without recurrent "suction" alarms and issue persisted despite the pump being in proper position and the presence of sufficient left ventricular volume. despite the low flow, the decision was made to keep the device implanted. it was noted also that the patient had clotting issues. there was no reported harm to the patient. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-04546 was provided in sections b1, b2, b5, d6a, d6b, h1, and h6. | The investigation into the reported low pump flows was completed. the device was returned for evaluation. analysis of the data logs showed the pump flow dropped following a sudden motor current spike. additionally, after the spike, the "suction" alarm triggered with low flow and non-pulsatile motor current waveforms. visual inspection of the pump revealed the presence of a large mass of biomaterial on and around the outflow cage and the impeller. a smaller mass of biomaterial was wrapped around the cannula. based on the available information, the root cause of the low pump flow was ingested biomaterial. this report is being filed as part of a retrospective review of historical records.
33357 | 1220648-2023-04547 | 18000330 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18000330&pc=OZD | 06/14/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 06/14/2022 | 10/24/2023 |  |  | 12/31/2023 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 |  |  | ABIOMED, INC. |  |  | US |  | 2023178369 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 21 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/14/2022 |  |  |  |  | The user facility reported that a 21-year-old male patient implanted with the impella cp for mechanical circulatory support developed thrombus. during removal of the device, echocardiogram results revealed a clot about 20-30 cm long in the descending aorta. previous echocardiogram results performed prior to explant results confirmed a clot visible under the impella catheter. vascular surgeon subsequently placed an aortic stent to push clot against the wall. "echo and angiogram did not show clot present. no clot was visible on the catheter." of note, the patient was reported to be hypercoagulopathic. | [5 empty] | The investigation into the reported thrombus was completed. the device was returned for evaluation. the pump was visually inspected, and no clot or abnormalities were observed. based on the available, the root cause of the issue was patient condition related since the patient was reported to be hypercoagulopathic. this report is being filed as part of a retrospective review of historical records.
33358 | 1220648-2023-04548 | 18000333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18000333&pc=OZD | 09/24/2023 | 01/31/2024 |  | 10/24/2023 |  |  | 09/24/2023 | 10/24/2023 |  |  | 06/30/2025 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024376514 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  |  | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported that after the impella cp was implanted, low pump flows occurred. staff tried to restart and reposition the impella, and they were unsuccessful. staff successfully replaced the impella. | [5 empty] | The investigation into the low pump flow issue has been completed. the data logs were not returned. however, the pump was returned for investigation. visual inspection found no issues on the pump. the pump ran without issue under bench test. the root cause of low flow was unable to be determined as the failure mode could not be reproduced and no data logs were returned for review. | The pump was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿
33359 | 1220648-2023-04549 | 18000348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18000348&pc=OZD | 12/02/2022 | 10/24/2023 |  | 10/24/2023 |  |  | 12/02/2022 | 10/24/2023 |  |  | 08/31/2024 |  | 10/24/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 |  |  | ABIOMED, INC. |  |  | US |  | 2023265347 |  | 0048-0045 |  | Yes |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Puncture/Hole | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/02/2022 |  |  |  |  | The user facility reported that a 52-year-old female patient implanted with the impella cp for mechanical circulatory developed a hole in the catheter of the device. the patient was reported to have tortuous anatomy with axillary insertion; however, the device was eventually inserted with some effort. afterwards, the medical staff noticed leaking at the location of the graft anastomosis. additional sutures were placed while the device was in the graft. after the pocket was closed, blood was noted to be leaking from the red impella plug by the grey/red junctions. no further information was provided. there was no reported harm to the patient. | [5 empty] | The investigation into the reported low pump flows was completed. the device was returned for evaluation. analysis of the data logs did not reveal any performance issues. the device was returned. the returned pump was visually inspected and a hole was observed in the catheter at the 29 cm mark. the cause of the bleeding from the red handle was blood ingress due to improper suturing resulting in a hole in the catheter. this report is being filed as part of a retrospective review of historical records.
33360 | 1220648-2023-04308 | 17984377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17984377&pc=OZD | 06/29/2023 | 10/23/2023 |  | 10/23/2023 |  |  | 06/29/2023 | 10/23/2023 |  |  | 04/30/2025 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 |  |  | ABIOMED, INC. |  |  | US |  | 2024333087 |  | 1000100 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/29/2023 |  |  |  |  | The user facility reported that a 44-year-old male patient implanted with the impella 5.5 for mechanical circulatory support developed a purge leak. blood was observed to be leaking from the red impella plug while the team was preparing to transfer the patient from the operating room table to the intensive care unit (icu) bed for transport. the physician was notified of the risks associated with electronics and blood exposure and elected to utilize bone wax and tegaderm to seal the area of leakage. there were no plans to exchange pump. no further information was provided. there was no reported harm to the patient. | [5 empty] | The investigation into the reported purge leak was completed. the device was not returned for evaluation. analysis of the data logs does not indicate any failure related to pump stop or optical signal caused by the blood in red handle. based on the available information, the root cause of the product damage issue was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33361 | 1220648-2023-04309 | 17984398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17984398&pc=OZD | 06/21/2023 | 01/11/2024 |  | 10/23/2023 |  |  | 06/21/2023 | 10/23/2023 |  |  | 12/31/2024 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 |  |  | ABIOMED, INC. |  |  | US |  | 2024315423 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/21/2023 |  |  |  |  | The user facility reported that a 52-year-old male patient implanted with the impella 5.5 for mechanical circulatory support developed thrombus. the patient was still intubated and sedated and had previously experienced multiple episodes of pulseless electrical activity (pea), where he self-converted. heparin was in the purge for treatment of left ventricular (lv) thrombus. no further information was provided. | [5 empty] | The investigation into the reported thrombus has been completed since the original report was filed. no product was returned. as per clinical lv thrombus observed in patient, but no thrombus was found on pump. the cause of the thrombus issue was most likely patient condition related due to lv thrombus observed in patient per clinical review. | The investigation into the reported thrombus was completed. the device was not received from the customer and therefore,an evaluation of the device was not possible. based on the available information a root cause to the reported event could not be determined. should additional information be received, this investigation will be updated accordingly. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33362 | 1220648-2023-04310 | 17984814 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17984814&pc=OZD | 03/13/2023 | 10/23/2023 |  | 10/23/2023 |  |  | 03/13/2023 | 10/23/2023 |  |  | 12/31/2024 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2024300418 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/13/2023 |  |  |  |  | The user facility reported a male patient was implanted with an impella 5.5 device for mechanical circulatory support. the clinical team reported that high purge pressure occurred about 7 days into impella support, after observing the purge pressure increase to above 1000 mmhg. the motor current was unaffected at the time. no purge line kinks were observed, and replacing the purge cassette did not resolve the issue. tpa was then administered into the purge line, which successfully resolved the issue several hours later. | [5 empty] | Only the pump cannula was returned. no damage or abnormalities were found on the segment of the impella that was returned. after reviewing the clinical information, it was determined that the cause of the high purge pressure was biomaterial occluding the purge gap since adding tpa to the purge resolved the issue. per the ifu: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ this report is being filed as part of a retrospective review of historical records.
33363 | 1220648-2023-04311 | 17984991 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17984991&pc=OZD | 09/23/2023 | 02/02/2024 |  | 10/23/2023 |  |  | 09/23/2023 | 10/23/2023 |  |  | 06/30/2025 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024372962 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had impella cp support ongoing when the team observed an access site bleed. the bleed was treated by femostop placement and 1 unit of red blood cells. the patient had the pressures drop and a suction alarm occurred, but the pump was not explanted. the pump did remain on for 4 days until it was weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
33364 | 1220648-2023-04312 | 17985061 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17985061&pc=OZD | 03/30/2023 | 10/23/2023 |  | 10/23/2023 |  |  | 03/30/2023 | 10/23/2023 |  |  | 11/30/2024 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2023281894 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/30/2023 |  |  |  |  | The user facility reported a patient of unknown gender and age was implanted with an impella 5.5 device for mechanical circulatory support. during support, a fluid leak was discovered between the purge filter and the red plug. it was noted that sodium bicarbonate was present in the purge at the time of the failure. as a result, the perfusionist had independently switched to a purge filter bypass. no harm done to the patient. | [5 empty] | The impella device was discarded by the customer and therefore,an evaluation of the device was not possible. upon conclusion of the investigation, a definitive cause for the sidearm leak could not be determined. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
33365 | 1220648-2023-04313 | 17985189 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17985189&pc=OZD | 09/23/2023 | 02/16/2024 |  | 10/23/2023 |  |  | 09/23/2023 | 10/23/2023 |  |  | 08/31/2024 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023272822 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Activation, Positioning or Separation Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported that despite the blue suture pad being hubbed out and angle matched, the patient developed significant sized hematoma post implant of the impella cpand dressing. patient received one unit of prbc's to increase hemoglobin, heparin was halted due to bleeding, and manual pressure was applied to the hematoma. additionally, staff noted that the bedside echo showed the impella cp was deep at 4.7 cm. the staff pulled back cautiously measuring 3.4 cm from aortic annulus. the patient was on impella cp support for 52.25 hours before the patient expired. the patient had a dnr. no allegations were made towards the impella cp for cause of death. | [5 empty] | The investigation into the access site bleeding ¿ major and the positioning issues has been completed. the device was not returned for investigation. per the given clinical information, the patient was vigorously moving during and after transfer from another hospital. the cause of the bleeding was high patient act values. the root cause of the positioning issue was patient movement. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows:¿assess access site for bleeding and hematoma.¿¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33366 | 1220648-2023-04314 | 17985347 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17985347&pc=OZD | 12/17/2021 | 10/23/2023 |  | 10/23/2023 |  |  | 12/17/2021 | 10/23/2023 |  |  | 07/31/2023 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022119313 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/17/2021 |  |  |  |  | The user facility reported a 55 year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was also reported that a ¿low purge pressure alarm¿ was triggered after replacing a leaky purge cassette. the new purge cassette was leaking at the yellow port. the purge bypass option was discussed. sodium bicarbonate was used in the purge solution. it was reported that there were no adverse events related to the impella device and no medical intervention was required. | [5 empty] | The investigation into the reported low purge pressure alarm and leaking at the yellow port has been completed. the medical center did not return any impella products for benchtop analysis nor any data logs for review, only the clinical report was evaluated. of note, sodium bicarbonate was used as a purge solution. the root cause of the luer leak/low pressure alarms was due to a damaged yellow luer. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33367 | 1220648-2023-04315 | 17985366 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17985366&pc=OZD | 12/21/2021 | 10/23/2023 |  | 10/23/2023 |  |  | 12/21/2021 | 10/23/2023 |  |  | 07/31/2023 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022115217 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/21/2021 |  |  |  |  | The user facility reported a 55 year-old male was implanted with an impella device for circulatory mechanical support. it was reported that due to a bicarbonate interaction, there was a cracked luer lock and a purge bypass was performed. there was no patient injury reported. | [5 empty] | The medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. of note, bicarbonate was used with device. the root cause of the luer lock crack/leak was due to a damaged yellow luer. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33368 | 1220648-2023-04316 | 17985421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17985421&pc=OZD | 12/24/2021 | 10/23/2023 |  | 10/23/2023 |  |  | 12/24/2021 | 10/23/2023 |  |  | 08/31/2023 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022136216 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  | White | Malfunction | No | Yes | Fluid/Blood Leak | Optical Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/24/2021 |  |  |  |  | The user facility reported a 67 year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that a purge fluid leak resolved with a bypass. the pump's placement signal also failed. the pump was removed and replaced. it was noted that bicarbonate was used during the purge. | [5 empty] | The pump was returned for evaluation which confirmed the optical signal issue on the pump. evaluation of the yellow luer confirmed leakage. the root cause of the sidearm leak was due to yellow luer damage as a result of bicarb weakening the plastic. the root cause of the ps issues could not be determined. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33369 | 1220648-2023-04318 | 17985452 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17985452&pc=OZD | 04/02/2023 | 10/07/2025 |  | 10/23/2023 |  |  | 04/02/2023 | 10/23/2023 |  |  | 01/31/2025 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024297543 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/02/2023 |  |  |  |  | The user facility reported a patient of unknown age and gender was implanted with an impella 5.5 device for mechanical circulatory support. during support, a fluid leak was noted from the purge filter. a purge bypass was subsequently completed and support with the implanted pump was maintained. there was no patient harm. | The complainant also reported that the placement signal failed without mobilizing the patient. as a result, neither the placement curve nor the left ventricle curve was visible on the screen. |  |  |  |  | The pump was returned; however, the sidearm was not attached and thus an evaluation related to the reported issue could not be completed. upon conclusion of the investigation, a definitive cause for the noted leak could not be established. per the ifu: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. placement signal not reliable alarm reproduced when plugged into aic. removing biomaterial from the sensor did not resolve the placement signal issue. optical continuity testing revealed fiber damage at the catheter kink next to the kink protector. therefore, it was determined that the cause of the optical signal issue was a damaged optical fiber line. all pertinent information available to abiomed has been submitted at this time. b5. updated to include optical signal issue description. h6. updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2023-04318. b7. other relevant history, including preexisting medical conditions updated. d3. updated. d6.a. / d6.b. updated. d10. concomitant medical products and therapy dates updated. f6. and f8. should have been left blank on the initial report. g1. manufacturing site name, email, and address updated.
33370 | 1220648-2023-04319 | 17985490 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17985490&pc=OZD | 12/31/2021 | 10/23/2023 |  | 10/23/2023 |  |  | 12/31/2021 | 10/23/2023 |  |  | 09/30/2023 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022141340 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Alarm System | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/31/2021 |  |  |  |  | The user facility reported a 62 year-old female was implanted with an impella device for mechanical circulatory support. it was reported that there were purge system open alarms after a cassette change. there was also leaking around the yellow luer and the staff was aware of the troubleshooting procedure of bypassing. there was 5%dextrose in bicarbonate in the purge. there was no patient injury reported. | [5 empty] | The investigation into the reported purge system leak has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated the cause of the purge leak was sidearm yellow luer damage due to bicarbonate in the purge. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33371 | 1220648-2023-04317 | 17985517 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17985517&pc=OZD | 12/27/2021 | 10/23/2023 |  | 10/23/2023 |  |  | 12/27/2021 | 10/23/2023 |  |  | 07/31/2023 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022130585 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Activation Failure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/27/2021 |  |  |  |  | The user facility reported a 81yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that after the pump was prepared with sodium bicarbonate and placed in the graft an "impella defective" alarm popped up. pump was immediately taken out and a new one prepped and inserted. there was no patient harm reported. | [5 empty] | The investigation for the pump stop has been completed. the data logs were reviewed and an impella defective alarm was observed. no pump stop occurred. pump never activated. the returned pump was run in a basin of water and the pump operated to specification. the cause of the impella defective alarm upon startup was most likely use related since the returned pump operated to specification. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33372 | 1220648-2023-04320 | 17985548 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17985548&pc=OZD | 04/08/2023 | 10/23/2023 |  | 10/23/2023 |  |  | 04/08/2023 | 10/23/2023 |  |  | 04/30/2023 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2022082653 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/08/2023 |  |  |  |  | The user facility reported a 55-year-old male patient was implanted with an impella 5.5 device for mechanical circulatory support. during support, the physician documented that the purge filter was defective which led to a drop in purge pressure. a purge bypass was successfully completed and support with the pump was maintained. there was no patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. the cause of the purge leak was not determined since product was not returned. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
33373 | 1220648-2023-04321 | 17985604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17985604&pc=OZD | 12/31/2021 | 10/23/2023 |  | 10/23/2023 |  |  | 12/31/2021 | 10/23/2023 |  |  | 03/31/2023 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022058031 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/31/2021 |  |  |  |  | The user facility reported a 65yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that during support, the pump developed a leak at the yellow luer after the purge solution was changed to sodium bicarbonate. a purge bypass was performed and the purge pressure and flows normalized. a few days later, the purge flows dropped which led to an pump stop. the pump was successfully explanted. there was no patient harm reported. | [5 empty] | The investigation into the leak at the yellow luer and the pump stop has been completed. the device nor the data logs were returned for evaluation. however, the clinical information was sufficient in determining the root cause, given the results of prior failure investigations. it was concluded that the cause of the purge leak was sidearm yellow luer damage due to sodium bicarbonate use. the cause of the pump stop was most likely blood ingress as a result of high purge pressure. the instruction for use for the related event indicates :¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33374 | 1220648-2023-04322 | 17985614 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17985614&pc=OZD | 06/04/2023 | 09/29/2025 |  | 10/23/2023 |  |  | 06/04/2023 | 10/23/2023 |  |  | 04/30/2025 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024332340 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 06/04/2023 |  |  |  |  | The user facility reported a 57-year-old female patient was implanted with an impella 5.5 device for mechanical circulatory support. us complainant reported that during pump prep, prior to insertion, there was an optical sensor failure alarm. they followed screen prompts and discussed with csc. they were unable to determine pump integrity, so they opted to prep a new pump. there was no known harm or injury to the patient. | [5 empty] | The investigation into the into the optical signal issue has been completed since the original report was filed. no product was returned for evaluation. data logs were returned, and optical system catheter/controller error observed at case start. it was found that the pump initial 5s parameters were trending snr low at zero, contrast decreasing, gain increasing to 2.59 and lamp at 100 with placement signal spike to 3000. later the 5s parameters and placement signal resolve within 2 minutes after pump plug connect. the cause of the optical signal failure was most likely an air gap from between the aic and the patient cable plug per clinical and log review. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Corrections that were made from the initial manufacturer device report number 1220648-2023-04322: d4 model number and unique identifier (udi) # corrected. f6 and f8 should have been left blank. g1 mdr reporting contact email.
33375 | 1220648-2023-04323 | 17985619 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17985619&pc=OZD | 01/10/2022 | 10/23/2023 |  | 10/23/2023 |  |  | 01/10/2022 | 10/23/2023 |  |  | 09/30/2023 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022149857 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/10/2022 |  |  |  |  | The user facility reported a 64 year-old female with acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was a crack in the yellow luer. there was no patient injury reported. | [5 empty] | The investigation into the reported purge system leak has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. of note, sodium bicarbonate was used as a purge solution. the cause of the purge leak was sidearm yellow luer damage. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33376 | 1220648-2023-04324 | 17985640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17985640&pc=OZD | 01/02/2022 | 10/23/2023 |  | 10/23/2023 |  |  | 01/02/2022 | 10/23/2023 |  |  | 09/30/2023 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022147149 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/02/2022 |  |  |  |  | The user facility reported a 52yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there were "purge system blocked" alarms. a visual inspection was performed and found a kinked purge line due to the patient laying on purge side arm. tissue plasminogen activator (tpa) was added to purge with no change in purge flow or purge pressure. a motor current rise occurred with pump stop. the pump was in use beyond the instructions for use. | [5 empty] | The investigation into the pump stop has been completed. the device was returned for evaluation. upon evaluation, blood was seen in the purge line indicating the pump stop was due to blood ingress. per the clinical account, the cause of the blood ingress is a result of the patient lying on the sidearm. therefore the root cause of the pump stop was due to blood ingress into the motor as a result of patient movement. instructions for use for the related event are as follows: ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33377 | 1220648-2023-04325 | 17985685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17985685&pc=OZD | 01/13/2022 | 10/23/2023 |  | 10/23/2023 |  |  | 01/13/2022 | 10/23/2023 |  |  | 08/31/2023 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022139013 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/13/2022 |  |  |  |  | The user facility reported a 77 year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there were purge pressure low alarms. the staff found that the patient¿s gown was wet and there was condensation within the purge sidearm retainer, and the connection was wet. the issue reportedly resolved when dropping the sidearm down at a 90-degree angle. as the purge pressure and flow was stable, the physician decided to manage the patient with the sidearm in this position. the sidearm was placed on an arm board, taped down, and allowed to hang at a 90-degree angle. the purge fluid was changed to 10% dextrose. although the purge pressure and flows were stable, the leak continued to occur intermittently. the pump support remained. there was no patient injury reported. | [5 empty] | The returned product was connected to a purge system and a purge leak at the sidearm filter was reproduced. analysis of the provided field data logs confirmed multiple low purge pressure alarms throughout the case. microscopic inspection revealed damage to the inlet neck of the purge sidearm filter. the root cause of the low purge pressure was a damaged purge filter. this damage most likely occurred due to excessive force during use with the retainer in place. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33378 | 1220648-2023-04326 | 17985726 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17985726&pc=OZD | 01/25/2022 | 10/23/2023 |  | 10/23/2023 |  |  | 01/25/2022 | 10/23/2023 |  |  | 08/31/2023 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022139031 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/25/2022 |  |  |  |  | The user facility reported a 85yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that 1 hour post explant of the pump, the patient developed neurological deficits and was diagnosed with a cerebral vascular accident (cva). the patient had a thrombectomy performed in the "pca". the patient's neurological function returned to baseline thereafter. there was no product problem reported. | [5 empty] | The investigation into the post explant cva has been completed. the device was not returned for evaluation. as the cva occurred post explant, the relationship between the cva and impella cannot be confirmed. the root cause of the cva could not be determined. the instructions for use for the related adverse event: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33379 | 1220648-2023-04327 | 17985735 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17985735&pc=OZD | 06/06/2023 | 10/23/2023 |  | 10/23/2023 |  |  | 06/06/2023 | 10/23/2023 |  |  | 04/30/2025 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2024339686 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 24 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/06/2023 |  |  |  |  | The user facility reported a 24-year-old male patient was implanted with an impella 5.5 device for mechanical circulatory support. during support, it was documented that high purge pressure alarms were triggered. it was verified that bicarbonate was in use at the time of the noted issue. the pump was subsequently explanted the following day and the patient remained supported with extra-corporeal membrane oxygenation (ecmo). there was no harm or injury and the patient. | [5 empty] | The device was returned and the investigation is ongoing. upon its completion, a supplemental report will be submitted with the finding. per the ifu: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
33380 | 1220648-2023-04328 | 17985780 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17985780&pc=OZD | 01/24/2022 | 10/23/2023 |  | 10/23/2023 |  |  | 01/24/2022 | 10/23/2023 |  |  | 10/31/2023 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022154370 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Increase in Pressure | Improper Flow or Infusion | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/24/2022 |  |  |  |  | The user facility reported a 64yo female was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was high purge pressure and low purge flow observed during case. bicarbonate temporarily fixed the problem and tissue plasminogen activator (tpa) was eventually needed. there was a clot observed in the outflow cage during explant. there was no patient harm reported. | [5 empty] | The investigation for the reported high purge pressure and low flows has been completed. the data logs were returned and confirm the high purge pressure condition. a gradual increase in purge pressure can be seen, this condition resolves toward the end of the case. the product was returned with the catheter cut. thrombus could be seen in the outflow and on the catheter. the root cause of the high purge pressure was determined to be biomaterial buildup as tpa was able to resolve the issue and no heparin was used in the purge for the first 4 days of support. the instructions for use provides details on how prevent thrombus formation in the pump. it states the following: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33381 | 1220648-2023-04329 | 17985820 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17985820&pc=OZD | 02/04/2022 | 10/23/2023 |  | 10/23/2023 |  |  | 02/04/2022 | 10/23/2023 |  |  | 09/30/2023 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED INC. |  |  | US |  | 2022141894 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/04/2022 |  |  |  |  | The user facility reported a 76yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the device was sitting toward the mitral valve causing incompetence of the mitral. attempts were made to reposition but the device would not sit properly in the left ventricle and was causing aortic insufficiency. after several attempts, the impella came back across the aortic valve, and was removed. the patient was placed on extracorporeal membrane oxygenation (ecmo). | [5 empty] | The investigation for the reported positioning issue has been completed. neither the product nor the data logs were returned. after reviewing the clinical information, it was determined that the cause of the positioning issues was most likely wireless insertion, which the impella is not indicated for the device. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33382 | 1220648-2023-04330 | 17985839 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17985839&pc=OZD | 02/26/2022 | 10/23/2023 |  | 10/23/2023 |  |  | 02/26/2022 | 10/23/2023 |  |  | 05/31/2023 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022088330 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Optical Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/26/2022 |  |  |  |  | The user facility reported 69yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that after 5 days in use, the pump lost a reliable placement signal. the team attempted troubleshooting and accidentally pulled the pump into the aorta and was unable to re-position the pump. the team medically managed the patient and replaced the pump. there was no patient harm reported. | [5 empty] | The investigation for the "placement signal low" alarms has been completed. data logs showed the sudden occurrence of "placement signal low" alarms was accompanied by a gradually decreasing placement signal waveform. there were frequent "suction" alarms before and after the "placement signal low" alarms with reduced motor current pulsatility. otherwise, all optical sensor metrics were within specification. after reviewing the clinical information, it was determined that the cause of the optical signal issue was most likely patient movement since the issue occurred immediately after the patient shifted. the patient's cardiac anatomy was likely a contributing factor. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33383 | 1220648-2023-04331 | 17985859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17985859&pc=OZD | 06/27/2023 | 01/24/2024 |  | 10/23/2023 |  |  | 06/27/2023 | 10/23/2023 |  |  | 10/31/2025 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024347080 |  | 1000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Prime | Activation Failure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/27/2023 |  |  |  |  | The user facility reported a 53-year-old female patient was implanted with an impella device for mechanical circulatory support. the complainant in japan had a 5.5 delivered for support of a patient who was in with a history of coronary artery disease (cad) and with a plan for coronary artery bypass grafting (cabg). the patient had extracorporeal membrane oxygenation (ecmo) and intra-aortic balloon pump (iabp). during priming, the 5.5 pump experienced an unspecified failure and was subsequently replaced by an impella cp device. the cp was successfully placed for support. no harm or injury noted. the pump was not noted to have stopped. | [5 empty] | The investigation into the reported pump did not start has been completed since the original report was filed. the device was not returned by the medical facility. the cause of the pump did not start issue cannot be determined since no data logs, no complaint device and nor sufficient data returned for investigation. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon conclusion of the investigation, a definitive cause for the reported issue could not be established. this report is being filed as part of a retrospective review of historical records.
33384 | 1220648-2023-04332 | 17985900 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17985900&pc=OZD | 03/07/2022 | 10/23/2023 |  | 10/23/2023 |  |  | 03/07/2022 | 10/23/2023 |  |  | 12/31/2023 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023172951 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/07/2022 |  |  |  |  | The user facility reported a 58yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that during insertion of the pump and torque of the catheter, the pump was kinked. the pump was removed and a different pump was placed. there was no patient harm reported. | [5 empty] | The investigation for the reported kinked pump has been completed. the device was not returned for evaluation. the clinical report provided sufficient information to determine root cause. therefore, the root cause of the delivery issue was due to the user technique. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33385 | 1220648-2023-04333 | 17986046 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17986046&pc=OZD | 03/10/2022 | 10/23/2023 |  | 10/23/2023 |  |  | 03/10/2022 | 10/23/2023 |  |  | 08/31/2023 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022132663 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/10/2022 |  |  |  |  | The user facility reported a 60 year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the cracked sidearm filter was leaking with the retainer in place. purge bypass performed. sodium bicarbonate was being used as the purge solution. there was no patient injury reported. | [5 empty] | The investigation into the reported leaking cracked filter has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. of note, sodium bicarbonate was used as a purge solution. the cause of the purge leak was sidearm filter damage. the cause of the filter damage could not be definitively determined. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33386 | 1220648-2023-04334 | 17986077 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17986077&pc=OZD | 03/10/2022 | 10/23/2023 |  | 10/23/2023 |  |  | 03/10/2022 | 10/23/2023 |  |  | 10/31/2023 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023159349 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/10/2022 |  |  |  |  | The user facility reported a 66 year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that purge solution was leaking from the red plug for the duration of therapy. there was no patient injury reported. | [5 empty] | The impella was returned with only the red pump handle and the purge sidearm intact; only a small segment of both the patient cable and the catheter remained. the red handle was opened, but definitive leaking was not observed. without the remaining parts of the pump attached, there is insufficient evidence to discover the source of the leaking. therefore, the cause of the purge leak was not determined since the impella was not returned intact. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33387 | 1220648-2023-04335 | 17986139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17986139&pc=OZD | 02/27/2022 | 10/23/2023 |  | 10/23/2023 |  |  | 02/27/2022 | 10/23/2023 |  |  | 10/31/2023 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022153149 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/27/2022 |  |  |  |  | The user facility reported a 43yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient had a significant coughing event that forced the impella out of the left ventricle and into the aorta. the patient was unharmed and deemed hemodynamically stable, so the pump was explanted. no damage or abnormalities were observed on the impella upon removal. the malposition caused no harm or injury. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, the clinical report provided sufficient information to determine the root cause of the reported issue. it was determined that the root cause of the positioning issue was due to patient movement. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33388 | 1220648-2023-04336 | 17986169 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17986169&pc=OZD | 03/21/2022 | 10/23/2023 |  | 10/23/2023 |  |  | 03/21/2022 | 10/23/2023 |  |  | 12/31/2023 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023176814 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/21/2022 |  |  |  |  | The user facility reported a 79yo male was implanted with an impella 5.5 device for mechanical circulatory support. an attempt was made to place a 5.5 via the graft but delivery was not possible. the pump would not pass the junction into the ascending aorta. the team removed the pump and found the pump to be thrombosed. the decision was made to insert an impella cp with success. there was no patient harm reported. | [5 empty] | The investigation for the reported thrombosis and delivery issues have been completed. the device nor data logs were returned for evaluation. however, the clinical report is sufficient to determine a root cause. therefore, the root cause of the thrombus was biomaterial buildup and the delivery issues was due to patient condition. the instructions for use provides suggestions on how to prevent thrombus formation into the device. it states :¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33389 | 1220648-2023-04337 | 17986256 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17986256&pc=OZD | 03/23/2022 | 10/23/2023 |  | 10/23/2023 |  |  | 03/23/2022 | 10/23/2023 |  |  | 10/31/2023 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022157461 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/23/2022 |  |  |  |  | The user facility reported a 45yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the impella 5.5 that had been used beyond its indicated use/duration, experienced a pump stop. troubleshooting was attempted by replacing the console but the decision was made to explant. upon explant on day 17, a thrombus was noted in the left ventricle. patient had been on anticoagulation with therapeutic labs for a few days. there was no patient harm reported. | [5 empty] | The investigation for the pump stop has been completed. the pump and the data logs were returned for evaluation. the field data log was reviewed and pump stop revolutions per minute (rpm) deviations alarm was confirmed at the end of the run. the returned pump was visually inspected and biomaterial was observed on the motor housing. the pump was activated in a basin of water and an immediate pump stop alarm 115 rpm deviations was reproduced. electrical resistances were measured and an open line was confirmed. the red handle was opened and a damaged was confirmed. the cause of the pump stop was an open electrical line from excessive force. the instructions for use as it relates to pump stop states the following: "if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33390 | 1220648-2023-04338 | 17986704 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17986704&pc=OZD | 03/29/2022 | 10/23/2023 |  | 10/23/2023 |  |  | 03/29/2022 | 10/23/2023 |  |  | 11/30/2023 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023168301 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/29/2022 |  |  |  |  | The user facility reported a 65 year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the purge sidearm was leaking. no further information was provided. there was no patient harm reported. | [5 empty] | The investigation into the reported leaking on the sidearm leak has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the root cause of the sidearm leak could not be determined since the device was not returned for evaluation. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33391 | 1220648-2023-04339 | 17986766 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17986766&pc=OZD | 03/22/2022 | 10/23/2023 |  | 10/23/2023 |  |  | 03/22/2022 | 10/23/2023 |  |  | 10/31/2023 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2021018079 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/22/2022 |  |  |  |  | The user facility reported a 53 year-old female was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was leaking at the clear yellow valve. the purge flow and purge pressure were in normal range while on bicarbonate. purge pressure alarms occurred and were due to a kinked and leaking sidearm. no other impella related issues. there was no patient harm reported. | [5 empty] | The investigation into the reported leaking on the yellow luer valve has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. sodium bicarbonate was used as a purge solution. the root cause of the yellow luer leak was due to damage as a result of use of bicarbonate. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33392 | 1220648-2023-04340 | 17986793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17986793&pc=OZD | 03/22/2022 | 09/21/2024 |  | 10/23/2023 |  |  | 03/22/2022 | 10/23/2023 |  |  | 11/30/2023 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023167076 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Pumping Stopped | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/22/2022 |  |  |  |  | The user facility reported a 31yo male was implanted with an impella 5.5 device for mechanical circulatory support. the patient was supported also by ecmo. the team had the 5.5 supporting for over a month. it was reported that the purge flows had been decreasing while running sodium bicarbonate. a pump stop occurred. a purge bypass was performed and they were able to get the pump restarted. the family later decided to withdraw care. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation for the pump stop issue has been completed. the pump nor the data logs were returned for evaluation. the clinical reported did not provide sufficient information to determine the root cause. therefore, the root cause of the pump stoppage could not be determined. instructions for use for the related event are as follows: ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b1, b2, b5, h1, and h6.
33393 | 1220648-2023-04341 | 17986844 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17986844&pc=OZD | 07/04/2023 | 10/23/2023 |  | 10/23/2023 |  |  | 07/04/2023 | 10/23/2023 |  |  | 04/30/2025 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024336766 |  | 1000100 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/04/2023 |  |  |  |  | The user facility reported a 76 year-old female was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that upon initiation of first impella 5.5, flow saturation was present. the pump was replaced without issue. there was no patient harm reported. | [5 empty] | The investigation for the reported flow saturation issue has been completed. the data logs were reviewed finding saturated flow following motor current spikes. additionally, at the same time the ¿impella position in aorta¿ alarm is occurring. the root cause is most likely biomaterial that is being ingested during this time, resulting in the saturated flow. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33394 | 1220648-2023-04342 | 17986891 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17986891&pc=OZD | 07/06/2023 | 10/23/2023 |  | 10/23/2023 |  |  | 07/06/2023 | 10/23/2023 |  |  | 05/31/2025 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024355533 |  | 1000100 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/06/2023 |  |  |  |  | The user facility reported 68yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the pump had come out of left ventricle position when attempted repositioning while in the operating suite. the pump could not recross. the pump was removed. there was no patient harm reported. | [5 empty] | The investigation for the positioning issue has been completed. a review of the clinical data confirmed pump was in aortic area (about 1 hour & 35 minutes into the case). pump then was stopped manually & disconnected from the console. product was not returned for review. based on clinical description & data analysis, the root cause of positioning issue was due to attempt to re-access the valve wirelessly. the wireless attempt goes against ifu recommendations. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33395 | 1220648-2023-04343 | 17986926 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17986926&pc=OZD | 02/14/2023 | 10/23/2023 |  | 10/23/2023 |  |  | 02/14/2023 | 10/23/2023 |  |  | 12/31/2024 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024292217 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/14/2023 |  |  |  |  | The user facility reported a 75-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported difficulty delivering the device due to significant resistance met while manipulating the device through the sterile sleeve. the pump was removed and replaced with no harm to the patient. | [5 empty] | Returned product had exceedingly tight repositioning unit, making it difficult to advance the unit along the catheter. additionally, a notable catheter kink was found at the 26-cm mark. repo unit was soaked to ensure blood was not contributing to issue, however the repo unit remained difficult to advance. the cath anchor was able to translate relative to the repo unit with no movement of the unit itself when attempting to advance. in conclusion, it was determined that the cause of the positioning issues was insufficient repo unit id. this report is being filed as part of a retrospective review of historical records.
33396 | 1220648-2023-04344 | 17987029 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17987029&pc=OZD | 06/07/2023 | 10/23/2023 |  | 10/23/2023 |  |  | 06/07/2023 | 10/23/2023 |  |  | 01/31/2025 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024294316 |  | 1000100 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/07/2023 |  |  |  |  | The user facility reported a 35-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella pump was inadvertently pulled out of the ventricle when the the patient went into a delirious state. as the patient's hemodynamics were improving, the decision was made to explant the pump. there was no patient harm resulting from the event. | [5 empty] | Root cause of the positioning issue was patient movement because the patient was in delirious state, where the patient pulled the catheter out of his ventricle completely. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. this report is being filed as part of a retrospective review of historical records.
33397 | 1220648-2023-04346 | 17987072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17987072&pc=OZD | 09/27/2023 | 10/31/2023 |  | 10/23/2023 |  |  | 09/27/2023 | 10/23/2023 |  |  | 03/31/2025 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024324376 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/27/2023 |  |  |  |  | The user facility reported a 46-year-old female with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the impella was started while 0.018 wire was still in place. staff was unable to restart the pump after multiple attempts and alternative controller. a new pump was used successfully. no harm to patient reported. | [5 empty] | The investigation into the pump did not start issue has been completed since the original report was filed. product was returned for evaluation. the cause of the pump did not start issue was use related with pump started several attempts with guidewire in place. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
33398 | 1220648-2023-04345 | 17987079 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17987079&pc=OZD | 09/26/2023 | 09/24/2024 |  | 10/23/2023 |  |  | 09/26/2023 | 10/23/2023 |  |  | 01/31/2025 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024317512 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/26/2023 |  |  |  |  | The japanese user facility reported a patient admitted in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the impella had a purge leak resulting in a persistent increase in purge flow rate and decrease in purge pressure. the impella was exchanged as there was concern that the pump would stop in the future. the new impella is operating without any issues. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." | The investigation for the purge leak has been completed. according to the clinical, it was observed that the purge flow rate was 30ml/hr and the purge pressure was 30mmhg. all purge lines were checked for leaks, and purge fluid was observed leaking from the base of the catheter plug and purge filter connection. the decision was made to replace the pump and no other issues occurred afterwards. no product was returned for an investigation. no data logs were returned for analysis. the most likely cause of the pump purge leak was damage to the filter-tubing joint. added: h6 component code 925 corrections to the initial mfr report: added: d6a/d6b. f6/f8 should have been left blank.
33399 | 1220648-2023-04347 | 17987092 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17987092&pc=OZD | 04/14/2022 | 10/23/2023 |  | 10/23/2023 |  |  | 04/14/2022 | 10/23/2023 |  |  | 01/31/2024 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023186077 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/14/2022 |  |  |  |  | The user facility reported a 71yo female was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that during insertion, the pump kinked and the device was removed and replaced. there was no patient harm reported. | [5 empty] | The investigation for the reported pump kink has been completed. the device was not returned for evaluation. however, clinical details provided are sufficient to determine a root cause. therefore, the root cause of the reported event was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
33400 | 1220648-2023-04348 | 17987132 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17987132&pc=OZD | 09/26/2023 | 01/18/2024 |  | 10/23/2023 |  |  | 09/26/2023 | 10/23/2023 |  |  | 05/31/2025 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024363567 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/26/2023 |  |  |  |  | The user facility reported a 75-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient had access site bleeding upon impella insertion. staff was unable to control the bleeding despite manual pressure and administration of protamine. the physician explanted the impella and replaced it with a new impella without any complication. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
33401 | 1220648-2023-04350 | 17987148 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17987148&pc=OZD | 09/23/2023 | 01/26/2024 |  | 10/23/2023 |  |  | 09/23/2023 | 10/23/2023 |  |  | 11/30/2024 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024297128 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/23/2023 |  |  |  |  | The user facility reported a 76-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced active access site bleeding. the patient was treated by re-dressing the groin, applying manual pressure, and one unit of red blood cells. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for investigation. per the provided clinical information, the patient was bleeding at access site around the sheath after gwru insertion with best practices. patient had calcified vessel condition. the root cause of the access site bleeding issue was most likely patient condition related due to patient having calcified vessels per clinical review.
33402 | 1220648-2023-04351 | 17987198 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17987198&pc=OZD | 05/02/2022 | 10/23/2023 |  | 10/23/2023 |  |  | 05/02/2022 | 10/23/2023 |  |  | 02/29/2024 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023188891 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/02/2022 |  |  |  |  | The user facility reported a 53 year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was a purge leak noted during a cassette change. a purge bypass was completed but the needle poked the clear tubing. there was no patient harm reported.. | [5 empty] | The device was returned and was visually inspected finding cracks in the yellow luer. the tubing between the filter and red handle was detached with a purge bypass connected. the clinical review found that the purge leak was noted during a cassette change. however, a needle poked the clear tubing during the purge bypass. it was further noted that the leak began when the patient was being turned and the leak occurred from the connection between the yellow luer to the reservoir. bicarbonate was used as the purge solution. the cause of the yellow luer purge leak was patient movement. this report is being filed as part of a retrospective review of historical records.
33403 | 1220648-2023-04353 | 17987207 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17987207&pc=PYX | 09/23/2023 | 02/02/2024 |  | 10/23/2023 |  |  | 09/23/2023 | 10/23/2023 |  |  | 01/31/2025 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024321259 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/23/2023 |  |  |  |  | The user facility reported a 64-year-old male with cardiomyopathy was implanted with an impella rp for mechanical circulatory support. the patient was also implanted with an impella 5.5 on the left side. the rp had low flows despite anticoagulation management. the rp was removed and replaced. the 2nd pump had same low flows. no lasting harm or injury but due to poor preload for the left sided 5.5 impella pump, the pump was explanted and patient placed on ecmo support instead. | [5 empty] | The investigation into the reported low pump flow has been completed since the original report was filed. upon investigation of the pump a clot was observed in the inflow cage and around the impeller. data logs confirm low pump flows and show a motor current spike and the placement signal trending upwards. the root cause of the low pump flow issue was determined to be biomaterial. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
33404 | 1220648-2023-04352 | 17987210 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17987210&pc=OZD | 05/09/2022 | 10/23/2023 |  | 10/23/2023 |  |  | 05/09/2022 | 10/23/2023 |  |  | 01/31/2024 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023186130 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/09/2022 |  |  |  |  | The user facility reported a 48 year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the pump was seen to have a purge leak on day 4 of support. the sidearm leak did not prompt pump replacement. the pump remains on for support and there is no harm or injury. | [5 empty] | The investigation into the sidearm leakage has been completed. the device for returned for evaluation. leak testing was performed and the leaking was reproduced. leaking could be seen from the broken filter. the root cause of the purge leak is damaged sidearm filter. this report is being filed as part of a retrospective review of historical records.
33405 | 1220648-2023-04349 | 17987222 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17987222&pc=OZD | 04/19/2022 | 10/23/2023 |  | 10/23/2023 |  |  | 04/19/2022 | 10/23/2023 |  |  | 02/29/2024 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023188861 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/20/2022 |  |  |  |  | The user facility reported a 77yo female was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there were flow and unreliable signal issues. the team made choice to explant the 5.5 and replace pump with a cp. there was no patient harm reported. | [5 empty] | The investigation for the flow saturation issues has been completed. data logs showed erratic, non-pulsatile motor current and high, non-pulsatile pump flow waveforms upon initial activation of the pump. there were also recurrent "impella position in aorta" alarms. the "impella stopped - motor current high" alarm triggered at the end of the case. in conclusion, given the nature of these metrics, it was determined that the cause of the saturated flow was most likely ingested biomaterial. this report is being filed as part of a retrospective review of historical records.
33406 | 1220648-2023-04354 | 17987225 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17987225&pc=OZD | 05/13/2022 | 10/23/2023 |  | 10/23/2023 |  |  | 05/13/2022 | 10/23/2023 |  |  | 02/29/2024 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023186188 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/13/2022 |  |  |  |  | The user facility reported 38yo male was implanted with an impella 5.5 device for mechanical circulatory support as a bridge to transplant. on day 19, high purge pressure developed and eventually a pump stop occurred. tissue plasminogen activator (tpa) in purge-purge system blocked alarm appeared. the pump was explanted and the field reported buildup on the pump. there was no patient harm reported. | [5 empty] | The investigation for the reported pump stop has been completed. the pump and purge cassette were returned for evaluation. an extreme kink was visualized on the catheter located on the connection to the red handle. a light test was performed and displayed fiber damage at the kink. a combination of biomaterial ingestion and the catheter kink contributed to the high purge pressure and pump stop. instructions for use for the related event are as follows: ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33407 | 1220648-2023-04355 | 17987272 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17987272&pc=OZD | 06/15/2023 | 10/23/2023 |  | 10/23/2023 |  |  | 06/15/2023 | 10/23/2023 |  |  | 04/30/2025 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024336867 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 21 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/15/2023 |  |  |  |  | The user facility reported a 21-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that upon attempting to re-advance the impella catheter following bypass cannulation, the catheter kinked and could not be repositioned. the pump was subsequently removed as a result of the noted issue. there was no patient harm. | [5 empty] | Pump was returned with catheter was cut off. visual inspection found kinks on catheter below the end cap. no other issue were found on the rest of the pump. the root cause of positioning issue was most likely due to use issue. this report is being filed as part of a retrospective review of historical records.
33408 | 1220648-2023-04356 | 17987461 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17987461&pc=OZD | 06/21/2023 | 01/24/2024 |  | 10/23/2023 |  |  | 06/21/2023 | 10/23/2023 |  |  | 01/31/2025 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024296581 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility had planned a minimally invasive aortic valve replacement (avr) for bicuspid atrioventricular (av) stenosis. after aortic valve replacement (avr)/mitral valve replacement (mvr) patient could not come off bypass. a sternotomy was performed, coronary artery bypass graft (cabg)x2 after vessel perforation and direct impella placement. angulation of impella catheter across aortic valve causing aortic insufficiency (ai) as leaflet tented open, the impella appeared against papillary. reposition attempted but surgeon concerned about ai and pulled back across valve. many more attempts to come off bypass with impella not across valve. then decision to place right sided pump, many difficulties with that insertion and eventual aborted. decided to use protek for va ecmo, the impella was aborted and graft sewn down. patient had 9.5 hours of pump time. the patient was sent to unit with open chest. | [5 empty] | The investigation into positioning issues has been completed. the impella product was not returned for investigation. as per clinical observations, the doctor was unable to cross the bicuspid aortic valve when repositioning the pump after it had been found too deeply into the ventricle. the cause of the positioning issue was most likely patient condition--diseased valve. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
33409 | 1220648-2023-04357 | 17987613 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17987613&pc=OZD | 06/22/2023 | 09/23/2025 |  | 10/23/2023 |  |  | 06/22/2023 | 10/23/2023 |  |  | 05/31/2025 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024344610 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that a positioning signal not reliable alarm occurred in a patient on impella support while the patient was laying flat. there was no harm to the patient as a result of this incident. | [5 empty] | The investigation into the reported "positioning issues" is still open at this time.the medical center did not return any impella products for benchtop analysis. upon completion of the investigation a final report will be forthcoming. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation into the reported "positioning issues" has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the data logs were returned and reviewed. per the clinical information provided and data log review, the root cause of the optical signal issue was not determined since product was not returned and insufficient clinical information was provided. the issue was either fiber break or an air gap. | H6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2023-04357. b7 other relevant history, including preexisting medical conditions updated. d6a updated. d10 concomitant medical products and therapy dates updated. g1 reporting contact email and reporting contact fax number updated.
33410 | 1220648-2023-04358 | 17987657 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17987657&pc=OZD | 07/04/2023 | 10/23/2023 |  | 10/23/2023 |  |  | 07/04/2023 | 10/23/2023 |  |  | 12/31/2024 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024315779 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/04/2023 |  |  |  |  | The user facility reported a 85-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was doing well and plans were to remove impella the next day. the patient had accidentally pulled the impella out of the eft ventricle¿(lv) into the aorta causing the automated impella controller (aic) to alarm the impella in the aorta. patient was stable with no changes in hemodynamics after this happened; physician removde the impella since they were planning on removing it the following day. patient did fine after removal and site was surgically closed achieving hemostasis. no known patient harm or injury. | [5 empty] | The medical center returned the pump for investigation into the pump's malpositioning. the anchor button functioned with no issues. in conclusion, it was determined that the cause of the positioning issues was the separation of the blue butterfly from the repo unit due to patient movement. this report is being filed as part of a retrospective review of historical records.
33411 | 1220648-2023-04359 | 17987727 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17987727&pc=OZD | 11/02/2022 | 03/05/2025 |  | 10/23/2023 |  |  | 11/02/2022 | 10/23/2023 |  |  | 06/30/2024 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023234462 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | Tachycardia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/02/2022 |  |  |  |  | The user facility reported a 48yo female with a st elevated mycoardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. the patient was in poor condition with runs of ventricular tachycardia (vt) and placed an intra-aortic balloon pump (iabp) prior to the 5.5. the 5.5 was not showing optimal placement signal and the optical sensor was questioned for damage. additionally the pump was not in a good position and leading to more vt. the team explanted and placed a new pump. there was no patient harm reported. | [5 empty] | The investigation for the reported positioning/optical signal issues and vt/vf is on-going at this time. no product was returned for investigation. upon completion of the investigation a final report will be forthcoming. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. based on the provided clinical information the cause of the positioning issue was most likely patient condition related. the first rt log (before cpr event and before pump is inserted into the body) shows all of the 5s parameters are within the normal ranges. the second rt log shows the 5s parameters remain within the normal range and the placement signal has some variations. during the time in which there are variations in the placement signal the cpr and repositioning of the pump both likely occur. no placement signal not reliable alarms are seen in the logs. this report is being filed as part of a retrospective review of historical records.
33412 | 1220648-2023-04360 | 17987841 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17987841&pc=OZD | 04/26/2022 | 10/23/2023 |  | 10/23/2023 |  |  | 04/26/2022 | 10/23/2023 |  |  | 12/31/2023 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023177906 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/26/2022 |  |  |  |  | The user facility reported that a 57-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced sudden low pump flows on the 8th day of support. the team attempted troubleshooting, and did not resolve the issue, so the console was swapped out and the issue was resolved. there was no reported harm to the patient. | [5 empty] | The investigation into the reported low pump flows was completed. the device was not returned for evaluation, but data logs were received. the console was analyzed, and no issues were found. the data logs were analyzed and show a motor current spike followed by low flows which is characteristic of ingested biomaterial. the root cause of the low pump flow was most likely ingested biomaterial. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33413 | 1220648-2023-04361 | 17987886 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17987886&pc=OZD | 04/21/2023 | 10/23/2023 |  | 10/23/2023 |  |  | 04/21/2023 | 10/23/2023 |  |  | 02/28/2025 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 |  |  | ABIOMED, INC. |  |  | US |  | 2024310558 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 26 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/21/2023 |  |  |  |  | The user facility reported that a 26-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced low pump flows after one month of support. the physician questioned the level of support the patient was being given. no further information was provided. there was no reported harm to the patient. | [5 empty] | The investigation into the reported low pump flows was completed. the pump and data stick were returned for evaluation. analysis of the data logs revealed alarms and low flows related to unsuitable pump position and suction event, which were likely related to ingested biomaterial. visual inspection of the pump revealed biomaterial on the impeller. functional testing of the pump failed until removal of the biomaterial. based on the available information, the root cause of the reported event was most likely due to ingested biomaterial. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33414 | 1220648-2023-04362 | 17988036 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17988036&pc=OZD | 09/24/2023 | 10/31/2023 |  | 10/23/2023 |  |  | 09/24/2023 | 10/23/2023 |  |  | 07/31/2025 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024372375 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | Material Deformation | Positioning Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/24/2023 |  |  |  |  | The user facility reported a 76-year-old male undergoing coronary artery bypass grafting was implanted with an impella 5.5 device for mechanical circulatory support. during support, the impella 5.5 had suction alarms that were likely due to right ventricle failure. impella was at p 4. additionally, the pump stopped due to the patient forcibly pulling the impella out. proximal portion of impella was pulled such that catheter stretched and centimeter markers were not visible. suture hub was still secured to skin and not bleeding. pump was removed. | [5 empty] | The investigation into the pump stop issue and the low pump flow issue has been completed since the original report was filed. product was returned for evaluation. the cause of the pump stop was an open electrical line due to excessive force applied by patient pulling the catheter and stretched, twisted per clinical, field investigation and log review. the cause of the low pump flow issue was patient condition related with persistent suction alarms at p-level higher than p-4 due to poor patient condition of premature ventricular contractions (pvc) and right ventricular failure (rvf), per clinical and log review. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
33415 | 1220648-2023-04363 | 17988205 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17988205&pc=OZD | 05/03/2023 | 10/23/2023 |  | 10/23/2023 |  |  | 05/03/2023 | 10/23/2023 |  |  | 03/31/2025 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED INC. |  |  | US |  | 2024323486 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Device Handling Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/03/2023 |  |  |  |  | The user facility in the eu reported a patient undergoing cardiac surgery was implanted with an impella 5.5 pump for mechanical circulatory support. bleeding was noted out of the red plug. bleeding did not require any intervention and pump was explanted after 13 days. | [5 empty] | The impella 5.5 pump was returned for investigation. a hole in the catheter was found near the 30cm marking on the catheter. blood was confirmed in the red handle. the cause of the product damage was blood ingress due to a hole in the catheter resulting from improper use. this report is being filed as part of a retrospective review of historical records.
33416 | 1220648-2023-04364 | 17988226 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17988226&pc=OZD | 06/10/2023 | 10/23/2023 |  | 10/23/2023 |  |  | 06/10/2023 | 10/23/2023 |  |  | 04/30/2025 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024338801 |  | 10000100 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [5 empty] | Male |  | Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/10/2023 |  |  |  |  | The user facility reported that a male patient implanted with the impella 5.5 for mechanical circulatory support experienced thrombus. there was clot noted on the inlet and outlet observed when the device was explanted. no further information was provided. there was no reported harm to the patient. the pump was explanted after 6 days of support. | [5 empty] | The investigation into the reported thrombus was completed. the device was not returned for evaluation, but images provided by the site confirmed the thrombus. analysis of data logs did not reveal any mc abnormality. based on the available information, the root cause of the thrombus issue was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33417 | 1220648-2023-04365 | 17988328 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17988328&pc=OZD | 07/05/2023 | 10/23/2023 |  | 10/23/2023 |  |  | 07/05/2023 | 10/23/2023 |  |  | 06/30/2024 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023247697 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/05/2023 |  |  |  |  | The user facility reported that a 62-year-old male implanted with the impella 5.5 for mechanical circulatory experienced three pump stops during the 50+ day support. during the pump replacement, staff noticed that the catheter appeared to be kinked. no further information was provided. there was no patient harm reported. | [5 empty] | The investigation into the pump stop was completed. the device was not returned for evaluation. based on the available information, the root cause of the pump stop issue was determined to be use-related due to the reported catheter kink near the kink protector. should additional information be received, this investigation will be updated accordingly. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33418 | 1220648-2023-04366 | 17988437 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17988437&pc=OZD | 12/30/2021 | 09/24/2024 |  | 10/23/2023 |  |  | 12/30/2021 | 10/23/2023 |  |  | 07/31/2023 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022145908 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Defective Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/30/2021 |  |  |  |  | The user facility reported a 44-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the red impella plug had bleeding upon insertion, bleeding had been resolved, and there was no mention of patient harm or injury requiring medical intervention. | It was reported heparin was withheld for the bleeding from the red impella plug as the patients prothrombin time/inr was high due to multi-organ failure. it was noted that the patient's care was withdrawn and the patient expired while on impella support. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The root cause of the blood leaking from the red plug could not be determined as product was not returned. the cause of the pump damage could not be determined. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-04366 was provided in sections b1, b2, b5, d6a, d6b, h1, and h6.
33419 | 1220648-2023-04367 | 17988966 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17988966&pc=OZD | 08/07/2021 | 10/23/2023 |  | 10/23/2023 |  |  | 08/07/2021 | 10/23/2023 |  |  | 05/31/2023 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022088357 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Crack | Fluid/Blood Leak | Decrease in Pressure |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/07/2021 |  |  |  |  | The user facility reported that a 65-year-old male patient was implanted with an impella 5.5 device for mechanical circulatory support. it was noticed during support that there was a crack in the yellow purge luer and that purge pressure (pp) was low. staff attempted to troubleshoot with possible purge bypass, but the outcome was not provided. bicarb was utilized as the purge solution. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge system leak was completed. the device was not returned for evaluation. based on the clinical report it can be confirmed that the cause of the purge system leak was due to damage of the purge luer with the use of sodium bicarbonate. per the ifu: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33420 | 1220648-2023-04368 | 17989619 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17989619&pc=OZD | 03/07/2022 | 10/23/2023 |  | 10/23/2023 |  |  | 03/07/2022 | 10/23/2023 |  |  | 10/31/2023 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022153090 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/07/2022 |  |  |  |  | The user facility reported a 68-year-old male bridging to a heart transplant was implanted with an impella device for mechanical circulatory support. the complainant reported observance of what appeared to be air bubbles after pump initiation. this anomaly resolved quickly and support was not interrupted. there was no harm to the patient because of this. the patient ultimately expired on support, however, there was no allegation of the impella being a factor. | [5 empty] | The cause of the presumed air embolism was not determined since product was not returned. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
33421 | 1220648-2023-04369 | 17989868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17989868&pc=OZD | 03/15/2023 | 10/23/2023 |  | 10/23/2023 |  |  | 03/15/2023 | 10/23/2023 |  |  | 08/31/2024 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023276920 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/15/2023 |  |  |  |  | The user facility reported a 72yo male was implanted with an impella 5.5 device for mechanical circulatory support. the medical team noted the anatomy/tortuousity of the patient presented challenge to pump delivery. it was reported that the pump could not be removed with direct access to subclavian. a partial opening of the chest was needed to remove the pump. there was no further harm reported. the removal took place after 4 days of successful support. | [5 empty] | The investigation for the reported removal issues has been completed. the clinical report provided sufficient details to determine the root cause. the root cause of the removal issue was determined to be patient condition as the clinical details specify that the patient's difficult anatomy caused additional resistance. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33422 | 1220648-2023-04370 | 17990064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17990064&pc=OZD | 08/30/2021 | 10/23/2023 |  | 10/23/2023 |  |  | 08/30/2021 | 10/23/2023 |  |  | 06/30/2023 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022102513 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/30/2021 |  |  |  |  | The user facility reported that a 64-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced low pump flow. shortly after starting pump and while ramping up the p-level, the impella position alarm sounded despite proper positioning as verified by echo and fluoroscopy. p-6 showed a flow of 4.5l with a maximum and minimum of both 4.5l and dampened motor current. the patient was taken for pump exchange due to concern pump failure. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported low pump flow was completed. the device was returned for evaluation. analysis of the data logs revealed a high motor current during ramp up and a motor current and pump flow shift upward while still on p4. impella in aorta alarms triggered after that point. since the staff confirmed the position via echo, the false position alarm was likely due to an interaction with the patient. additionally, the returned device was visually inspected, and a cannula kink was observed. no biomaterial was observed. the most likely root cause of the low pump flow or suction or non-pulsatile motor current was due to a kinked cannula. this report is being filed as part of a retrospective review of historical records.
33423 | 1220648-2023-04371 | 17990256 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17990256&pc=OZD | 08/06/2021 | 10/23/2023 |  | 10/23/2023 |  |  | 08/06/2021 | 10/23/2023 |  |  | 05/31/2023 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022085192 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Other | Health Professional | User facility | Company representation |  |  | 49 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | Yes | Initial submission |  |  |  |  | 08/06/2021 |  |  |  |  | The user facility reported a 49-year-old male patient was implanted with an impella 5.5 pump for mechanical circulatory support. the staff noted that purge flow rates were twice as high upon the patient's return\ from the operating room (or). the impella 5.5 purge reservoir/filter was found to be cracked and leaking" from the purge filter area inside the purge protector. flows were within normal limits of 2-30 and no alarms were triggered. purge bypass was discussed, however it was later noted that the patient was weaned and the pump was explanted. there was no harm reported to the patient. | [5 empty] | The investigation into the reported purge leak was completed. the device was not returned for evaluation. based on the conclusion of the investigation, the root cause of the purge system leak was most likely damage to the purge sidearm. per the ifu: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
33424 | 1220648-2023-04372 | 17990553 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17990553&pc=OZD | 09/06/2021 | 10/23/2023 |  | 10/23/2023 |  |  | 09/06/2021 | 10/23/2023 |  |  | 06/30/2023 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022102507 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/06/2021 |  |  |  |  | The user facility reported that a 70-year-old male implanted with the impella 5.5 pump experienced high purge pressure. the representative reported that the high purge pressure alarm was delayed. the patient was treated with tpa for several hours. subsequent minimum and maximum flows were equal, and the pump was later replaced. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported high purge pressure was completed. analysis of the returned data logs confirmed purge system blocked alarms. functional testing of the pump showed the purge pressure to be above 1000 mmhg and purge flow <1 ml/hr. visual inspection of the catheter reveal blood ingress in the purge lumen. as the data logs from the beginning of the case were not returned and no major abnormalities were found with the pump, the root cause of the high purge pressure was undetermined. this report is being filed as part of a retrospective review of historical records.
33425 | 1220648-2023-04373 | 17990931 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17990931&pc=OZD | 08/11/2021 | 10/23/2023 |  | 10/23/2023 |  |  | 08/11/2021 | 10/23/2023 |  |  | 03/31/2023 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022063763 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Other | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | Yes | Initial submission |  |  |  |  | 08/11/2021 |  |  |  |  | The user facility reported a 61-year-old male patient was implanted with an impella 5.5 pump for mechanical circulatory support. a "purge pressure low" alarm was triggered during support and the purge side arm was noted to be leaking. a purge bypass was recommended and instructions were sent. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge system leak was completed. the device was not returned for evaluation. based on the clinical report it was concluded that the cause of the purge system leak was due to damage of the purge luer with the use of sodium bicarbonate. the root cause of the purge leak was sidearm yellow luer damage. this report is being filed as part of a retrospective review of historical records.
33426 | 1220648-2023-04374 | 17991004 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17991004&pc=OZD | 09/10/2021 | 10/23/2023 |  | 10/23/2023 |  |  | 09/10/2021 | 10/23/2023 |  |  | 06/30/2023 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022109092 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Break | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/10/2021 |  |  |  |  | The user facility reported that a 74-year-old female patient implanted with the impella 5.5 for mechanical circulatory support experienced a purge sidearm break after the scrub technician attempted to place it on the purge cover. the physician completed a purge bypass with an 18-gauge angiocath. the alarms immediately stopped, and no further issues were noted. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge sidearm break/ purge leak was completed. the device was not returned for evaluation. based on the available information, the root cause of the purge sidearm break/ purge leak was sidearm bond damage due to the user. this report is being filed as part of a retrospective review of historical records.
33427 | 1220648-2023-04375 | 17991130 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17991130&pc=OZD | 06/02/2023 | 02/28/2024 |  | 10/23/2023 |  |  | 06/02/2023 | 10/23/2023 |  |  | 12/31/2024 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 |  |  | ABIOMED, INC. |  |  | US |  | 2024304183 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/02/2023 |  |  |  |  | The user facility reported 66-year-old male implanted with an impella 5.5 device for mechanical circulatory support experienced high purge pressure and a drop in purge flows. there were no kinks noted and the pump was adequately primed with sodium bicarbonate. the purge tubing was subsequently exchanged and primed with 2 mg of alteplase in 50 ml of sterile water and tissue plasminogen activator (tpa) was initiated. after a few hours of tpa, purge flows increased to 3.5 and purge pressure decreased to 713. no further information was provided. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported high purge pressure has been completed since the original report was filed. the pump was not returned for investigation. data logs show a 7-minute rise in purge pressure for 5 hours. purge cassette replacement results were unsuccessful. the purge pressure rise gradually decreased and returned to the normal range after infusing tissue plasminogen activator (tpa). the root cause of the high purge pressure issue was most likely biomaterial since the purge pressure returned to the normal range after the infusion of tissue plasminogen activator. | The device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. per the ifu: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ this report is being filed as part of a retrospective review of historical records.
33428 | 1220648-2023-04376 | 17991289 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17991289&pc=OZD | 05/19/2023 | 10/23/2023 |  | 10/23/2023 |  |  | 05/19/2023 | 10/23/2023 |  |  | 02/28/2025 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2024307220 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/19/2023 |  |  |  |  | The user facility reported a 75-year-old male patient was implanted with an impella 5.5 device for mechanical circulatory support. on the second day of support, the pump had 2 temporary, self resolving, pump stops on after a notable increase in purge pressure (pp). the stops caused no harm or injury. the purge pressure (pp) was noted to prompt discussions of the use of lytics in the purge to resolve the issue. no lasting harm or injury. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the pump stop was not determined since product/data was not returned. the cause of the high purge pressure was not determined since product/data was not returned. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
33429 | 1220648-2023-04377 | 17991541 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17991541&pc=OZD | 04/27/2023 | 01/24/2024 |  | 10/23/2023 |  |  | 04/27/2023 | 10/23/2023 |  |  | 12/31/2024 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMART ASSIST SET, EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2024308271 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 44 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/27/2023 |  |  |  |  | The user facility reported a 44-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. no unloading from pump; suspicion of thrombus; patient on ecmella therapy. patient stable. | [5 empty] | The investigation into the low pump flow has been completed since the original report was filed. product was returned for evaluation. there was a kink found near the inflow cage of the pump. the root cause of the low pump flow issue was a kinked cannula. the failure mode will be monitored and trended. | The impella device was received from the customer and the investigation is ongoing. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
33430 | 1220648-2023-04378 | 17992729 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17992729&pc=OZD | 09/24/2023 | 09/29/2025 |  | 10/23/2023 |  |  |  | 10/23/2023 |  |  | 05/31/2025 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024345523 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | Device Sensing Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that there were persistent suction alarms and placement signal after performing a purge cassette/bag change for the impella 5.5 support. the physician performed echo at bedside showing pump to be shallow at 3.8 from aorta valve annulus to inlet area. the patient¿s left ventricle is small. to troubleshoot the surgeon slightly advanced impella approximate 1-2 cm. the physician adjusted placement signal which stopped alarms. it was further reported that the patient experienced gi bleeding prompting a colonoscopy and clipping x3. gi removed all sources of anticoagulation due to melena and low hemoglobin. a total of 19 units of prbc, 2 ffp¿s, and 1 platelet were given for the gi and nasal bleeding. the patient only remains on antiplatelet therapy at this time and is hemodynamically stable. the impella was removed the following day due to suction and placement signal issues. | [5 empty] | The investigation into the optical signal issue and the vascular damage - peripheral vessel issue has been completed since the original report was submitted. the device was not returned for investigation. due to limited clinical details, the root cause of the vascular damage - peripheral vessels cannot be determined. the root cause of the optical signal issue is most likely due to sensor wear. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the product was not returned for analysis. the data logs were evaluated and all 5s parameters were within specification. the placement signal can be seen drifting for 12 hours until it spikes to 3000. this pattern is indicative of sensor wear. the root cause of the optical signal issue is most likely due to sensor wear. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial. manufacturer device report number 1220648-2023-04378. b7 other relevant history, including preexisting medical conditions updated. g1 reporting contact email updated.
33431 | 1220648-2023-04379 | 17992733 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17992733&pc=OZD | 09/24/2023 | 12/12/2023 |  | 10/23/2023 |  |  |  | 10/23/2023 |  |  | 07/31/2025 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024387277 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 56-year-old female was implanted with an impella cp for mechanical circulatory support. however, an hour after implant, the decision was made to explant the pump due to there being no flow past sheath. ischemia concerns prompted the premature explant. | [5 empty] | The investigation into the patient's limb ischemia has been completed. no product was returned. as per clinical, patient was experiencing ischemia concerns with no flow past impella sheath and pump was explanted. the cause of the limb ischemia issue was most likely patient condition with diseased/small vessels related with ischemia confirmed only on impella leg per clinical. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33432 | 1220648-2023-04380 | 17992769 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17992769&pc=OZD | 09/23/2023 | 02/02/2024 |  | 10/23/2023 |  |  |  | 10/23/2023 |  |  | 08/31/2025 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024386744 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Alarm System | Biocompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/23/2023 |  |  |  |  | The user facility reported a 57-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the impella had an active aorta alarm and the physician decided to explant based on hemodynamic stability. anticoagulation was stopped roughly 45 minutes pre removal. the device was weaned off and removed. it was noted that the proximal aspect of cannula just distal to the outlet cage had some biomaterial wrapped around the circumference. | [5 empty] | The investigation for the reported removal issues has been completed. the impella device was returned for evaluation. upon visual inspection, there was a cannula kink observed around the impella. a thrombus was observed at the inlet cage. there were no decreases in flow observed during impella support and flows were within specification ranges. the root cause of the thrombus issue was patient condition related. the cannula kink may have been due to the impella removal process. d8/d9 the original manufacturer device report had no selected to the question, device returned to manufacturer. the device was returned at that time. yes to this question has been selected on this report. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
33433 | 1220648-2023-04381 | 17992774 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17992774&pc=OZD | 09/23/2023 | 02/02/2024 |  | 10/23/2023 |  |  | 09/23/2023 | 10/23/2023 |  |  | 05/31/2025 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024368927 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/23/2023 |  |  |  |  | The user facility reported a 70-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the physician experienced difficulty advancing catheter once already across the aortic valve. flows were not enough to support the patient. during repositioning to try to improve flows, a kink in the nitinol cannula was observed. when readvanced back into ventricle, kink went away but flows remained lower than needed. the impella was replaced with a new impella cp without issues. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow issue has been completed. the product was returned for review. data confirmed the low flow when pump started & remained to the rest of the case that triggered auto suction response and suction alarms. kinked cannula found 14 mm from of and cannula bonding site. pump ran without issue under bench test. the root cause of low flow was due to a kinked cannula.
33434 | 1220648-2023-04382 | 17992781 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17992781&pc=OZD | 09/24/2023 | 10/23/2023 |  | 10/23/2023 |  |  |  | 10/23/2023 |  |  | 07/31/2025 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024389258 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Inaccurate Delivery | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/24/2023 |  |  |  |  | The user facility reported an 82-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. upon implanting the impella, the repositioning sheath was in place, bleeding was seen from distal end of blue suture hub. upon investigation, it was found that the blue suture hub was loose around the white plastic portion near the wire re-access port. staff could not control the bleeding from this area, there was an attempt to push plastic portion into the blue suture hub, but it was unsuccessful. the decision was made to replace the impella. | [5 empty] | The investigation into the delivery issue and access site bleeding has been completed. product was returned and evaluated. per the clinical information provided, it seems that the blue butterfly was inserted into the patient which is against the impella ifu. the root cause of the access site bleeding issue was use related to improper technique. the failure modes will be monitored and trended. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33435 | 1220648-2023-04383 | 17992783 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17992783&pc=OZD | 09/25/2023 | 09/26/2025 |  | 10/23/2023 |  |  |  | 10/23/2023 |  |  |  |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1338986 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 09/25/2023 |  |  |  |  | The user facility reported a 65-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the nurse lost the placement signal on the automated impella controller (aic). the aic was exchanged with a different aic with no issues. there was no harm to the patient as a result of this incident. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | Corrections to the initial mfr report: d2-common device name updated d4-updated model number f6/f8 should have been left blank. g1 mdr reporting contact email h6 impact code 4629 added, component code changed to 765. | The investigation of automated impella controller (aic) - controller failure (red alarm) has been completed. aic - controller failure (red alarm): there was no evidence of ¿controller failure¿ in the log. it is most likely that the loss of the placement signal was misreported as ¿controller failure. aic - loss of opm data: the root cause of the controller error and loss of placement signal is due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic software operated as designed.
33436 | 1220648-2023-04384 | 17993177 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17993177&pc=OZD | 09/23/2023 | 08/02/2024 |  | 10/23/2023 |  |  | 09/23/2023 | 10/23/2023 |  |  | 07/31/2025 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024388010 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/23/2023 |  |  |  |  | It was further reported that the patient expired while on support. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that the use of the device cannot conclusively be associated with the outcome. | The user facility reported a 67-year-old male (race unknown) in cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was transferred to the tertiary center and, upon transfer, the team treated the patient for an access site bleed. the team gave multiple blood products. site bleed had a fem-stop applied as well. support was maintained with the implanted pump. |  |  |  |  | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- b5 mso review, h6 component code 925 in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33437 | 1220648-2023-04385 | 17993182 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17993182&pc=OZD | 06/05/2023 | 02/28/2024 |  | 10/23/2023 |  |  | 06/05/2023 | 10/23/2023 |  |  | 04/30/2025 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024330825 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/05/2023 |  |  |  |  | The user facility reported that a 45-year-old female patient implanted with the impella 5.5 for mechanical circulatory support experienced high purge pressure. the nurse attempted to troubleshoot per protocol but was unable to resolve the alarms. tissue plasminogen activator (tpa) protocol was subsequently initiated. no further information was provided. there was no reported harm to the patient. | [5 empty] | The investigation into the limb ischemia ¿ reversible has been completed. the device and data were not returned for investigation. the purge pressure rise gradually decreased and returned to the normal range after infusing tissue plasminogen activator (tpa). the likely cause of the high purge pressure issue was biomaterial, as the purge pressure returned to normal after infusing tissue plasminogen activator. | The device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33438 | 1220648-2023-04386 | 17993185 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17993185&pc=OZD | 06/02/2023 | 09/19/2025 |  | 10/23/2023 |  |  | 06/09/2023 | 10/23/2023 |  |  | 01/31/2025 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 |  |  | ABIOMED, INC. |  |  | US |  | 2024297939 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 06/09/2023 |  |  |  |  | The user facility reported 72-year-old male was implanted with an impella 5.5 device for mechanical circulatory support experienced high purge pressure issues. on the morning of the scheduled explant, the patient experienced high purge pressure and purge system blocked alarms. there was no kinking of the purge tubing, and the remainder of the catheter was free of obstruction. the alarm would clear when the yellow luer was disconnected from the purge system, leading the medical staff to suspect an obstruction somewhere downstream of the purge connection, within the catheter itself. the device was subsequently explanted. visual inspection of the pump did not reveal any obvious signs of clot at the outlet of the device. there was no reported harm to the patient. | The complainant reported an ongoing placement signal issue during impella support. it was noted that false impella in aorta alarms would sound, forcing the care team to reduce support to lower than p-4 to clear the alarm. each time they cleared the alarm, the staff would then return the support level to p-8. support continued this way until the morning of a scheduled explant, at which point a purge system blocked alarm sounded coinciding with a high purge pressure alarm. the staff moved forward with the scheduled explant following the noted alarm. |  |  |  |  | The investigation into the optical signal issue and the high purge pressure issues has been completed since the initial report was submitted. the product was not returned for investigation. based on the data log review, the root cause of the high purge pressure is most likely due to ingested biomaterial. the root cause of the optical signal issue is most likely due to the patient¿s condition. | The device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the product was not returned for investigation. the data logs indicate a small mc spike, a gradual rise in purge pressure with a decrease in purge flow. based on the data log review, the root cause of the high purge pressure is most likely due to ingested biomaterial. the optical parameters were reviewed. there was no "placement signal unreliable" alarm triggered. the gain and the light are within spectrum. the snr changes in accordance with p-level changes. the root cause of the optical signal issue is most likely due to the patients condition. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
33439 | 1220648-2023-04387 | 17993199 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17993199&pc=OZD | 09/05/2022 | 10/23/2023 |  | 10/23/2023 |  |  | 09/05/2022 | 10/23/2023 |  |  | 05/31/2024 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023217184 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Difficult to Advance | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/05/2022 |  |  |  |  | The user facility reported a 62-year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there were impella inlet and outlet in aorta alarms. the pump was in the aorta on p1. attempts were made to recross the valve, however a foreign body appeared to be obstructing the path on tee. the decision was made to explant the pump and, upon removal, clots were observed in the inflow and outflow. the pump was successfully replaced with a new impella 5.5 through the same graft. there was no harm done to the patient. | [5 empty] | The investigation for the reported positioning issues has been completed. the data log confirms impella position in aorta alarms. no product was returned for analysis. based on the provided clinical information, the cause of the positioning issue was most likely use related. this report is being filed as part of a retrospective review of historical records.
33440 | 1220648-2023-04388 | 17993269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17993269&pc=OZD | 04/20/2023 | 09/19/2025 |  | 10/23/2023 |  |  | 04/20/2023 | 10/23/2023 |  |  | 12/13/2024 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024311328 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/20/2023 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. roughly four days into support, a thrombectomy was required in the patient's right axillary artery. the pump was explanted the following day and a thrombus was observed at the inlet and outlet of the device. the patient was stable following the noted event. | The complainant also reported impella 5.5 issues with the positioning signal. impella position confirmed with echocardiography. p-level switches between p0 and p9. later, the physician repositioned the pump 3 cm deeper in the left ventricle, and after that, all waves and measurements were better and stable recommendations for transesophageal echocardiography were also given. |  |  |  |  | The returned pump was visually inspected and biomaterial was observed on the inlet and outlet. upon conclusion of the investigation, the cause of the thrombus in outflow on explant was patient condition related. this report is being filed as part of a retrospective review of historical records. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the returned pump was visually inspected, and biomaterial was observed on the inlet and outlet. a cannula kink was observed. the cause of the false position in aorta alarm was most likely patient condition related. software version unknown due to no data logs returned. all pertinent information available to abiomed has been submitted at this time. b1. product problem updated. b5. updated to include optical signal issue description. h6. updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2023-04388. d3, d3 manufacturer name, city and state corrected. d6a / d6b missing from initial report, updated. d10, concomitant medical products and therapy dates updated. g1 reporting contact email updated. f6 and f8 erroneously entered in initial report. h6 component code updated for pump code.
33441 | 1220648-2023-04389 | 17993283 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17993283&pc=OZD | 05/26/2023 | 02/16/2024 |  | 10/23/2023 |  |  | 05/26/2023 | 10/23/2023 |  |  | 03/31/2025 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 |  |  | ABIOMED, INC. |  |  | US |  | 2024321735 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/26/2023 |  |  |  |  | The user facility reported 64-year-old male patient was implanted with an impella 5.5 device for mechanical circulatory support. during support, it was noted that high purge pressure was encountered. the pump was subsequently replaced as a result of the noted issue. there was no reported harm to the patient. | [5 empty] | The device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. per the ifu: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the high purge pressure has been completed. the device was not returned for investigation. it was reported that the pump was explanted after reported several high purge pressures at the end of case. the cause of the high purge pressure issue was most likely biomaterial based on data log review.
33442 | 1220648-2023-04390 | 17993293 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17993293&pc=OZD | 05/15/2023 | 10/23/2023 |  | 10/23/2023 |  |  | 05/15/2023 | 10/23/2023 |  |  | 03/31/2025 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMART ASSIST SET, EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2024329251 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/15/2023 |  |  |  |  | The user facility reported a 58-year-old male was implanted with an impella 5.5 device for mechanical circulatory support. during support, a decrease in purge flow (pf) and increase in purge pressure (pp) was encountered. as a result of the noted issue, the pump was removed and another impella 5.5 pump was implanted for ongoing support. there was no patient harm as a result of the failure. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon conclusion of the investigation, the cause of the high purge pressure was not determined since neither product nor data logs were returned. per the ifu: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ this report is being filed as part of a retrospective review of historical records.
33443 | 1220648-2023-04391 | 17993306 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17993306&pc=OZD | 05/03/2023 | 10/23/2023 |  | 10/23/2023 |  |  | 05/03/2023 | 10/23/2023 |  |  | 01/31/2025 |  | 10/23/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2024307198 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/03/2023 |  |  |  |  | The user facility reported a 65-year-old male was implanted with an impella device for mechanical circulatory support. during support, it was noted that low purge flows and high purge pressure was encountered. the customer was advised on potential options to treat the issue or to wean the patient due to the potential of a pump stop. however, the staff opted not to do lysis or exchange the pump. support was maintained with the implanted pump with no noted harm to the patient. | [5 empty] | The pump was returned with the catheter cut. the pump was attached to the aic and purge cassette, and was purged easily though the length of catheter that remained. the catheter was then cut close to the motor housing to test the remaining portion of the catheter, and purge fluid was pushed through easily with a syringe. the blockage had been isolated to the motor. the cannula was then removed and biomaterial was observed in the purge gap. the cause of the high purge pressure was biomaterial. per the ifu: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ this report is being filed as part of a retrospective review of historical records.
33444 | 1220648-2023-04305 | 17982084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17982084&pc=OZD | 09/22/2023 | 02/12/2024 |  | 10/21/2023 |  |  | 09/22/2023 | 10/21/2023 |  |  | 08/31/2025 |  | 10/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024385363 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/22/2023 |  |  |  |  | The user facility reported a 79-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device and an impella 5.5 (bipella) for mechanical circulatory support. the patient experienced hemolysis with bipella support. the patient¿s prognosis was bleak and bipella was used as a last resort to manage cardiogenic shock. | [5 empty] | The investigation into the hemolysis issue has been completed. the device was not returned for investigation. based on the clinical information provided, the root cause of the hemolysis was most likely due to improper positioning as stated in the clinical review of the impella pigtail was wrapped on the pap muscle. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
33445 | 1220648-2023-04208 | 17982087 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17982087&pc=OZD | 07/03/2021 | 10/19/2023 |  | 10/21/2023 |  |  | 07/03/2021 | 10/21/2023 |  |  | 02/28/2023 |  | 10/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022054940 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/03/2021 |  |  |  |  | The user facility reported a 34-year-old male was implanted with an impella 5.5 as a bridge to heart and kidney transplant. the nurse reported increase in purge flow, reduced purge pressure and visible leaking from filter in y-connector. after 56 days of support the purge line cracked and was successfully troubleshot. the pump remained on for support. no lasting harm or injury. the pump was explanted after 72 days. | [5 empty] | The investigation into the reported purge leak has been completed. the medical center returned the product. the root cause of the increased purge flow and reduced purge pressure was the cracks in the y connector. the root cause of the cracks in the y connector was product malfunction. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33446 | 1220648-2023-04065 | 17982088 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17982088&pc=OZD | 10/04/2021 | 10/17/2023 |  | 10/21/2023 |  |  | 10/04/2021 | 10/21/2023 |  |  | 07/31/2023 |  | 10/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022112655 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/04/2021 |  |  |  |  | The user facility reported a 71-year-old male patient was implanted with an impella 5.5 device for mechanical circulatory support. during support, it was noted that a fluid leak developed from a crack on the yellow luer of the purge system. the cassette was changed to troubleshoot the issue, but the leak remained. a purge filter bypass was subsequently performed to temporarily resolve the noted leak and the pump was explanted later that day. there was no patient harm as a result of the failure. | [5 empty] | The device was returned and an evaluation was completed. functional testing of the device verified a fluid leak from the noted crack during use. based on the provided information, the root cause for the noted crack was determined to be the use of bicarbonate in the purge which weakened the plastic on the luer. should additional relevant information become available, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records.
33447 | 1220648-2023-04066 | 17982089 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17982089&pc=OZD | 10/28/2021 | 10/17/2023 |  | 10/21/2023 |  |  | 10/28/2021 | 10/21/2023 |  |  | 05/31/2023 |  | 10/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022091914 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/28/2021 |  |  |  |  | The user facility reported that a 38-year-old female patient implanted with the impella 5.5 for mechanical circulatory support began to experience lower flows than expected in the operating room (or). the flow rate gradually decreased from 4 l/min to 3.2 l/min with constant "suction" alarms. the pump was confirmed to be in the proper position via echo, and repositioning did not resolve the issue. the clinicians observed clots on and the cannula via echo, and clotting was observed in the inferior vena cava (ivc). ultimately, the decision was made to exchange the pump. no further information was provided. there was no reported harm to the patient. | [5 empty] | The investigation into the reported low pump flows was completed. the device was returned for evaluation. visual inspection of the pump revealed significant amounts of biomaterial on and around the impeller. no other abnormalities were noted. functional testing of the pump was able to reproduce the suction alarms and low flows. based on the available information, the most likely root cause of the low pump flows was ingested biomaterial. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33448 | 1220648-2023-04209 | 17982090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17982090&pc=OZD | 07/12/2021 | 10/19/2023 |  | 10/21/2023 |  |  | 07/12/2021 | 10/21/2023 |  |  | 03/31/2023 |  | 10/21/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022056267 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/12/2021 |  |  |  |  | The user facility reported a 64 year-old male was implanted with an impella 5.5 device for mechanical circulatory support for bridge to transplant. it was reported that after a purge cassette change, a leak was noted at the yellow luer lock or filter area. sodium bicarbonate was running in the purge. a purge bypass procedure was successfully completed. there was no patient harm reported. | [5 empty] | The investigation for the purge leak has been completed. neither the product nor the data logs were returned for investigation. however, the clinical information was sufficient in determining the root cause, given the results of prior failure investigations. it was concluded that the cause of the purge leak was sidearm yellow luer damage due to sodium bicarbonate use. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33449 | 1220648-2023-04236 | 17974752 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17974752&pc=OZD | 09/22/2023 | 06/03/2024 |  | 10/20/2023 |  |  | 09/22/2023 | 10/20/2023 |  |  | 06/30/2025 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024378385 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Pacing Inadequately | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/22/2023 |  |  |  |  | The user facility reported a 78-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the impella was inserted and advanced into the left ventricle. the wire was removed. the impella started at p-auto with flows of 3.5 lpm. shortly after the impella support was initiated, the patient went into complete heart block requiring temporary pacemaker placement. the patient was transported to icu with a temporary pacemaker. | [5 empty] | The investigation into the vf/vt/af/at (complete heart block) issues has been completed since the original report was submitted. the pump was not returned for investigation. as the necessary information was not provided, the root cause of the vt/vf was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
33450 | 1220648-2023-04237 | 17975230 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17975230&pc=OZD | 09/21/2023 | 01/15/2024 |  | 10/20/2023 |  |  | 09/21/2023 | 10/20/2023 |  |  | 03/31/2025 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024324572 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the impella 5.5 inlet cage completely separated from nitinol cannula during explant of device after a successful wean. inlet cage was able to be removed from conduit without further surgical intervention. no reported adverse effect to patient. the device was not saved for investigation. the support had been a successful 4 day support with the 5.5 after escalation from the impella cp. | [5 empty] | The impella device was not received from the customer as it was discarded. therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿avoid manual compression of the inlet and outlet areas of the cannula assembly.¿ | The investigation into the reported product damage (detached inflow/outflow - peripheral vessel) has been completed since the original report was filed. pump was not returned for investigation. according to clinical records, the inlet cage was detached from the cannula while removing the impella. the cause of the product damage-detached inlet cage issue was not determined since product was not returned and sufficient clinical information was not provided.
33451 | 1220648-2023-04238 | 17975258 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17975258&pc=OZD | 09/20/2023 | 01/17/2024 |  | 10/20/2023 |  |  | 09/21/2023 | 10/20/2023 |  |  | 07/31/2025 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024387961 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/21/2023 |  |  |  |  | The user facility reported a 78-year-old female who had just undergone cardiac surgery was implanted with an impella cp device for mechanical circulatory support. during support, the patient had limb ischemia. the patient had a very thick hyper dynamic ventricle that was confirmed on echo. fem fem bypass was added due to distal perfusion ischemia and was now perfusing well. it was reported 5 days later that the right leg had positive posterior tibialis pulse and was warm to touch, but the leg was mottled from knee to toe. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported limb ischemia has been completed. the device was not returned for evaluation. however, clinical information provided was sufficient to determine the root cause of the reported issue. as per clinical, limb ischemia observed on impella leg and fem-fem bypass done resolving ischemia. the cause of the limb ischemia issue was most likely patient condition with fem-fem bypass resolving ischemia based on clinical review.
33452 | 1220648-2023-04239 | 17975270 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17975270&pc=OZD | 09/21/2023 | 02/08/2024 |  | 10/20/2023 |  |  | 09/21/2023 | 10/20/2023 |  |  | 06/30/2025 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024373676 |  | 0048-0045 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that while attempting to reposition, the impella cp was pulled out of the ventricle. a decision was made to place a new impella. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve¿ | The investigation into the positioning issues has been completed. the device was not returned for investigation. data logs confirmed the failure mode. the logs showed about 8 hours & 50 minutes into support, pump position in aorta that triggered alarm impella position in aorta. pump then was disconnected from the console. based on clinical description and data analysis, the root cause of positioning issue was most likely due to use related.
33453 | 1220648-2023-04240 | 17975286 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17975286&pc=PYX | 09/22/2023 | 02/05/2024 |  | 10/20/2023 |  |  | 09/22/2023 | 10/20/2023 |  |  | 01/31/2025 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024305350 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/22/2023 |  |  |  |  | The user facility reported a 65-year-old male in non-st elevated myocardial infarction was implanted with an impella rp device for mechanical circulatory support. during support, the pump had reduced flows and there was an alarm stating purge system blockage indicating a likely large clot in the cannula. the rp was successfully replaced. | [5 empty] | The investigation into the reported low pump flow has been completed since the original report was filed. the impella was returned for evaluation. upon investigation of the pump large clots were found in the inflow and outflow cages. there was also a kink in the cannula and kinking damage to the catheter by the motor housing. data logs show a high motor current spike followed by low flow and suction alarms. placement signal trends upwards during the low pump flow. the root cause of the low pump flow was determined to be biomaterial. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
33454 | 1220648-2023-04241 | 17975442 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17975442&pc=OZD | 09/20/2023 | 10/20/2025 |  | 10/20/2023 |  |  | 09/20/2023 | 10/20/2023 |  |  |  |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1538036 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | Activation Failure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 09/20/2023 |  |  |  |  | The user facility in the eu reported the automated impella controller (aic) had problems during preparation. start procedure could not be completed. a back up aic was put into place. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
33455 | 1220648-2023-04242 | 17975492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17975492&pc=OZD | 02/14/2023 | 10/20/2023 |  | 10/20/2023 |  |  | 09/21/2023 | 10/20/2023 |  |  |  |  | 10/20/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, JP |  |  | ABIOMED, INC |  |  | US |  | 1316937 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Activation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 09/21/2023 |  |  |  |  | The user facility reported a 54-year-old male patient was implanted with an impella cp device for mechanical circulatory support. during support, it was noted that the utilized aic displayed a controller error alarm. the specified issue found during a daily inspection of the console was noted to be an optical sensor problem. however, upon a follow up evaluation of the returned device, it was discovered that the cause of the overarching issue was a defective pbm board. there was no noted harm resulting from the failure. | [5 empty] | The device was returned and an evaluation was completed. upon conclusion of the investigation, it was determined that the controller error alarm was caused by a loss of communication to the epm due to a loss of power from the pbm board. should additional relevant information become available, a supplemental report will be submitted.
33456 | 1220648-2023-04243 | 17975755 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17975755&pc=OZD | 09/25/2023 | 06/20/2024 |  | 10/20/2023 |  |  | 09/25/2023 | 10/20/2023 |  |  | 06/30/2025 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024376311 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  |  |  | Malfunction | No | Yes | Material Deformation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/25/2023 |  |  |  |  | The user facility reported a 79 -year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the 0.018¿ guidewire was malformed when it came out of the hoop. the guidewire had a large curve on it. the physician did not feel comfortable to use the wire with the patient. the guidewire was not used. no patient harm was reported. | [5 empty] | The investigation for the product damage has been completed. the impella device was not returned for evaluation. therefore, without sufficient clinical details, the root cause of the reported issues product damage could not be determined. added- d6a/d6b. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
33457 | 1220648-2023-04244 | 17975890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17975890&pc=OZD | 07/14/2022 | 10/20/2023 |  | 10/20/2023 |  |  | 07/14/2022 | 10/20/2023 |  |  | 03/31/2024 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023202602 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/14/2022 |  |  |  |  | The user facility reported a 72-year-old male patient was scheduled to be implanted with an impella 5.5 device for mechanical circulatory support. during setup, the aic displayed a placement signal not reliable alarm. the decision was subsequently made to not implant the noted pump and a second pump was then placed for support. no harm or injury from delay. | [5 empty] | The data logs were reviewed and the unreliable placement signal alarm was observed. the returned product was connected to the aic; no errors were observed. to confirm that the root cause was most likely misuse, the returned product was connected to the snr test bench and the observed snr graphs and 5s values were within spec. upon conclusion of the investigation, a definitive cause for the noted issue could not be established. this report is being filed as part of a retrospective review of historical records.
33458 | 1220648-2023-04245 | 17976175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17976175&pc=OZD | 09/20/2021 | 10/20/2023 |  | 10/20/2023 |  |  | 09/20/2021 | 10/20/2023 |  |  | 06/30/2023 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022100754 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/20/2021 |  |  |  |  | The user facility reported a 58-year-old male was implanted with an impella device for mechanical circulatory support. it was reported that the purge sidearm cracked despite the sidearm retainer being in place. the sidearm was cut off for a bypass as a result. bicarbonate was being used in the purge. no reports of adverse events were reported. | [5 empty] | The impella 5.5 was returned for analysis however the sidearm was not. the root cause of the purge system leak was determined to most likely be a crack in the yellow lure. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33459 | 1220648-2023-04246 | 17976203 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17976203&pc=OZD | 08/29/2021 | 10/20/2023 |  | 10/20/2023 |  |  | 08/29/2021 | 10/20/2023 |  |  | 03/31/2023 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022061915 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 55 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | Yes | Initial submission |  |  |  |  | 08/29/2021 |  |  |  |  | A medwatch was received that reported a 55-year-old male was implanted with an impella device as a bridge to transplant. it was reported that a leak was noted on the connector of the purge line purge cassette after performing recommended troubleshooting the purge pressure and flow reestablished. there was no patient harm reported as a result of the issue. | [5 empty] | The device was not returned for evaluation. clinical review was performed finding that bicarbonate was used in the purge. the bypass technique was used. the cause of the purge leak was sidearm yellow luer damage as a result of bicarb weakening the plastic of the yellow luer. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33460 | 1220648-2023-04247 | 17976261 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17976261&pc=OZD | 10/08/2021 | 10/20/2023 |  | 10/20/2023 |  |  | 10/08/2021 | 10/20/2023 |  |  | 06/30/2023 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022091893 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/08/2021 |  |  |  |  | The user facility reported a 82yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that during removal of a pump, the physician encountered resistance. a clot was noted in the inlet. additionally, the sterile sleeve was detached from the distal lock closest to the red impella plug. there was no patient harm reported and the pump remained on for support of over 234 hours. | [5 empty] | The investigation for the reported sleeve damage and pump thrombus has been completed. evaluation of the returned product confirmed that the sterile sleeve was detached from the distal lock. a kink in the cannula was noted proximal to the outflow cage. a clot was noted in the inlet. the data logs provided for reviewed and showed no high motor current alarms. multiple instances of 'preventing retrograde flow' alarms were noted. the root cause of the thrombus noted on explant was most likely due to patient condition. the root cause of the damage to the sterile sleeve was most likely excessive force as it was reported that the physician had difficulty with resistance during the pump's removal. the instructions for use provides details on how prevent thrombus formation in the pump. it states the following: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33461 | 1220648-2023-04248 | 17976290 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17976290&pc=OZD | 08/25/2023 | 02/06/2025 |  | 10/20/2023 |  |  | 09/21/2023 | 10/20/2023 |  |  |  |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1134032 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The us based hospital biomed department reported upon a broken purge disc on an automated impella controller (aic). the instructions for use were followed and aic was removed from service for a backup and repair. there was no patient use at the time and no harm or injury possible. | [5 empty] | The console was received for evaluation. however, the evaluation has not yet been completed. upon investigation completion, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. according to the device analysis the cause of the issue was a defective component. the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-04248. b3: date of event. d2: common device name. e2: should have been selected as no. g1: reporting contact email revised on manufacturer device report in accordance with updated procedures. g1: reporting contact fax number missing. h6: component code was missing.
33462 | 1220648-2023-04249 | 17976465 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17976465&pc=OZD | 09/20/2022 | 10/20/2023 |  | 10/20/2023 |  |  | 09/20/2022 | 10/20/2023 |  |  | 11/30/2023 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2023161218 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/20/2022 |  |  |  |  | The user facility reported a 66-year-old male patient was implanted with an impella device for mechanical circulatory support. during support, a fluid leak was identified at the yellow luer connection. a purge bypass was successfully completed to resolve the leak. no harm done to the patient. | [5 empty] | The product was returned and the leak was confirmed to be located at the yellow luer. a visual inspection identified a crack in the component. upon conclusion of the investigation, the root cause of the purge leak was determined to be a damaged yellow luer. per the ifu: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
33463 | 1220648-2023-04250 | 17976675 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17976675&pc=OZD | 10/21/2022 | 10/20/2023 |  | 10/20/2023 |  |  | 10/21/2022 | 10/20/2023 |  |  | 04/30/2024 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2023215895 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Advance | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/21/2022 |  |  |  |  | The user facility reported a 52-year-old male was implanted with an impella device for mechanical circulatory support. during attempted delivery, it was noted that difficulties were encountered as a result of the guidewire kinking/looping. the pump subsequently kinked at the junction between the motor and catheter shaft resulting in potential damage to the sensor. after a second unsuccessful attempt to deliver the pump, the decision was made to replace the device. the patient was considered to be stable following the event. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. after reviewing the clinical information, it was discussed that the cause of the delivery issues was potentially related to the damaged guidewire; however, there is insufficient evidence to ascertain a definitive cause. this report is being filed as part of a retrospective review of historical records.
33464 | 1220648-2023-04251 | 17976678 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17976678&pc=OZD | 10/13/2021 | 10/20/2023 |  | 10/20/2023 |  |  | 10/13/2021 | 10/20/2023 |  |  | 07/31/2023 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022113198 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 24 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/13/2021 |  |  |  |  | The user facility reported a 24yo male was implanted with an impella 5.5 device for mechanical circulatory support as a bridge to heart transplant. it was reported that during a difficult insertion, the pump kinked. the pump was pulled back and replaced and support continued to the transplant. there was no patient harm reported. | [5 empty] | The investigation for the kinked catheter has been completed. the product was returned and a kink in the catheter was confirmed. the root cause of the catheter kink was determined to be delivery issues due to patient condition. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33465 | 1220648-2023-04252 | 17976721 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17976721&pc=OZD | 01/24/2023 | 10/20/2023 |  | 10/20/2023 |  |  | 01/24/2023 | 10/20/2023 |  |  | 10/31/2024 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023268876 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a failed delivery of the impella 5.5 device past the anastomosis. the impella 5.5 insertion was aborted and an impella cp was used instead. there was no harm to the patient. | [5 empty] | The investigation into the reported delivery issues has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that there were numerous attempts made to pass impella 5.5 through anastomosis, all unsuccessful. catheter removed from sheath where kink became apparent. the physician elected to swap for new pump. the patient reportedly had tortuous vasculature and excessive force was noted to have been used. the root cause of the delivery issue was determined to be patient condition as the patient had tortuous vasculature.
33466 | 1220648-2023-04253 | 17976820 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17976820&pc=OZD | 10/17/2021 | 10/20/2023 |  | 10/20/2023 |  |  | 10/17/2021 | 10/20/2023 |  |  | 04/30/2023 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022080671 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Break | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/17/2021 |  |  |  |  | The user facility reported a 65 year-old female was implanted with an impella device for bridge to possible transplant. it was reported that the purge system cracked at the yellow connector while working on a purge bypass. there was no patient harm reported. the pump remained on for 3 days. | [5 empty] | The investigation into the reported crack in the yellow connector of the purge system has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. during review, it was found that sodium bicarbonate was used as a purge solution and the sidearm retainer was used along with the 3-point fixation technique. a review of the data logs found multiple purge pressure low and purge system open alarms. no further abnormal purge alarms were captured for the remainder of the case. the root cause of the low purge pressure was most likely a damaged yellow luer. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33467 | 1220648-2023-04254 | 17976840 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17976840&pc=OZD | 10/17/2021 | 10/20/2023 |  | 10/20/2023 |  |  | 10/17/2021 | 10/20/2023 |  |  | 07/31/2023 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022126631 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/17/2021 |  |  |  |  | The user facility reported a 67 year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the purge filter cracked. the physician performed a bypass. no reports of adverse events were reported. | [5 empty] | The device was returned for evaluation. during inspection, it was confirmed that a purge sidearm bypass was performed. a crack in the yellow luer of the purge sidearm was confirmed. the root cause of the purge leak/damage was yellow luer damage. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33468 | 1220648-2023-04256 | 17976865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17976865&pc=OZD | 11/24/2021 | 10/20/2023 |  | 10/20/2023 |  |  | 11/24/2021 | 10/20/2023 |  |  | 03/31/2023 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022067492 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/24/2021 |  |  |  |  | The user facility reported a 66 year-old female was implanted with an impella device for bridge to decision for transplant. it was reported that during a purge cassette change, the yellow luer cracked on impella purge sidearm. as a result, purge open alarms due to leaking of purge solution occurred and a purge sidearm bypass had to be completed to maintain function. there was no patient harm reported. | [5 empty] | The device was returned for evaluation. microscopic inspection of the purge sidearm revealed a crack on the yellow luer at the weld line. sodium bicarbonate was used as a purge solution. the root cause of the reported cracked luer and the subsequent purge system open alarms was determined to be a damaged yellow luer. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33469 | 1220648-2023-04255 | 17976929 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17976929&pc=OZD | 10/12/2021 | 10/20/2023 |  | 10/20/2023 |  |  | 10/12/2021 | 10/20/2023 |  |  | 06/30/2023 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022098178 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/12/2021 |  |  |  |  | The user facility reported a 46 year-old female was implanted with an impella 5.5 device for bridge to transplant. it was reported that the device cracked after the patient rolled over on it. the device was replaced. no reports of adverse events were reported. | [5 empty] | The investigation into the reported cracked device has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the cause of the purge leak/crack was most likely sidearm bond damage. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33470 | 1220648-2023-04258 | 17977015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17977015&pc=OZD | 11/28/2021 | 10/20/2023 |  | 10/20/2023 |  |  | 11/28/2021 | 10/20/2023 |  |  | 08/31/2023 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022135034 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/28/2021 |  |  |  |  | The user facility reported a 65yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that while boosting the patient up in bed, the nurse grabbed under the patient arm and then noticed the console alarming. a transesophageal echo was performed and showed the pump pulled into the aorta. attempts were made to cross aortic valve. however, the pump was removed completely in the operating room a few hours later. there was no patient harm reported. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the positioning issue was use related. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33471 | 1220648-2023-04259 | 17977038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17977038&pc=OZD | 12/17/2021 | 10/20/2023 |  | 10/20/2023 |  |  | 12/17/2021 | 10/20/2023 |  |  | 08/31/2023 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022136249 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/17/2021 |  |  |  |  | The user facility reported a 73yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the device was accidentally pulled from the left ventricle. the pump was replaced. there was no patient harm reported. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the improper positioning was use related. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33472 | 1220648-2023-04260 | 17977124 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17977124&pc=OZD | 10/29/2022 | 10/20/2023 |  | 10/20/2023 |  |  | 10/29/2022 | 10/20/2023 |  |  | 04/30/2024 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2023230536 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/29/2022 |  |  |  |  | The user facility reported a 72-year-old male patient was implanted with an impella device for mechanical circulatory support. after a purge bag change, fluid was seen dripping from the junction of the purge system and pump. the fluid leak coincided with a drop in purge pressure and a purge filter bypass was subsequently completed to resolve the noted issue. there was no harm to the patient. | [5 empty] | The returned sidearm was visually inspected and a crack was observed in the yellow luer. a syringe with colored fluid was connected to the sidearm, light pressure was applied, and a leak was reproduced from the yellow luer. the cause of the purge leak was sidearm yellow luer damage. no bicarbonate reported in the patient case or the patient update details. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
33473 | 1220648-2023-04257 | 17977145 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17977145&pc=OZD | 12/10/2021 | 10/20/2023 |  | 10/20/2023 |  |  | 12/10/2021 | 10/20/2023 |  |  | 08/31/2023 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022124046 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/10/2021 |  |  |  |  | The user facility reported a 74yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that it was difficult to explant the pump. the graft was cut open and clots could be seen. the physician suspected that the clotting was preventing the catheter from moving. | [5 empty] | The investigation for the difficulty in removal of the pump has been completed. the device was returned and the catheter was able to move without any issues however all the blood had dried. in the graft a significant amount of blood could be seen. the root cause of the difficult explant was most likely a result of significant blood clotting in the graft. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33474 | 1220648-2023-04261 | 17977308 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17977308&pc=OZD | 10/12/2022 | 10/20/2023 |  | 10/20/2023 |  |  | 10/12/2022 | 10/20/2023 |  |  | 06/30/2024 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2023249971 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | Material Puncture/Hole | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/12/2022 |  |  |  |  | The user facility reported a 50-year-old male patient was implanted with an impella device for mechanical circulatory support. during support, drops of blood were seen coming from the red plug. after discussing the noted observation with the physician, the decision was made to maintain support with the implanted pump. no intervention was required and patient support continued without complication. | [5 empty] | The pump was returned and a visual inspection verified the presence of blood in the red handle. further inspection identified a hole in the catheter from which the leak originated. upon conclusion of the investigation, the leak was confirmed and the cause was traced to a hole in the catheter. this report is being filed as part of a retrospective review of historical records.
33475 | 1220648-2023-04262 | 17977374 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17977374&pc=PYX | 09/22/2023 | 08/21/2024 |  | 10/20/2023 |  |  | 09/22/2023 | 10/20/2023 |  |  | 07/31/2024 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023232114 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient's cardiogenic shock worsened overnight and the decision was made to place the impella rp. continuous renal replacement therapy (crrt) initiated for metabolic acidosis (ph 7.26 post intervention, now 7.32 after initiation). the urine was red and labs were hemolyzed. transesophageal echocardiogram (tee) revealed loss of anterior wall motion and the left ventricular ejection fraction (lvef) was 10% down from 45%. decision was therefore made to place an impella cp and support patient with bipella. | It was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the hemolysis has been completed. according to the clinical, the patient began producing red colored urine and hemolysis was confirmed. later the same day, tee revealed loss of anterior wall motion so the decision was made to place an impella cp device as well. in an attempt to resolve the issues the pumps were repositioned. however, the hemolysis remained so the decision was made to remove both pumps. after the pumps were removed hemolysis still remained. the patient expired on support shortly afterward. no product was returned for an investigation. no data logs were returned for analysis. the cause of the hemolysis was not determined because no product or logs were returned and not enough clinical information was given. the instructions for use for the impella rp smartassist system with automated impella controller states the following: section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, cardiac or vascular injury (including ventricular perforation)¿. section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions including catheter position, pre-existing medical conditions, and small left ventricular volumes may also play a role in patient susceptibility to hemolysis.¿ added- b5 mso review, h6 component code 925 in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
33476 | 1220648-2023-04263 | 17977528 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17977528&pc=OZD | 11/23/2022 | 10/20/2023 |  | 10/20/2023 |  |  | 11/23/2022 | 10/20/2023 |  |  | 07/31/2024 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2023237913 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 40 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/23/2022 |  |  |  |  | The user facility reported a 40-year-old male patient was implanted with an impella device for mechanical circulatory support. during support, blood was seen coming from the red plug. as a result of the noted issue, the pump was replaced for ongoing support. there was no reported harm to the patient. | [5 empty] | The device was returned and an evaluation was completed. in regards to the reported leak, a visual inspection was performed and a hole in the catheter was identified near the motor housing. based on the provided information, the cause of the product damage was blood ingress due to a hole in the catheter resulting from improper use. this report is being filed as part of a retrospective review of historical records.
33477 | 1220648-2023-04264 | 17977603 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17977603&pc=OZD | 12/23/2022 | 10/20/2023 |  | 10/20/2023 |  |  | 12/23/2022 | 10/20/2023 |  |  | 04/30/2024 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023236588 |  | 1000211 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/23/2022 |  |  |  |  | The user facility reported a 81-year-old male patient was implanted with an impella device for mechanical circulatory support. the complainant reported a pump failure alarm following initiation of support. the white cable was reseated, the alarm disappeared, and the issue resolved. despite successful restart of the pump, the physician opted to replace the device. there was no harm to the patient. | [5 empty] | The pump was returned for evaluation. the log shows the pump unplugged during repositioning of the pump. when the pump was plugged back in, impella defective alarms occurred. a visual inspection of the pump and a visual inspection inside the red handle did not reveal any abnormalities. the pump was plugged into the automated impella controller (aic) and run and the impella defective alarm was not reproduced. the pump was unplugged while running and plugged back into the console and the alarm was still not reproduced. the cause of the positioning issue was use related. this report is being filed as part of a retrospective review of historical records.
33478 | 1220648-2023-04265 | 17978007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17978007&pc=OZD | 02/07/2023 | 10/20/2023 |  | 10/20/2023 |  |  | 02/07/2023 | 10/20/2023 |  |  | 11/30/2024 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2024287570 |  | 0550-0002 |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 40 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/07/2023 |  |  |  |  | The user facility reported a 40-year-old male patient was implanted with an impella 5.5 device for mechanical circulatory support. during support, a purge leak was identified and the purge cassette was replaced to troubleshoot the issue. however, the leak persisted resulting in completion of a purge bypass. an increase in purge pressure was noted, but support was maintained with the implanted pump. there was no patient harm resulting from the failure. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. a cause for the reported leak could not be established. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
33479 | 1220648-2023-04266 | 17978055 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17978055&pc=OZD | 02/01/2023 | 10/20/2023 |  | 10/20/2023 |  |  | 02/01/2023 | 10/20/2023 |  |  | 12/31/2024 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2024289222 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/01/2023 |  |  |  |  | The user facility reported a 60-year-old male patient was implanted with an impella device for mechanical circulatory support. during support, high purge pressure/purge flow low alarms were triggered. lysis treatment was attempted but the purge pressure remained high and the decision was made to explant the pump. the patient was transferred to a left ventricular assist device (lvad) for ongoing support. there was no patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon conclusion of the investigation, a definitive cause for the noted failure could not be determined. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ this report is being filed as part of a retrospective review of historical records.
33480 | 1220648-2023-04267 | 17978263 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17978263&pc=OZD | 02/24/2023 | 10/20/2023 |  | 10/20/2023 |  |  | 02/24/2023 | 10/20/2023 |  |  | 12/31/2024 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2024300416 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/24/2023 |  |  |  |  | The user facility reported a 53-year-old male patient was implanted with an impella device for mechanical circulatory support. during support, the purge flow dropped and the aic alarmed low purge flow/high purge pressure. lysis was initiated and purge flow recovered. g5 with nabic was subsequently administered and support was maintained. after three days, the purge flow went back to initial value of about 10ml/h. there was no patient harm. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, the root cause of the high purge pressure was unable to be determined. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ this report is being filed as part of a retrospective review of historical records.
33481 | 1220648-2023-04268 | 17978296 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17978296&pc=OZD | 04/05/2022 | 10/20/2023 |  | 10/20/2023 |  |  | 04/05/2022 | 10/20/2023 |  |  | 09/30/2023 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022151102 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/05/2022 |  |  |  |  | The user facility reported a 58yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the impella got accidentally explanted as patient stepped on the white cable. the device was replaced with a intraortic balloon pump. there was no patient harm reported. | [5 empty] | The investigation for the positioning issue has been completed. the device was not returned for evaluation. however, the clinical report provide sufficient details to determine a root cause. therefore, the root cause of the positioning issue was patient stepping on the patient cable pulling the pump out of the ventricle. the instructions for use provides details on handling the device to prevent issue with pump usage. it states "¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33482 | 1220648-2023-04269 | 17979188 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17979188&pc=OZD | 03/14/2023 | 10/20/2023 |  | 10/20/2023 |  |  | 03/14/2023 | 10/20/2023 |  |  | 12/31/2024 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024301277 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/14/2023 |  |  |  |  | The user facility reported a 74-year-old male patinet was implanted with an impella 5.5 device for mechanical circulatory support. the us complainant had a 5.5 being prepped and exhibited a position not reliable alarm. the team made choice to replace both the pump and the automated impella controller(aic). | [5 empty] | The device was returned and an evaluation was completed. no external damage or abnormalities were found on the returned pump. when the pump was connected to the aic, the ps spiked out of range and the placement signal not reliable alarm reproduced. the optical 5s parameters of the returned pump were out of specification. optical continuity testing revealed optical fiber damage within the red handle at a significant optical lumen kink against the post proximal to the patient cable. upon conclusion of the investigation, the cause was determined to be an out-of-box damaged optical fiber line. this report is being filed as part of a retrospective review of historical records.
33483 | 1220648-2023-04270 | 17979523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17979523&pc=OZD | 03/15/2023 | 10/20/2023 |  | 10/20/2023 |  |  | 03/15/2023 | 10/20/2023 |  |  | 11/30/2024 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2023282276 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/15/2023 |  |  |  |  | The user facility reported a 61-year-old male patient was implanted with an impella device for mechanical circulatory support. during support, the aic displayed a purge system blocked alarm. the purge flow at the time of the alarm was .5 ml/h and the purge pressure was 1226 mmhg. blood was also visible on the purge side arm. a purge bypass was completed, but the issue remained. due to the persistent issue, the decision was made to replace the pump. another pump was successfully inserted for ongoing support. there was no patient harm. | [5 empty] | The device was returned for evaluation. no data logs were returned. the sidearm was not returned. the returned pump was attached to a purge cassette with a purge bypass and the high purge pressure was reproduced on a console. biomaterial was observed in the purge gap. the pump was soaked in water with tergazyme overnight, but the high purge pressure still did not resolve. no kinks were noted. upon conclusion of the investigation, the cause of the high purge pressure was biomaterial. this report is being filed as part of a retrospective review of historical records.
33484 | 1220648-2023-04271 | 17980252 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17980252&pc=OZD | 09/20/2023 | 01/18/2024 |  | 10/20/2023 |  |  | 09/20/2023 | 10/20/2023 |  |  |  |  | 10/20/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 1161391 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Male |  |  |  | Malfunction | No | Yes | Air/Gas in Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 09/20/2023 |  |  |  |  | The user facility reported a patient (age, gender and race unknown) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the team exchanged the purge fluid bag and the automated impella controller (aic) repeatedly displayed an air in the system alarm. the team utilized the de-airing procedure without success. the team then set up the back up aic and it malfunctioned with an error of the purge system. the team switched back to the initial aic and the issue was resolved. the patient remained stable throughout this event. the initial aic was used successfully. | [5 empty] | The investigation into the aic - purge issue ¿ other has been completed since the original report was filed. the console was returned for investigation. upon console return, the entire purge assembly was checked for any signs of abnormality but there were no obvious malfunctions. a pump and purge were then run with no problem and the purge failures from the field were not replicated. a pump and purge were then run with a kinked purge line and the failures seen in the field were reproduced. data logs were reviewed which found that the console was having multiple intermittent low purge pressure and air in purge alarms as the case wizard bag change procedure was conducted. eventually, these alarms resolved and the console was able to operate without failure for ten more hours. the failures seen in the field were most likely caused by a kinked purge line. | Correction d4: a correction was made to the lot number and serial number. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
33485 | 1220648-2023-04272 | 17980339 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17980339&pc=OZD | 09/20/2023 | 02/26/2024 |  | 10/20/2023 |  |  | 09/20/2023 | 10/20/2023 |  |  | 06/30/2025 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024372882 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/20/2023 |  |  |  |  | The user facility reported a 63-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the physician used the right femoral artery as the access point with 14 fr sheath and repo sheath. the right femoral artery spasmed and there was no distal flow with the repo sheath in. an antegrade stick was made for distal perfusion. | [5 empty] | The investigation into the limb ischemia reversible issue has been completed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33486 | 1220648-2023-04273 | 17980424 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17980424&pc=OZD | 09/20/2023 | 02/06/2024 |  | 10/20/2023 |  |  | 09/20/2023 | 10/20/2023 |  |  | 06/30/2025 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024373896 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 18 YR |  | Female |  |  | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an 18-year-old female was implanted with an impella device for mechanical circulatory support. during support, it was noted that the impella cp stopped several times. on the third attempt the impella restarted. the decision was made to explant the impella. the patient was not harmed. | [5 empty] | The investigation for the reported pump stop has been completed. the impella pump was not returned for evaluation. however, data logs were reviewed which confirmed there was a pump stop. however, without the device for evaluation, the root cause of the pump stop was unable to be determined as limited clinical details were provided. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
33487 | 1220648-2023-04275 | 17980447 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17980447&pc=OZD | 09/21/2023 | 09/19/2024 |  | 10/20/2023 |  |  | 09/21/2023 | 10/20/2023 |  |  | 09/30/2024 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023274092 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Improper Flow or Infusion | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 09/21/2023 |  |  |  |  | The user facility reported a 34-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the waveforms and impella flows exhibited a classic sign for clot with flows increasing initially and then plateauing at an exceedingly low rate. many alarms were present including "impella in ventricle," "placement signal low," and "suction." successful insertion of second catheter. | Procedural outcome noted that the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The investigation into the low pump flow issue has been completed since the original report was filed. the device was returned for investigation. data logs showed about 10 minutes into support, pump position was in ventricle that triggered impella position in ventricle and suction alarms. next, pump was run for 15 more minutes then was stopped manually and disconnected from the console. visual inspection found no issues on the pump. pump ran without issue under bench test. the root cause of positioning issue was most likely use related. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-04275 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
33488 | 1220648-2023-04274 | 17980451 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17980451&pc=OZD | 09/15/2023 | 10/20/2023 |  | 10/20/2023 |  |  | 09/21/2023 | 10/20/2023 |  |  |  |  | 10/20/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1324024 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 09/21/2023 |  |  |  |  | The user facility biomedical engineer reported their automated impella controller (aic) displayed a battery failure alarm (red alarm). the aic is coming back for investigation. there was no patient involvement. | [5 empty] | The investigation into the reported automated impella controller battery failure has been completed. data analysis: there was a battery failure alarm (transport connection blown/ missing ) due to low pack voltage . additionally, the lt logs showed that battery 1 had a cell imbalance on cell 3. device analysis: the failure mode was reproduced on an engineering bench. the battery 1 pack voltage was measured to be 0 v rather than the expected 16 v, and the battery lcd indicator was blank (fully discharged). bat-test confirmed battery 1 cell 3 was outputting 300mv less than the other cells in battery 1 which triggered the cell imbalance condition. root cause: the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure. the failure modes will be monitored and trended.
33489 | 1220648-2023-04276 | 17980453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17980453&pc=OZD | 09/22/2023 | 02/01/2024 |  | 10/20/2023 |  |  | 09/22/2023 | 10/20/2023 |  |  | 08/31/2025 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024385080 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 56-year-old female was implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 was noted to have migrated to the aorta following suction alarms. echo was performed to check the position and pump was attempted to be pulled back. however, attempts were unsuccessful, leading to the pump being explanted and new pump being placed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported positioning issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. data confirmed pump was in aortic area (about 6.75 hours into the case). pump then was stopped manually & disconnected from the console. the root cause of positioning issue was due to attempt to re-cross aortic valve wirelessly. the impella is not indicated for wireless re-access per ifu 10003049. .
33490 | 1220648-2023-04277 | 17980470 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17980470&pc=OZD | 08/31/2023 | 07/17/2024 |  | 10/20/2023 |  |  | 09/20/2023 | 10/20/2023 |  |  | 06/30/2025 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024373698 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/20/2023 |  |  |  |  | The user facility reported a 68-year-old female in anterior st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. after removal of the device, the patient had a post closure groin complication. patient was brought down to cath lab for explant, and perclose x2 was used. closure appeared to go well. patient was brought up to inpatient unit, where a hematoma was found, and hemoglobin dropped. the patient's heartrate began to increase and blood pressure dropped, causing the patient to be sent to icu. icu was able to get bleeding under control. it is unknown if blood products were given. | [5 empty] | The investigation for the access site bleed and removal issues has been completed. according to the clinical, five hours after beginning impella cp support, the patient had post closure groin complications. the patient was brought down to cath lab for explant, and two perclose were used. although the closure appeared to go well, after the patient was brought to an inpatient unit a hematoma was found and their hgb dropped. the patient's hr then began to increase and their bp dropped, causing them to be sent to icu where the bleeding was resolved. no product was returned for a visual inspection. data log analysis was not necessary for this complaint. the cause of the minor access site bleeding was not determined because no product was returned and not enough clinical information was given. the instructions for use referenced in the initial report is for the impella cp with smart assist system in the following sections: section: general patient care considerations section: entering cardiac output section: potential adverse events (united states), which now includes the statement "cardiac or vascular injury (including ventricular perforation).¿ added-h6 component code 925 added-d6a/d6b f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33491 | 1220648-2023-04278 | 17980518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17980518&pc=OZD | 09/20/2023 | 02/16/2024 |  | 10/20/2023 |  |  | 09/20/2023 | 10/20/2023 |  |  | 04/30/2025 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024345733 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 88 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/20/2023 |  |  |  |  | The user facility reported an 88-year-old male who had cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the patient was bleeding within the bronchus (due to intubation), pcps puncture site, and impella puncture site. platelet (thrombocyte) count level decreased to 3 and hb6. the doctor said that since the patient had a tendency to bleed, the patient would be managed without heparin in the purge. blood products were administered. the patient's coagulation improved. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed. the device was not returned for investigation. based on clinical description, the root cause of access site bleeding was most likely due to patient condition of a tendency to bleed. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33492 | 1220648-2023-04279 | 17980530 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17980530&pc=OZD | 09/20/2023 | 04/16/2024 |  | 10/20/2023 |  |  | 09/20/2023 | 10/20/2023 |  |  | 08/31/2024 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023260864 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/20/2023 |  |  |  |  | The user facility reported a 75-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the impella site continued to bleed after repositioning sheath was positioned requiring the patient to go to icu with fem stop in place. patient continued to require fem stop causing leg ischemia. the impella was explanted 1.5 hours earlier than planned. | [5 empty] | The investigation into the access site bleeding ¿ major and the limb ischemia ¿ reversible has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia and the access site bleeding issues was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related events are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33493 | 1220648-2023-04280 | 17980541 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17980541&pc=OZD | 09/20/2023 | 02/20/2024 |  | 10/20/2023 |  |  | 09/20/2023 | 10/20/2023 |  |  | 02/28/2025 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024337453 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp delivered to support a transcatheter aortic valve replacement (tavr) patient. the pump failed to deliver through the 1st 14fr sheath. the team replaced the sheath with a new 14fr and delivery was accomplished. the pump had low flows/suction and the team believed perhaps there was a damage made to the pump in the insertion at the sensor. | [5 empty] | The investigation into the delivery issues- anatomy and the low pump flow issues has been completed. the device was not returned for investigation. per the clinical information provided, the patient had a newly implanted tavr before impella use. the doctor encountered difficulty inserting the impella using a non-impella sheath, which was resolved by using the sheath recommended for impella. the cause of the difficulty inserting/removing pump through introducer issue was related to the use of a non-recommended sheath. the cause of the low pump issue was most likely biomaterial based on data log review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
33494 | 1220648-2023-04281 | 17980557 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17980557&pc=OZD | 09/22/2023 | 01/22/2024 |  | 10/20/2023 |  |  | 09/22/2023 | 10/20/2023 |  |  |  |  | 10/20/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1329519 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the purge disk for an automated impella controller (aic) had completely broken off. there was no known harm or injury resulting from the damage. | [5 empty] | The device was returned and the investigation is ongoing. upon completion of the investigation, a supplemental report will be submitted. | The investigation for the console purge disc/flag issues has been completed. the console and the logs have been returned for evaluation. the logs were reviewed but were not related to the reported issue. the console was evaluated and confirmed a broken purge retainer. clinical information provided stated that the purge disk had completely broken off. the root cause of the reported issue was a broken purge retainer.
33495 | 1220648-2023-04282 | 17980599 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17980599&pc=OZD | 09/21/2023 | 04/16/2024 |  | 10/20/2023 |  |  | 09/21/2023 | 10/20/2023 |  |  | 06/30/2025 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024368532 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A patient of unknown age and gender received hemodynamic support from an impella cp smartassist heart pump due to acute myocardial infarction/cardiogenic shock. after inserting the impella cp, the heart pump initially delivered a pump flow of 3 l/min, which then suddenly dropped to 0.1 l/min. to rule out suction as a cause, the correct positioning of the pump in the left ventricle was checked and found to be good. therefore, the initial impella cp was replaced with a new impella cp. with the new impella cp heart pump, the percutaneous coronary intervention (pci) procedure could be carried out without any problems. the patient was subsequently stable and was transferred to the intensive care unit. | [5 empty] | The investigation for the reported low pump flow has been completed. the pump was returned for investigation. the impeller blades were damaged. there was no instance of biomaterial or any foreign material ingested. the logs were not returned for investigation. the root cause of the low pump flow is due to a damaged impeller, however the cause of the damage is unknown since limited clinical information was provided. added- d9 device return date, h6 problem code 2954 | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
33496 | 1220648-2023-04283 | 17980630 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17980630&pc=OZD | 10/28/2020 | 02/28/2024 |  | 10/20/2023 |  |  | 09/21/2023 | 10/20/2023 |  |  | 08/31/2022 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1500208 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Other | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Nerve Damage | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The medical center in (b)(6), 2020 had an impella cp support of 36 hours that was weaned and explanted. the patient had the support for deterioration after coronary artery bypass grafting (cabg). the patient had left or on iabp support, but was escalated the cp. later in (b)(6), 2023 there was a publication in a local newspaper that was found sharing the patient had been harmed and suffered lasting limb issues. the impella access limb had been restrained too tightly and the patient developed "permanent nerve injury" with foot drop. the peroneal nerve was compromised. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported limb ischemia has been completed since the original report was filed. no product was returned for review. without further clinical details surrounding the patient's condition and the product returned for analysis, the root cause of the nerve injury could not be determined.
33497 | 1220648-2023-04284 | 17980665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17980665&pc=OZD | 09/22/2023 | 10/20/2023 |  | 10/20/2023 |  |  | 09/22/2023 | 10/20/2023 |  |  |  |  | 10/20/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, JP |  |  | ABIOMED, INC. |  |  | US |  | 1379919 |  | 0048-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Failure of Device to Self-Test | Application Program Freezes, Becomes Nonfunctional | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 09/22/2023 |  |  |  |  | The user facility reported while using the automated impella controller (aic) for training, an alarm stating "self-diagnosis failed." the aic was not operational. the staff turned the power on and off several times and the problem persisted. the aic was to be returned for investigation. | [5 empty] | The investigation into the aic-self check error has been completed. data analysis: the reported failure has been confirmed. logs show the aic was booted ten times on the complaint date. device analysis: the console was tested after arriving in field service. upon visual inspection, it was confirmed the partition boot manager (pbm) board was corroded on both sides. the root cause for the system self-diagnosis failure was due to pbm corrosion. the failure modes will be monitored and trended.
33498 | 1220648-2023-04285 | 17980678 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17980678&pc=OZD | 09/21/2023 | 06/27/2024 |  | 10/20/2023 |  |  | 09/21/2023 | 10/20/2023 |  |  | 01/31/2024 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023182467 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/21/2023 |  |  |  |  | Additional information: the patient expired on impella support. per abiomed medical safety office review: there is no association between impella use and outcome. patient's peri-procedural condition was poor. | The user facility reported a 62-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, there was an alarm of high purge pressure and purge system blocked. it is unknown if the impella 5.5 was replaced at this time. it is noted that there was no patient harm. |  |  |  |  | The investigation into the high purge pressure issue has been completed since the original report was submitted. the device was not returned for investigation. the cause of the high purge pressure was not determined since product/data was not returned. added b5 and b6. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6 and f8 removed. h6 medical device problem code, corrected from 2954 to 1491. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
33499 | 1220648-2023-04286 | 17980691 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17980691&pc=OZD | 09/21/2023 | 02/07/2025 |  | 10/20/2023 |  |  | 09/21/2023 | 10/20/2023 |  |  |  |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1084353 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 50-year-old female was implanted with an impella 5.5 device for mechanical circulatory support. the automated impella controller (aic) displayed ¿impella position unknown¿/ active error alarm alarms. a bedside ultrasound confirmed placement. therefore, nurse was advised to swap the automated impella controller (aic). there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ | The investigation for the controller failure-red alarm has been completed. available console data logs didn't include data from the day of occurrence. review of available logs didn't reveal any observable anomaly with the console. the console was returned for evaluation. visual inspection internal circuitry of the console didn't reveal any observable anomaly with the console. a know good pump and purge line was run with console and no observable anomaly noted. the root cause for the reported controller error could not be determined due to insufficient information and the failure not being reproducible. added d9 device return date corrections to the initial mfr report: d2: updated common device name. d4: updated udi number. g1: mdr reporting contact email.
33500 | 1220648-2023-04287 | 17980693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17980693&pc=OZD | 09/20/2023 | 03/19/2024 |  | 10/20/2023 |  |  | 09/21/2023 | 10/20/2023 |  |  |  |  | 10/20/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1145540 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Alarm System | No Display/Image | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 09/21/2023 |  |  |  |  | The user facility planned to do an automated impella controller (aic) to aic transfer. when the new aic was turned on and set-up started, an alarm appeared stating "change purge cassette". there was no cassette in the aic. then, the screen froze. the team restarted the aic and had no further issues. there was no patient harm related to this event. | [5 empty] | The investigation into the aic - display issue has been completed since the original report was filed. the console was returned and tested after arriving in fs. the failure mode was reproduced. the root cause of the screen freeze was an aic software issue. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
33501 | 1220648-2023-04288 | 17980712 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17980712&pc=OZD | 09/21/2023 | 02/16/2024 |  | 10/20/2023 |  |  | 09/21/2023 | 10/20/2023 |  |  | 12/31/2024 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024305871 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/21/2023 |  |  |  |  | The user facility reported a 62-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support as a bridge to left ventricular assistance device (lvad). on the way to the icu, the patient developed a massive hematoma. the patient was moved back to the or where the physician did an axillary cutdown. vascular repaired anastomosis and fixed arterial leaking. one unit of packed red blood cells was given. the impella 5.5 was removed and the patient received the lvad. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed. the device was not returned for investigation. based on clinical description, the root cause of access site bleeding was most likely due to patient movement. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33502 | 1220648-2023-04298 | 17980937 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17980937&pc=OZD | 09/22/2023 | 03/07/2024 |  | 10/20/2023 |  |  | 09/22/2023 | 10/20/2023 |  |  | 07/31/2025 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024379792 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the physician decided to place intra-aortic balloon pump therapy (iabp) and remove impella cp. the reason for this was that the repo sheath was completely occlusive with no distal flow after an angiogram was taken through the re-access port. the physician was uncomfortable in placing an antegrade stick so impella was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible has been completed since the original report was filed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
33503 | 1220648-2023-04300 | 17980938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17980938&pc=OZD | 09/21/2023 | 01/24/2024 |  | 10/20/2023 |  |  | 09/22/2023 | 10/20/2023 |  |  |  |  | 10/20/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | ABIOMED, INC. |  |  | US |  | 1138132 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 09/22/2023 |  |  |  |  | The user facility reported a patient (unknown age, race, gender) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the automated impella controller (aic) displayed an alarm stating "purge disc not recognized". a backup aic was put into place. there was no patient harm. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation into the console purge disc/flag issue has been completed. the logs were returned for evaluation. the logs revealed that the purge disc not detected alarm occurred several times. the console was tested and the issue with the purge disc not being able to be recognized was reproduced. visual inspection of the console found that the purge retainer flag was broken which would result in a ¿purge disc not detected¿ alarm. the root cause of this issue was a broken purge pressure sensor flag.
33504 | 1220648-2023-04299 | 17980939 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17980939&pc=OZD | 09/22/2023 | 07/30/2024 |  | 10/20/2023 |  |  | 09/22/2023 | 10/20/2023 |  |  | 04/30/2025 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024355677 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/22/2023 |  |  |  |  | The user facility reported a 63-year-old male was implanted with an impella cp device for mechanical circulatory support. when the peel away sheath was removed from patient to suture device into place, patient became hostile and started flailing, causing access site to bleed and hematoma to form. site became bigger, and vascular team was called to ccl. it was the team's decision to remove device to repair the vessel. closure was successful. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the major bleed was patient movement as the bleeding began when the patient became combative. the instructions for use for the impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added-h6 component code 925 added-d6a/d6b. f6/f8 should have been left blank in the initial report.
33505 | 1220648-2023-04301 | 17980942 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17980942&pc=OZD | 09/21/2023 | 06/27/2024 |  | 10/20/2023 |  |  | 09/22/2023 | 10/20/2023 |  |  | 08/31/2025 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024386699 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/22/2023 |  |  |  |  | The user facility reported a 54-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. during support, a retroperitoneal bleed was found, vascular cutdown performed in hopes of removing cp and upgrading to 5.5. however axillary vessels were too small. md elected to place the patient on ecmo and leave groin cp in place. vascular surgery repaired insertion site. however the patient required multiple blood products overnight. | [5 empty] | The investigation into the vascular damage - peripheral vessel issue has been completed since the original report was submitted. the product was not returned. due to limited clinical details, the root cause of the vascular damage - peripheral vessels could not be determined. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted report incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
33506 | 1220648-2023-04297 | 17980954 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17980954&pc=OZD | 09/21/2023 | 10/20/2023 |  | 10/20/2023 |  |  | 09/21/2023 | 10/20/2023 |  |  |  |  | 10/20/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1103170 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) had a frozen screen. there was no reported patient harm. | [5 empty] | The investigation into the reported display issue has been completed. data analysis shows the purge cassette was reinserted 11 times. every time upon cassette insertion, the console ic6073 received ¿purge disc not detected - alarm,¿, which made the aic not pass through the cassette and disc insertion screen, which confirms the reported failure frozen screen. the aic console was also tested, and the issue was able to be reproduced on the bench. upon visual inspection, the missing disc detect flag was confirmed and a crack in the retainer was noted. the cause of the issue was determined to be a defective component. the failure modes will be monitored and trended.
33507 | 1220648-2023-04303 | 17980966 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17980966&pc=PYX | 09/22/2023 | 04/16/2025 |  | 10/20/2023 |  |  | 09/22/2023 | 10/20/2023 |  |  | 07/31/2025 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024385507 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | Company representation |  |  |  | 57 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 09/22/2023 |  |  |  |  | The user facility reported a 57-year-old male was implanted with an impella rp device for mechanical circulatory support. the peel-away sheath stayed in for the duration of support. after small but continuous oozing from peel-away sheath, decision made to remove peel-away and insert 15f sheath. after peeling away the sheath, mattress suture tightened with significant oozing noted and eventually new mattress suture placed for hemostasis. the oozing became so significant that left groin access was obtained to visualize right groin and assess for significant bleed. arterial bleed was not noted under fluoro, instead it was a much less significant venous bleed than expected. given the continuous oozing from site (tr/high central venous pressure/fluid overload without dialysis today likely contributing to oozing), the mattress suture was removed, and a deeper mattress suture was placed. this final mattress suture stopped the significant oozing from site and the patient returned to ccu. | Additional information received on 04/07/2025 via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding the patient that endured r groin bleeding with a "definite" relationship to the impella device used: "received pt in bed s/p rt femoral rp impella placement, p6 site oozing pressure bag applied. 9/24am right groin continue to bleed when patient coughs dressing changed and pressure bag placed. 9/24 evening: bleeding looks like contained @ this time, with pressure dressing¿. the patient recovered with no sequelae. |  |  |  |  | The investigation into the access site bleeding ¿ major issue has been completed. the device was not returned for investigation. based on clinical description, the root cause of access site bleeding was most likely due to anticoagulation management related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | B5 updated with additional information regarding the previously reported access site bleeding. instructions for use as follows: impella rp with smart assist system with automated impella controller. section: warnings and cautions. ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2023-04303. d6a and d6b revised on the initial manufacturer device report number 1220648-2023-04303 in accordance with updated procedures. f6/f8-should have been left blank initial manufacturer device report number 1220648-2023-04303. g1 reporting contact email revised on manufacturer device report 1220648-2023-04303 in accordance with updated procedures. g2 report source; study was inadvertently omitted on the initial manufacturer device report number 1220648-2023-04303. g4 pma/510(k)number was erroneously entered on the initial manufacturer device report number 1220648-2023-04303. h6 component code revised from the initial manufacturer device report number 1220648-2023-04303 in accordance with updated procedures. h10 instructions for use revised from the initial manufacturer device report number 1220648-2023-04303 in accordance with updated procedures.
33508 | 1220648-2023-04302 | 17980973 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17980973&pc=OZD | 09/22/2023 | 01/12/2024 |  | 10/20/2023 |  |  | 09/22/2023 | 10/20/2023 |  |  | 07/31/2025 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024378448 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had an impella 5.5 support ongoing for over 3 weeks when the access site had to have a hematoma washout. the patient remains on the support while awaiting transplant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for benchtop evaluation and investigation. per the clinical information provided, an impella surgical site washout was conducted due to the hematoma. the cause of the access site bleeding issue was not determined since product was not returned and sufficient clinical information was not provided.
33509 | 1220648-2023-04304 | 17980980 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17980980&pc=OZD | 09/22/2023 | 07/24/2024 |  | 10/20/2023 |  |  | 09/22/2023 | 10/20/2023 |  |  | 04/30/2025 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024355065 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 75-year-old female was implanted with an impella cp device for mechanical circulatory support. patient developed a large hematoma at impella site on the left side and femstop in place. heparin was paused for a while due to access site bleeding and the hematoma stabilized. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for access site bleeding/hematoma has been completed. the impella device was not returned for evaluation. the root cause of the major access site bleed was anticoagulation management because the bleeding was resolved by pausing heparin. d6a/d6b added h6 component code added 925 the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-04304: g1 reporting contact fax number missing.
33510 | 1220648-2023-04306 | 17981108 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17981108&pc=OZD | 09/14/2023 | 06/18/2024 |  | 10/20/2023 |  |  | 09/14/2023 | 10/20/2023 |  |  | 02/28/2025 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024337865 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | No | Risk Manager | Other | Health Professional | User facility | Company representation |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Deformation | [6 empty] | Cardiac Arrest | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/14/2023 |  |  |  |  | The user facility med watch report states a 64-year-old female in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support as a bailout measure. after implantation, the patient started to decompensate and coded. chest compressions were performed and the pump folded on itself during the compressions. vascular was consulted but cardiologist was able to slowly pull back and after several attempts the pump unfolded at the common iliac and eventually removed. shortly after removal of impella patient coded again. despite several rounds of cpr the team was unable to achieve return of spontaneous circulation and patient expired. no allegations were made towards the impella for cause of death. | The patient expired while on support. the outcome was reviewed by abiomed's medical safety officer and it was determined that there was not enough information to associate the use of the device with the patient's outcome. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the product damage has been completed. no product was returned for analysis. the clinical description stated that upon placement, the patient coded once again requiring chest compressions. the pump most likely folded on itself during chest compressions. the root cause of the product damage - cannula kink is most likely due to use. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 code 4114 was reported incorrectly on the previously submitted report. a new code has been added to type of investigation codes.
33511 | 1220648-2023-04307 | 17981248 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17981248&pc=OZD | 09/10/2022 | 10/20/2023 |  | 10/20/2023 |  |  | 09/10/2022 | 10/20/2023 |  |  | 06/30/2024 |  | 10/20/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023233413 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 52-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella 5.5 device for mechanical circulatory support. the complainant had an impella 5.5 support ongoing for a total of 21 days when during the support there were positional alarms for pump in aorta. flat motor current and uncoupled waveforms with impella in aorta alarms. echo performed showing inlet cage well inside the left ventricle (lv). surgeon repositioned multiple times with no change in mc, waveforms or flow. | [5 empty] | The investigation into the reported saturated flow has been completed. product was returned for investigation. the device was visually inspected and found biomaterial entraining the impeller. the investigation determined that the root cause of the issue was biomaterial. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
33512 | 1220648-2023-04198 | 17964555 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17964555&pc=OZD | 07/04/2021 | 10/19/2023 |  | 10/19/2023 |  |  | 07/04/2021 | 10/19/2023 |  |  | 04/30/2023 |  | 10/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022074306 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Device Sensing Problem | Material Integrity Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/04/2021 |  |  |  |  | The user facility reported a 70-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. during initial prep, the axillary kit sheath had notable ridges along the outside. a new axillary kit was subsequently utilized for impella implant. during placement, the pump had to be torqued to position mid ventricle, which caused a lot of torque in power cable. device unplugged to get torque out and replugged in. when ramping up flows, got a pump stop. device restarted and increased to p9. at p9, flows were 6.0 (mean) lpm and lv waveform was above position signal waveform. contacted csc and recommended replacing pump. new 5.5 opened and inserted without issues. no issues with flows on new device. there was no patient harm resulting from the noted issues. | [5 empty] | The returned introducers were visually inspected and the first introducer sheath was found to have scoreline grooves while the second introducer sheath scorelines were smooth. the scoreline observation was determined to be visual only and there was no weakness to the sheath wall and the sheath was not from the same lot of suspect sheaths that had weak sheath walls. the associated pump was not returned and an evaluation could not be completed.the observation made by the field was determined to be not related to the function of the sheath. the cause of the placement signal issue was not determined since the pump was not returned for evaluation. this report is being filed as part of a retrospective review of historical records.
33513 | 1220648-2023-04199 | 17965006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17965006&pc=OZD | 07/26/2021 | 10/19/2023 |  | 10/19/2023 |  |  | 07/26/2021 | 10/19/2023 |  |  | 04/30/2023 |  | 10/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2022093843 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/26/2021 |  |  |  |  | The user facility reported a 55-year-old male patient was implanted with an impella device for mechanical circulatory support. during support, a leak was observed from the purge arm filter. the pump was subsequently explanted as a result of the noted issue. the patient was stable with no report of harm or medical intervention. | [5 empty] | The product was returned for evaluation and bone wax was present on the purge side arm filter. upon removal of the bone wax, a leak from a crack in the filter was confirmed. the provided data logs confirmed low purge pressure alarms triggered on 26jul2021 at 20:11. the root cause of the leak in the purge side arm filter was most likely use related as it was reported that the catheter was not fixed properly, despite retraining. per the ifu: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
33514 | 1220648-2023-04200 | 17965147 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17965147&pc=OZD | 09/21/2023 | 04/22/2024 |  | 10/19/2023 |  |  | 09/21/2023 | 10/19/2023 |  |  | 05/31/2025 |  | 10/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024369916 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/21/2023 |  |  |  |  | The user facility reported an 81-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. during support, there were low pump flows on the impella cp. the pump behaved like it ingested clot, but activated clotting time was over 330. the pump was repositioned multiple times. staff bolused fluid and placed patient on a saline drip, still motor current kept getting flat. aortic and left ventricle placement signals continued to separate and continued to get reduced flows and suction events. the patient was stable following percutaneous coronary intervention so the impella cp was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the low pump flow and access site bleeding has been completed. the impella device was not returned for evaluation. data logs reported low pump flow from the start of the case while operating at p-level above p6. no motor current spike or trending placement sign was reported during the case. the cause of the issue was not determined since product was not returned and sufficient clinical information was not provided. added- h6 impact code 4628.
33515 | 1220648-2023-04201 | 17965238 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17965238&pc=OZD | 09/21/2023 | 12/01/2025 |  | 10/19/2023 |  |  | 09/21/2023 | 10/19/2023 |  |  |  |  | 10/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1503160 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 09/21/2023 |  |  |  |  | The user facility reported a 36-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. cp support ongoing when there were purge alarms and the purge disc was found to be broken. the purge disc holder came off the automated impella controller (aic). the team successfully set up the back up aic. there was no patient harm. | [5 empty] | Corrections to the initial mfr report: g1 mdr reporting contact email, d2 device name has been updated, d4 model number has been updated, d9 device return date added, f6 and f8 should have been left blank, h6 type of investigation code added. | Data analysis: on sep 9, 2023, during support, the console displayed ¿purge pressure low ¿ alarm¿ (event #431) a couple of times. the user then performed an incomplete cassette & bag change procedure, which likely indicates the procedure was left incomplete due to a broken retainer. this aligns with the reported failure mode device analysis: the console was tested after arriving at field service (fs). visual inspection confirmed a broken retainer and flag. root cause: the root cause of the purge pressure low was a defective purge retainer. h3 updated. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
33516 | 1220648-2023-04202 | 17965419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17965419&pc=OZD | 08/12/2021 | 10/19/2023 |  | 10/19/2023 |  |  | 08/12/2021 | 10/19/2023 |  |  | 04/30/2023 |  | 10/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2022092174 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/12/2021 |  |  |  |  | The user facility reported a 61-year-old male patient was implanted with an impella device for mechanical circulatory support. during support, the purge sidearm reportedly broke. a purge sidearm bypass was completed and support was maintained. there was no patient harm. | [5 empty] | The device was returned and an evaluation was completed. in regards to the report of a broken purge sidearm, a review of the provided data logs noted a decrease in purge pressure coinciding with an increase in purge flow on 10aug2021. this triggered purge pressure low alarms. the purge sidearm was returned in pieces with a needle for bypass in the purge lumen of the pump handle. upon conclusion of the investigation, the root cause of the purge sidearm damage was bond damage after patient movement.this report is being filed as part of a retrospective review of historical records.
33517 | 1220648-2023-04203 | 17965429 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17965429&pc=OZD | 05/03/2023 | 10/19/2023 |  | 10/19/2023 |  |  | 05/03/2023 | 10/19/2023 |  |  | 12/31/2024 |  | 10/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 |  |  | ABIOMED, INC. |  |  | US |  | 2024315825 |  | 1000100 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/03/2023 |  |  |  |  | The user facility reported that a 62-year-old female patient implanted with the impella 5.5 for mechanical circulatory support. the 5.5 was placed as care and support needed to be escalated from an impella cp. after over 6 days the team had a had a suspicion there was clot on the catheter. no further information was provided. | [5 empty] | The investigation into the reported thrombosis was completed. the device was not returned for evaluation. based on the available information, the root cause of the issue was not determined. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33518 | 1220648-2023-04204 | 17965492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17965492&pc=OZD | 05/26/2023 | 10/19/2023 |  | 10/19/2023 |  |  | 05/26/2023 | 10/19/2023 |  |  | 04/30/2025 |  | 10/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024333504 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/26/2023 |  |  |  |  | The user facility reported that a 50-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced low pump flows and suction alarms after the device was placed. echocardiogram showed optimal device placement, volume status and no detectable clot. the alarms were unresolved despite all troubleshooting measures. later in the support the plasma free hemoglobin (pfhgb) increased which is indicative of hemolysis, and the p-level was decreased from p-8 to p-6. the device was subsequently explanted. upon explant, no clot was visible on the device. | [5 empty] | The investigation into the reported low pump flows was completed. the device was not returned for evaluation. analysis of the data logs did not align completely with the clinical account. the data logs showed that the pump was initially running on p5, and a few suction alarms were noted, with period of pump speed adjustments. based on the log analysis, the slightly low flow and suction observed at the beginning of the case was resolved with a decrease of p level. without the returned device, the root cause of the low flow cannot be definitively determined. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33519 | 1220648-2023-04205 | 17965578 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17965578&pc=OZD | 05/26/2023 | 10/19/2023 |  | 10/19/2023 |  |  | 05/26/2023 | 10/19/2023 |  |  | 03/31/2025 |  | 10/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024327577 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/26/2023 |  |  |  |  | The user facility reported that a 66-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced a bend/kink and damage. there were no delivery or positioning issues reported. no vessel conditions were reported. there was no reported harm to the patient. the kinked pump was removed after just 20minutes of support and was replaced. | [5 empty] | The investigation into the reported catheter kink was completed. the device was returned for evaluation. visual inspection of the device confirmed a catheter kink near the motor housing. analysis of the data logs do not show any abnormalities. based on the available, the root cause of the product damage was a catheter kink by the motor housing from excessive force/a use related issue. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33520 | 1220648-2023-04207 | 17965600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17965600&pc=OZD | 04/27/2023 | 10/19/2023 |  | 10/19/2023 |  |  | 04/27/2023 | 10/19/2023 |  |  | 01/31/2025 |  | 10/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 |  |  | ABIOMED, INC. |  |  | US |  | 2024310011 |  | 1000100 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/27/2023 |  |  |  |  | The user facility reported that a 71-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced positioning issues and intermittent impella in aorta and suction alarms. the patient was subsequently taken for pump explant after under 3 hours of support. there was no harm reported to the patient. | [5 empty] | The investigation into the reported pump position issue was completed. the device was not returned for evaluation. based on clinical description, the root cause of positioning issue was most likely use- related. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33521 | 1220648-2023-04206 | 17965610 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17965610&pc=OZD | 11/08/2021 | 10/19/2023 |  | 10/19/2023 |  |  | 11/08/2021 | 10/19/2023 |  |  | 04/30/2023 |  | 10/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2022074227 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/08/2021 |  |  |  |  | The user facility reported a 67-year-old female patient was implanted with an impella device for mechanical circulatory support. the complainant reported a damaged purge system filter during support. a purge system bypass was performed and support continued with no harm to the patient. | [5 empty] | The field lt log was reviewed and low purge pressure alarm was confirmed 19 hours into support. purge returned to normal after the event and remained within specification through the remainder of the run. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon conclusion of the investigation, the cause of the purge leak was likely sidearm filter damage from ipa interaction; however, the product was not returned to confirm. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
33522 | 1220648-2023-04210 | 17965792 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17965792&pc=OZD | 07/10/2021 | 10/19/2023 |  | 10/19/2023 |  |  | 07/10/2021 | 10/19/2023 |  |  | 03/31/2023 |  | 10/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022067490 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/10/2021 |  |  |  |  | The user facility reported a 66 year-old female in cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the purge system was leaking which required a pump replacement after 2 weeks of support. there was no patient harm as a result of the issue. | [5 empty] | The investigation into the reported purge system leak has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. however, based upon previous reported issues of the same failure, the most likely cause of the purge leak was sidearm bond damage. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33523 | 1220648-2023-04211 | 17965840 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17965840&pc=OZD | 07/23/2021 | 10/19/2023 |  | 10/19/2023 |  |  | 07/23/2021 | 10/19/2023 |  |  | 05/31/2023 |  | 10/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022088119 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/23/2021 |  |  |  |  | The user facility reported a 60 year-old male with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was a low purge pressure alarm and leaking was found at the yellow luer of the purge sidearm. purge bypass performed. there was no patient harm as a result of the issue. | [5 empty] | The impella device was returned for evaluation. upon visual inspection, a crack was seen in the sidearm of the yellow luer. no other damage was found to the sidearm. the cause of the purge leak was sidearm yellow luer damage. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33524 | 1220648-2023-04212 | 17965884 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17965884&pc=OZD | 07/30/2021 | 10/19/2023 |  | 10/19/2023 |  |  | 07/30/2021 | 10/19/2023 |  |  | 06/30/2023 |  | 10/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022092871 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  |  | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/30/2021 |  |  |  |  | The user facility reported a 57 year old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the pump came out of position. there was no harm or injury noted. the pump was replaced and support proceeded on the second pump. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the positioning issues was determined to be patient movement as the patient coughing caused the impella to pull back across the valve. wireless reaccess is not permitted per the ifu. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33525 | 1220648-2023-04213 | 17965965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17965965&pc=OZD | 08/08/2021 | 10/19/2023 |  | 10/19/2023 |  |  | 08/08/2021 | 10/19/2023 |  |  | 05/31/2023 |  | 10/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022088145 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Difficult or Delayed Positioning | Material Deformation | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/08/2021 |  |  |  |  | The user facility reported a 55yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the pump was unable to pass the anatomy due to multiple kinks on cannula and catheter, a new 5.5 pump was used, and there was no harm to the patient. | [5 empty] | The investigation for the kinked pump has been completed. the device was not returned for evaluation. however, clinical report provides sufficient details to determine the root cause of the reported event. the cannula kinks were most likely due to patient condition. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33526 | 1220648-2023-04214 | 17966099 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17966099&pc=OZD | 01/21/2022 | 10/19/2023 |  | 10/19/2023 |  |  | 01/21/2022 | 10/19/2023 |  |  | 03/31/2023 |  | 10/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2022073836 |  | 0550-0001 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 46 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/21/2022 |  |  |  |  | The user facility reported a 46-year-old female patient was implanted with an impella device for mechanical circulatory support. the complainant reported a pump stop occurring during support that was successfully restarted. support continued with no harm to the patient. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. the root cause of the reported pump pause was not able to be determined through this investigation as no product or data logs were provided. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
33527 | 1220648-2023-04215 | 17966210 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17966210&pc=OZD | 02/23/2022 | 10/19/2023 |  | 10/19/2023 |  |  | 02/23/2022 | 10/19/2023 |  |  | 11/30/2023 |  | 10/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023164773 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/23/2022 |  |  |  |  | The user facility reported a 67-year-old male patient was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported a placement signal failure during support. the pump was removed and replaced with no harm to the patient. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon conclusion of the investigation, the cause of the optical placement signal issue could not be determined. this report is being filed as part of a retrospective review of historical records.
33528 | 1220648-2023-04216 | 17966472 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17966472&pc=OZD | 07/08/2021 | 10/19/2023 |  | 10/19/2023 |  |  | 07/08/2021 | 10/19/2023 |  |  | 04/30/2021 |  | 10/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022070047 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 50 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | Pressure Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/08/2021 |  |  |  |  | The user facility reported a 50-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the us complainant had a 5.5 supporting a patient for 33 days when the purge system had to be bypassed due to a leak. the pump still supports the patient despite the bypass without harm or injury. | [5 empty] | The device has been returned for evaluation. visual inspection of the returned product revealed the purge sidearm had been removed and a purge bypass was put in place. additionally, there was a significant catheter kink next to the kink protector of the red pump handle. no other damage or abnormalities were found.the cause of the purge leak was most likely sidearm yellow luer damage due to sodium bicarbonate use. this report is being filed as part of a retrospective review of historical records.
33529 | 1220648-2023-04217 | 17966486 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17966486&pc=OZD | 04/17/2022 | 10/19/2023 |  | 10/19/2023 |  |  | 04/17/2022 | 10/19/2023 |  |  | 01/31/2023 |  | 10/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022036509 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 25 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/17/2022 |  |  |  |  | The user facility reported a 25-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported an impella in aorta alarm with flattened motor current/waveform. imagery found what appeared to be a clot lodged within the device's inlet. further information obtained found systemic heparin had been discontinued previously. the pump was removed and the patient was bridged to (left ventricular assist device) lvad. | [5 empty] | The pump was returned and an evaluation was completed. upon visual inspection of the device, a clot was observed in the outflow. upon conclusion of the investigation, the noted issue was attributed to the patient's condition. this report is being filed as part of a retrospective review of historical records.
33530 | 1220648-2023-04218 | 17966495 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17966495&pc=OZD | 08/09/2021 | 10/19/2023 |  | 10/19/2023 |  |  | 08/09/2021 | 10/19/2023 |  |  | 09/30/2022 |  | 10/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1499125 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 66 YR |  | Male |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/09/2021 |  |  |  |  | The user facility reported a 65 year-old male with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported the side arm was broken by the retainer clip. new 5.5 opened and used. there was no patient harm as a result of the issue. | [5 empty] | The device was returned for review and confirmed to have the sidearm split in half. based on the clinical review, the cause of the sidearm damage was a sidearm bond damage after improper technique. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33531 | 1220648-2023-04219 | 17966605 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17966605&pc=OZD | 08/14/2021 | 10/19/2023 |  | 10/19/2023 |  |  | 08/14/2021 | 10/19/2023 |  |  | 06/30/2023 |  | 10/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022097958 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/14/2021 |  |  |  |  | The user facility reported a 44yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that during the support the pump came out of the ventricle and could not be re-advanced past the valve. the pump was removed and replaced with a new impella. there was no patient harm reported. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the positioning issue was use related that the impella device is not indicated for wireless re-access. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33532 | 1220648-2023-04220 | 17966618 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17966618&pc=OZD | 08/16/2021 | 10/19/2023 |  | 10/19/2023 |  |  | 08/16/2021 | 10/19/2023 |  |  | 04/30/2023 |  | 10/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022076860 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Difficult or Delayed Positioning | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/16/2021 |  |  |  |  | The user facility reported a 71 yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that due to anatomy, the catheter kinked and they were unable to advance. a new pump was used. there was no patient harm reported. | [5 empty] | The investigation for the kinked catheter has been completed. the device was not returned for evaluation. however, the clinical report has sufficient details to determine the root cause of the reported event. the root cause of the delivery issues is patient's underlying native condition, due to the severe calcification of the innominate artery. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33533 | 1220648-2023-04221 | 17966741 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17966741&pc=OZD | 08/22/2021 | 10/19/2023 |  | 10/19/2023 |  |  | 08/22/2021 | 10/19/2023 |  |  | 04/30/2023 |  | 10/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022102543 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 60 YR |  | Female |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/22/2021 |  |  |  |  | The user facility reported a 60-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that there was a leak at a purge filter. a purge filter bypass with needle was completed successfully. there was no patient harm as a result of the issue. | [5 empty] | Evaluation of the returned product confirmed that a purge bypass had been performed. the yellow luer was returned separate from the purge sidearm and was observed to be cracked in half at the weld line. leak reproduction was attempted on the intact purge filter and blue reservoir; however, no leak was observed. the root cause of the purge sidearm leak was determined to be a damaged yellow luer. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33534 | 1220648-2023-04222 | 17966945 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17966945&pc=OZD | 04/27/2022 | 10/19/2023 |  | 10/19/2023 |  |  | 04/27/2022 | 10/19/2023 |  |  | 12/31/2023 |  | 10/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2023171782 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | Material Split, Cut or Torn | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/27/2022 |  |  |  |  | The user facility reported a 48-year-old female patient was implanted with an impella device for mechanical circulatory support. during support, a tear in the sterile sleeve was identified. a sterile bandage was applied and support with the pump continued. roughly two days later, a fluid leak was identified in the purge system. purge reservoir placement and a filter change resolved the noted leak. the pump was noted to have been used beyond indication. there was no harm or injury/infection noted as a result of the failure. | [5 empty] | The pump was returned and the damage to the sterile sleeve and the side arm was confirmed. the sidearm was broken at the purge tubing in-between the kink protector and the start of the retainer. the most likely root cause for the sidearm damage was purge line damage due to patient movement. this is a use related issue as the device was not fixed in a safe location. the damage to the sterile sleeve was also determined to be most likely a result of a use issue as it was reportedly torn off. per the ifu: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
33535 | 1220648-2023-04223 | 17967121 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17967121&pc=OZD | 09/09/2021 | 10/19/2023 |  | 10/19/2023 |  |  | 09/09/2021 | 10/19/2023 |  |  | 06/30/2023 |  | 10/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022100786 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 59 YR |  | Female |  |  |  | Malfunction | Yes | Yes | Device Fell | [6 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/09/2021 |  |  |  |  | The user facility reported a 59yo female was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the console was dropped and the pump stopped. the white cord was stuck in place and pliers were used to remove it. there were no visual issues with the pin. the plug was tried in the new console, but it would not recognize a pump being plugged in. patient returned to the operating room for removal and placement of the new impella. there was no patient harm reported. additionally there was a report from the nurse on the second day of support the patient had some bleeding at site and had to go back to or. team had made choice to withhold heparin in purge. | [5 empty] | The investigation for the reported pump stop has been completed. the device was not returned for evaluation. however, the information provided in the clinical is sufficient in determining root cause. therefore, the cause of the pump stop was most likely broken pins from the aic being dropped. the investigation into the access site bleeding is ongoing at this time. upon completion of the investigation, a final report will be forthcoming. the instructions for use provides details on how to manage a pump stop : ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33536 | 1220648-2023-04224 | 17967165 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17967165&pc=OZD | 09/10/2021 | 10/19/2023 |  | 10/19/2023 |  |  | 09/10/2021 | 10/19/2023 |  |  | 06/30/2023 |  | 10/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022106248 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/10/2021 |  |  |  |  | The user facility reported a 53-year-old male was implanted with an impella device for mechanical circulatory support. it was reported that while in the operating room, there was a "purge system open" alarm on the console. it was found that the purge flow was 30ml/hr and the purge pressure was 0mmhg. the customer then found a leak coming from the purge sidearm. the purge sidearm was cut off and a blunt needle repair was used to get through case. there was no patient harm reported as a result of the issue. | [5 empty] | The purge sidearm was returned for analysis and a leak was confirmed. data logs were also provided which showed 'purge system open' and 'purge pressure low' alarms. the root cause of the leak was a crack in the purge sidearm filter. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33537 | 1220648-2023-04225 | 17967221 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17967221&pc=OZD | 09/11/2021 | 10/19/2023 |  | 10/19/2023 |  |  | 09/11/2021 | 10/19/2023 |  |  | 06/30/2023 |  | 10/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022092875 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 55 YR |  | Female |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/11/2021 |  |  |  |  | The user facility reported a 55-year-old female was implanted with an impella device for mechanical circulatory support. it was reported that the staff observed a cracked purge arm on impella 5.5 and an ¿open purge alarm¿ was displayed. purge bypass performed with mitigation of open purge alarm. purge pressure and purge flows returned to normal. pump function and flows did not appear to be impacted. the patient was supported throughout. physician elected to replace affected pump with new 5.5. no reports of adverse events were reported. | [5 empty] | The returned purge sidearm was visually inspected and a crack on the male portion of the yellow luer was observed. of note, sodium bicarbonate was used as a purge solution. therefore, the root cause of the reported purge sidearm crack/leak was determined to be a damaged yellow luer. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33538 | 1220648-2023-04226 | 17967259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17967259&pc=OZD | 05/26/2022 | 10/19/2023 |  | 10/19/2023 |  |  | 05/26/2022 | 10/19/2023 |  |  | 11/30/2023 |  | 10/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023166227 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/26/2022 |  |  |  |  | The user facility reported a 66-year-old female patient was implanted with an impella 5.5 device for mechanical circulatory support. upon insertion of the pump, a placement signal not reliable alarm appeared on the console. the pump was subsequently removed and an impella cp pump was implanted for ongoing support. there was no patient harm. | [5 empty] | The device was returned and an evaluation was completed. the data logs showed the raw position signal (ps) measuring around 50 mmhg and the displayed measuring 500 mmhg during the case. the pump was initially connected to the automated impella controller (aic) upon return and the placement signal measured 400/400 mmhg. 5s parameters measured normal. pump was run in tergazyme and the ps slowly dropped to 11/8 mmhg. upon conclusion of the investigation, the root cause of the optical signal issue was determined to be most likely due to biomaterial. this report is being filed as part of a retrospective review of historical records.
33539 | 1220648-2023-04227 | 17967314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17967314&pc=OZD | 09/16/2021 | 10/19/2023 |  | 10/19/2023 |  |  | 09/16/2021 | 10/19/2023 |  |  | 11/30/2022 |  | 10/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022022666 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/16/2021 |  |  |  |  | The user facility reported a 32yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported there was a kink that was discovered once the pump was removed, according to the site, the kink was likely from operator hands when trying to insert the pump into the graft. to remedy issue another 5.5 pump was used, and there was no mention of harm to the patient. | [5 empty] | The investigation for the kinked catheter has been completed. the device was not returned for evaluation. however, the clinical reported provided sufficient details to determine the root cause. the root cause of the reported delivery issues was determined to be the reported catheter kink. based on the provided information, the catheter kink was most likely caused by excessive force during pump insertion. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33540 | 1220648-2023-04229 | 17967466 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17967466&pc=OZD | 06/08/2022 | 10/19/2023 |  | 10/19/2023 |  |  | 06/08/2022 | 10/19/2023 |  |  | 01/31/2024 |  | 10/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023186129 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/08/2022 |  |  |  |  | The user facility reported a 56-year-old male patient was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported a purge leak from the device's side arm (yellow luer). a bypass configuration was completed without any harm to the patient and support continued. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. as the product was not returned, a definitive cause for the issue could not be established. however, a clinical review of the provided information, determined a potential cause for the leak to be sidearm yellow luer damage due to the use of sodium bicarbonate in the purge solution. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
33541 | 1220648-2023-04228 | 17967471 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17967471&pc=OZD | 08/18/2021 | 10/19/2023 |  | 10/19/2023 |  |  | 08/18/2021 | 10/19/2023 |  |  | 06/30/2022 |  | 10/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1498647 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 59 YR |  | Male |  |  |  | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/18/2021 |  |  |  |  | The user facility reported a 59-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the purge side arm cracked while in the retainer. the 3-point fixation was not in the correct place . a side arm bypass done and the pump was removed the next morning. there was no patient harm as a result of the issue. | [5 empty] | The purge side arm and protector were returned for evaluation. upon evaluation, a crack in the filter was confirmed. the root cause of the purge filter crack was due side arm filter damage as a result of the 3 point fixation not being in the correct place. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33542 | 1220648-2023-04230 | 17967669 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17967669&pc=OZD | 07/24/2021 | 10/19/2023 |  | 10/19/2023 |  |  | 07/24/2021 | 10/19/2023 |  |  | 06/30/2023 |  | 10/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022092850 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  |  | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/24/2021 |  |  |  |  | The user facility reported a 56 year old male patient was implanted with an impella 5.5 pump for mechanical circulatory support. during support, it was noted that the purge sidearm was leaking at the yellow luer. purge bypass was completed which resulted in purge pressure and flow returning to normal ranges. sodium bicarbonate was used in the purge solution. there were no reports of adverse events, medical intervention or harm to the patient. | [5 empty] | The investigation into the reported purge leak was completed. the device was not returned for evaluation. based on a clinical review of the provided information, the root cause of the purge leak was most likely a damaged yellow luer due to the use of sodium bicarbonate in the purge system. this report is being filed as part of a retrospective review of historical records.
33543 | 1220648-2023-04231 | 17967756 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17967756&pc=OZD | 09/20/2021 | 10/19/2023 |  | 10/19/2023 |  |  | 09/20/2021 | 10/19/2023 |  |  | 06/30/2023 |  | 10/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022102419 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/20/2021 |  |  |  |  | The user facility reported a 67yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the pump moved back into the aorta. repositioning was performed the following day, but the patient had to have echo imaging to successfully support. the pump had been left overnight in the aorta. the pump remained on til the durable lvad was placed. | [5 empty] | The investigation for the reported positioning issue has been completed. no product or logs were returned for this investigation. however, the clinical report provided sufficient detail to determine the root cause. the root cause of the positioning issues was determined to be use related due to patient movement as the impella 5.5 was pulled back into the aorta accidently while the patient was being rolled. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33544 | 1220648-2023-04232 | 17969751 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17969751&pc=OZD | 09/20/2023 | 08/16/2024 |  | 10/19/2023 |  |  | 09/20/2023 | 10/19/2023 |  |  | 03/31/2025 |  | 10/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024347216 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that while performing the impella cp implant, the staff was unable to advance repositioning sheath after placing device. the device prolapsed in the ventricle when attempting to advance repositioning sheath. the impella was removed and replaced. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation for the delivery issues has been completed. product was not returned for review. the root cause of delivery issue was determined to be use related because the catheter was kinked during advancement through repositioning sheath after placement. added- d6a/d6b.
33545 | 1220648-2023-04233 | 17969810 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17969810&pc=OZD | 09/20/2023 | 02/02/2024 |  | 10/19/2023 |  |  | 09/20/2023 | 10/19/2023 |  |  | 06/30/2025 |  | 10/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024374919 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | Yes | Yes | Pumping Stopped | [6 empty] | Thrombocytopenia | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that patient on impella cp support was in the process of being weaned when there was a pump stop. impella cp was explanted and no new pump was required. the pump had supported for 15 days when it stopped. | In addition to the originally reported pump stop, information was received via abiomed's loqi database on 2023-10-06 that the patient endured thrombocytopenia prior to the pump stop (on 2023-09-11). platelet delivery was prompted. |  |  |  |  | The investigation for the pump stop and thrombocytopenia has been completed. the pump was not returned for investigation. data logs show several motor current spikes with pump stops occurring at the end of the case. additionally, the data logs indicate an hpp event during which there was a sudden three-minute rise in purge pressure coinciding with the first pump stop. according to clinical notes, the doctor reported a drop in platelet count, and the pump failed to restart after the third attempt. the cause of the pump stop issue was not determined since product was not returned and insufficient clinical information was not provided. the cause of the thrombocytopenia issue was most likely patient condition related and determined to be unknown relation to impella due to insufficient clinical information. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | Updated b(5), g(4) with new information received 2023-10-06. upon completion of our investigation, a supplemental mdr will be filed.
33546 | 1220648-2023-04234 | 17969819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17969819&pc=OZD | 09/20/2023 | 10/31/2023 |  | 10/19/2023 |  |  | 09/20/2023 | 10/19/2023 |  |  | 06/30/2025 |  | 10/19/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024377971 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during placement of impella cp, the patient¿s blood pressure dropped, and low flows were experienced. the impella cp was removed, and an aorta-gram showed an aortic dissection at a right turn in the aorta and heavy calcification. balloon tamponade was done for 20 minutes and dissection resolved. | [5 empty] | The investigation into the vascular damage issue has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. there was heavy calcification at the site of the aortic dissection and no resistance was noted during pump insertion. the cause of the vascular damage was most likely patient condition related due to the patient's cad, calcification, and tortuosity. the failure mode will be monitored and trended. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33547 | 1220648-2023-04235 | 17969987 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17969987&pc=OZD | 05/26/2023 | 10/19/2023 |  | 10/19/2023 |  |  | 05/26/2023 | 10/19/2023 |  |  | 12/31/2024 |  | 10/19/2023 |  | TEMPORARY NON-ROLLER TYPE CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024322790 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 47-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant had an impella cp support ongoing for six days when the pump stopped. the stopped pump was due to an underlying coagulopathy and clot formation of unknown cause/etiology. the pump was explanted and not replaced, as the patient was stable with ecmo alone. | [5 empty] | The investigation into the reported "pump stop" has been completed. product was returned for investigation. the device was visually inspected and found a cannula kink and also, biomaterial was found in the purge gap. there were no other abnormalities found. as per clinical investigation, the patient had an impella cp support ongoing for six days when the pump stopped. the stopped pump was due to an underlying coagulopathy and clot formation of unknown cause/etiology. the pump was explanted and not replaced, as the patient was stable with ecmo alone. the investigation determined that the root cause of the pump stop issue was biomaterial found on the outlet. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
33548 | 1220648-2023-04124 | 17956130 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17956130&pc=OZD | 01/08/2022 | 10/18/2023 |  | 10/18/2023 |  |  | 01/08/2022 | 10/18/2023 |  |  | 09/30/2023 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022146546 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  |  | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/08/2022 |  |  |  |  | The user facility reported that a 58-year-old female patient was implanted with an impella 5.5 pump for mechanical circulatory support. following a change in purge tubing, it was noted that a fluid leak was visible from the purge sidearm. it was noted that sodium bicarbonate was used in the purge solution throughout support. a purge bypass was subsequently completed to resolve the noted leak. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge leak was completed. the device was not returned for evaluation. the clinical information was reviewed. the root cause of the purge leak was sidearm yellow luer damage due to sodium bicarbonate in the purge solution. this report is being filed as part of a retrospective review of historical records.
33549 | 1220648-2023-04125 | 17956738 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17956738&pc=OZD | 01/10/2022 | 10/18/2023 |  | 10/18/2023 |  |  | 01/10/2022 | 10/18/2023 |  |  | 08/31/2023 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022138995 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/10/2022 |  |  |  |  | The user facility reported that a 47-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced low pump flows and "impella position in aorta" alarms. the alarm was reportedly still active following repositioning in addition to the low pump flows. the pump was scheduled for explant the following day. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported low pump flows was completed. the pump and data stick were returned for evaluation. analysis of the data logs confirmed impella position in aorta alarms on (b)(6) 2022 which persisted for the remainder of the case. of note, an overall loss of motor current pulsatility was observed at the onset of the alarms and pump flows were slightly below normal limits. the pump was visually inspected, and no abnormalities were noted. based on the available information, the root cause of the suspected low pump flow was believed to be related to the patient condition; however a definitive cause could not be established. this report is being filed as part of a retrospective review of historical records.
33550 | 1220648-2023-04126 | 17956862 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17956862&pc=OZD | 01/23/2022 | 10/18/2023 |  | 10/18/2023 |  |  | 01/23/2022 | 10/18/2023 |  |  | 09/30/2023 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022143118 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/23/2022 |  |  |  |  | The user facility reported that a 56-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced immediate flow saturation when the pump was turned on. a separate impella 5.5 was successfully implanted. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported saturated flow was completed. the pump was returned for evaluation. visual inspection revealed biomaterial around the base and main body of the impeller. functional testing of initially failed, but the pump was eventually able to start after the biomaterial was partially removed/dissolved. visual inspection of the pump following the run revealed a small amount of biomaterial still present under the outflow cage. the root cause of the flow saturation was most likely the observed ingested biomaterial applying resistance to the pump rotor. this report is being filed as part of a retrospective review of historical records.
33551 | 1220648-2023-04127 | 17957005 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17957005&pc=OZD | 01/23/2022 | 10/18/2023 |  | 10/18/2023 |  |  | 01/23/2022 | 10/18/2023 |  |  | 03/31/2023 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022056270 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Consumer | Health Professional |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/23/2022 |  |  |  |  | The user facility reported that a 62-year-old female patient implanted with the impella 5.5 for mechanical circulatory support experienced a purge leak. the patient initially used heparin in the purge but changed to sodium bicarbonate for eight hours. prior to to the change, staff noted a buildup of dextrose at the yellow-to-yellow connection. a purge bypass was performed without issue and pump pressure and flows returned to normal. despite the temporary resolution, the decision was still made to replace the pump. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge leak was completed. the device was not returned for evaluation. the clinical information was reviewed. the root cause of the purge leak was sidearm yellow luer damage due to sodium bicarbonate in the purge solution. per the ifu: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
33552 | 1220648-2023-04128 | 17957030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17957030&pc=OZD | 10/26/2022 | 10/18/2023 |  | 10/18/2023 |  |  | 10/26/2022 | 10/18/2023 |  |  | 06/24/2024 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023239092 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/26/2022 |  |  |  |  | The user facility reported that a 69-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced high purge pressure on the 10th day of support. heparin or sodium bicarbonate were removed at the time of the high purge pressure. once added back into the system the high purge pressure resolved. no alarms sounded. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported high purge pressure was completed. the pump and data stick were returned for evaluation. functional testing of the pump passed. analysis of the data logs confirmed the high purge pressure and the removal of heparin/bicarb from the purge prior to the increase in purge pressure. the root cause of the high purge pressure was most likely biomaterial buildup as it was resolved by adding bicarb/heparin back into the purge solution. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33553 | 1220648-2023-04129 | 17957083 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17957083&pc=OZD | 10/05/2022 | 09/20/2024 |  | 10/18/2023 |  |  | 10/05/2022 | 10/18/2023 |  |  | 02/29/2024 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023186216 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Fluid/Blood Leak | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/05/2022 |  |  |  |  | The user facility reported that a 38-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced a purge leak. an "open purge alarm" sounded as a result of the breakdown of the purge reservoir on the device. a needle bypass was performed to keep prevent pump failure. a leak was also reported from the purge sidearm reservoir. no retainer was in use. of note, the patient utilized a 12" extension set and sodium bicarbonate was reported to be used in the purge solution. there were no reports of harm to the patient. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The investigation into the reported purge leak was completed. the device was not returned for evaluation. the root cause of the purge leak could not conclusively be determined as the product was not returned. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-04129 was provided in sections b2, b5, h1, and h6.
33554 | 1220648-2023-04130 | 17957127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17957127&pc=OZD | 01/21/2022 | 10/18/2023 |  | 10/18/2023 |  |  | 01/21/2022 | 10/18/2023 |  |  | 10/31/2023 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022153886 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Consumer | Health Professional |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/21/2022 |  |  |  |  | The user facility reported that a 53-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced pump saturation/purge pressure high upon insertion. the pump had been pulled back across the left ventricle (lv) to reposition. it was removed from the body and reinserted. while attempting to increase the flow, the lv signal showed pump saturation and the purge system blocked alarm sounded. the purge pressure was high and there was pump saturation, with ao ps below lv ps. the pump started; however, all impella flows (if), including max/min/mean values, were 6.2 l. the physician replaced the device and support was initiated without issue. there were no adverse events to the patient and no medical intervention was required. additional information provided reported that the patient had not been adequately anticoagulated for weeks. the patient was previously on va ecmo and an impella cp. it was reported that it was possible that the device was pulled out of the packaging without lifting the number 2 tab first. the dextrose concentration in the purge was d5w. | [5 empty] | The investigation into the reported high purge pressure and pump saturation was completed. the pump, purge cassette, and data logs were returned for evaluation. analysis of the data logs confirmed a saturated pump and high purge pressure. visual inspection of the pump noted coagulated blood around the purge gap. the pump operated as expected during functional testing. the root cause of the high purge pressure was most likely the lack of anticoagulation related to patient management. this report is being filed as part of a retrospective review of historical records.
33555 | 1220648-2023-04131 | 17957159 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17957159&pc=OZD | 11/08/2022 | 09/20/2024 |  | 10/18/2023 |  |  | 11/08/2022 | 10/18/2023 |  |  | 06/30/2024 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023230454 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  | White | Death | Yes | Yes | Restricted Flow rate | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/08/2022 |  |  |  |  | The user facility reported that a 72-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced low pump flows and worsening respiratory status on the 15th day of support. there was a lack of decompression in the left ventricle seen on echo. the clinical team suspected a clot in device. the patient was subsequently sent for device exchange and a clot was confirmed clot in inflow. no further information was provided. there was no reported harm to the patient. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the reported low pump flows was completed. the device was not returned for evaluation, but data logs were received. analysis of the data logs were revealed multiple suction alarms on the last day of support. based on the available information and the report of a clot visualized in the inflow; the root cause of the low flow was most likely due to biomaterial in the inflow. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-04131 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
33556 | 1220648-2023-04132 | 17957198 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17957198&pc=OZD | 01/23/2022 | 09/26/2024 |  | 10/18/2023 |  |  | 01/23/2022 | 10/18/2023 |  |  | 06/30/2023 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022098180 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/23/2022 |  |  |  |  | The user facility reported that a 79-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced a purge sidearm leak. heparin was in the purge solution. a purge bypass and subsequent pump exchange were performed without issue. there were no reports of harm to the patient. | It was additionally reported there was bleeding complications in the operating room with the impella 5.5 and the patient received multiple blood products. the procedure outcome for the patient noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The investigation into the reported purge sidearm leak was completed. the pump, purge cassette, and data stick were returned for evaluation. based on the available information, the root cause of the purge leak was sidearm filter damage. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-04132 was provided in sections b1, b2, b5, d6a, d6b, h1, and h6.
33557 | 1220648-2023-04133 | 17957203 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17957203&pc=OZD | 11/21/2022 | 10/18/2023 |  | 10/18/2023 |  |  | 11/21/2022 | 10/18/2023 |  |  | 08/31/2023 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022125940 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/21/2022 |  |  |  |  | The user facility reported that a 48-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced thrombus. when the physician was explanting the device at bedside, after removal, he noticed a clot on the catheter inflow and on the motor housing near the graft insertion site. no further information was provided. there was no reported harm to the patient. | [5 empty] | The investigation into the reported thrombus was completed. the device was not returned for evaluation. no logs were returned and there are no logs available on constellation. based on the available information, the root cause of the thrombus was not determined since product. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33558 | 1220648-2023-04135 | 17957303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17957303&pc=OZD | 11/21/2022 | 10/18/2023 |  | 10/18/2023 |  |  | 11/21/2022 | 10/18/2023 |  |  | 09/30/2024 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023259141 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/21/2022 |  |  |  |  | The user facility reported that a 66-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced a purge pressure increase several minutes after starting the pump, with purge pressure rapidly increasing. visual inspection did not reveal any kinks in the lines. the purge line was flushed, and the purge cassette changed. the tissue plasminogen activator (tpa) protocol was subsequently initiated as precaution for suspected clots blocking the purge. the decision was made to replace the pump to minimize risk to the patient. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported high purge pressure was completed. the pump, purge cassette, and data logs were returned for evaluation. analysis of the data logs confirmed high purge pressure about ten minutes after initial pump activation. the purge pressure remained high and pump flow waveform were saturated throughout support. visual inspection of the pump and purge cassette revealed no external damage or abnormalities. functional testing of the purge cassette was within normal limits with no purge related alarms. visual inspection of the motor house revealed no evidence of blood ingress; however, there was evidence of tpa clearing out the purge line. the root cause of the high purge pressure was ingested biomaterial. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33559 | 1220648-2023-04136 | 17957326 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17957326&pc=OZD | 12/06/2022 | 10/18/2023 |  | 10/18/2023 |  |  | 12/06/2022 | 10/18/2023 |  |  | 09/30/2024 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023258412 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Activation, Positioning or Separation Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/06/2022 |  |  |  |  | The user facility reported that a 39-year-old male patient implanted with the impella 5.5 for mechanical circulatory support developed hemolysis six days into support. the plasma free hemoglobin (pfhs) levels were elevated. at that time, the clinicians also noted that the pump had "flipped" in the left ventricle (lv) and was touching the mitral valve. the pump was subsequently successfully repositioned and there were no further reports of hemolysis related issues. | [5 empty] | The investigation into the reported hemolysis was completed. the device was returned for evaluation. analysis of the data logs revealed numerous "impella position unknown" alarms on the day of the reported hemolysis, accompanied by a minor loss in placement signal (ps) waveform pulsatility. based on the available information, the root cause of the hemolysis was most likely improper positioning of the device. impella 5.5 ifu: "evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33560 | 1220648-2023-04138 | 17957442 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17957442&pc=OZD | 01/23/2023 | 10/18/2023 |  | 10/18/2023 |  |  | 01/23/2023 | 10/18/2023 |  |  | 11/30/2024 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023282851 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  |  | White | Malfunction | No | Yes | Material Integrity Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/23/2023 |  |  |  |  | The user facility reported that a 54-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced sterile sleeve damage. the clinical team reported that the sterile sleeve became "undone" overnight. due to concerns about the sterility of the catheter, the decision was made to exchange the device. there was no reported harm to the patient. | [5 empty] | The investigation into the reported pump damage was completed. the device was returned for evaluation. visual inspection revealed there was damage on the distal end of the sterile sleeve. the tuohy was unscrewed and came apart, causing the sleeve to become unsecured from the catheter. based on the available information, the root cause of the product damage was use related. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33561 | 1220648-2023-04140 | 17957566 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17957566&pc=OZD | 01/20/2023 | 09/20/2024 |  | 10/18/2023 |  |  | 01/20/2023 | 10/18/2023 |  |  | 11/30/2024 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023278424 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Material Integrity Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/20/2023 |  |  |  |  | The user facility reported that a 61-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced damage to the sterile sleeve when a clinical team member tore the sterile sleeve while attempting to reposition the device. the pump's performance was not affected. the device was not removed due to this issue. there was no reported harm to the patient. | The procedural outcome noted that the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the reported pump damage was completed. no product was returned for benchtop analysis. after reviewing the clinical information, it was determined that the cause of the issue was most likely a damaged sterile sleeve due to improper handling of the device during repositioning. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-04140 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
33562 | 1220648-2023-04141 | 17957633 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17957633&pc=OZD | 01/27/2023 | 10/18/2023 |  | 10/18/2023 |  |  | 01/27/2023 | 10/18/2023 |  |  | 10/31/2024 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023268038 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/27/2023 |  |  |  |  | The user facility reported that a 76-year-old male patient implanted with the impella 5.5 for mechanical circulatory support developed thrombus. upon, explant, the clinicians observed a clot on the device. the pump performed with no issues throughout support. there was no reported harm to the patient. | [5 empty] | The investigation into the reported thrombus was completed. the device was returned for evaluation. visual inspection of the returned pump revealed the presence of biomaterial on the impeller within the cannula. based on the available information, the root cause of the reported thrombus ingestion was due to patient condition related. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33563 | 1220648-2023-04143 | 17957691 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17957691&pc=OZD | 01/27/2022 | 10/18/2023 |  | 10/18/2023 |  |  | 01/27/2022 | 10/18/2023 |  |  | 09/30/2023 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022151097 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 16 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/27/2022 |  |  |  |  | The user facility reported that a 16-year-old male patient was implanted with the impella 5.5 for mechanical circulatory support. four days into support, flow saturation was observed from the pump. upon explant, the motor was noted to be encased in thrombus. the patient was subsequently implanted with a left ventricular assist device (lvad) to bridge to heart transplant. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported saturated flow was completed. the device was returned for evaluation. visual inspection of the pump confirmed a large thrombus encasing the motor and outflow. the root cause of the inaccurate high flow resulting in saturated flow was ingested biomaterial applying resistance to the pump rotor. this report is being filed as part of a retrospective review of historical records.
33564 | 1220648-2023-04142 | 17957700 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17957700&pc=OZD | 02/04/2023 | 10/18/2023 |  | 10/18/2023 |  |  | 02/04/2023 | 10/18/2023 |  |  | 10/31/2024 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023269566 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  |  |  | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/04/2023 |  |  |  |  | The user facility reported that a 52-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced a tear in the sterile sleeve by the 3-point fixation clip. per their protocol, the physician cleaned the area with iodine and placed a tegaderm over the site. there was no reported harm to the patient. | [5 empty] | The investigation into the reported sterile sleeve tear was completed. the device was not returned for evaluation. based on the available information, the root cause of the product damage to the sterile sleeve was not determined. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33565 | 1220648-2023-04144 | 17957734 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17957734&pc=OZD | 02/08/2022 | 10/18/2023 |  | 10/18/2023 |  |  | 02/08/2022 | 10/18/2023 |  |  | 09/30/2023 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022140779 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility |  |  |  | 58 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/08/2022 |  |  |  |  | The user facility reported that a 58-year-old male patient was implanted with the impella 5.5 for mechanical circulatory support. on the second day of support, a fluid leak was discovered at the purge sidearm. a purge bypass was performed to resolve the issue. no further information was provided. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge leak was completed. the device was not returned for evaluation. the clinical information was reviewed. the root cause of the purge leak was not determined as neither the product nor sufficient clinical information was provided. this report is being filed as part of a retrospective review of historical records.
33566 | 1220648-2023-04145 | 17957747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17957747&pc=OZD | 02/05/2023 | 10/18/2023 |  | 10/18/2023 |  |  | 02/05/2023 | 10/18/2023 |  |  | 07/31/2024 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022155779 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/05/2023 |  |  |  |  | The user facility reported that a 36-year-old male patient implanted with the impella 5.5 for mechanical circulatory support developed thrombus. once the device was removed it was noted to have a large clot coming from the outlet cage. the device was explanted, and a heartmate was placed without issue. there was no reported harm to the patient. | [5 empty] | The investigation into the reported thrombus was completed. the device was returned for evaluation. based on the available information and analysis the root cause of the thrombus on the pump upon explant was related to patient condition. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33567 | 1220648-2023-04147 | 17957833 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17957833&pc=OZD | 02/19/2022 | 10/18/2023 |  | 10/18/2023 |  |  | 02/19/2022 | 10/18/2023 |  |  | 10/31/2023 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022157646 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/19/2022 |  |  |  |  | The user facility reported a 57-year-old male patient was implanted with an impella 5.5 device for mechanical circulatory support. during support, the staff noted the purge sidearm had fluid leaking out during manipulation of the patient's automated impella controller (aic). heparin was in the purge. the pump was subsequently explanted that same day. | [5 empty] | The investigation into the reported purge leak was completed. the device was not returned for evaluation. the clinical information was reviewed. the root cause of the purge leak was not determined as neither the product nor sufficient clinical information was provided. per the ifu: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
33568 | 1220648-2023-04146 | 17957834 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17957834&pc=OZD | 10/28/2022 | 10/18/2023 |  | 10/18/2023 |  |  | 10/28/2022 | 10/18/2023 |  |  | 06/30/2024 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023230439 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/28/2022 |  |  |  |  | The user facility reported that a 30-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced access site bleeding in the axillary pocket. the patient was subsequently transfused 1 unit of packed red blood cells (prbc) and a jackson pratt (jp) drain was placed prior to closing pocket. additionally, the patient was administered protamine and heparin was stopped for bleeding management. it was reported that the bleeding was not attributed to impella. | [5 empty] | The investigation into the reported access site bleeding was completed. the device was not returned for evaluation. no data logs were returned and there were no logs available. based on the available information, the root cause of the access site bleeding was most likely related to patient condition related. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33569 | 1220648-2023-04134 | 17957862 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17957862&pc=OZD | 11/21/2022 | 10/18/2023 |  | 10/18/2023 |  |  | 11/21/2022 | 10/18/2023 |  |  | 04/30/2024 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023212679 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/21/2022 |  |  |  |  | The user facility reported that a 46-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced a pump stop/failure to start. the device would not start despite multiple attempts and exchanging bthe automated impella controller (aic). another device was subsequently implanted without issue. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported pump stop/ unable to start was completed. the device was returned for evaluation. functional testing of the pump was able to duplicate the high motor current pump stop alarm and controller failure alarm. ct scan of the motor revealed a motor cable line disconnect from the motor pcb with insufficient molding epoxy observed around the soldering area. the root cause of the pump unable to start was insufficient delo on the motor pcb motor cable solder joints resulting in an open electrical line. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33570 | 1220648-2023-04148 | 17957934 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17957934&pc=OZD | 02/22/2023 | 10/18/2023 |  | 10/18/2023 |  |  | 02/22/2023 | 10/18/2023 |  |  | 11/30/2024 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023277505 |  | 1000100 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/22/2023 |  |  |  |  | The user facility reported that a 64-year-old male patient implanted with the impella 5.5 for mechanical circulatory support had a "tear" in the impella catheter near the red handle. the wiring within the catheter was exposed at the tear. to avoid any potential performance issue, the decision was made to replace the pump with another impella 5.5. there was no reported harm to the patient. | [5 empty] | The investigation into the reported sterile sleeve damage was completed. the device was not returned for evaluation. based on the available information, the root cause of the catheter damage was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33571 | 1220648-2023-04137 | 17957946 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17957946&pc=OZD | 01/20/2023 | 09/20/2024 |  | 10/18/2023 |  |  | 01/20/2023 | 10/18/2023 |  |  | 10/31/2024 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023274497 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Material Integrity Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/20/2023 |  |  |  |  | The user facility reported that a 67-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced damage to the sterile after placement of the device and weaning off the intra-aortic balloon pump (iabp). the sterile sleeve was noted to have completely detached from the device locking mechanism. a suture and tegaderm were used to secure the sleeve back onto the locking mechanism. of note, the patient was reported to have tortuosity making placement difficult. there was no reported harm to the patient. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The investigation into the reported pump damage was completed. without the returned product, the cause of the damaged sterile sleeve could not be determined. this report is being filed as part of a retrospective review of historical records. the failure mode will be monitored and trended. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type.
33572 | 1220648-2023-04149 | 17957961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17957961&pc=OZD | 02/20/2022 | 10/18/2023 |  | 10/18/2023 |  |  | 02/20/2022 | 10/18/2023 |  |  | 09/30/2023 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022146604 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Suction Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/20/2022 |  |  |  |  | The user facility reported that a 64-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced low pump flows and continuous suction immediately after start-up. the pump was explanted as a result and upon explant a clot was seen in the cannula. there was no reported harm to the patient. | [5 empty] | The investigation into the reported low pump flows was completed. the device was not returned for evaluation. based on the information available for clinical review, the root cause of the low pump flows was most likely ingested biomaterial. this report is being filed as part of a retrospective review of historical records.
33573 | 1220648-2023-04139 | 17957999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17957999&pc=OZD | 12/04/2022 | 10/18/2023 |  | 10/18/2023 |  |  | 12/04/2022 | 10/18/2023 |  |  | 07/31/2024 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023251500 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/04/2022 |  |  |  |  | The user facility reported that a 57-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced sterile sleeve damage. on the 29th day of support staff observed that the repositioning sleeve was torn during pump repositioning. blood cultures were positive from an undetermined source. the catheter was sutured at 53 and 54 cm. the plan was to wean the patient and explant due to repeat positive blood cultures. of note, the positive blood cultures were not attributed to the device, or the sleeve being torn. no further information was provided. | [5 empty] | The investigation into the reported damage was completed. the device was not returned for evaluation. based on the provided clinical information, the cause of the product damage was most likely a damaged sterile sleeve. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33574 | 1220648-2023-04150 | 17958021 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17958021&pc=OZD | 02/27/2023 | 10/18/2023 |  | 10/18/2023 |  |  | 02/27/2023 | 10/18/2023 |  |  | 12/31/2024 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023284001 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem | [5 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/27/2023 |  |  |  |  | The user facility reported that a 59-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced low pump flows after placing the device. the device was subsequently explanted due to suspected clot which was confirmed upon explant. the physician elected not to place another device. there were no reported of harm to the patient. | [5 empty] | The investigation into the reported low pump flows was completed. the device was not returned for evaluation. based on the available information, the root cause of the low pump flow was most likely biomaterial since a clot was observed on inlet after explant. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33575 | 1220648-2023-04151 | 17958080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17958080&pc=OZD | 02/22/2022 | 10/18/2023 |  | 10/18/2023 |  |  | 02/22/2022 | 10/18/2023 |  |  | 07/31/2023 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022121442 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/22/2022 |  |  |  |  | The user facility reported that a 59-year-old male patient was implanted with the impella 5.5 for mechanical circulatory support. during support, the sterile sleeve of the device disconnected when the patient was moving from the bed to a chair, and while ambulating. the constant movement required periodic repositioning of the device which put stress on the connection between the locking mechanism and the sleeve causing the separation. the physician was able to push the sleeve back up and twist the tuohy-borst so that the sterile sleeve was held in the proper place; however, there appeared to be contamination of the sleeve. no further information was provided. there was no reported harm to the patient. | [5 empty] | The investigation into the reported sterile sleeve damage was completed. based on the available information and no device return, the root cause of the sterile sleeve damage was undetermined. this report is being filed as part of a retrospective review of historical records.
33576 | 1220648-2023-04152 | 17958217 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17958217&pc=OZD | 03/04/2023 | 10/18/2023 |  | 10/18/2023 |  |  | 03/04/2023 | 10/18/2023 |  |  | 10/31/2024 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023266024 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/04/2023 |  |  |  |  | The user facility reported that a 62-year-old female patient implanted with the impella 5.5 for mechanical circulatory support experienced positioning issues. the impella 5.5 catheter could not be delivered over the 0.018¿ impella wire so the team had to use the 0.027¿ wire. finally, and with much difficulty, after thirty minutes the aortic valve was successfully crossed. then, the physician inadvertently pulled the impella back into the aorta and was unable to get it back across the valve. as a result, the device was removed, and a different physician placed a new impella cp via the groin. no further information was provided. there was no harm reported to the patient. | [5 empty] | The investigation into the reported pump positioning issue was completed. the device was returned for evaluation. visual inspection of the pump revealed a significant cannula kink near the outflow cage and a significant catheter kink near the motor housing. after reviewing the clinical information, it was determined that the root cause of the positioning issues was wireless re-access; the impella device is not indicated for wireless re-access per ifu. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33577 | 1220648-2023-04154 | 17958301 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17958301&pc=OZD | 03/08/2023 | 10/18/2023 |  | 10/18/2023 |  |  | 03/08/2023 | 10/18/2023 |  |  | 12/31/2024 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 |  |  | ABIOMED, INC. |  |  | US |  | 2024293946 |  | 1000100 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Malfunction | No | Yes | Material Puncture/Hole | Activation, Positioning or Separation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/08/2023 |  |  |  |  | The user facility reported that a 47-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced positioning issues. the left ventricle (lv) signal was reported to be choppy. an echo showed that the device inlet was trapped in a mitral valve leaflet. while positioning it, the device got pulled back into the aorta and the sterile sleeve ripped. the pump was explanted and replaced. there was no harm reported to the patient. | [5 empty] | The investigation into the reported pump position issue was completed. the device was not returned for evaluation. the analysis of the data logs revealed that there are no positioning related alarms seen on the day of the reported positioning issue. based on the available information, the root cause of the positioning issue was most likely use related based on the provided clinical information. the cause of the product damage was a damaged sterile sleeve likely caused by excessive force. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33578 | 1220648-2023-04153 | 17958408 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17958408&pc=OZD | 01/28/2022 | 10/18/2023 |  | 10/18/2023 |  |  | 01/28/2022 | 10/18/2023 |  |  | 10/31/2023 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022155800 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/28/2022 |  |  |  |  | The user facility reported that a 62-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced low pump flows due to a clot. the pump was subsequently exchanged without issue. of note, the patient's pre-operative act was 218. 3000 units of heparin were administered and the act after insertion increased to 287. there was no reported harm to the patient. | [5 empty] | The investigation into the reported low pump flows was completed. the device was returned for evaluation. evaluation of the returned pump found a clot around the impeller. the root cause of the low flows was ingested biomaterial. this report is being filed as part of a retrospective review of historical records.
33579 | 1220648-2023-04155 | 17958437 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17958437&pc=OZD | 03/01/2022 | 10/18/2023 |  | 10/18/2023 |  |  | 03/01/2022 | 10/18/2023 |  |  | 08/31/2023 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022133321 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/01/2022 |  |  |  |  | The user facility reported that a 37-year-old male patient was implanted with the impella 5.5 for mechanical circulatory support. one day into support, purge fluid was seen leaking out of the catheter/ red sidearm. hospital staff monitored the purge flow, pressure and motor current. support continued for 2 more weeks. no further information was provided. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge leak was completed. the device was not returned for evaluation. the clinical information was reviewed. the root cause of the purge leak was not determined as neither the product nor sufficient clinical information was provided. per the ifu: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
33580 | 1220648-2023-04156 | 17958448 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17958448&pc=OZD | 03/13/2023 | 10/18/2023 |  | 10/18/2023 |  |  | 03/13/2023 | 10/18/2023 |  |  | 12/31/2024 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 |  |  | ABIOMED, INC. |  |  | US |  | 2024294372 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Open or Close | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/13/2023 |  |  |  |  | The user facility reported that a 54-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced positioning issues. the blue lock on the catheter was unable to unlock. no other issues were noted, and the pump was in good position. no further information was reported. there was no harm reported to the patient. | [5 empty] | The investigation into the reported pump position issue was completed. the device was not returned for evaluation. based on the available information, the root cause of inability to unlock the cap lock could not be determined. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33581 | 1220648-2023-04157 | 17958458 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17958458&pc=OZD | 03/18/2022 | 09/19/2024 |  | 10/18/2023 |  |  | 03/18/2022 | 10/18/2023 |  |  | 11/30/2023 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023169831 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | Insufficient Information | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/18/2022 |  |  |  |  | The user facility reported that a 59-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced high purge pressure and low purge flow two days into support. it was also noted that no alarm was triggered despite the high purge pressure. despite the noted issues, support continued until wean and explant. no further information was provided. there were no reports of harm to the patient. | The procedural outcome noted that the patient expired. medical safety review of the event noted the use of the device cannot conclusively be associated with the (death) outcome. |  |  |  |  | The investigation into the reported high purge pressure and low purge flow was completed. the device was not returned for evaluation. after review of the clinical information, it was determined that the cause of the high purge pressure was undetermined as neither the product nor the data logs were returned. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. as noted in the impella instructions for use: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ b.1, b.5 and h.6 revised due to medical safety review of file since the initial manufacturer device report number was submitted. b.7 revised as information was inadvertently omitted on the initial manufacturer device report number. d.4 revised unique identifier (udi) # as it was previously entered incorrectly on the initial manufacturer device report number. d.6a and d.6b revised from the initial submission of manufacturer device report number in accordance with updated procedures. f.6 and f.8 dates were reported on manufacturer device report number and should not have been. h.4 revised device manufacture date as it was previously entered incorrectly on the initial manufacturer device report number. h.10 additional manufacturer narrative instructions for use were inadvertently omitted from the previously submitted manufacturer device report number and were added.
33582 | 1220648-2023-04158 | 17958987 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17958987&pc=OZD | 03/22/2022 | 10/18/2023 |  | 10/18/2023 |  |  | 03/22/2022 | 10/18/2023 |  |  | 06/30/2023 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022097934 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/22/2022 |  |  |  |  | The user facility reported that a 58-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced positioning issues with their device. the device was pulled back into the aorta by the patient who had been recovering from extended sedation. an echo was used to confirm that the pump was in the aorta. the patient was stable so the pump was put on p-2 and left in the aorta until it could be removed. there was no harm reported to the patient. | [5 empty] | The investigation into the reported pump position issue was completed. the device was not returned for evaluation. based on the clinical information received, the root cause of the positioning issues was determined to be patient movement as the patient's hand got caught in the three-point fixation. this report is being filed as part of a retrospective review of historical records.
33583 | 1220648-2023-04160 | 17959078 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17959078&pc=OZD | 04/04/2022 | 10/18/2023 |  | 10/18/2023 |  |  | 04/04/2022 | 10/18/2023 |  |  | 12/31/2023 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023177746 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/04/2022 |  |  |  |  | The user facility reported that a 67-year-old male patient was implanted with the impella 5.5 for mechanical circulatory support. during support, purge fluid was observed to be leaking from the purge sidearm, but the exact location was not clearly specified. sodium bicarbonate was used in the purge solution. a purge sidearm bypass was performed to resolved the issue. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge leak was completed. the device was not returned for evaluation. the root cause of the purge leak was not determined since no product was returned. this report is being filed as part of a retrospective review of historical records.
33584 | 1220648-2023-04159 | 17959227 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17959227&pc=OZD | 03/31/2022 | 10/18/2023 |  | 10/18/2023 |  |  | 03/31/2022 | 10/18/2023 |  |  | 10/31/2023 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022153877 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/31/2022 |  |  |  |  | The user facility reported that a 66-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced positioning issues with their device. the device displaced into the aorta after repositioning of the pump was attempted. the impella was unable to re-cross the valve and had to be explanted as a result. there was no harm reported to the patient. | [5 empty] | The investigation into the reported pump positioning issue was completed. the device was returned for evaluation. there were no documented issues related to returned device found during the investigation. based on the available information, the root cause of the positioning issue was determined to be use related per the clinical details. this report is being filed as part of a retrospective review of historical records.
33585 | 1220648-2023-04161 | 17959239 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17959239&pc=OZD | 04/08/2022 | 10/18/2023 |  | 10/18/2023 |  |  | 04/08/2022 | 10/18/2023 |  |  | 01/31/2024 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023186113 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/08/2022 |  |  |  |  | The user facility reported that a 56-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced thrombus. as the physician was advancing the device, he pulled back on the 0.018 impella wire. an attempt was made to buddy wire the device into the left ventricle, but this was unsuccessful. the physician removed the pump to rewire. the wire was again, back in the left ventricle. when the physician inspected the pump prior to loading it on the 0.018 wire, he "thought he saw some clot" and elected to use another device. no further information was provided. | [5 empty] | The investigation into the reported thrombus was completed. the device was returned for evaluation. visual inspection was completed, and no biomaterial was observed. based on the available clinical information, the root cause of the issue was most likely biomaterial; however, there was no confirmed issue with the pump. this report is being filed as part of a retrospective review of historical records.
33586 | 1220648-2023-04162 | 17959295 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17959295&pc=OZD | 04/09/2022 | 10/18/2023 |  | 10/18/2023 |  |  | 04/09/2022 | 10/18/2023 |  |  | 07/31/2023 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022119986 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/09/2022 |  |  |  |  | The user facility reported that a 61-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced high purge pressure issues. it was noted that a purge solution of d5w with 0.45% sodium chloride was erroneously used initially during support. purge pressure started rising shortly after and purge flow decreased. the purge fluid was swapped out for d5 solution with heparin and the purge pressure returned to normal ranges. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported high purge pressure and low purge flow was completed. the device was not returned for evaluation. after review of the clinical information, the root cause was determined to be biomaterial buildup due to use of improper purge solution. this report is being filed as part of a retrospective review of historical records.
33587 | 1220648-2023-04163 | 17959328 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17959328&pc=OZD | 04/11/2022 | 10/18/2023 |  | 10/18/2023 |  |  | 04/11/2022 | 10/18/2023 |  |  | 12/31/2023 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023177803 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/11/2022 |  |  |  |  | The user facility reported that a 58-year-old male patient implanted with the impella 5.5 for mechanical circulatory support received suction alarms and lower flows (diastolic 0.0 l/min) with their device. the patient was transferred to another facility and flows improved to p-7 without suction. no further information was provided. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported low pump flows was completed. the device was not returned for evaluation. based on the available information, the root cause of the reported event is undetermined. this report is being filed as part of a retrospective review of historical records.
33588 | 1220648-2023-04055 | 17959366 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17959366&pc=OZD | 09/18/2023 | 02/26/2024 |  | 10/18/2023 |  |  | 09/18/2023 | 10/18/2023 |  |  | 04/30/2025 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024351029 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/18/2023 |  |  |  |  | The user facility reported an 83-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient had a small common femoral artery and that vessel size could have been the cause (per medical team) for limited flow down the leg where the impella was inserted. an external bypass was done to remediate this issue and palpable pulse returned to the patient¿s leg. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported limb ischemia has been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
33589 | 1220648-2023-04164 | 17959367 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17959367&pc=OZD | 04/13/2022 | 10/18/2023 |  | 10/18/2023 |  |  | 04/13/2022 | 10/18/2023 |  |  | 11/30/2023 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023169727 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Break | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/13/2022 |  |  |  |  | The user facility reported that 65-year-old male patient implanted with the impella 5.5 for mechanical circulatory support was undergoing a purge cassette change when the purge sidearm broke. a purge sidearm bypass was performed without issue. heparin was used in the purge solution throughout support. no further information was provided. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge leak was completed. the device was returned for evaluation. visual inspection of the device confirmed the purge sidearm was broken into two pieces between the pressure reservoir and the check valve. the root cause of the purge leak was sidearm bond damage due to the use of excessive force when exchanging the purge cassette. this report is being filed as part of a retrospective review of historical records.
33590 | 1220648-2023-04165 | 17959464 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17959464&pc=OZD | 04/18/2022 | 10/18/2023 |  | 10/18/2023 |  |  | 04/18/2022 | 10/18/2023 |  |  | 12/31/2023 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023171550 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/18/2022 |  |  |  |  | The user facility reported that a 53-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced purge flows <0.5 ml/h, and pressures up to 1590mmhg. tissue plasminogen activator (tpa) was administered for approximately 6 hours without success. the pump was subsequently replaced. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge system blocked was completed. the pump and purge cassette were returned for evaluation. analysis of data logs confirmed the reported event. the pump was visually inspected, and blood was observed in the sidearm. biomaterial was also observed in the motor. the root cause of the issue was biomaterial. this report is being filed as part of a retrospective review of historical records.
33591 | 1220648-2023-04166 | 17959611 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17959611&pc=OZD | 09/19/2023 | 12/01/2025 |  | 10/18/2023 |  |  | 09/19/2023 | 10/18/2023 |  |  |  |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1136132 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a battery failure on the automated impella controller (aic). the aic was removed and there was no reported patient harm. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | Corrections to the initial mfr report: g1 mdr reporting contact email. d2 device name has been updated. h6 type of investigation code added.
33592 | 1220648-2023-04058 | 17959687 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17959687&pc=OZD | 09/08/2022 | 10/16/2023 |  | 10/18/2023 |  |  | 09/08/2022 | 10/18/2023 |  |  | 06/30/2024 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023234456 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/08/2022 |  |  |  |  | The user facility reported a 50-year-old male patient was implanted with an impella 5.5 device for mechanical circulatory support. placement signal issues and low flow was observed directly after insertion. the pump was replaced successfully to resolve the noted issue. no harm done to the patient | [5 empty] | The product was returned and the optical sensor was confirmed to be operating within specification. the data logs confirm a high placement signal. the 5s parameters show the sensor was functioning. the root cause of the placement signal issue was most likely patient condition. this report is being filed as part of a retrospective review of historical records.
33593 | 1220648-2023-04059 | 17959690 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17959690&pc=OZD | 09/16/2023 | 10/16/2023 |  | 10/18/2023 |  |  | 09/16/2023 | 10/18/2023 |  |  | 07/31/2025 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024373503 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Power Up | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/16/2023 |  |  |  |  | The user facility reported a 74-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support for escalation of therapy. the impella 5.5 advanced over the wire into left ventricle (lv). pump was inadvertently started with wire in place. immediate impella stopped alarm occurred. pump removed. the patient was supported chemically while new 5.5 was prepped. new 5.5 placed without issue. or team educated while in room and will do follow-up education with md and staff to ensure all understand pump cannot be started while wire is in place. | [5 empty] | The investigation into the pump that did not start has been completed. impella 5.5 was returned for evaluation. as per clinical, pump was started with guidewire in place leading to pump stop alarm and the pump was explanted. visually inspected the return pump using borescope and found the guidewire marks to be present around the impeller. the pump was run with guidewire in place during the case causing interference with the impeller. no other abnormalities were found. the cause of the pump did not start issue was use related due to guidewire interference with impeller per field investigation and clinical review. the failure modes will be monitored and trended.
33594 | 1220648-2023-04056 | 17959691 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17959691&pc=OZD | 03/16/2023 | 10/16/2023 |  | 10/18/2023 |  |  | 03/16/2023 | 10/18/2023 |  |  | 11/30/2024 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2023283741 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/16/2023 |  |  |  |  | The user facility reported a 60-year-old male patient was implanted with an impella device for mechanical circulatory support. one day after replacing purge solution using the biocarbonate protocol, the purge pressure increased till 1062mmhg and purge flow decreased to 0.5ml/h. then the yellow luer started to leak from the joint with reservoir and the filter. soon after blood appeared in the filter. the pump was subsequently replaced as a result of the noted issue. there was no patient harm. | [5 empty] | The returned pump was visually inspected and blood ingress was observed in the filter and in a small portion of the purge cassette tubing next to the yellow luer. the retainer was removed, the sidearm was fluid pressure tested using a syringe and a leak from the reservoir weld seal was reproduced at a high pressure. the purge cassette was attempted to run through de-air but the tubing was blocked. the location of the tubing blockage was identified to be a knot tied in the tubing beyond the purge disc. the cause of the high purge pressure was most likely a kinked purge line in the purge cassette tubing. refer to (b)(4) for additional pattern details. the cause of the purge leak was an insufficient sidearm reservoir membrane radial seal.
33595 | 1220648-2023-04057 | 17959692 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17959692&pc=OZD | 09/08/2022 | 10/13/2025 |  | 10/18/2023 |  |  | 09/08/2022 | 10/18/2023 |  |  | 05/31/2024 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023223543 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/08/2022 |  |  |  |  | The user facility reported a 69-year-old male patient was implanted with an impella 5.5 device for mechanical circulatory support. the us complainant had a 5.5 failed delivery. the team observed native anatomic limitations. when the 5.5 failed delivery a cp was used instead, which also failed delivery. the case was aborted due to length of procedure and blood loss. one unit prbc was given in the cath lab to treat the noted blood loss. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
33596 | 1220648-2023-04167 | 17959719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17959719&pc=OZD | 04/27/2022 | 10/18/2023 |  | 10/18/2023 |  |  | 04/27/2022 | 10/18/2023 |  |  | 10/31/2023 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022156702 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/27/2022 |  |  |  |  | The user facility reported that a 71-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced a purge leak. a purge sidearm crack was noted when moving the patient to a different bed. a purge bypass technique was performed without issue. sodium bicarbonate was used in the purge. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge leak was completed. the device was not returned for evaluation. an image provided by the site indicated a leak at the yellow luer, which was likely a result of sodium bicarbonate weakening the plastic. the root cause of the purge leak was sidearm yellow luer damage as a result of bicarb. this report is being filed as part of a retrospective review of historical records.
33597 | 1220648-2023-04168 | 17959734 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17959734&pc=OZD | 03/19/2023 | 10/18/2023 |  | 10/18/2023 |  |  | 03/19/2023 | 10/18/2023 |  |  | 04/30/2023 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022076926 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Leak/Splash | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/19/2023 |  |  |  |  | The user facility reported that a 37-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced a purge leak and that the purge sidearm was cracked. a purge bypass was completed; however, the device was subsequently replaced following the bypass. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge leak was completed. the pump was returned with the bypass needle in place. the sidearm was not returned. based on the available information, the root cause of the purge leak was due to a leaking sidearm, but the exact root cause of the leak cannot be determined as the sidearm wasn't returned for review. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33598 | 1220648-2023-04169 | 17959791 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17959791&pc=OZD | 02/17/2023 | 10/18/2023 |  | 10/18/2023 |  |  | 02/17/2023 | 10/18/2023 |  |  | 06/30/2024 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023234681 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Contamination | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/17/2023 |  |  |  |  | The user facility reported that a 51-year-old male patient implanted with the impella 5.5 for mechanical circulatory support. the impella purge side-arm clear casing was contaminated with feces. the casing was opened and soiled area was cleaned. infectious disease physician cleared to continue impella support after cleaning site. there was no reported harm to the patient. | [5 empty] | The investigation into the reported positioning issues was completed. the device was not returned for evaluation. based on the available information, the root cause of the product damage, specifically contamination of the purge side-arm clear retainer with feces, was use related. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33599 | 1220648-2023-04170 | 17959911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17959911&pc=OZD | 03/20/2023 | 10/18/2023 |  | 10/18/2023 |  |  | 03/20/2023 | 10/18/2023 |  |  | 12/31/2024 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024292193 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/20/2023 |  |  |  |  | The user facility reported that a 67-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced a purge leak. while attempting to extubate the patient, he became agitated and completely pulled the purge line from the red impella plug. the alarm sounded due to the purge line being removed and the clinical staff was unable to correct it. the staff was informed that if and when the pump stopped, they would need to manage the patient with chemical support. the patient was subsequently taken to the operating room (or) to replace the existing pump with a new one. no further information was provided. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge leak was completed. the device was not returned for evaluation. based on the available information, the root cause of the purge leak was use related and related to the patient pulling the purge line from the red pump plug. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33600 | 1220648-2023-04171 | 17959978 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17959978&pc=OZD | 06/30/2023 | 10/18/2023 |  | 10/18/2023 |  |  | 06/30/2023 | 10/18/2023 |  |  | 04/30/2025 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024333093 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/30/2023 |  |  |  |  | The user facility reported a 61-year-old male patient was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the pump stopped while the patient was up walking around. the pump was restarted but when he stood up, it stopped again. the pump was removed several hours later for a scheduled heart transplant. | [5 empty] | The investigation for the reported pump stops has been completed. no product was returned for evaluation. as per clinical, pump stops seemed to have occurred due to patient movement. data logs were returned, and several 119 pump stop alarms were observed at the end of the case. the cause of the pump stop issue was an open electrical line due to patient movement based on the clinical and data log review. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
33601 | 1220648-2023-04172 | 17960028 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17960028&pc=OZD | 06/30/2023 | 10/18/2023 |  | 10/18/2023 |  |  | 06/30/2023 | 10/18/2023 |  |  | 03/31/2025 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024321650 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/30/2023 |  |  |  |  | The user facility reported 48yo male was implanted with an impella 5.5 device for mechanical circulatory support. the clinical team reported that the catheter anchor failed to translate along the catheter when the lock was disengaged, resulting in the inability to reposition the impella. the clinicians attempted to resolve the issue by using rotaglide lubricant, however it proved ineffective. the decision was then made to explant the pump. there was no patient harm reported. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was returned for evaluation. upon evaluation, significant kinking and bunching of the catheter was observed between the 48 and 50-cm marks. additionally, the catheter was bent at the 46-cm mark. the cath anchor functioned with no issues and was able to move along the catheter normally, except at the areas of the catheter damage. the top repo unit nub was displaced within its hole. in conclusion, it was determined that the cause of the positioning issues was a cath anchor use issue; excessive force was applied when manipulating the repo unit. this report is being filed as part of a retrospective review of historical records.
33602 | 1220648-2023-04173 | 17960447 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17960447&pc=OZD | 09/06/2022 | 10/18/2023 |  | 10/18/2023 |  |  | 09/06/2022 | 10/18/2023 |  |  | 05/31/2024 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023229020 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/06/2022 |  |  |  |  | The user facility reported a 55-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported a failure to deliver the device past the patient's mid-axillary stenosis. upon inspection of the device, it was noted that a kink was present where the catheter attaches to the cannula. an impella cp was subsequently used for patient support. no harm to the patient as a result of inability to deliver device. | [5 empty] | The device was returned and an evaluation was completed. a visual inspection verified that a kink was present on the catheter where the catheter attaches to the cannula. based on the provided information, it was noted that the cause for the reported kinking was believed to be a steep insertion angle during delivery. this report is being filed as part of a retrospective review of historical records.
33603 | 1220648-2023-04174 | 17961001 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17961001&pc=OZD | 12/09/2022 | 10/18/2023 |  | 10/18/2023 |  |  | 12/09/2022 | 10/18/2023 |  |  | 08/31/2024 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2023258084 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Advance | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/09/2022 |  |  |  |  | The user facility reported a 75-year-old female undergoing cardiac surgery was presented for attempted impella escalation for mechanical circulatory support. upon attempted delivery, resistance was met and the pump was unable to advance through the patient's vasculature. a deformation was eventually noted on the catheter shaft close to the motor housing and the decision was made to abort the escalation. support continued with the original pump in the femoral artery. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. based on the clinical account, the root cause of the delivery issue was patient condition, as an old graft was in place. this report is being filed as part of a retrospective review of historical records.
33604 | 1220648-2023-04175 | 17961086 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17961086&pc=OZD | 09/08/2023 | 10/09/2024 |  | 10/18/2023 |  |  | 09/21/2023 | 10/18/2023 |  |  | 07/31/2025 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024375562 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 32 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Nerve Damage | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/21/2023 |  |  |  |  | Per loqi registry communication: a 32-year-old male with cardiomyopathy was escalated to an impella 5.5 for mechanical circulatory support. the patient had to go back to the or for a washout of right axillary wound and peripheral nerve decompression due to the position of impella 5.5 causing neuropraxia to the patient's right arm. per documentation, the patient had a neurology consultation where they recommended a ct of head and neck. ct results no hematoma or injury to c5-6 nerve root. patient went to or for exploration of wound and no hematoma or nerve compression was found during this procedure. impella 5.5 was still repositioned during this time, but notes washout was "uneventful." | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related events are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation of vascular damage - peripheral vessel, peripheral nerve injury, and positioning issues has been completed. the impella device was not returned for evaluation. the most likely cause of the positioning issues was use related based on the clinical details regarding that the impella was placed in an improper position. the root cause of the vascular damage - peripheral vessels and peripheral nerve injury cannot be determined as insufficient clinical details were provided. the following have been reported incorrectly or missing from manufacturer device report. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. h6 codes 1942 and 3038 were reported incorrectly. new codes were added to health effect - clinical code, health effect - impact code, and component code.
33605 | 1220648-2023-04176 | 17961183 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17961183&pc=OZD | 04/01/2023 | 10/18/2023 |  | 10/18/2023 |  |  | 04/01/2023 | 10/18/2023 |  |  | 02/28/2023 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024308543 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/01/2023 |  |  |  |  | The user facility reported that a 67-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced positioning issues. the device was having difficulty getting any flow due to its position. transesophageal echocardiogram (tee) results showed the device was against the inferior wall and under a pap. while trying to reposition, the device popped back into the aorta and was unable to be advanced. a new device was placed. no further information was provided. there was no harm reported to the patient. | [5 empty] | The investigation into the reported pump position issue was completed. the device was not returned for evaluation. base on the clinical account the root cause of the issue was use related because impella popped back while attempting repositioning and was unable to be advanced. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33606 | 1220648-2023-04177 | 17961197 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17961197&pc=OZD | 12/22/2022 | 10/18/2023 |  | 10/18/2023 |  |  | 01/22/2023 | 10/18/2023 |  |  | 09/30/2024 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2023262100 |  | 0550-0008 |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/22/2023 |  |  |  |  | The user facility reported a 69-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. during support, it was noted that the purge filter for the pump became damaged. the pump was subsequently removed as a result of the noted issue. no harm done to the patient. | [5 empty] | The impella device was not returned by the customer and therefore, evaluation of the device was not possible. the cause of the purge leak was not determined since neither the product nor sufficient clinical information were provided. per the ifu: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
33607 | 1220648-2023-04178 | 17961222 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17961222&pc=OZD | 04/01/2023 | 10/18/2023 |  | 10/18/2023 |  |  | 04/01/2023 | 10/18/2023 |  |  | 11/30/2024 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023279014 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Difficult to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/01/2023 |  |  |  |  | The user facility reported that a 40-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced catheter anchor issues. the physician had a hard time initially advancing the catheter anchor and locking it over the catheter. the clinical representative explained the appropriate process to the physician. the catheter was eventually placed; however, with a lot of resistance. the clinical representative inspected the locking mechanism for damage but was unable to determine if it came that way or if the damage had been due to the physician's technique. there was no reported harm to the patient. | [5 empty] | The investigation into the reported catheter anchor issue was completed. the device was returned for evaluation. functional testing of the device passed. based on the available information, the root cause of the catheter anchor issue is due to a use issue as the catheter anchor was able to move along the catheter without any issues within benchtop testing at abiomed. there was no damage observed on the returned pump. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33608 | 1220648-2023-04179 | 17961302 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17961302&pc=OZD | 04/17/2023 | 10/18/2023 |  | 10/18/2023 |  |  | 04/17/2023 | 10/18/2023 |  |  | 03/31/2025 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024319202 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Power Up | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/17/2023 |  |  |  |  | The user facility reported that a 61-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced a pump did not start issue. after the device was inserted, positioned & wire removed, the flows attempted to initiate. a high, erratic motor current was noted then the pump stop error occurred. it was also reported that there was an impella defective alarm. multiple attempts to restart the pump were unsuccessful. while obtaining a second automated impella controller (aic), the physician noted improved left ventricular function and decided against placing a new pump. the surgery was completed surgery with minimal pressor support. there was no reported harm to the patient. | [5 empty] | The investigation into the reported pump did not start issue was completed. the device was returned for evaluation. analysis of the data logs revealed that following insertion, there was an immediate motor current (mc) spike, indicating an open line in all three motor phases. the pump was electrically tested and the open line was confirmed. based on the available information, the root cause of the pump being unable to start was an open line due to a tensile pull/excessive force on the catheter which stretched the motor cables to the point of breakage. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33609 | 1220648-2023-04180 | 17961331 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17961331&pc=OZD | 02/07/2023 | 10/18/2023 |  | 10/18/2023 |  |  | 02/07/2023 | 10/18/2023 |  |  | 02/29/2024 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023191196 |  | 1000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/07/2023 |  |  |  |  | The user facility reported a 54-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the pump shifted out of the left ventricle during impella support. due to the noted issue, the pump was removed and replaced with an impella cp pump. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. although it is unknown why the pump dislodged, the staff attempted to re-access the left ventricle (lv) wirelessly, which is a contraindication per ifu 0550-9002. as a result, the cause for the noted positioning issue was attributed to user error. this report is being filed as part of a retrospective review of historical records.
33610 | 1220648-2023-04181 | 17961346 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17961346&pc=OZD | 04/15/2023 | 10/18/2023 |  | 10/18/2023 |  |  | 04/15/2023 | 10/18/2023 |  |  |  |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023264103 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/15/2023 |  |  |  |  | The user facility reported that a 69-year-old female patient implanted with the impella 5.5 for mechanical circulatory support experienced a pump stop. the device was placed using the appropriate size graft. the team noted when cross clamping after direct placement to place the device into surgical mode on the automated impella controller (aic), the pump would not go to p-0, and allow for surgical mode to be an option. after multiple attempts to place into surgical mode, the device was successfully exchanged for a new one. there was no reported harm to the patient. | [5 empty] | The investigation into the reported pump stop was completed. the device was returned for evaluation. functional testing of the pump passed. upon further evaluation it was noted that the site had navigated to the ¿settings/service¿ menu, but surgical mode was never selected. based on the available information, the most likely root cause was use error. impella 5.5 ifu. "maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33611 | 1220648-2023-04182 | 17961453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17961453&pc=OZD | 04/10/2023 | 10/18/2023 |  | 10/18/2023 |  |  | 04/10/2023 | 10/18/2023 |  |  | 11/30/2024 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024286372 |  | 1000211 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/10/2023 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. during support, the patient developed heparin induced thrombocytopenia (hit). following diagnosis, heparin was removed from the purge fluid and the patient was switched to argatroban. although there was no notable malfunction of the pump, blood data suggested that a blood clot had formed at an unknown location. support continued with the implanted pump despite the condition. | [5 empty] | The device was returned for evaluation. upon visual inspection of the returned pump, a thrombus was confirmed on the impeller. no other abnormalities were observed. upon conclusion of the investigation, the cause of the thrombus was determined to be related to the patient condition based on the provided clinical information. this report is being filed as part of a retrospective review of historical records.
33612 | 1220648-2023-04183 | 17961921 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17961921&pc=OZD | 04/15/2023 | 10/18/2023 |  | 10/18/2023 |  |  | 04/15/2023 | 10/18/2023 |  |  | 12/31/2024 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2024291860 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 19 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/15/2023 |  |  |  |  | The user facility reported that a 19-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced a purge leak. the device began leaking through the purge reservoir/filter and low purge flow alarm continuously sounded. the purge cassette was changed without resolution. the staff identified that the source of the leaking was the purge cassette, a purge bypass was successfully completed with 19ga needle and y connected from the purge cassette. the purge bypass was secured to the arm board. no further information was provided. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge leak was completed. the pump and purge cassette were returned for evaluation. the pump was returned without all parts intact, they had cut a portion & a purge bypass needle was connected to the purge lumen. the catheter was also cut off below the motor housing. no other issues were found on the rest of the pump. based on the available information, the root cause of the purge leak was the damaged sidearm of impella. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33613 | 1220648-2023-04184 | 17961964 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17961964&pc=OZD | 04/18/2023 | 10/18/2023 |  | 10/18/2023 |  |  |  | 10/18/2023 |  |  | 12/31/2024 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024307429 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/18/2023 |  |  |  |  | The user facility reported that a 55-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced thrombus. clot observed in lv noted on echo. clinical staff observed clot attached to device at explant. of note, anticoagulation had not begun for 4-5 days. the patient was also diagnosed with hit (heparin induced thrombocytopenia). no further information was provided. | [5 empty] | The investigation into the reported thrombus was completed. the device was not returned for evaluation. based on clinical description, the patient was diagnosed with hit (heparin induced thrombocytopenia) resulting in a delay of using anticoagulation for the first 4-5 days. therefore, the root cause of the thrombus was most likely due to anticoagulation management issue. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33614 | 1220648-2023-04185 | 17962048 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17962048&pc=OZD | 04/24/2023 | 10/18/2023 |  | 10/18/2023 |  |  | 04/24/2023 | 10/18/2023 |  |  | 07/31/2024 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023254975 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/24/2023 |  |  |  |  | The user facility reported that a 60-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced a purge leak on day 22 of support. the next day, the automated impella controller (aic) alarmed for low purge pressure. the leak at purge sidearm was considerably large. the purge cassette was subsequently changed without resolution of the alarm. a purge arm bypass was performed which was successful in resolving the alarm and leak issues. of note, due to patient's abdominal surgeries and wounds from this hospital admission, stat locks were not placed in usual 3-point fixation across the patient¿s abdomen. instead, they were placed down his right abdomen and secured with 2 stat locks. a similar method was used after the purge bypass with arm board in place and patient education was provided regarding importance of keeping arm board from being pulled or tucked underneath patient. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge leak was completed. the device was not returned for evaluation. based on the available information, the root cause of the issue was a leak from the sidearm, but the exact location of the leak was unknown. the sidearm was reported to have been cleaned and placed with improper fixation. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33615 | 1220648-2023-04186 | 17962111 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17962111&pc=OZD | 08/17/2023 | 03/08/2024 |  | 10/18/2023 |  |  |  | 10/18/2023 |  |  | 05/31/2025 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024361093 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 62 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/18/2023 |  |  |  |  | A long-term outcome and quality indicator (loqi) impella registry notification was received 2023-09-18 indicating a patient enduring ventricular arrhythmia (svt and vt runs) during impella support. the patient had presented with acute myocardial infarction and cardiogenic shock for hemodynamic support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the ventricular arrhythmia (svt and vt runs) issue has been completed since the original report was filed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined.
33616 | 1220648-2023-04187 | 17962335 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17962335&pc=OZD | 08/25/2023 | 06/12/2024 |  | 10/18/2023 |  |  | 10/18/2023 | 10/18/2023 |  |  | 07/31/2025 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024370578 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Renal Impairment | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/18/2023 |  |  |  |  | Long-term outcome and quality indicator (loqi) impella registry notification received (b)(6) 2023 indicating this patient to have endured acute renal insufficiency and atrial arrhythmia with a possible relationship to the impella device: "nephrology diagnosed subject with acute renal insufficiency" noted along with "post operative a-fib treated with amiodarone bolus." the patient presented for hemodynamic support via the impella 5.5 device. | There is no association between impella use and outcome. |  |  |  |  | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions- including catheter position, pre-existing medical conditions, and small left ventricular volumes- may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the renal insufficiency and the vf/vt/af/at (atrial arrhythmia) issues has been completed since the original report was submitted. the device was not returned for investigation. the cause of the renal failure was not determined because no product was returned, and not enough clinical information was given. as the necessary information was not provided, the root cause of the vt/vf was not determined.
33617 | 1220648-2023-04188 | 17962365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17962365&pc=OZD | 09/19/2023 | 01/25/2024 |  | 10/18/2023 |  |  | 09/19/2023 | 10/18/2023 |  |  | 07/31/2025 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024379796 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the physician had difficulty removing impella 5.5. originally was being viewed under echo guidance. the physician requested fluoro imaging and was able to pull the catheter across the anastomosis safely upon explant impella was unraveled exposing electrical cable. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation into the removal issues has been completed since the original report was filed. the product was returned for investigation. there was clear sign of excessive force used leading to catheter damage and exposing of electrical cable. per the investigation and clinical information provided, the root cause of the removal issue was use related to excessive force.
33618 | 1220648-2023-04189 | 17962367 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17962367&pc=OZD | 09/19/2023 | 02/07/2025 |  | 10/18/2023 |  |  | 09/19/2023 | 10/18/2023 |  |  |  |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1504376 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the automated impella controller (aic) had a battery communication failure (red alarm). | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation for the console (aic) battery failure-red alarm has been completed. after reviewing the data logs, the reported battery failure issue was confirmed. there were numerous instances of battery failure alarms (transport connection blown/missing) (event set #360) observed as well as battery communication failure on the reported complaint date. the console was returned for evaluation. the battery failure issue was able to be reproduced. results of running the batttest utility revealed that cell 4 in battery1 had a voltage that was significantly less than the rest of the cells in the battery. lcd screen of battery one was blank. the reported battery failure issue was determined to be caused by internal battery cell imbalance. it is the root cause for both battery failure alarm and battery communication error alarm.
33619 | 1220648-2023-04190 | 17962368 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17962368&pc=OZD | 09/18/2023 | 01/23/2024 |  | 10/18/2023 |  |  | 09/19/2023 | 10/18/2023 |  |  | 08/31/2025 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024385000 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 85-year-old female noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the complainant reported the impella cp was removed and the patient continued on ecmo. gradually, over the next few days the patient's distal pulses became more faint and eventually disappeared. the patient was taken back to the or for a lower extremity fasciotomy and thrombectomy. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reversible limb ischemia has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. the root cause of the limb ischemia was patient condition related, and unknown relation to the impella because it occurred 3 days after removal. the failure mode will be monitored and trended.
33620 | 1220648-2023-04191 | 17962377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17962377&pc=OZD | 09/19/2023 | 01/31/2024 |  | 10/18/2023 |  |  | 09/19/2023 | 10/18/2023 |  |  | 08/31/2025 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024384514 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported had placed an impella cp for support of the stemi 59-year-old male patient. the patient was being transferred to tertiary center by helicopter/medflight when due to user error the cord was broken and the pump stopped. the patient was returned to the lab for removal and replacement of a new pump. the second pump was successfully placed. there was no lasting harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the reported pump stop has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible; only the clinical report was evaluated. data logs showed pump ran without issue during the case, then suddenly was disconnected from the console while running. based on clinical description & data analysis, the root cause of pump stop was due to use related.
33621 | 1220648-2023-04192 | 17962378 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17962378&pc=OZD | 09/19/2023 | 08/02/2024 |  | 10/18/2023 |  |  | 09/19/2023 | 10/18/2023 |  |  | 07/31/2025 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024387986 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/19/2023 |  |  |  |  | The user facility reported a 71-year-old male in anterior acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support prior to percutaneous coronary intervention. the support was successful, but there was an access site hematoma observed. the team made choice to explant the pump but leave the vessel accessed for possible repair. the repair was determined to be necessary when the vascular surgeon observed superficial femoral artery/profunda dissection. the vessel was treated and the patient survived. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ | The investigation of access site bleeding and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. instructions for use as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention: section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ d6a/d6b added h6 component code added the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-04192 f6/f8-should have been left blank.
33622 | 1220648-2023-04193 | 17962381 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17962381&pc=OZD | 09/19/2023 | 11/17/2025 |  | 10/18/2023 |  |  | 09/19/2023 | 10/18/2023 |  |  |  |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1597593 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 09/19/2023 |  |  |  |  | The abiomed field team was at the user facility for cath lab training. when the automated impella controller (aic) was set up, the aic alarmed "controller error" (yellow alarm). this aic had recently been in service for about three weeks on an impella 5.5 patient. the console had been sent with the patient to an outside facility for transfer and was recently returned. the aic will be returned for investigation. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | Corrections to the initial mfr report: g1: mdr reporting contact email. d2: device name has been updated. d4: model number has been updated. f6: and f8 should have been left blank. h6: type of investigation code added.
33623 | 1220648-2023-04194 | 17962383 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17962383&pc=OZD | 09/19/2023 | 02/21/2024 |  | 10/18/2023 |  |  | 09/19/2023 | 10/18/2023 |  |  | 04/30/2025 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024348662 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility had a 14fr sheath introducer placed for support and placement of the impella cp pump for a aortic regurgitation (ar)/aortic stenosis (as) 64-year-old male patient. the patient had a bleed from the access site. two units of blood were given to deliver back to the patient. additionally the leg had an antegrade sheath to provide flow to the limb. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major and the limb ischemia ¿ irreversible has been completed. the device was not returned for investigation. the cause of the bleeding was most likely congestive liver failure, patient condition related. the cause of the major and the limb ischemia was not determined.
33624 | 1220648-2023-04195 | 17962484 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17962484&pc=OZD | 09/18/2023 | 01/31/2024 |  | 10/18/2023 |  |  | 09/18/2023 | 10/18/2023 |  |  | 07/31/2025 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024379736 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/18/2023 |  |  |  |  | The user facility reported a 65-year-old female post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. on day 10 of the support, the high purge pressure alarms prompted the team to troubleshoot by adding tissue plasminogen activator to the purge. the tpa was added but, later the 5.5 pump stopped. the pump was explanted as a result of the failure. there was no noted harm due to the reported issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the reported high purge pressure and pump stop has been completed since the original report was filed. the product and data logs were not returned for review. the root cause of pump stop and high purge pressure were unable to be determined as limited clinical details were provided and no product was returned for review.
33625 | 1220648-2023-04196 | 17962491 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17962491&pc=OZD | 09/18/2023 | 02/05/2024 |  | 10/18/2023 |  |  | 09/18/2023 | 10/18/2023 |  |  | 04/30/2025 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024353376 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 72-year-old male patient was implanted with impella cp prior to percutaneous coronary intervention (pci). the physician left the peel away sheath in the right femoral artery as the impella cp was going to be removed in a short time. the patient had peripheral artery disease (pad) in the right leg that limited the outflow of the antegrade sheath, and the antegrade sheath clotted off which made the right femoral retrograde sheath clotted off. the impella was successfully removed per plan. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the thrombus issue has been completed. the product was not returned for benchtop investigation. data showed mc spikes followed by low flow when pump started and remained to the rest of the case that triggered auto suction response and suction alarms. based on clinical description & data review, the root cause of low flow was most likely due to biomaterial ingestion.
33626 | 1220648-2023-04197 | 17962514 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17962514&pc=OZD | 09/18/2023 | 01/31/2024 |  | 10/18/2023 |  |  | 09/18/2023 | 10/18/2023 |  |  | 03/31/2025 |  | 10/18/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024341756 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a 41-year-old male patient was implanted with an impella cp pump for mechanical circulatory support. during support, persistent suction was noted and the decision was made to replace the pump in case a clot had been ingested. a new impella was implanted for ongoing support. there was no patient harm due to the noted issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow issue has been completed. the pump was not returned for investigation. the data logs showed suction alarm during the entire case run with low pump flow. there were no signs related to improper position or biomaterial interaction. as per clinical details, the pump was explanted and replaced with another cp pump, yet the suction alarms persisted. the suction alarms resolved with the second pump after the physician administered fluids to the patient. the cause of the low pump flow issue was most likely patient condition related to low blood volume since the pump flow was restored after the given fluid volume, based on given clinical details.
33627 | 1220648-2023-04067 | 17949081 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17949081&pc=OZD | 10/29/2021 | 10/17/2023 |  | 10/17/2023 |  |  | 10/29/2021 | 10/17/2023 |  |  | 07/31/2023 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022113226 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/29/2021 |  |  |  |  | The user facility reported that a 68-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced pump positioning issues. the device came fully out of the left ventricle, and it was noted that the centimeter mark on the catheter did not change. the patient was taken to the operating room (or) where the surgeon attempted to reposition multiple times via transesophageal echocardiogram (tee) unsuccessfully. a new device was subsequently implanted. additional information indicated that the tuohy-borst was properly locked, and the 3-point fixation technique was utilized; however, it was also noted that the patient was ambulatory and walking. there was no reported harm to the patient. | [5 empty] | The investigation into the reported positioning issue was completed. the device was not returned for evaluation. based on the information that the patient was ambulatory at the time, the root cause of the reported positioning issues was most likely related to patient movement. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33628 | 1220648-2023-04068 | 17949258 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17949258&pc=OZD | 10/22/2021 | 10/17/2023 |  | 10/17/2023 |  |  | 10/22/2021 | 10/17/2023 |  |  | 07/31/2023 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022115214 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/22/2021 |  |  |  |  | The user facility reported that a 44-year-old female patient implanted with the impella 5.5 for mechanical circulatory was decannulated to be taken off extracorporeal membrane oxygenation (ecmo). approximately an hour later, the patient went into hypoxic arrest. cpr was performed and the patient was recannulated for ecmo. following this event, the impella 5.5 flows were unable to increase above 2l/min and continuous suction and positioning alarms were triggered. the patient had a patent foramen ovale (pfo) and was also having poor right heart function. the physician suspected that a clot had traveled through the pfo and was ingested by the device. the patient's volume and pump position were checked and found to be adequate. the pump continued to sufficiently support the patient and was left implanted until the case concluded. no further information was provided. | [5 empty] | The investigation into the reported low pump flow was completed. the device was not returned for evaluation. analysis of the data logs confirmed the drop in flow accompanied with a loss of pulsatility in the motor current. multiple suction and positioning alarms were triggered from the time of the hypoxic arrest to the end of the case. based on the available information, the most likely root cause of the suction and low flow was determined to be a clot in the cannula or outflow restricting blood flow to the impeller. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33629 | 1220648-2023-04069 | 17949313 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17949313&pc=OZD | 05/04/2022 | 09/22/2024 |  | 10/17/2023 |  |  | 05/04/2022 | 10/17/2023 |  |  | 01/31/2024 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023181451 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  |  | White | Injury | Yes | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/04/2022 |  |  |  |  | The user facility reported that a 68-year-old female patient implanted with the impella 5.5 for mechanical circulatory support experienced high purge pressure within the first hour of support. no alarms sounded. troubleshooting was performed however the specifics were not provided. the high purge pressure eventually resolved, and the pump was able to keep supporting the patient. there were no reports of harm to the patient. | The procedural outcome noted that the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The investigation into the reported high purge pressure was completed. the device was not returned for evaluation. the root cause of the high purge pressure was unable to be determined as no product or data was returned for analysis. data for this pump is not available. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-04069 was provided in sections b1, b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding.
33630 | 1220648-2023-04070 | 17949330 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17949330&pc=OZD | 05/08/2022 | 10/17/2023 |  | 10/17/2023 |  |  | 05/08/2022 | 10/17/2023 |  |  | 12/31/2023 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023174130 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/08/2022 |  |  |  |  | The user facility reported that a 54-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced a purge leak due to a purge sidearm crack. a purge bypass was performed without issue and the pump was replaced later that day. the sidearm was discarded so the part of the sidearm that resulted in the leak was unknown. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge sidearm leak issue was completed. the pump was returned for evaluation, but the sidearm was discarded at the site. based on the available information, the root cause of the purge sidearm leak was not determined since sufficient product was not returned. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33631 | 1220648-2023-04071 | 17949337 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17949337&pc=OZD | 11/12/2021 | 10/17/2023 |  | 10/17/2023 |  |  | 11/12/2021 | 10/17/2023 |  |  | 07/31/2023 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022113550 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/12/2021 |  |  |  |  | The user facility reported that a 38-year-old male implanted with the impella 5.5 pump experienced a purge system blocked alarm. the purge flow was reportedly less than 0.5 ml/hr and purge pressure was 1075mmhg with a mean motor current of 195 on p-4 despite the administration of tissue plasminogen activator (tpa). additional information indicated that the purge solution had sodium bicarbonate 25meq/l with d5w. the pump was ultimately replaced with a new one. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge system blocked alarm was completed. the device and data logs were returned for evaluation. analysis of the data logs confirmed a sharp increase in purge pressure and purge system blocked alarms. visual inspection of the pump revealed blood under the impeller but no catheter kinks. functional testing of the pump revealed normal functionality. the root cause of the sharp increase in purge pressure and purge system blocked alarms was potentially a kinked purge line, however this could not be confirmed. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33632 | 1220648-2023-04072 | 17949362 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17949362&pc=OZD | 05/09/2022 | 10/17/2023 |  | 10/17/2023 |  |  | 05/09/2022 | 10/17/2023 |  |  | 08/31/2023 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022125362 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/09/2022 |  |  |  |  | The user facility reported that a 36-year-old male patient implanted with the impella 5.5 for mechanical circulatory support developed thrombus. upon removal of the device, it was found to have what appears to be adhesion on one window of the inlet and on the propeller inside the outlet. no further information was provided. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported thrombus was completed. the device was returned for evaluation. visual inspection of the pump found biomaterial on two struts of the inlet cage and in the outflow cage. functional testing of the device passed. based on the available information, the root cause of the reported event is biomaterial/patient condition. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33633 | 1220648-2023-04073 | 17949399 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17949399&pc=OZD | 05/11/2022 | 10/17/2023 |  | 10/17/2023 |  |  | 05/11/2022 | 10/17/2023 |  |  | 01/31/2024 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023184021 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/11/2022 |  |  |  |  | The user facility reported that a 52-year-old female patient implanted with the impella 5.5 for mechanical circulatory support experienced a purge leak and low purge pressure per the automated impella controller (aic). the clinical staff then noticed the purge filter was cracked. no further information was provided. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge leak and low purge pressure issue was completed. the device was not returned for evaluation. based on the available information, the root cause of the low purge pressure was the cracked purge filter. the root cause of the purge leak was due to sidearm filter damage. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33634 | 1220648-2023-04074 | 17949469 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17949469&pc=OZD | 11/16/2021 | 10/17/2023 |  | 10/17/2023 |  |  | 11/16/2021 | 10/17/2023 |  |  | 08/31/2023 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022135024 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/16/2021 |  |  |  |  | The user facility reported that a 57-year-old male implanted with the impella 5.5 pump experienced a purge flow of less than 0.5 ml/hr. staff attempted to troubleshoot without resolution. the purge bag was changed and de-aired several times. the patient was subsequently taken for pump exchange due to concern for pump stop. post exchange acts were reported to be normal with 25 u/ml of heparin in the purge. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported low purge flow and high purge pressure was completed. the pump and data logs were returned for evaluation. analysis of the data logs revealed normal purge flow through pump delivery; however, when the pump was activated, purge flow dropped. a high motor current event also occurred at the activation of impella. functional testing of the pump reproduced the high purge pressure and visual inspection revealed biomaterial in the purge gap. the biomaterial was most likely ingested upon pump activation correlating with the high motor current event. the root cause of the high purge pressure was ingested biomaterial occluding the purge gap. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33635 | 1220648-2023-04075 | 17949495 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17949495&pc=OZD | 05/13/2022 | 10/17/2023 |  | 10/17/2023 |  |  | 05/13/2022 | 10/17/2023 |  |  | 02/29/2024 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023193677 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/13/2022 |  |  |  |  | The user facility reported that a 62-year-old female patient implanted with the impella 5.5 for mechanical circulatory support experienced a purge leak. on the 9th day of support a purge sidearm leak was reported. the location of the leak was the yellow luer. sodium bicarbonate was used in the purge solution. purge bypass instructions were provided. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge leak was completed. the device was not returned for evaluation. the root cause of the purge leak was due to a crack in the yellow luer. this is most likely a result of the use of sodium bicarbonate. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended. per the ifu: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol."
33636 | 1220648-2023-04107 | 17949620 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17949620&pc=OZD | 07/26/2022 | 10/17/2023 |  | 10/17/2023 |  |  | 07/26/2022 | 10/17/2023 |  |  | 02/29/2024 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023196465 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/26/2022 |  |  |  |  | The user facility reported that a 57-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced a purge leak. the device had been in place for 54 days and developed a crack. although a crack was noted, the location was unknown. a new device placed. no further details were provided. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge leak was completed. the device was not returned for evaluation. based on the available information and no product returned for evaluation, the root cause for the reported purge leak was undetermined. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33637 | 1220648-2023-04108 | 17949641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17949641&pc=OZD | 08/02/2022 | 10/17/2023 |  | 10/17/2023 |  |  | 08/02/2022 | 10/17/2023 |  |  | 03/31/2024 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023203546 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/02/2022 |  |  |  |  | The user facility reported that a 46-year-old female patient implanted with the impella 5.5 for mechanical circulatory support experienced a purge leak. during a routine purge cassette change, the yellow luer was noted to be cracked. a purge bypass was subsequently performed with assistance from clinical support. of note, sodium bicarbonate had been used in the purge. the patient was successfully weaned off support two days after the purge bypass and the pump was discarded. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge leak was completed. the device was not returned for evaluation. based on the clinical information provided, the root cause of the purge leak was most likely sidearm yellow luer damage. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33638 | 1220648-2023-04105 | 17949674 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17949674&pc=OZD | 09/20/2023 | 09/11/2024 |  | 10/17/2023 |  |  | 09/20/2023 | 10/17/2023 |  |  |  |  | 10/17/2023 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1666954 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Alarm System | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 09/20/2023 |  |  |  |  | The user facility reported a 73-year-old male in st elevated myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. during support, the automated impella controller (aic) had a "controller error" alarm. the aic was exchanged without difficulty. there was no patient harm. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console (aic) yellow alarm has been completed. data logs confirmed the reported failure mode given there was a controller error alarm (opm comm crc error ¿ event set # 1045) triggered during the clinical procedure. real time (rt) logs confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) are represented as -16384 values due to the crc error. the console was returned for evaluation. during functional evaluation, an optical bench fw communication protocol anomaly occurred, resulting in misalignment of data communication packets received by epc the controller error and loss of placement signal. the aic sw operated as designed. added- h6 impact code 4629 d2-corrected common device name. d4-corrected model number. f6/f8 should have been left blank in the initial report.
33639 | 1220648-2023-04103 | 17949687 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17949687&pc=OZD | 09/17/2023 | 04/29/2024 |  | 10/17/2023 |  |  | 09/17/2023 | 10/17/2023 |  |  | 02/28/2025 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024324108 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Insufficient Flow or Under Infusion | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the impella cp's flows were not adequate on max support due to high systemic vascular resistance. furthermore, the patient began to have limb ischemia on support. milrinone was added and the patient did regain pulsatility with initiation of the milrinone. additionally, the patient was experiencing hemolysis. plasma free hemoglobin was elevated with dark red blood in their foley/urine output. the impella cp was removed and the replaced with the impella 5.5. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported issues has been completed since the original report was submitted. the device was not returned for investigation. the cause of the hemolysis issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the low pump flow issue was most likely due to the patient¿s condition as max flows were not adequate due to the patient's high svr. as all the necessary information and returned device were not provided, the root cause of the limb ischemia and the vt/vf was not determined. adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ to conform to updated procedures, sections d6 and h10 have been revised with updated information.
33640 | 1220648-2023-04102 | 17949777 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17949777&pc=OZD | 09/17/2023 | 02/21/2024 |  | 10/17/2023 |  |  | 09/17/2023 | 10/17/2023 |  |  | 06/30/2025 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024373882 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp stopped alarm occurred with rapid spike in motor current. attempts were made to restart the pump multiple times without success. systemic heparin was held for supratherapeutic act and heparin 50 u/ml remained in purge. the pump was ultimately replaced as patient became unstable off support. the physician replacing the pump noted a large clot stuck on pump inlet. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop and the thrombus issues has been completed since the original report was filed. the device was returned and tested after arriving in fs. visually inspected and biomaterial observed in the inlet, pigtail of the pump. the cause of the pump stop was due to bearing wear failure within design life per field investigation and data log review. however, the cause of the thrombus issue was not determined as the necessary clinical information was not provided.
33641 | 1220648-2023-04101 | 17949798 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17949798&pc=OZD | 12/01/2021 | 10/17/2023 |  | 10/17/2023 |  |  | 12/01/2021 | 10/17/2023 |  |  | 07/31/2023 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022117716 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/01/2021 |  |  |  |  | The user facility reported that a 63-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced low pump flows. the physician did not believe the device was accurately unloading and calculating the flow rate. no further information was provided. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported low pump flows was completed. the pump and purge cassette were returned for evaluation. visual inspection of the pump revealed a large amount of biomaterial around the impeller in the cannula. following the investigation, the biomaterial was determined to be the root cause of the pump not unloading properly and lower flow than expected flows. this report is being filed as part of a retrospective review of historical records.
33642 | 1220648-2023-04106 | 17949888 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17949888&pc=OZD | 12/12/2021 | 10/17/2023 |  | 10/17/2023 |  |  | 12/12/2021 | 10/17/2023 |  |  | 10/31/2022 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2021016460 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/12/2021 |  |  |  |  | The user facility reported that a 51-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced ectopy. the physician attempted to reposition the device with echo guidance but was unable to pull it back across the valve. the device was unable to re-cross the valve without a wire. no further information was provided. there was no harm reported to the patient. | [5 empty] | The investigation into the reported pump positioning issue was completed. the device was not returned for evaluation. the clinical information was reviewed and based on this information, the root cause of the positioning issue was determined to be user related coinciding with wireless re-access attempts. wireless re-access is not permitted per ifu 0550-9002. this report is being filed as part of a retrospective review of historical records.
33643 | 1220648-2023-04076 | 17949911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17949911&pc=OZD | 05/26/2022 | 10/17/2023 |  | 10/17/2023 |  |  | 05/26/2022 | 10/17/2023 |  |  | 02/29/2024 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023197391 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/26/2022 |  |  |  |  | The user facility reported that a 54-year-old male patient implanted with the impella 5.5 for mechanical circulatory support. at the explant after 9 days, there was an observation made of thrombus material at the pump motor. no further information was provided. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported thrombus was completed. the device was returned for evaluation. visual inspection of the pump found no biomaterial in the cannula or on the inlet cage. a picture of the inflow cage was provided by the field and showed no evidence of biomaterial. based on the analysis, there was no issue with the pump. the device functioned/operated to specification. the root cause of the reported event is undetermined. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33644 | 1220648-2023-04094 | 17950134 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17950134&pc=OZD | 09/18/2023 | 11/03/2025 |  | 10/17/2023 |  |  | 09/18/2023 | 10/17/2023 |  |  |  |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1314172 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | Optical Problem | Noise, Audible |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the fan is making a loud noise and aic is giving an optical error. there was no reported patient harm. this malfunction was noted by the biomedical department. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: console log only contained entries from (b)(6) 2023 (reported day of event was (b)(6) 2023). there was evidence of optical placement signal issues (events 1036 / signal invalid), but no entries relevant to fan operation. rtlog data indicates that optical parameters during first pump connection were consistent with air gap (very low snr, weak optical signal) and improved after pump was disconnected and reconnected. device analysis: console was tested in danvers, but fan noise was not reproduced. analysis of the ic5244 service records indicated that after last pm on (b)(6) 2022, the console had aic sw v11.1 installed when left abiomed facility, but current logs present on the device show aic sw v12.2, indicating that a field software upgrade must have been performed. the report of fan noise might be related to observations made during console¿s field software upgrade (that took place on or prior to (b)(6) 2023), which is known to temporarily affect fan operation during pbm firmware re-flash. although the initial psnr condition was mostly resolved after pump was reconnected (consistent with an air gap contributing to initial problems during case start), the optical parameters (particularly snr) suggested some underlying issues with optical bench. during testing of the optical system, inconsistent performance of optical bench had been observed, failing "5s" specification depending on test cavity length: test cavity test plug pump simulator. cavity length: 14333 nm. snr: 7101. contrast: 66.2%. lamp: 93.6%. gain: 1.33. mano: 770.5 mmhg. light: 2.66 v cavity length: 15886 nm. snr: 459. contrast: 19.1%. lamp: 59.7%. gain: 1.35. mano: 770.7 mmhg. light: 2.65 v cavity length: 15769 nm. snr: 1451. contrast: 30.1%. lamp: 65.9%. gain: 1.35. mano: 770.7 mmhg. light: 2.67 v. pass fail fail. monitoring of the optical bench ccd output showed noise in the analog signal, which caused low snr, depending on location of the fringe pattern, respectively to the location of noisy part of the baseline. console inspection also revealed some contamination of the blue receptacle fiber which, however, was unlikely to contribute to the psnr condition (1036). while testing aic with rlm connected to a test network, it's been discovered that there was partition switching taking place during rlm boot-up, indicating a corrupted microsd card. inspection of the backstrap cable revealed misaligned miniusb connector and compromised insulation on usb wires, which might have contributed to microsd card corruption. root cause: the root cause of the reported fan noise was most likely a misinterpretation of a momentarily high fan speed during aic sw upgrade (performed in the field on 2023-09-15). the reported issue of ¿aic giving an optical error¿ has been confirmed in device logs, and was result of low snr due to noisy optical bench. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h3 updated. h6 updated. d2 common device name updated. g1 reporting contact email and reporting contact fax number updated. | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
33645 | 1220648-2023-04092 | 17950174 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17950174&pc=OZD | 09/18/2023 | 11/01/2023 |  | 10/17/2023 |  |  | 09/18/2023 | 10/17/2023 |  |  | 06/30/2025 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024368528 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/18/2023 |  |  |  |  | The user facility reported a patient (unknown age, gender and race) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after approximately one hour, the pump was in the wrong position in the aorta. the physician declined assistance from abiomed representative and explanted the impella cp. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed since the original report was filed. the medical center did not return the impella device. no benchtop analysis was possible. it was reported that the root cause of the positioning issue is most likely due to use as the md was offered support with positioning from csc but declined. there was nothing reported about the patient¿s condition or any malfunctions with the device. there were insufficient clinical details, no product or data logs returned.
33646 | 1220648-2023-04090 | 17950178 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17950178&pc=OZD | 09/18/2023 | 11/06/2023 |  | 10/17/2023 |  |  | 09/18/2023 | 10/17/2023 |  |  |  |  | 10/17/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1377267 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  |  |  | Malfunction | No | Yes | Defective Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a console and impella cp support ongoing after the first pump was damaged in the surgical cross clamp. the second pump was delivered and shortly there after there was a yellow controller alarm that prompted the team to replace the console. no lasting harm or injury from the interruption in the console/pump support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - controller error (yellow alarm) has been completed since the original report was filed. product was returned for evaluation. it was reported that the console was tested on an engineering lab bench and the failure mode was not reproduced while running an optical pump simulator. there was no issue with the console hardware. the root cause of the controller error was due to an aic software issue.
33647 | 1220648-2023-04097 | 17950190 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17950190&pc=OZD | 09/19/2023 | 12/11/2023 |  | 10/17/2023 |  |  | 09/19/2023 | 10/17/2023 |  |  | 05/31/2025 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024376995 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/19/2023 |  |  |  |  | The user facility reported a 57-year-old female with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the impella 5.5 pump was placed via the axillary and support is ongoing, though due to access site bleed and other sites/systems (gi bleed, nasal) heparin is being withheld. the pump has been supporting the patient for over one month to date. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported access site bleeding has been completed since the original report was filed. no product was returned. as per clinical, bleeding from access site around the sheath and heparin was stopped to achieve hemostasis, other information are unknown. the cause of the access site bleeding issue was not determined since product was not returned and due to insufficient clinical information.
33648 | 1220648-2023-04077 | 17950242 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17950242&pc=OZD | 05/29/2022 | 10/17/2023 |  | 10/17/2023 |  |  | 05/29/2022 | 10/17/2023 |  |  |  |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023197495 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/29/2022 |  |  |  |  | The user facility reported that a 61-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced positioning issues. the clinical team reported that the device was inadvertently pulled out of the left ventricle during support. subsequent multiple attempts were made to re-insert the device into the left ventricle; however, the clinicians were unable to successfully cross the aortic valve and reaccess the ventricle. the decision was then made to explant the pump and replace it with another device. there was no harm reported to the patient. | [5 empty] | The investigation into the reported pump position issue was completed. the device was not returned for evaluation. after reviewing the clinical information, it was determined that the cause of the positioning issue was user attempted wireless re-access. the impella device is not indicated for wireless re-access per ifu 0550-9002. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33649 | 1220648-2023-04109 | 17950262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17950262&pc=OZD | 07/05/2022 | 09/25/2024 |  | 10/17/2023 |  |  | 08/22/2022 | 10/17/2023 |  |  | 04/30/2024 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023211247 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Fluid/Blood Leak | Adverse Event Without Identified Device or Use Problem | [5 empty] | Stroke/CVA | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/22/2022 |  |  |  |  | The user facility reported that a 48-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced a purge leak. the team started sodium bicarbonate through the purge since the patient suffered a stroke. after more the 14 days on sodium bicarbonate purge, the purge sidearm yellow luer began to leak. no alarms were present, and the purge flows/pressures were within normal range. the team did not want to replace the pump, so the impella ran, while leaking, for the next couple of weeks. there were no reports of harm to the patient. | In addition, it was reported that on (b)(6) 2022 heparin was removed from purge solution due to concern for stroke. the patient was sent for a ct scan and a subdural hematoma noted. neurosurgery declined craniotomy due to need for anticoagulation. the patient was alert and lethargic but does not have left sided function. a craniotomy was performed, (b)(6) 2022, and heparin was continued to be withheld due to the head bleed. on (b)(6) 2022 it was noted that care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the reported purge leak was completed. the device was not returned for evaluation. the root cause of the purge leak was most likely sidearm yellow luer damage as reported by the clinical team. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-04109 was provided in sections b1, b2, b5, d6a, d6b, h1, and h6.
33650 | 1220648-2023-04091 | 17950274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17950274&pc=OZD | 09/18/2023 | 12/03/2025 |  | 10/17/2023 |  |  | 09/18/2023 | 10/17/2023 |  |  |  |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1097391 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The us complainant had an automated impella controller (aic) controller with a purge alarm. the vad coordinator noted there was a prompt for disconnection of the luer after purge in air alarm. there was no reported patient harm. | [5 empty] | The impella console was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | D.9 device return date was added. a.2 age should have been left blank in accordance with updated procedures on the initial manufacturer device report number 1220648-2023-04091 as it is unknown. a.3 unknown should have been selected on the initial manufacturer device report number 1220648-2023-04091 as it is unknown. d.1 and d.2 brand name and common device name were revised in accordance with updated procedures since the initial manufacturer device report number 1220648-2023-04091 was submitted. d.4 model number was revised in accordance with updated procedures since the initial manufacturer device report number 1220648-2023-04091 was submitted. d.4 revised catalog number, lot number, unique identifier (udi) # and serial number as they were reported incorrectly on the initial manufacturer device report number 1220648-2023-04091. g.1 revised mdr reporting contact email since the initial submission of manufacturer device report number 1220648-2023-04091 in accordance with updated procedures. added reporting contact fax number in accordance with updated procedures since the initial manufacturer device report number 1220648-2023-04091 was submitted. added manufacturer site fax number as it was inadvertently omitted from the initial manufacturer device report number 1220648-2023-04091. h.4 revised as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2023-04091. h.6 code 3221 was reported incorrectly on the initial manufacturer device report number 1220648-2023-04091 in accordance with updated procedures.
33651 | 1220648-2023-04080 | 17950372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17950372&pc=OZD | 06/13/2022 | 10/17/2023 |  | 10/17/2023 |  |  | 06/13/2022 | 10/17/2023 |  |  | 02/29/2024 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023197679 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/13/2022 |  |  |  |  | The user facility reported that a 53-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced positioning issues. the patient experienced multiple impella in aorta alarms in the cath lab. the pump was advanced by the physician, and the alarms self-corrected. no further issues were reported after repositioning. there was no harm reported to the patient. | [5 empty] | The investigation into the reported pump position issue was completed. the device was not returned for evaluation. based on the available information, the positioning issue was most likely due to improper use or technique. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33652 | 1220648-2023-04078 | 17950385 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17950385&pc=OZD | 06/07/2022 | 10/17/2023 |  | 10/17/2023 |  |  | 06/07/2022 | 10/17/2023 |  |  | 10/31/2024 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023181475 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/07/2022 |  |  |  |  | The user facility reported that a 72-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced a purge leak. a retainer was placed on the purge sidearm right after the pump was in proper position. two hours later, a leak from the purge sidearm was observed. the purge cassette was replaced, but the leak did not resolve. no cleaning products were reported to have been used on the sidearm. a purge bypass was performed successfully, and the purge flow and rate were restored to previous baseline settings. the etiology of leak remains known. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge leak was completed. the device was returned for evaluation. functional testing confirmed fluid leaking out of the top of the purge reservoir membrane cap. the leak was noted to be coming through a tear in the blue membrane at the left of the center hole of the housing cap. the root cause of the purge leak was sidearm reservoir damage. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended. per the ifu: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol."
33653 | 1220648-2023-04079 | 17950408 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17950408&pc=OZD | 06/08/2022 | 10/17/2023 |  | 10/17/2023 |  |  | 06/08/2022 | 10/17/2023 |  |  | 02/29/2024 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023196017 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Suction Problem | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/08/2022 |  |  |  |  | The user facility reported that a 60-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced low pump flows and suction alarms. upon increasing flows during initial insertion, flows initially, successfully increased to p8/avg 4.5lpm; however, minutes after start-up, flows dropped, and suction alarms were triggered. transesophageal echocardiogram (tee) imaging assesses left ventricular as sufficient and accurate positioning was confirmed. the physician attempted to reposition the device and increase flows, but suction events continued, and the pump was explanted as a result. act at insertion was >300 sec with heparin. visual inspection of the device upon explantation, revealed no clots noted externally on the pump/cannula. there was no reported harm to the patient. | [5 empty] | The investigation into the reported low pump flows and suction alarms was completed. the device was returned for evaluation. analysis of the data logs showed a sudden decrease in flows and confirmed the suction alarms. visual inspection of the pump revealed thrombus on the impeller. the root cause of the low flow was determined to be biomaterial. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33654 | 1220648-2023-04095 | 17950680 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17950680&pc=OZD | 03/15/2023 | 10/17/2023 |  | 10/17/2023 |  |  | 03/15/2023 | 10/17/2023 |  |  | 06/30/2024 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023248461 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Tamponade | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/15/2023 |  |  |  |  | The user facility reported a 56yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient had tamponade and was taken back to the operating room to clean out the clot. the chest was left open for a few hours. | [5 empty] | The investigation for the reported tamponade has been completed. after reviewing the clinical information, it was determined that the cardiac tamponade was unrelated to the impella. there was no relation of the event to the device. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33655 | 1220648-2023-04085 | 17950698 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17950698&pc=OZD | 11/27/2021 | 10/17/2023 |  | 10/17/2023 |  |  | 11/27/2021 | 10/17/2023 |  |  | 06/30/2023 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022098174 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/27/2021 |  |  |  |  | The user facility reported that a 62-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced sudden low pump flows and suction alarms. the patient was taken for pump exchange. upon explant, a large clot was observed, and the pump was sent to pathology. there was no reported harm to the patient. | [5 empty] | The investigation into the reported low pump flows was completed. the device was not returned for evaluation. based on the available information, the most likely root cause of the low pump flows was ingested biomaterial. this report is being filed as part of a retrospective review of historical records.
33656 | 1220648-2023-04099 | 17950770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17950770&pc=OZD | 09/18/2023 | 02/02/2024 |  | 10/17/2023 |  |  | 09/18/2023 | 10/17/2023 |  |  | 06/30/2025 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024364338 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the impella 5.5 access site had bleeding access site complications. heparin was halted and manual pressure was held. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ | The investigation into the reported, access site bleeding has been completed since the original report was filed. product was not returned for review. as per clinical patient had bleeding at access site and packing of the site resolved bleeding. all other information were unknown. the cause of the access site bleeding issue was not determined due to insufficient clinical information and since product was not returned.
33657 | 1220648-2023-04088 | 17950813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17950813&pc=OZD | 07/28/2022 | 10/17/2023 |  | 10/17/2023 |  |  | 07/28/2022 | 10/17/2023 |  |  | 10/31/2023 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023161413 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/28/2022 |  |  |  |  | The user facility reported that after insertion of the impella 5.5 for a 79-year-old male patient, the pump started having purge system issues with flows steadily decreasing and pressures steadily rising. staff troubleshot with measures available. the patient was unable to have tpa for the purge pressure resolution. due to increased need to perform troubleshooting and concern for pump longevity, pump performance was set at lower flows and the patient was weaned for removal before a pump stop occurred. no further information was provided. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge pressure issues was completed. the device and data logs were returned for evaluation. analysis of the data logs revealed a slight increase in purge pressure to around 800 mmhg, and then a decrease back to 500 mmhg after the start/stop method. functional testing of the pump revealed that the purge pressure remained around 500 mmhg at all p levels. the root cause of the high purge pressure is most likely biomaterial. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33658 | 1220648-2023-04084 | 17951093 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17951093&pc=OZD | 05/29/2022 | 10/17/2023 |  | 10/17/2023 |  |  | 05/29/2022 | 10/17/2023 |  |  | 12/31/2023 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023175334 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/29/2022 |  |  |  |  | The user facility reported that a 39-year-old female patient implanted with the impella 5.5 for mechanical circulatory support experienced pump positioning issues. the clinical team stated that the "pump came back into aorta unable to be advanced after multiple attempts" the plan was to take the patient for pump exchange. no further information was provided. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported positioning issue was completed. based on the available information, the root cause of the reported event was undetermined. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33659 | 1220648-2023-04098 | 17951137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17951137&pc=OZD | 11/26/2021 | 10/17/2023 |  | 10/17/2023 |  |  | 11/26/2021 | 10/17/2023 |  |  | 09/30/2023 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022141282 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/26/2021 |  |  |  |  | The user facility reported that a 65-year-old male patient was implanted with the impella 5.5 for mechanical circulatory support. following implant, it was noted that the pump experienced saturated flows and the mc was flat. the pump was replaced as a precaution. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported saturated flow was completed. the device was not returned for evaluation. the root cause of the saturated flow was unable to be determined since product was not returned. this report is being filed as part of a retrospective review of historical records.
33660 | 1220648-2023-04100 | 17951154 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17951154&pc=OZD | 04/14/2023 | 10/17/2023 |  | 10/17/2023 |  |  | 04/14/2023 | 10/17/2023 |  |  | 01/31/2025 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024293924 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/14/2023 |  |  |  |  | The user facility reported that a 29-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced positioning issues. the clinicians reported that they were unable to advance the repositioning unit along the catheter. they stated that pressing the catheter anchor button failed to release the anchor, even though it functioned normally at the beginning of support. there was no reported harm to the patient. | [5 empty] | The investigation into the reported pump positioning issues was completed. the device was not returned for evaluation. based on the available information, the root cause of the reported event was not determined. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33661 | 1220648-2023-04081 | 17951178 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17951178&pc=OZD | 06/28/2022 | 10/17/2023 |  | 10/17/2023 |  |  | 06/28/2022 | 10/17/2023 |  |  | 03/31/2024 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023209651 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Puncture/Hole | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/28/2022 |  |  |  |  | The user facility reported that a 51-year-old male patient implanted with the impella 5.5 for mechanical circulatory support observed leaking out of the red pump handle within two hours of support. as a result, the decision was made to exchange the device. upon removal, the clinicians noted a significant kink on the catheter near the motor housing. no further information was provided. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported malfunction was completed. the device was returned for analysis. visual inspection of the pump revealed two puncture holes in the catheter between the 33-cm and 34-cm marks. additionally, the catheter was significantly kinked near the motor housing. dried blood was found in the red pump handle. based on the available information, it was determined that the root cause of the issue was blood ingress due to a hole in the catheter due to improper use. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33662 | 1220648-2023-04089 | 17951338 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17951338&pc=OZD | 11/29/2021 | 10/17/2023 |  | 10/17/2023 |  |  | 11/29/2021 | 10/17/2023 |  |  | 08/31/2023 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022126009 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | Fracture | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/29/2021 |  |  |  |  | The user facility reported a 66-year-old male patient was implanted with an impella 5.5 device for mechanical circulatory support. during support, it was noted that the yellow luer of the purge sidearm broke after a purge cassette change. purge fluid leaked from the luer and the "purge system open" alarm triggered. a purge bypass was performed to successfully resolve the issue. sodium bicarbonate was used in the purge solution throughout support. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge sidearm break and purge leak was complete. the device was not returned for evaluation. after review of the clinical information, it was determined that the cause of the purge leak was sidearm yellow luer damage due to the use of sodium bicarbonate in the purge solution. per the ifu: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
33663 | 1220648-2023-04112 | 17951392 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17951392&pc=OZD | 10/05/2022 | 10/17/2023 |  | 10/17/2023 |  |  | 10/05/2022 | 10/17/2023 |  |  | 02/29/2024 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023194890 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/05/2022 |  |  |  |  | The user facility reported that a 53-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced a purge leak. the yellow luer became cracked while changing the purge cassette. the purge fluid is sodium bicarbonate. no further information was provided. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge leak was completed. the device was not returned for evaluation. based on the available information, the root cause of the reported event was most likely due to sidearm yellow luer damage. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33664 | 1220648-2023-04086 | 17951417 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17951417&pc=OZD | 06/30/2022 | 10/17/2023 |  | 10/17/2023 |  |  | 06/30/2022 | 10/17/2023 |  |  | 01/31/2024 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023183679 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/30/2022 |  |  |  |  | The user facility reported that a 40-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced a purge leak at the purge sidearm. the physician changed the pump without issue. of note, it was reported that the patient may have rolled and caused damage. no further information was provided. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge leak was completed. the pump and purge cassette were returned for evaluation. visal inspection revealed a broken filter to tubing joint. based on the clinical account and the device evaluation, the root cause of the break resulting in the purge leak was excessive force. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended. per the ifu: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.".
33665 | 1220648-2023-04082 | 17951483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17951483&pc=OZD | 11/24/2021 | 10/17/2023 |  | 10/17/2023 |  |  | 11/24/2021 | 10/17/2023 |  |  | 06/30/2023 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022096677 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/24/2021 |  |  |  |  | The user facility reported that a 66-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced low pump flows midway through support. the patient decompensated and began to receive inadequate hemodynamic support, despite the impella being in proper position and all pump metrics being within normal limits. notably, the patient presented with considerably low end-diastolic pressure (edp) and ejection fraction (ef) values. the patient was then placed on venous- arterial extracorporeal membrane oxygenation (va ecmo) for additional support with the impella. the pump remained implanted until the patient was successfully weaned for transplant. there was no reported harm to the patient. | [5 empty] | The investigation into the reported low pump flows was completed. the device was returned for evaluation. analysis of the data logs did not reveal any performance issues. the device was returned with the entire catheter cut off. no other damage or abnormalities were noted. after reviewing the clinical information, it was discussed that the root cause of the low pump flow could potentially be the patient's condition, however a definitive cause could not be established. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33666 | 1220648-2023-04113 | 17951557 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17951557&pc=OZD | 10/11/2022 | 10/17/2023 |  | 10/17/2023 |  |  | 10/11/2022 | 10/17/2023 |  |  | 06/30/2024 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023232051 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/11/2022 |  |  |  |  | The user facility reported that a 54-year-old female patient implanted with the impella 5.5 for mechanical circulatory support had a clot noted in the cannula post explant. no further information was provided. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported thrombus was completed. the pump and purge cassette were returned for evaluation. analysis of data logs revealed suction triggered throughout support. visual inspection of the pump revealed biomaterial on the outflow and the inflow cage. the root cause of the reported event was biomaterial. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33667 | 1220648-2023-04093 | 17951774 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17951774&pc=OZD | 09/17/2023 | 12/27/2024 |  | 10/17/2023 |  |  | 09/17/2023 | 10/17/2023 |  |  |  |  | 10/17/2023 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1249508 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Communication or Transmission Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant in the eu reported that the automated impella controller (aic) did not recognize the inserted impella cardiac pump catheter. there was no known harm or injury. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation for the console "pump not detected" has been completed. data logs from the console confirm impella defective alarm being sent. logs show an occasional communication error with the impellatronics (epm circuitry) each time a pump is connected. the console was returned for evaluation. when the console was run with a known good pump, there was no observable anomaly. the root cause could not be determined due to not being reproducible. impellatronics board was replaced as a precaution. corrections to the initial mfr report: d2 common device name updated. g1 mdr reporting contact email updated.
33668 | 1220648-2023-04104 | 17952075 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17952075&pc=OZD | 12/09/2021 | 10/17/2023 |  | 10/17/2023 |  |  | 12/09/2021 | 10/17/2023 |  |  | 08/31/2023 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022136270 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/09/2021 |  |  |  |  | The user facility reported that a 62-year-old female patient implanted with the impella 5.5 for mechanical circulatory support experienced "purge system blocked" alarms coinciding with a purge flow of <.5, and pp of 1105 after leaving the cath lab. the tpa protocol was initiated and initially effective; however, the system started to fail again. on day 7 of support, the patient was subsequently taken for pump exchange. no further information was provided. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge system blocked alarms was completed. the pump was returned for evaluation. visual inspection revealed biomaterial in the purge gap. the pump catheter had been cut so the full purge line could not be tested. functional testing did not confirm any blockages. the root cause of the high purge pressure was ingested biomaterial occluding the purge gap. per the ifu: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ this report is being filed as part of a retrospective review of historical records.
33669 | 1220648-2023-04083 | 17952109 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17952109&pc=OZD | 06/22/2022 | 10/17/2023 |  | 10/17/2023 |  |  | 06/22/2022 | 10/17/2023 |  |  | 03/31/2024 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023205523 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/22/2022 |  |  |  |  | The user facility reported that a 69-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced positioning issues. the cable became wrapped around the patient's leg overnight and was pulled back into the aorta with the patient's movement. the patient remained hemodynamically stable; however, multiple attempts to readvance the device across the aortic valve were unsuccessful. as a result, the device was explanted, and an intra-arterial balloon pump (iabp) was placed. there was no harm reported to the patient. | [5 empty] | The investigation into the reported pump positioning issue was completed. no data or device were returned for evaluation. based on the information received, the root cause of the positioning issues was determined to be patient movement. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33670 | 1220648-2023-04114 | 17952128 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17952128&pc=OZD | 12/20/2021 | 10/17/2023 |  | 10/17/2023 |  |  | 12/20/2021 | 10/17/2023 |  |  |  |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 1468544 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/20/2021 |  |  |  |  | The user facility reported that a 61-year-old female patient implanted with the impella 5.5 for mechanical circulatory support experienced "purge system blocked" alarms after priming the pump. the purge cassette was changed without resolution. the pump was subsequently replaced. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge system blocked alarm was completed. the pump was returned for evaluation. functional testing of the pump indicated that it had not completed the case start. the case start was completed successfully in the lab and the pump functioned without issue. upon conclusion of the investigation, it was discussed that a kink in the purge line could result in the reported purge blockage; however, a definitive cause for the failure could not be determined. this report is being filed as part of a retrospective review of historical records.
33671 | 1220648-2023-04111 | 17952252 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17952252&pc=OZD | 09/28/2022 | 10/17/2023 |  | 10/17/2023 |  |  | 09/28/2022 | 10/17/2023 |  |  | 03/31/2024 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023219710 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/28/2022 |  |  |  |  | The user facility reported that a 72-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced a purge leak. the nurse called for assistance because she changed the purge cassette twice and there was leaking from the sidearm. purge bypass instructions were sent, and a purge bypass was performed without issue. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported purge leak was completed. the device was not returned for evaluation. based on the available information, the root cause of the purge leak was unable to be determined as no product was returned for analysis. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33672 | 1220648-2023-04115 | 17952271 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17952271&pc=OZD | 12/30/2021 | 09/19/2024 |  | 10/17/2023 |  |  | 12/30/2021 | 10/17/2023 |  |  | 08/31/2023 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022136203 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/30/2021 |  |  |  |  | The user facility reported that a 53-year-old male patient was implanted with an impella 5.5 for mechanical circulatory support. during support, the pump position moved from 54cm to 50cm and came back across the aortic valve and could not be re-advanced through the valve. the pump was removed that same day. | Additional information noted the patient expired on support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the reported positioning issue was completed. the device was not returned for evaluation. based on the information received, the most likley root cause of the pump positioning issue was use related as the pump position was moved 4cm and resultantly out of the lv. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6.
33673 | 1220648-2023-04116 | 17952425 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17952425&pc=OZD | 09/21/2022 | 10/17/2023 |  | 10/17/2023 |  |  | 09/21/2022 | 10/17/2023 |  |  |  |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023225617 |  |  |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/21/2022 |  |  |  |  | The user facility reported that a 47-year-old male patient implanted with the impella 5.5 for mechanical circulatory support exchanged the device due to high purge pressure and suspected clotting. the patient was also heparin-induced thrombocytopenia (hit) positive and had been on an argatroban drip for anticoagulation. purge fluid was d5 there were no reports of harm to the patient. | [5 empty] | The investigation into the reported high purge pressure was completed. the device was not returned for evaluation. analysis of the data logs revealed a gradual increase in purge pressure rose and the lack of heparin in the purge. the clinical information was also reviewed. based on the information received, the root cause of the high purge pressure was most likely biomaterial. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33674 | 1220648-2023-04118 | 17952556 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17952556&pc=OZD | 09/20/2023 | 01/23/2024 |  | 10/17/2023 |  |  | 09/20/2023 | 10/17/2023 |  |  | 04/30/2024 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023219348 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Deformation | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/20/2023 |  |  |  |  | The user facility reported a 62-year-old female in cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella cp had been placed under fluoroscopy guidance successfully. while removing the peel-away sheath, the impella was pulled back into the aorta. the pigtail was pushed up against the valve and collapsed and prolapsed on itself. the provider was able to snare the pigtail and straighten the catheter so that it could removed without incident or injury. the cp was replaced successfully. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues and the product damage (cannula kink) has been completed. the device was not returned for investigation. per the clinical information provided, the root cause of the cannula kink and the positioning issues is due to use; while the peel away introducer was removed, the pump was pulled back into the aorta. pigtail was pushed up against the valve and collapsed and prolapsed on itself.
33675 | 1220648-2023-04119 | 17952579 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17952579&pc=OZD | 09/20/2023 | 12/08/2023 |  | 10/17/2023 |  |  | 09/20/2023 | 10/17/2023 |  |  | 04/30/2025 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024345677 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/20/2023 |  |  |  |  | The user facility in japan reported a 52-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient had a ct scan that showed cerebral edema, but no bleeding. the heparin in the purge liquid was changed to meiron. a following ct scan confirmed a cerebral hemorrhage. the relationship between impella and cerebral hemorrhage is unknown. even after completing percutaneous coronary intervention and being admitted to the icu, disseminated intravascular coagulation (dic) was present and the coagulation system was prolonged ( activated partial thromboplastin time unknown). after 25.5 hours of impella support, the patient expired. the relationship between impella and cause of death is unknown. the cause of death is unknown, but active treatment was stopped and palliative treatment was performed due to cerebral hemorrhage, and the patient expired because circulation could not be maintained. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported stroke/cva been completed since the original report was filed. on (b)(6) 2023 heparin was stopped being used in the purge and then the hemorrhagic stoke occurred on (b)(6) 2023. there was no product or data logs returned. the root cause of the stroke/cva is most likely due to the patients condition and the relationship to the impella is unknown.
33676 | 1220648-2023-04110 | 17952645 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17952645&pc=OZD | 09/23/2022 | 09/15/2024 |  | 10/17/2023 |  |  | 09/23/2022 | 10/17/2023 |  |  | 02/29/2024 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic |  | Death | Yes | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | Asystole | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/23/2022 |  |  |  |  | The user facility reported that a 68-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced high purge pressure within three hours of support. tissue plasminogen activator (tpa) was added to the purge which was successful in increasing the purge flow and decreasing purge pressure. of note, the patient was reported to have a pre-existing history of severe clotting issues pre-insertion. there were no reports of harm to the patient. | It was reported prior to the tissue plasminogen activator (tpa) being added to the purge, the patient was decompensating with the pump having intermittent suction alarms with persistent purge pressure high alarms. after the tpa protocol resolved the high purge pressure alarms, the patient was asystolic. it was noted the patient's care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the reported high purge pressure was completed. the device was not returned for evaluation. the data logs were reviewed and show a gradual increase in purge pressure leading to high purge pressure alarms. the logs also show the purge pressure decreasing as the tpa began to work. the root cause of the high purge pressure was most likely biomaterial as tpa was able to resolve the issue. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-04110 was provided.
33677 | 1220648-2023-04120 | 17952769 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17952769&pc=OZD | 09/18/2023 | 01/24/2024 |  | 10/17/2023 |  |  | 09/18/2023 | 10/17/2023 |  |  |  |  | 10/17/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1574621 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Communication or Transmission Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | A 66-year-old female patient was treated for acute myocardial infarction/cardiogenic shock and was scheduled to receive hemodynamic support with an impella cp smartassist heart pump. during the preparation of the impella cp heart pump, an error occurred in the automated impella controller (aic), which prevented the preparation of the impella catheter from being completed. the aic was replaced by the backup controller. the impella cp could then be prepared and introduced without any further problems. the patient suffered no harm as a result of the incident. | [5 empty] | The investigation into the reported pump not detected issue has been completed. the automated impella controller (aic) controller was received for analysis. console logs from the reported day of event are consistent with complaint. the cause of the aic issue was a defective impellatronics board, combined with the failure of software mitigation mechanism. the failure modes will be monitored and trended. | The investigation into the pump not detected and the aic - purge issue - other has been completed since the original report was filed. product was returned for evaluation. it was reported that the cause of the aic issue was a defective impellatronics board, combined with the failure of software mitigation mechanism and the aic - purge issue - other was most likely use related.
33678 | 1220648-2023-04122 | 17952881 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17952881&pc=OZD | 09/18/2023 | 01/30/2024 |  | 10/17/2023 |  |  | 09/18/2023 | 10/17/2023 |  |  | 05/31/2025 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024365524 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient was having gross hemolysis while on impella support. after confirming pump positioning volume was given. plasma free hemoglobin levels were high, but did improve with volume resuscitation but left ventricular thrombus showed on tte echo. the following day the pump ingested a clot, and the pump was removed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions including catheter position, pre-existing medical conditions, and small left ventricular volumes may also play a role in patient susceptibility to hemolysis.¿ | The investigation has been completed. the pump was not returned for review. data logs showed after the pump started, mc spiked followed low flow that triggered auto suction response and suction alarms. this event remained to the rest of the case. based on clinical description and data analysis, the root cause of hemolysis was most likely due to biomaterial ingestion.
33679 | 1220648-2023-04096 | 17952990 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17952990&pc=OZD | 09/18/2023 | 03/28/2025 |  | 10/17/2023 |  |  | 09/18/2023 | 10/17/2023 |  |  |  |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1320564 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 09/18/2023 |  |  |  |  | The user facility's biomedical engineer reported their automated impella controller (aic) had a repeated error message. it is unknown what the error message was as the medical staff did not document the alarm for abiomed team. a loaner was sent to initiate the return for investigation. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) ¿ other has been completed. the console logs confirmed that multiple impella defective alarms were issued when the pump was plugged in to the console on the date of complaint. the alarms are all due to errors reading from the eeprom of the pump. visual inspection of the consoles internal circuitry did not reveal any malfunctions. when a known good pump and purge was connected to the console, the console was able to operate within specification. (no observable anomaly). the root cause for the impella defective alarms could not be determined due to not being able to reproduce. a.1 revised since the initial manufacturer device report number 1220648-2023-04096 was submitted in accordance with updated procedures. a.3 revised since the initial manufacturer device report number 1220648-2023-04096 was submitted in accordance with updated procedures. d.2 revised common device name since the initial manufacturer device report number 1220648-2023-04096 was submitted in accordance with updated procedures. d.4 revised model number since the initial manufacturer device report number 1220648-2023-04096 was submitted in accordance with updated procedures.. d.9 revised device returned selection as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2023-04096. e.1 added fax number as it was inadvertently omitted from the initial manufacturer device report number 1220648-2023-04096. e.2 revised since the initial manufacturer device report number 1220648-2023-04096 was submitted in accordance with updated procedures. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2023-04096 and should not have been. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2023-04096 was submitted in accordance with updated procedures.
33680 | 1220648-2023-04087 | 17953109 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17953109&pc=OZD | 07/11/2022 | 10/17/2023 |  | 10/17/2023 |  |  | 07/11/2022 | 10/17/2023 |  |  |  |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023208299 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 17 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/11/2022 |  |  |  |  | The user facility reported that a 17-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced thrombus. after removal of the device for heart transplant, significant clot burden was found in the cannula. the pump ran with excellent flows and without problems up to explant. there was also clot formed noted within the ventricle of the native heart post-removal. no further information was provided. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported thrombus was completed. the device was not returned for evaluation. analysis of the field logs did not reveal any high motor current event indicative of clot ingestion or suction alarms while the pump was operational. the pump flow was within specification throughout support. additionally, returned field images confirmed biomaterial in the outlet and inlet. based on the available information, the root cause of the reported event is biomaterial pulled into the pump during explant. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33681 | 1220648-2023-04117 | 17953281 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17953281&pc=OZD | 09/18/2023 | 08/02/2024 |  | 10/17/2023 |  |  | 09/18/2023 | 10/17/2023 |  |  | 05/31/2025 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024358089 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A patient in the eu of unknown age received hemodynamic support from an impella cp pump due to acute myocardial infarction/cardiogenic shock. a ¿purge pressure high¿ alarm and a purge line leak have been reported. the impella cp was removed and replaced with a new impella cp. the patient suffered no harm as a result of the incident. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation of purge leak ¿ pump and high purge pressure has been completed. the product was returned and investigated. data logs show a gradual 2-hour rise in purge pressure, followed by pump flow saturation and pump stoppage, with several high motor current alarms. biomaterial was observed in purge gap. a crack was found in the filter, causing purge fluid leakage. further investigation revealed blood in the purge line. the pump stopped due to damage to the sidearm filter, which allowed blood to ingress into the motor housing. this also caused the high purge pressure. the pump stopped due to damage to the sidearm filter, which allowed blood to ingress into the motor housing. this also caused the high purge pressure. d6a/d6b added. d9 date device returned to manufacturer added.
33682 | 1220648-2023-04121 | 17953496 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17953496&pc=OZD | 05/25/2023 | 10/17/2023 |  | 10/17/2023 |  |  | 05/25/2023 | 10/17/2023 |  |  | 01/31/2025 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024331189 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Activation, Positioning or Separation Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that had an impella cp that was caught under the papillary and was being repositioned but, the pump came across the valve and was unable be re-deliver. the patient additionally had signs of hemolysis. the impella cp was explanted and replaced without any harm or injury. | [5 empty] | The investigation into the reported positioning and possible hemolysis has been completed since the original report was filed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. based on clinical description, the root cause of positioning issue was most likely due to use issue and the root cause of hemolysis was unable to be determined as no product was returned for review. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33683 | 1220648-2023-04123 | 17954213 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17954213&pc=OZD | 09/19/2023 | 04/29/2024 |  | 10/17/2023 |  |  | 09/19/2023 | 10/17/2023 |  |  | 07/31/2025 |  | 10/17/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024377574 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/19/2023 |  |  |  |  | The user facility reported a 57-year-old male (race unknown) with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. on day eight of support, the pump stopped and exhibited in aorta/malposition alarms. the team was able to restart the pump and the alarms were resolved. the patient remains on impella support one month after the alarms. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation for the reported positioning issues and the pump stop has been completed. the pump was not returned for investigation. data logs show a motor current spike and suction alarm, followed by saturated pump flow and a pump stop-motor high current alarm. the pump was successfully restarted after the third pump stop and ran without issue. the cause of the temporary pump stop issue was most likely biomaterial based on data log review. there is a likelihood that biomaterial interaction caused the saturated flow and miscalculation in the placement signal, resulting in a false pump position related issue. to conform to updated procedures, section d6 has revised with updated information.
33684 | 1220648-2023-04045 | 17940151 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17940151&pc=OZD | 09/16/2023 | 02/05/2024 |  | 10/16/2023 |  |  | 09/17/2023 | 10/16/2023 |  |  | 08/31/2025 |  | 10/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024382958 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/17/2023 |  |  |  |  | The user facility reported a 31-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the physician decided on escalation of care and the impella cp was removed and an impella 5.5 was placed. upon impella 5.5 removal, a clot was discovered in the inlet cage. the nurse stated the patient was subtherapeutic. heparin was in purge and was running systematically. the team reported the patient suffered an embolic stroke. the patient had an intra-aortic balloon pump placed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the stroke/cva issue has been completed. the device was not returned for investigation. the clinical stated there was a clot observed on the inflow cage upon explant and the data logs show a small spike in motor current followed by decreased flows. though the product was not returned, the root cause of the stroke is most likely due to ingested biomaterial.
33685 | 1220648-2023-04046 | 17940220 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17940220&pc=OZD | 03/21/2022 | 10/16/2023 |  | 10/16/2023 |  |  | 03/21/2022 | 10/16/2023 |  |  | 11/30/2023 |  | 10/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023169301 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Sensing Intermittently | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/21/2022 |  |  |  |  | The user facility reported a 72-year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was documented that the persistent signal issues were encountered and that the optical sensor was not reliable. as a result, the pump was explanted and a new device was placed for ongoing support. there was no patient harm or injury as a result of the noted issue. | [5 empty] | The pump was returned and an evaluation was completed. a review of the data logs indicated that the sensor was initially working. after 15 minutes of support, the raw ps drops to -32768, lamp increases to 100%, and snr drops to zero. product was returned and showed the sensor working without issue, with 5s parameters within specification. the patient cable was slightly unplugged from the optical receptacle, and the 5s parameters in the data logs mimic those from the case. thus, the root cause of the ps issues is due to the patient cable becoming unplugged. this report is being filed as part of a retrospective review of historical records.
33686 | 1220648-2023-04047 | 17940281 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17940281&pc=OZD | 09/17/2023 | 03/26/2024 |  | 10/16/2023 |  |  | 09/17/2023 | 10/16/2023 |  |  | 04/30/2025 |  | 10/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - EU |  |  | ABIOMED, INC. |  |  | US |  | 2024346214 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Improper Flow or Infusion | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/17/2023 |  |  |  |  | The user facility reported a 65-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after a check of the motor current, several spikes occurred around 20 min before doctor's call. after request of anticoagulation status, the physician confirmed that activated partial thromboplastin time out of range for the whole day before his arrival. reduction of flow up to 0.0 lpm. probable clot formation was presumed, then confirmed after pump removal. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported low pump flow and suspected thrombus has been completed. the impella device was returned for evaluation. visual inspection found biomaterial wrapped around the impeller. data logs show a spike in motor current and active suction alarm preceded by low pump flow during the start of the case. the cause of the low pump flow issue was most likely biomaterial, based on data log review. the instructions for use states ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. it also states ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿
33687 | 1220648-2023-04048 | 17940461 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17940461&pc=OZD | 04/19/2022 | 10/16/2023 |  | 10/16/2023 |  |  | 04/19/2022 | 10/16/2023 |  |  | 02/28/2023 |  | 10/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2022059818 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | Fracture | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/19/2022 |  |  |  |  | The user facility reported a 56-year-old male patient was implanted with an impella device for mechanical circulatory support. the complainant reported the purge filter broke during cpr on the second day of support. a purge side arm bypass was performed with no harm to the patient. | [5 empty] | The device was returned and an evaluation was completed. visual inspection of the returned pump confirmed a sidearm bypass was in place. the purge sidearm itself was not returned. otherwise, no damage or abnormalities were found. when the pump was connected to the aic, it started normally and its performance was within specification. after reviewing the clinical information, it was determined that the cause of the purge leak was most likely damage to the sidearm filter during cpr. this report is being filed as part of a retrospective review of historical records.
33688 | 1220648-2023-04049 | 17940667 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17940667&pc=OZD | 08/02/2022 | 10/16/2023 |  | 10/16/2023 |  |  | 08/02/2022 | 10/16/2023 |  |  | 04/30/2024 |  | 10/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2023222286 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Advance | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/02/2022 |  |  |  |  | The user facility reported a 71-year-old male patient was implanted with an impella device for mechanical circulatory support. very difficult patient anatomy led to kinks at catheter shaft. after changing to a new pump it was possible place the new pump without any problems. patient stable on support, no harm or injury inflicted. | [5 empty] | The pump was returned and an evaluation was completed. a visual inspection verified that the catheter on the pump was kinked. from the clinical review, the root cause of this compliant was determined to be related to the patient's condition/anatomy. this report is being filed as part of a retrospective review of historical records.
33689 | 1220648-2023-04050 | 17941107 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17941107&pc=OZD | 09/18/2023 | 01/25/2024 |  | 10/16/2023 |  |  | 09/18/2023 | 10/16/2023 |  |  | 04/30/2025 |  | 10/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - EU |  |  | ABIOMED, INC. |  |  | US |  | 2024345085 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/18/2023 |  |  |  |  | The user facility in the eu reported a 73-year-old male (race unknown) noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient had significant bleeding beside the repo-sheath. there was prolonged manual compression. the impella cp was explanted on the same day prematurely. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed. the device was not returned for investigation. based on clinical description, the root cause of access site bleeding was most likely due to anticoagulation management related.
33690 | 1220648-2023-04051 | 17941183 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17941183&pc=OZD | 12/21/2022 | 10/16/2023 |  | 10/16/2023 |  |  | 12/21/2022 | 10/16/2023 |  |  | 11/30/2024 |  | 10/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023273602 |  | 0048-0045 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/21/2022 |  |  |  |  | The user facility reported a 46 year old male was implanted with an impella cp device for mechanical circulatory support as a bridge for heart transplant. it was reported that there was a suspicion of hemolysis. the patient eventually received a heart transplant. | [5 empty] | The investigation for the reported hemolysis issue has been completed. the cause of the hemolysis and its relation to the impella were not determined since the product and data logs were not determined. the instructions for use has details on the handling of hemolysis. it states ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33691 | 1220648-2023-04052 | 17941231 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17941231&pc=OZD | 09/18/2023 | 01/16/2024 |  | 10/16/2023 |  |  | 09/18/2023 | 10/16/2023 |  |  | 07/31/2025 |  | 10/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024377012 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 18 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/18/2023 |  |  |  |  | The user facility reported an 28-year-old male with non ischemic cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, the purge fluid was leaking around the purge filter and cable junction. the purge flows were unchanged and purge pressures were stable. the clinicians were advised to replace the pump if purge flows continually increase or purge pressures drastically decrease. the patient remained on impella support for eight additional days. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation into the reported purge leak has been completed since the original report was filed. the product was not returned for review. the root cause of the purge leak was unable to be determined as limited clinical details were provided and no product was returned for review.
33692 | 1220648-2023-04053 | 17941255 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17941255&pc=OZD | 09/19/2023 | 02/26/2024 |  | 10/16/2023 |  |  | 09/19/2023 | 10/16/2023 |  |  | 03/31/2025 |  | 10/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2024333100 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/19/2023 |  |  |  |  | The user facility in japan reported a patient (age and gender unknown) noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the impella cp support was ongoing when, on day two, post-op there was an observation made of limb ischemia. the ischemia was treated by pump explant. the pulses had been lost and the patient had no known peripheral artery disease. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported limb ischemia has been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
33693 | 1220648-2023-04054 | 17941420 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17941420&pc=OZD | 09/13/2023 | 09/11/2024 |  | 10/16/2023 |  |  | 09/13/2023 | 10/16/2023 |  |  |  |  | 10/16/2023 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1414254 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 09/19/2023 |  |  |  |  | The user facility's biomedical engineer reported a battery failure on their automated impella controller (aic). the aic will be returned for investigation. there was no patient involvement. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation for the console (aic) battery failure-red alarm has been completed. the data logs confirmed the reported failure. there was a battery failure alarm (transport connection blown/ missing ¿ event set #360) due to low pack voltage. the batt test utility revealed that both batteries were fully depleted and were therefore unable to be charged or communicate with the pbm (console¿s charging circuitry). the console was returned for evaluation. upon functional testing, the batteries were fully depleted. when replaced with new batteries, the issue resolved. the console¿s battery charging circuitry was confirmed to be working as expected. the most likely cause of the batteries being fully depleted is that the console was not connected to ac power once the batteries reached a low capacity, and therefore the batteries were not prevented from fully discharging. d2-corrected common device name. d4-corrected model number. e2-health professional changed to no. f6/f8 should have been left blank in the initial report.
33694 | 1220648-2023-04060 | 17944434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17944434&pc=OZD | 09/16/2023 | 08/08/2024 |  | 10/16/2023 |  |  |  | 10/16/2023 |  |  | 06/30/2025 |  | 10/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024385787 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/16/2023 |  |  |  |  | The user facility reported a 67-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella insertion site was oozing requiring multiple dressing changes. hemoglobin dropped from 15.9 to 14.8. a decision was made to give 500 of albumin but the oozing continued so stiches were added to the insertion site. | It was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that the use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | B.5 additional information was received, and abiomed's medical safety officer review was completed. the investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely use issue since the bleeding was resolved with an additional stitch. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d.6a and d.6b revised from the initial submission of manufacture device report number 1220648-2023-04060 in accordance with updated procedures. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2023-04060 and should not have been. h.6 added code 925 as added since the initial manufacture device report number 1220648-2023-04060 was submitted in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacture device report number 1220648-2023-04060 in accordance with updated procedures.
33695 | 1220648-2023-04061 | 17946555 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17946555&pc=OZD | 09/16/2023 | 01/21/2025 |  | 10/16/2023 |  |  |  | 10/16/2023 |  |  |  |  | 10/16/2023 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1455741 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Power Up | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant in japan reported a aic device startup failure. when the staff turned on the power, the led light turned green, but the screen remained black and froze. staff tried to shut it down by holding down the power button, but it wouldn't turn off. staff tried forcefully shutting down by turning off the battery power and restarting it, but the same problem occurred and staff was unable to use the aic. while in storage, it was plugged into an ac outlet. there was no reported patient harm. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation for the console (aic) power on issues has been completed. data logs were reviewed from the console show that the date and time of the console was incorrect and that the console underwent a system self check failure after no communication from the cpld process. the console was returned for evaluation. during functional testing, inserted the power plug into the outlet and turned on the power, but the monitor did not work. after reassembling it, it started up and confirmed a self-check error. confirm that sau_motor_1 is n.a using fwcheck. after plugging in, the power led did not glow amber. the led on the power module also did not light up, and it was confirmed that the power supply (0042-3007) has malfunctioned. both fuses were blown, they were replaced the power supply with a different one, the blown fuse did not occur again. after replacing the power supply, confirm that the power module led lights up. power was turned on but found a self-check error. the root cause for the system self check failure was a defective impellatronics board. the root cause for the power on issues was a defective power supply unit. corrections to the initial mfr report: d2 updated common device name g1 mdr reporting contact email.
33696 | 1220648-2023-04062 | 17946558 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17946558&pc=OZD | 09/18/2023 | 01/23/2024 |  | 10/16/2023 |  |  | 09/18/2023 | 10/16/2023 |  |  | 06/30/2025 |  | 10/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024371256 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/18/2023 |  |  |  |  | The user facility reported a 79-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the impella cp was being inserted and shortly after the cross clamp of the aorta, there was an unexpected pump stop. the catheter was positioned shallow in the ventricle due to papillary muscle. it was unclear if the cross clamp directly led to the pump stop. multiple unsuccessful attempts were made to restart the impella. the pump was explanted and replaced. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | Since the medical center did not return the impella device, no benchtop analysis was possible. based on the clinical info provided, the data logs were reviewed and showed multiple high mc stops, the cause of the pump stop cannot be determined since no complaint device returned.
33697 | 1220648-2023-04063 | 17946560 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17946560&pc=OZD | 09/16/2023 | 03/18/2024 |  | 10/16/2023 |  |  | 09/18/2023 | 10/16/2023 |  |  |  |  | 10/16/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, EU |  |  | ABIOMED, INC. |  |  | US |  | 1204669 |  | 0042-0000-EU |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 09/18/2023 |  |  |  |  | The user facility reported a 76-year-old male in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. during set-up of the purge system for the procedure, the automated impella controller (aic) did not recognize the disc. the aic was successfully replaced with a backup aic. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the console was returned and tested after arriving in fs. the root cause of this issue could not be determined. purge disc not detected alarms were unable to be reproduced. hairline fracture found on left side of retainer during physical inspection of console. unrelated to the reported product malfunction.
33698 | 1220648-2023-04064 | 17946605 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17946605&pc=OZD | 09/18/2023 | 02/16/2024 |  | 10/16/2023 |  |  |  | 10/16/2023 |  |  | 08/31/2025 |  | 10/16/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024382315 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/18/2023 |  |  |  |  | The user facility reported a 59-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, the patient developed a hematoma at the right axillary access site. the patient was sent to the or for evacuation. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed. the device was not returned for investigation. the root cause of the access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review.
33699 | 1035166-2023-00113 | 17936782 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17936782&pc=DYB | 03/03/2023 | 08/23/2023 |  | 10/14/2023 |  |  | 09/20/2023 | 10/14/2023 |  |  | 10/01/2025 |  | 10/14/2023 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3006 |  |  | OSCOR INC. |  |  | US |  | DP-17198 |  | 0052-3006 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | 48 YR |  | Male |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | We have been notified of a complaint involving an oscor 23 fr introducer (material #0052-3006, lot # dp-17198). patient was suffering from cardiomyopathy and during implant of an impella 5.5 pump, blood was seen at the side of the hemostatic valve. upon removal of the introducer, it was noted that the component was cracked. no blood products were required, and a second axillary kit was available for use as needed. no medical intervention was required and there was no delay in the procedure. patient is a 48-year-old, male and is in stable condition. device was discarded. | [5 empty] | Conclusion not yet available, evaluation in process. a follow-up will be submitted as soon as the investigation is complete. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | The following sections were updated in follow-up 1. no product was provided for analysis. procedure completed successfully with different device. there was no adverse event with the patient reported. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. no product was returned to oscor inc. for evaluation; however, a complaint notification was provided. without the return of the actual device in question for evaluation, oscor inc. is unable to determine the exact cause for this incident. at this time, it is not possible to assign a definitive root cause for the event as reported. the complaint is non-verifiable as the product was not returned for evaluation. no corrections or corrective actions will be taken. device was not returned for analysis and review of the manufacturing documents did not reveal any discrepancies that would have contributed to this event. this complaint is non-verifiable. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may have not been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
33700 | 1220648-2023-04028 | 17931155 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17931155&pc=OZD | 09/05/2022 | 10/13/2023 |  | 10/13/2023 |  |  | 09/05/2022 | 10/13/2023 |  |  | 03/31/2024 |  | 10/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023199538 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/05/2022 |  |  |  |  | The user facility reported a 51yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the pump was falling out of the ventricle after the patient sneezed. the pump could not be readvanced, therefore, was replaced. there was no harm to the patient. | [5 empty] | The investigation for the positioning issue has been completed. the device was not returned for evaluation. however, the clinical report provided sufficient information to determine the root cause. therefore, the cause of the positioning issue was patient movement from the patient sneezing. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33701 | 1220648-2023-04029 | 17933087 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17933087&pc=OZD | 02/07/2023 | 10/13/2023 |  | 10/13/2023 |  |  | 02/07/2023 | 10/13/2023 |  |  | 12/31/2024 |  | 10/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024288216 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Male |  |  | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/07/2023 |  |  |  |  | The user facility reported a male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the pump was moving toward the mitral valve. repositioning was attempted to move impella but met with resistance against the mitral valve. the access site was explored and a graft twist was noted. the chest was opened at the top two sternal wires and found that the sternum closure had pinched the graft. impella was explanted and the graft was closed. there was no patient harm reported. | [5 empty] | No product was returned. only the clinical report was analyzed. the cause of the positioning issue was use related due to sternal wires pinching the graft. the failure mode will be monitored and trended.
33702 | 1220648-2023-04030 | 17933172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17933172&pc=OZD | 11/30/2022 | 10/13/2023 |  | 10/13/2023 |  |  | 11/30/2022 | 10/13/2023 |  |  | 09/30/2024 |  | 10/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023257465 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/30/2022 |  |  |  |  | The user facility reported a 77yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that during support, a pump stop alarm was triggered. a cable between the sterile sleeve and proximal to the red impella plug has been severed. the pump was removed and replaced. there was no patient harm reported. | [5 empty] | The investigation for the reported pump stop has been completed. the device was not returned for evaluation. however, the based on the clinical report, sufficient details provided can be used to determine the root cause. therefore, the cause of the pump stop was an electrical open most likely from excessive force. instructions for use for the related event are as follows: ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33703 | 1220648-2023-04031 | 17933537 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17933537&pc=OZD | 09/16/2023 | 01/29/2024 |  | 10/13/2023 |  |  | 09/16/2023 | 10/13/2023 |  |  | 06/30/2025 |  | 10/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024377403 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/16/2023 |  |  |  |  | The user facility reported a 71-year-old female with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. support was ongoing when there were signs of hemolysis despite being in appropriate pump position and having the preload appropriate. the team attempted to troubleshoot by adding tissue plasminogen activator to the purge, but found instead that when this failed, there was a need to replace the pump. a second new impella cp was successfully placed. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. the data logs were not available to review. the root cause of the hemolysis issue cannot be determined due to insufficient data returned. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
33704 | 1220648-2023-04032 | 17933751 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17933751&pc=OZD | 10/16/2022 | 10/13/2023 |  | 10/13/2023 |  |  | 10/16/2022 | 10/13/2023 |  |  | 04/30/2024 |  | 10/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023214593 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/16/2022 |  |  |  |  | The user facility reported a 64yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that an impella defective alarm occurred during impella support. the alarm appeared after support was initiated at level p2. the console was swapped in an attempt to troubleshoot the issue but the failure remained. as a result, the pump was subsequently removed and a new impella 5.5 pump was implanted without further issue. | [5 empty] | The investigation for the reported defective alarm has been completed. data logs showed that the "impella defective" alarm triggered twice in quick succession after the pump was started and ran for several seconds. no damage or abnormalities were found on the returned pump. when the pump was connected to an aic, it started and ran with no issues or alarms. therefore, it was determined that the cause of the pump not being detected was use related. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33705 | 1220648-2023-04033 | 17934478 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17934478&pc=OZD | 09/15/2023 | 01/26/2024 |  | 10/13/2023 |  |  | 09/15/2023 | 10/13/2023 |  |  | 03/31/2025 |  | 10/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2024334322 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 42 YR |  | Male |  | Non Hispanic | Asian | Malfunction | Yes | Yes | Pumping Stopped | [6 empty] | Non specific EKG/ECG Changes | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/15/2023 |  |  |  |  | The user facility in japan reported a 42-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. shortly after the cp pump started, there was a pump alarm and the pump stopped. the impella cp was replaced and ecmo was added due to the malfunction. | [5 empty] | The root cause of the pump stop/ did not start issue was not determined since product was returned for investigation but the pump stop could not be reproduced. also, no logs were returned for investigation. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
33706 | 1220648-2023-04034 | 17934749 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17934749&pc=OZD | 07/14/2023 | 10/13/2023 |  | 10/13/2023 |  |  | 09/20/2023 | 10/13/2023 |  |  | 03/31/2025 |  | 10/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024346134 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/13/2023 |  |  |  |  | The user facility reported an 82-year-old male with left ventricle dysfunction was to be implanted with an impella cp device for mechanical circulatory support. the impella pump was not inserted due to optical sensor failure. a new device was set up and implanted after the first pump failed. there was no adverse event to the patient as a result of this incident. | [5 empty] | The investigation into the optical system error alarm has been completed. pump was not returned for investigation. the data logs show an optical system error alarm, along with the air gap trend. based on the logs, it is evident that when the pump was unplugged and re-plugged, the 5s parameters were within specification. the cause of the optical signal failure was most likely an air gap from between the aic and the patient cable plug based on data log review. the failure modes will be monitored and trended.
33707 | 1220648-2023-04035 | 17936450 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17936450&pc=OZD | 04/19/2023 | 08/02/2024 |  | 10/13/2023 |  |  | 09/14/2023 | 10/13/2023 |  |  | 12/31/2024 |  | 10/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024315995 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | Tachycardia | [10 empty] | Required Intervention | [28 empty] | Initial submission | Followup |  |  |  | 09/14/2023 |  |  |  |  | The decision was made to withdraw care and the patient expired. per medical safety officer review use of the device cannot conclusively be associated with the outcome. | Long-term outcome and quality indicator (loqi) impella registry notifications were received on 2023-09-14 reporting a 77 year old male patient enduring limb ischemia prompting warming of the patient in an attempt to resolve. it was noted that a distal perfusion catheter was placed, however, it's unclear if this was proactive or reactive in nature. the patient also endured hemolysis that prompted a blood transfusion. lastly, the patient sustained vt that prompted a cardioversion and repositioning of the pump. |  |  |  |  | The investigation of limb ischemia, hemolysis, and vf/vt has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b5 updated with additional information . h6 component code added the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-04035. d4 model number. in accordance with updated procedures, section d6 has been revised from the initial submission. f6/f8-should have been left blank . g1 reporting contact fax number missing. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
33708 | 1220648-2023-04036 | 17936486 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17936486&pc=OZD | 09/13/2023 | 01/26/2024 |  | 10/13/2023 |  |  | 09/13/2023 | 10/13/2023 |  |  |  |  | 10/13/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1106952 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Activation Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that performing software update on automated impella controller (aic). after turning it back on, the battery failure (red alarm) came up. restarting several times didn¿t resolve the issue. | [5 empty] | The automated impella controller (aic) device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - battery failure (red alarm) has been completed since the original report was filed. the console was returned for investigation. the failure mode was reproduced on an engineering lab bench. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure, a known pattern failure.
33709 | 1220648-2023-04037 | 17936487 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17936487&pc=OZD | 09/13/2023 | 04/19/2024 |  | 10/13/2023 |  |  | 09/13/2023 | 10/13/2023 |  |  | 04/30/2025 |  | 10/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024353479 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/13/2023 |  |  |  |  | The user facility reported an 83-year-old female noted as high risk percutaneous coronary intervention (hrpci) was implanted with an impella cp device for mechanical circulatory support. impella turned on auto and suction on p2. activated clotting time (act) not obtained before pump turned on. act 237 after pump in body. patient had small aortic arch possible suction due to patient anatomy. impella protected pci supported throughout case on p-9 with 2.6 l/min of flow with suction per physician order. device explanted without complication. there was no patient harm, | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow issue has been completed since the original report was submitted. the device was not returned for investigation. the cause of the low pump flow issue was not determined since product was not returned and sufficient clinical information was not provided.
33710 | 1220648-2023-04038 | 17936488 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17936488&pc=OZD | 09/13/2023 | 01/22/2024 |  | 10/13/2023 |  |  | 09/13/2023 | 10/13/2023 |  |  |  |  | 10/13/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, EU |  |  | ABIOMED, INC. |  |  | US |  | 1354144 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Failure of Device to Self-Test | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 09/13/2023 |  |  |  |  | The user facility reported their automated impella controller (aic) started up with a blue screen. this is a system self-check error. the console to be returned for investigation and repair. this event occurred during set-up and there was no patient involvement. | [5 empty] | The investigation into the aic - system self-check error has been completed since the original report was filed. the console was returned for investigation. it was determined that the cause of the aic issue was contamination of a communication line on the pbm. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
33711 | 1220648-2023-04039 | 17936491 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17936491&pc=OZD | 09/13/2023 | 01/29/2024 |  | 10/13/2023 |  |  | 09/14/2023 | 10/13/2023 |  |  | 04/30/2025 |  | 10/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2024334308 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/14/2023 |  |  |  |  | The user facility reported a patient (unknown age or gender) in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella cp support was ongoing when the pump was explanted after six days. upon the explant the team observed a large clot burden at the pump cannula and inlet. the pump was explanted and the patient suffered a cerebral infarct/cardio vascular accident. there were no problems with the pump operation, a thrombectomy was performed. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. product and data were not returned for review. the root cause of cerebral hemorrhage was most likely due to patient condition and unknown relation to impella. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33712 | 1220648-2023-04040 | 17936493 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17936493&pc=OZD | 09/14/2023 | 01/24/2024 |  | 10/13/2023 |  |  | 09/14/2023 | 10/13/2023 |  |  | 06/30/2025 |  | 10/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024375584 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 73-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that there were high purge pressure alarms overnight leading to pump failure. multiple attempts to restart the pump failed. the patient was brought to the ccl and impella was removed. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation into the high purge pressure and the pump stop issues has been completed since the original report was filed. product was returned for evaluation. data logs confirmed high purge pressure and high purge pressure alarms. there was a severe kink in the purge line of the catheter with blood ingress into the purge line as well as a hole in the catheter. additionally, there was blood ingress in the purge gap. it was determined that the root cause of the pump stop was a kinked purge line at the 52 cm catheter marking.
33713 | 1220648-2023-04041 | 17936494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17936494&pc=OZD | 09/14/2023 | 01/23/2024 |  | 10/13/2023 |  |  | 09/14/2023 | 10/13/2023 |  |  | 06/30/2025 |  | 10/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024372300 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Female |  |  |  | Malfunction | No | Yes | Material Deformation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/14/2023 |  |  |  |  | The user facility reported a female (unknown age and race) noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the physician noted there was a visible kink in the cannula approx. 1 cm below the outlet on fluoroscopy. the physician tried to reposition to relieve kink. inlet and outlet appeared to be in appropriate position in proximity to aortic valve. physician stated that there was no suction and continued throughout the case despite dropping p levels, all the way to p2. volume status did not seem to be the problem given the patients presentation. pump was explanted and replaced with another with no further issues. there was no patient harm. | [5 empty] | The investigation into the product damage (cannula kink) and the low pump flow has been completed since the original report was filed. pump was returned for investigation. cannula kink was observed near motor. also, doctor reported reposition of the pump did not solve the cannula kink. data logs shows the low pump flow with suction alarm during the case run. the cause of the low pump flow issue was a kinked cannula. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
33714 | 1220648-2023-04042 | 17936495 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17936495&pc=OZD | 09/14/2023 | 02/09/2024 |  | 10/13/2023 |  |  | 09/14/2023 | 10/13/2023 |  |  | 03/31/2025 |  | 10/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024327461 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/14/2023 |  |  |  |  | The user facility reported a 30-year-old female with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the impella 5.5 was initially attempted in the right axillary artery, but the patient had oozing from around the graft at anastomosis and extra stitching and suturing were done. the impella was placed in the left axillary artery. day three of support and the patient was transferred from hospital bed to a transport cot with no issues. during automated impella controller (aic) to aic transfer, the white impella plug was unplugged from one aic and plugged into the new aic, immediately, "defective impella catheter" alert happened. troubleshooting was unsuccessful. the impella and purge were not running at this time. the patient was taken to the operation room and the impella was removed and a new 5.5 was successfully placed. the patient remained stable. | [5 empty] | The investigation into the access site bleeding ¿ major and the pump stop issues has been completed since the original report was filed. the device was returned for investigation. imc logs confirmed event # 3 was triggered when pump was connected to the console. pump was connected to 2 consoles to reproduce the failure mode. both consoles could not detect the pump & triggered impella defective alarms. visual inspection found a hole on the catheter at mark 30 & blood inside the pump plug. the cause of the pump not detected was blood ingress into red pump plug due to a hole in the catheter resulting from improper use. the cause of the bleeding was use related since the bleeding occurred from an insufficient stitching of the graft to the vessel. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
33715 | 1220648-2023-04043 | 17936524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17936524&pc=OZD | 09/15/2023 | 01/23/2024 |  | 10/13/2023 |  |  | 09/15/2023 | 10/13/2023 |  |  | 07/18/2023 |  | 10/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - EU |  |  | ABIOMED, INC. |  |  | US |  | 2024363538 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Power Up | Failure to Prime | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/15/2023 |  |  |  |  | The user facility reported a patient (unknown age, gender and race) noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the automated impella controller (aic) did not start set-up/priming with the impella cp pump. the team changed to a different aic but the problem was not resolved. the team changed to a different impella cp and the aic was recognizing and set-up priming worked without any problems. there was no patient harm. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | The investigation into the priming issue has been completed since the original report was filed. pump and purge cassette was returned for investigation. data logs were not provided. during test run, pump and purge cassette were primed without any issue. no other physical abnormalities were observed. there was no issue found during the case. the device functioned/operated to specification.
33716 | 1220648-2023-04044 | 17936530 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17936530&pc=OZD | 09/15/2023 | 02/14/2024 |  | 10/13/2023 |  |  | 09/15/2023 | 10/13/2023 |  |  | 07/31/2025 |  | 10/13/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024378412 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | Thrombosis/Thrombus | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/15/2023 |  |  |  |  | The user facility reported a 66-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the patient is heparin-induced thrombocytopenia positive. the patient developed a hematoma and bleeding at the access site. sandbags for pressure did not completely stop the oozing. hematoma washout and start anticoagulation planned. no blood products required. clots found and removed during hematoma washout. | [5 empty] | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for investigation. according to the clinical evaluation, the cause of the access site bleeding issue was high patient act values since the decision to halt the anticoagulant to control the bleeding. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33717 | 1220648-2023-04003 | 17918720 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17918720&pc=OZD | 09/13/2023 | 09/17/2024 |  | 10/12/2023 |  |  | 09/13/2023 | 10/12/2023 |  |  |  |  | 10/12/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1205057 |  | 0042-0000 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Activation Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Additional information noted patient care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 65-year-old male patient was implanted with an impella 5.5 device for mechanical circulatory support. during support, it was noted that the aic experienced a controller failure alarm. the automated impella controller (aic) was switched to a back up and the issue was resolved. there was no patient harm as a result of the failure. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-04003 was provided. | The automated impella controller (aic) device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
33718 | 1220648-2023-04004 | 17918747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17918747&pc=OZD | 09/13/2023 | 01/23/2024 |  | 10/12/2023 |  |  | 09/13/2023 | 10/12/2023 |  |  | 06/30/2025 |  | 10/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024370701 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 75-year-old male patient underwent a high-risk percutaneous coronary intervention (hrpci) and received hemodynamic support from an impella cp smartassist heart pump. the peel away introducer sheath was damaged / split at the tip, causing excessive bleeding during the case. due to vascular injury, the attending physician decided to replace the entire impella catheter with a new device on the opposite femoral artery. initial femoral access site has been closed using an angio-seal device. operation was then completed without further problems. | [5 empty] | The investigation into vascular damage-peripheral vessel and introducer damage-mechanical has been completed. no product was returned for evaluation. the root cause of the vascular damage - peripheral vessel cannot be determined due to insufficient clinical information. the root cause of the introducer damage-mechanical is most likely due to a sheath tip split but it cannot be confirmed if it was due to usage or a malfunction. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33719 | 1220648-2023-04005 | 17918940 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17918940&pc=OZD | 09/13/2023 | 01/22/2024 |  | 10/12/2023 |  |  | 09/13/2023 | 10/12/2023 |  |  | 06/30/2025 |  | 10/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024383197 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Insert | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A patient of unknown age received hemodynamic support from an impella cp smartassist heart pump during high-risk percutaneous coronary intervention (hrpci). after the pci procedure, the peel-away introducer sheath was removed. the repositioning sheath could not be placed. obesity has been reported as a contributing factor. the impella cp heart pump was subsequently removed and the access site was surgically closed. the patient received catecholamines and extracorporeal membrane oxygenation (va-ecmo) support. the patient was subsequently stable. | [5 empty] | The investigation into the vascular damage has been completed since the original report was filed. the pump was returned for investigation. as per clinical, there was no introducer damage, but patient had vascular damage during repo unit insertion due to adiposity patient condition. no other abnormalities were found. the cause of the vascular damage - peripheral vessel issue was patient condition related due to obesity (adiposity) leading to resistance during repo unit insertion. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33720 | 1220648-2023-04006 | 17919051 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17919051&pc=OZD | 09/14/2023 | 10/12/2023 |  | 10/12/2023 |  |  | 09/14/2023 | 10/12/2023 |  |  | 09/30/2023 |  | 10/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022142501 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 65-year-old female noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. during support, it was noted that a fluid leak developed at the purge sidearm. the sidearm leak was treated by tape being applied over a crack on the reservoir. ecpella support remained ongoing despite the noted leak. there was no patient harm as a result of the reported issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. although it was noted that the leak had originated from a crack in the sidearm, a definitive cause for the fracture could not be established. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
33721 | 1220648-2023-04007 | 17919082 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17919082&pc=OZD | 09/14/2023 | 02/02/2024 |  | 10/12/2023 |  |  | 09/14/2023 | 10/12/2023 |  |  | 08/30/2025 |  | 10/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024381524 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Infusion or Flow Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 75-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant had a cp support that was ongoing with suction/low flows. the team treated the low flows and access site bleed with one unit of blood infused and three vials of albumin. the team wished to upgrade to a 5.5 support but vessel sizing not appropriate. patient remained on cp support. | [5 empty] | The investigation into the low pump flow and the access site bleeding has been completed. the device was not returned for investigation. the data logs indicated several suction alarms, with the placement signal trending downward and the low pump flow was observed at the end of the case. the cause of the low pump flow issue was most likely patient condition related, based on data log review, and provided clinical details. the cause of the access site bleeding issue was not determined since the product was not returned and sufficient clinical information was not provided. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33722 | 1220648-2023-04008 | 17919124 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17919124&pc=OZD | 09/15/2023 | 02/05/2025 |  | 10/12/2023 |  |  | 09/15/2023 | 10/12/2023 |  |  |  |  | 10/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1230742 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 67-year-old male patient was implanted with an impella 5.5 device for mechanical circulatory support. during support, it was noted that the aic failed to detect installed purge discs after many attempts and multiple purge cassettes. the aic was switched to a back up console. there was no patient harm as a result of the failure. | [5 empty] | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. according to the data and device analysis the root cause of the issue was a missing purge flag and cracked purge retainer. d2 common device name revised on manufacturer device report 1220648-2023-04008 in accordance with updated procedures g1 reporting contact email revised on manufacturer device report 1220648-2023-04008 in accordance with updated procedures g1 reporting contact fax number missing n manufacturer device report 1220648-2023-04008. | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
33723 | 1220648-2023-04009 | 17919176 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17919176&pc=OZD | 09/15/2023 | 01/25/2024 |  | 10/12/2023 |  |  | 09/15/2023 | 10/12/2023 |  |  | 03/31/2025 |  | 10/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP JP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024326121 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 26 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 26-year-old male patient was presented for hemodynamic support via impella cp implant. during insertion of the pump, an echocardiogram confirmed that the impella pigtail had penetrated the left ventricle which required surgical repair. the impella was removed. | [5 empty] | The investigation into the ventricle perforation issue has been completed. the device was not returned for investigation. based on clinical description, the root cause of ventricle perforation was most likely due to pump was in improper position. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33724 | 1220648-2023-04010 | 17919184 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17919184&pc=OZD | 07/13/2022 | 10/12/2023 |  | 10/12/2023 |  |  | 07/13/2022 | 10/12/2023 |  |  | 04/30/2024 |  | 10/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023213205 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/13/2022 |  |  |  |  | The user facility reported a 58yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that while on support, when the surgical site/anastamosis failed. there was blood loss. the team made the choice to explant the impella. the patient was hemodynamically stable without impella support. | [5 empty] | The investigation for the reported access site bleed has been completed. the cause of the access site bleeding was not able to be determined since neither the product nor sufficient clinical information was provided. the instructions for use provides details on prevention of trauma and bleeding with the catheter. it states ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33725 | 1220648-2023-04011 | 17919226 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17919226&pc=OZD | 08/06/2021 | 10/12/2023 |  | 10/12/2023 |  |  | 08/06/2021 | 10/12/2023 |  |  | 03/31/2023 |  | 10/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022064511 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/06/2021 |  |  |  |  | It was reported that during insertion. there was bleeding through the graft around the sheath due to the patient's anatomy. the 5.5 pump was removed and an impella cp was placed. the patient received two units of blood for blood loss during the attempted insertion. | [5 empty] | The investigation for the reported access site bleeding and the delivery issues has been completed. the site reported that there was difficulty inserting impella 5.5 due to small patient anatomy. additionally, the site reported bleeding around the sheath through the graft during insertion. the decision was made to insert impella cp instead. the root cause of the delivery issues was related to patient anatomy. a root cause for the reported bleeding could not be determined through this evaluation as no product was returned for evaluation. the instructions for use provides suggestions on preventing trauma to insertion site. it states ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33726 | 1220648-2023-04012 | 17920621 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17920621&pc=OZD | 09/13/2023 | 01/30/2024 |  | 10/12/2023 |  |  | 09/13/2023 | 10/12/2023 |  |  | 04/30/2025 |  | 10/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024351287 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/13/2023 |  |  |  |  | The user facility reported a 71-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the physician was using a single access and the diagnostic catheter felt tight and was not able to torque. the physician downsize to a 6fr sheath and began working on the coronary access. during this time suction alarms started, and the pump performance level setting(p-level) dropped. the impella was repositioned, but the p level was unable to increase. the decision was made to switch out the impella. the new impella was successfully implanted and there was no harm to the patient. | [5 empty] | The investigation into the low pump flow issue has been completed. the pump was returned for investigation. data logs confirmed the failure mode with multiple mc spikes follow flow dropped to (b)(4) that triggered suction response and suction alarms. the pump then was stopped manually and disconnected from the console. the impeller blades were visually inspected and found the impeller blades were damaged. no other issues were found on the rest of the pump. the root cause of low flow was a fractured impeller blade. the fracture was characteristic of an interaction with another device but that was unable to be definitively determined since fluoroscopic imaging and/or additional clinical details were not provided. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
33727 | 1220648-2023-04013 | 17920687 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17920687&pc=OZD | 09/13/2023 | 02/26/2024 |  | 10/12/2023 |  |  | 09/13/2023 | 10/12/2023 |  |  | 05/31/2025 |  | 10/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024358538 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/13/2023 |  |  |  |  | The user facility reported a 67-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, a distal perfusion cannula was placed for lack of distal flow. | [5 empty] | The investigation into the reported limb ischemia has been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33728 | 1220648-2023-04014 | 17923098 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17923098&pc=PYX | 09/13/2023 | 02/22/2024 |  | 10/12/2023 |  |  | 09/13/2023 | 10/12/2023 |  |  | 04/30/2025 |  | 10/12/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024337967 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/13/2023 |  |  |  |  | The user facility reported a 65-year-old female with right heart failure was implanted with an impella rp device for mechanical circulatory support. during support, it was noted that the patient developed hemolysis and the pump was subsequently explanted. | [5 empty] | The investigation into the reported hemolysis been completed since the original report was filed. the cause of the hemolysis cannot be determined since the complaint device not returned for investigation and insufficient clinical data returned. data logs reviewed: no mc spikes observed. no positioning issue occurred. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
33729 | 1220648-2023-04015 | 17923232 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17923232&pc=OZD | 09/13/2023 | 02/02/2024 |  | 10/12/2023 |  |  | 09/13/2023 | 10/12/2023 |  |  | 06/30/2025 |  | 10/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024357260 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Non Hispanic | White | Malfunction | Yes | Yes | Restricted Flow rate | Increase in Pressure | Pumping Stopped |  |  |  |  | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/13/2023 |  |  |  |  | The user facility reported a 48-year-old female with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support as a bridge to transplant. the impella was experiencing "high purge pressure/purge system blocked" alarms. tissue plasminogen activator was hung, however the alarm did not resolve and purge pressure high/purge system blocked alarms persisted. the following day, it was reported there was a pump stop with "impella stopped motor current high" alarms. staff explanted the impella and a large clot was observed on the outlet cage. also, an unknown left ventricle (lv) thrombus was identified immediately following explant. decision was made to treat lv thrombus with thrombolytics. the implant of a new 5.5 was aborted. | [5 empty] | The investigation into the reported high purge pressure, pump stop and thrombus has been completed since the original report was filed. impella was returned with the pump catheter cut near the motor housing side. the pump was visually inspected and found biomaterial to be ingested in the outflow, purge gap, impeller and the inflow. catheter was stripped near the motor housing from 20 cm marking and observed blood to be ingressed in the purge line. no other abnormalities were found during field investigation. data logs were reviewed, and purge pressure rise was observed with likely biomaterial ingestion with saturated flow along with motor current spike at the end. pump stop with high motor current alarm was observed. the cause of the high purge pressure issue was most likely biomaterial based on the field investigation and data log review. the cause of the pump stop was due to blood ingress. the cause of the blood ingress was unknown since the pump was returned with catheter cut open on the motor side. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. per the ifu: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
33730 | 1220648-2023-04017 | 17925627 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17925627&pc=OZD | 05/16/2023 | 10/12/2023 |  | 10/12/2023 |  |  | 05/16/2023 | 10/12/2023 |  |  | 03/31/2025 |  | 10/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2024323947 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/16/2023 |  |  |  |  | The user facility in the eu reported a 60-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. due to cross clamp with the clamp, a small lesion occurred. this was a use error, as the wrong clamp was used for clamping aorta during surgical procedure. a minor surgical intervention (sutures) resolved the issue. | [5 empty] | Based on the clinical account, the root cause of the patient injury was due to a use issue, as the md used a hard clamp on the aorta causing a lesion. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
33731 | 1220648-2023-04016 | 17925637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17925637&pc=OZD | 06/02/2023 | 10/12/2023 |  | 10/12/2023 |  |  | 06/02/2023 | 10/12/2023 |  |  | 01/31/2025 |  | 10/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024331513 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/02/2023 |  |  |  |  | The user facility reported a 74-year-old female with cardiomyopathy undergoing surgery was implanted with an impella cp for mechanical circulatory support. after two days of support, the impella cp pump stopped. . the cp stopped after the pump was cross-clamped for the surgical procedure for the mitral valve. the pump would not restart post cross-clamp. the pump was removed but not noted to be replaced. no lasting harm or injury from the stop. | [5 empty] | The investigation into the "pump stop" has been completed. the device was returned for evaluation. returned complaint cp pump was visually inspected and found to have biomaterial observed around impella. pump connected to automated impella controller (aic) and pump stop alarm reproduced. pump soaked to remove the biomaterial. biomaterial removed and pump run in water successfully without stopping within specific p-level and flow rate. root cause of the pump stop issue was biomaterial ingested due to patient condition related and after removing it the pump worked perfect within specific flow range. the failure modes will be monitored and trended. instructions for use for pump stop event: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
33732 | 1220648-2023-04018 | 17925813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17925813&pc=OZD | 09/14/2023 | 02/26/2024 |  | 10/12/2023 |  |  | 09/14/2023 | 10/12/2023 |  |  | 06/30/2025 |  | 10/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024368587 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 31 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/14/2023 |  |  |  |  | The user facility in the eu reported a 31-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. after the impella was implanted by an experienced implanter without any problems, it has been noticed in the icu that the leg is underperfused. the patient is prone to dissections and has some in other places besides the one on the leg. since the vessel was too small for a 5.5, a new cp was implanted via the axillary artery. | [5 empty] | The investigation into the limb ischemia reversible issue has been completed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33733 | 1220648-2023-04019 | 17925963 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17925963&pc=OZD | 09/14/2023 | 02/26/2024 |  | 10/12/2023 |  |  | 09/14/2023 | 10/12/2023 |  |  | 06/30/2025 |  | 10/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024373884 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 20 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/14/2023 |  |  |  |  | The user facility reported a 20-year-old female (race unknown) in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient's vessel was too small in diameter for the 14fr and the team was concerned about limb ischemia. the team placed a 5fr sheath on left groin for fem-fem bypass. | [5 empty] | The investigation into the reported limb ischemia has been completed since the original report was filed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33734 | 1220648-2023-04020 | 17926014 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17926014&pc=OZD | 09/14/2023 | 10/12/2023 |  | 10/12/2023 |  |  | 09/14/2023 | 10/12/2023 |  |  | 05/31/2025 |  | 10/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024363604 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Material Deformation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/14/2023 |  |  |  |  | The user facility reported a 65-year-old female in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the cannula kinked in the patient's aorta. patient had a tortuous anatomy. the impella was replaced. there was no harm to the patient. | [5 empty] | The investigation into the kinked cannula has been completed. no product was returned. as per clinical evaluation, impella cannula kinked after insertion with tortuous vessel condition observed around aorta earlier with imaging. the cause of the product damage - cannula kink issue was use related due to improper technique of advancing pump for insertion with tortuous vessel conditions observed. the failure modes will be monitored and trended.
33735 | 1220648-2023-04024 | 17926231 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17926231&pc=OZD | 09/10/2023 | 09/22/2025 |  | 10/12/2023 |  |  | 09/13/2023 | 10/12/2023 |  |  | 05/31/2025 |  | 10/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024354026 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 09/13/2023 |  |  |  |  | The user facility reported a 77-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the md noted the optical signal to be less reliable. the pump position was evaluated and the team adjusted position. the pump could then not be repositioned across the valve and was explanted and replaced. no lasting harm as the replacement pump was placed without any problems. | [5 empty] | The investigation into the reported positioning issues has been completed since the original report was filed. the returned complaint device visually inspected. biomaterial was ingested around the impeller. the optical sensor was tested and found to be within specification range. the data logs were reviewed and mc spike occurred followed by impella in aorta alarms. the root cause of the positioning issue was use related because the pump was pulled back across the valve. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the root cause of false impella in aorta alarm was the ingestion of biomaterial. all pertinent information available to abiomed has been submitted at this time. d6a/d6b updated with additional information. f6 and f8 should have been left blank from the initial manufacturer device report number 1220648-2023-04024. g1 updated with correct mdr reporting contact email.
33736 | 1220648-2023-04025 | 17926242 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17926242&pc=OZD | 09/13/2023 | 03/28/2024 |  | 10/12/2023 |  |  | 09/13/2023 | 10/12/2023 |  |  | 03/31/2025 |  | 10/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP JP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024366046 |  | 0048-0024 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 64-year-old male patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the medical team had an impella cp stop during the support. the pump was unable to be restarted despite all typical troubleshooting measures. the pump was explanted and replaced when patient condition deteriorated without support. ecmo was additionally placed. | [5 empty] | The investigation for the reported pump stop has been completed. the pump was returned for evaluation. visual inspection found blood ingress in the purge lumen and other related components. high purge pressures (pp) provided during bench testing. data logs showed high pp immediately during case start and remained for the rest of the case with purge system blocked and high pp alarms. logs also show about 16 minutes from the pump starting, there was a high motor current pump stop that could not be restarted. the root cause of the temporary pump stop was blood ingress due to a purge blockage from damage to the reservoir to filter joint. a data trend chart was pulled and this is the only instance of a temporary pump stop caused from a purge blockage of the reservoir to joint filter due to a product malfunction. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿.
33737 | 1220648-2023-04021 | 17926252 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17926252&pc=OZD | 04/16/2023 | 10/12/2023 |  | 10/12/2023 |  |  | 04/16/2023 | 10/12/2023 |  |  | 11/30/2024 |  | 10/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2023278791 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/16/2023 |  |  |  |  | The user facility reported a 58-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. after a red alarm for high purge pressure/low purge flow at the 7th day of support team decides to replace the purge cassette. after some hours (approximately 7) same problem so the physicians decide to replace the impella catheter. they found large clot inside the inflow cage. patient stable. | [5 empty] | The returned pump was inspected for the cause of the low pump flow and biomaterial was found near the impeller and occluding the inflow cage. the cause of the low pump flow was biomaterial. this report is being filed as part of a retrospective review of historical records.
33738 | 1220648-2023-04023 | 17926732 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17926732&pc=OZD | 09/12/2023 | 04/26/2024 |  | 10/12/2023 |  |  | 09/12/2023 | 10/12/2023 |  |  | 06/30/2025 |  | 10/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024364322 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/12/2023 |  |  |  |  | The user facility reported a 51-year-old male (unknown race) was escalated from an impella cp to an impella 5.5. during support, the team noticed a purge leak at the yellow luer lock connection to purge side arm. the cassette was exchanged but the leak continued. is was small leak. pump flows (pf) and pump pressure (pp) have not been affected. the team will be monitoring closely for any changes to pf/pp. | [5 empty] | The investigation into the purge leak ¿ pump issue has been completed since the original report was filed. the pump was returned, tested, and evaluated. the failure could not be reproduced. the root cause of the purge leak - pump could not be found as the leak could not be reproduced. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
33739 | 1220648-2023-04026 | 17927084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17927084&pc=OZD | 06/07/2023 | 10/12/2023 |  | 10/12/2023 |  |  | 06/07/2023 | 10/12/2023 |  |  | 01/31/2025 |  | 10/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - EU |  |  | ABIOMED, INC. |  |  | US |  | 2024316618 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/07/2023 |  |  |  |  | The user facility reported a 61-year-old female with heart transplant rejection was implanted with an impella cp for mechanical circulatory support. the impella cp was supporting a patient for 22 days when the pump stopped and was re-started. the team was utilizing the cp beyond indication and was alerted that the pump may fail. the pump remained on for support despite the issue. | [5 empty] | The investigation into the pump stop has been completed. product & data were not returned for review. the pump stop occurred on day 22 which is 14 days over the 8 day design life of the impella cp optical. the root cause of the pump stop was unable to be determined as limited clinical details were provided and no product was returned for review. the failure modes will be monitored and trended. instructions for use for pump stop: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
33740 | 1220648-2023-04022 | 17927088 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17927088&pc=PYX | 09/12/2023 | 10/12/2023 |  | 10/12/2023 |  |  | 09/12/2023 | 10/12/2023 |  |  | 06/30/2025 |  | 10/12/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024366098 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Material Deformation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/12/2023 |  |  |  |  | The user facility reported a 35-year-old female post cardiotomy cardiogenic shock and with low cardiac output syndrome was implanted with an impella rp for mechanical circulatory support. the impella rp delivered with difficulty via the 23fr sheath. the team observed the rp to have kinked during attempts. the rp was replaced successfully. | [5 empty] | The investigation into the delivery issues with kinked catheter has been completed. the product was returned and evaluated. the cause of the delivery issues was use related due to improper technique. the failure modes will be monitored and trended.
33741 | 1220648-2023-04027 | 17927264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17927264&pc=OZD | 06/10/2023 | 11/19/2024 |  | 10/12/2023 |  |  | 06/10/2023 | 10/12/2023 |  |  | 12/31/2024 |  | 10/12/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024315410 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/10/2023 |  |  |  |  | The user facility reported a 58-year-old female in advanced heart failure was implanted with an impella device for mechanical circulatory support. the pump was inserted, however, the md was unable to reposition because the catheter would not unlock when the button was being pressed. there was no harm to the patient as a result of this incident. | [5 empty] | The cause of positioning issue cannot be determined since the complaint device not retuned for investigation. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The cause of positioning issue cannot be determined since the complaint device not retuned for investigation the impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
33742 | 1220648-2023-03984 | 17911573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17911573&pc=OZD | 11/16/2022 | 10/11/2023 |  | 10/11/2023 |  |  | 11/16/2022 | 10/11/2023 |  |  | 02/29/2024 |  | 10/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023191053 |  | 1000211 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/16/2022 |  |  |  |  | The user facility reported a 66-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant in japan had a 5.5 support in which had a thrombosis at subrenal artery. there was no harm or injury or allegation made against the impella contributing to the thrombosis. there was heparin-induced thrombocytopenia¿hit+ concerns. the impella had flow drops during the support. | [5 empty] | The left log show suction events and preventing retrograde flow alarms during the end of the case. a visual inspection of the pump did not reveal any abnormalities. no kinks or biomaterial was observed. the pump was attached to a purge cassette and an aic and was run. the pump operated to specification and the low pump flow was not reproduced. the cause of the low pump flow was patient condition related. this report is being filed as part of a retrospective review of historical records.
33743 | 1220648-2023-03985 | 17911591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17911591&pc=OZD | 11/01/2022 | 10/11/2023 |  | 10/11/2023 |  |  | 11/01/2022 | 10/11/2023 |  |  | 06/30/2024 |  | 10/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2023243430 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 1 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/01/2022 |  |  |  |  | The user facility reported an unspecified patient age/gender/race admitted in cardiogenic shock was implanted with an impella device for mechanical circulatory support. after implanting and starting the pump, the purge pressure increased immediately. a thrombus was seen in the echo. pump has successfully been replaced. no harm done to the patient. | [5 empty] | The returned pump was visually inspected and biomaterial was observed on the impeller and blocking the purge gap. the cause of the issue was biomaterial. patient was reported to be predisposed to thrombus formation. this report is being filed as part of a retrospective review of historical records.
33744 | 1220648-2023-03986 | 17913222 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17913222&pc=OZD | 11/26/2022 | 10/11/2023 |  | 10/11/2023 |  |  | 11/26/2022 | 10/11/2023 |  |  | 04/30/2024 |  | 10/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 PUMP SET JP |  |  | ABIOMED, INC. |  |  | US |  | 2023223251 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission |  |  |  |  | 11/26/2022 |  |  |  |  | The user facility in japan reported a 65-year-old male admitted in cardiogenic shock was implanted with an impella device for mechanical circulatory support. the complainant reported a positioning issue during impella support. it was documented that the pump was pulled back into the aorta as a result of the patient moving. support was maintained throughout the night until the pump could be removed the following morning. there was no harm to the patient as a result of the noted issue. | [5 empty] | Based on the clinical account, the root cause of the positioning issues was that the tuohy borst was not tightened. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. this report is being filed as part of a retrospective review of historical records.
33745 | 1220648-2023-03987 | 17913284 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17913284&pc=OZD | 09/11/2023 | 06/30/2025 |  | 10/11/2023 |  |  | 09/11/2023 | 10/11/2023 |  |  |  |  | 10/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1130671 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Activation Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | A patient of unknown age and gender received hemodynamic support from an impella cp smartassist heart pump due to acute myocardial infarction/cardiogenic shock. over a period of several hours, the automated impella controller displayed several short ¿controller error¿ alarms. the users decided against using the backup controller because removal of the impella heart pump was already planned. the patient was successfully weaned from impella support and suffered no harm from the event. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - controller failure (red alarm) has been completed. data analysis: the logs confirm multiple issuances of the controller failure (pud sw corruption) alarm; device analysis: the failure was reproduced in the lab. extraction of the pud firmware revealed that it had been corrupted. replacing the pud resolved the issue. conclusion: the console was confirmed for the reported aic - controller failure (red alarm). the root cause of the controller failure alarm was due to pud fw corruption. the cause of the pud fw corruption was not determined. d2 revised common device name since the initial manufacturer device report number 1220648-2023-03987 was submitted in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2023-03987 in accordance with updated procedures.
33746 | 1220648-2023-03988 | 17913840 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17913840&pc=OZD | 09/12/2023 | 11/20/2024 |  | 10/11/2023 |  |  | 09/12/2023 | 10/11/2023 |  |  |  |  | 10/11/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1114578 |  | 0042-0010 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Rupture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an automated impella controller (aic) was rupturing the purge discs of installed cassettes. the aic was swapped to resolve the noted issue. there was no patient harm. | [5 empty] | Upon review of the provided information, it was determined that the correct aic serial number involved in this event is ic1883. as such, this is a duplicate report of 1220648-2024-21968. please see 1220648-2024-21968. for all further investigation / reporting related to this event. | The automated impella controller (aic) device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
33747 | 1220648-2023-03989 | 17913997 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17913997&pc=OZD | 09/12/2023 | 01/29/2024 |  | 10/11/2023 |  |  | 09/12/2023 | 10/11/2023 |  |  | 05/31/2025 |  | 10/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024371339 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 71-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. following implant, slow distal flow was noted in the impella leg. an external bypass was performed to provide blood flow to the sfa and distal leg. upon eventual explant of the pump, it was documented that the two pre-closes failed and a balloon tamponade was utilized. a contrast injection of the left femoral artery (lfa) was taken, and a perforation was noted. the surgeon was subsequently called for a thrombectomy of the sfa and closure. surgical wound closure was successful. | [5 empty] | The investigation into the vascular damage - peripheral vessel and the limb ischemia ¿ reversible has been completed. the device was not returned for benchtop investigation. based on clinical description, the root cause of limb ischemia and the of vascular damage was most likely due to patient condition. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33748 | 1220648-2023-03990 | 17914454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17914454&pc=OZD | 05/15/2023 | 10/11/2023 |  | 10/11/2023 |  |  | 05/15/2023 | 10/11/2023 |  |  | 01/31/2025 |  | 10/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024330986 |  | 0048-0045 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The team placed an impella cp to support a patient down in the community with a myocardial infarction. the complainant reported the purge sidearm cracked. the remainder of sidearm was removed. the pump support was for just 28 hours. | [5 empty] | The investigation into the reported events has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the cause of the purge leak- pump issue was not determined since the product was not returned for evaluation. the failure modes will be monitored and trended. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir."
33749 | 1220648-2023-03992 | 17914540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17914540&pc=OZD | 05/10/2023 | 10/11/2023 |  | 10/11/2023 |  |  | 05/10/2023 | 10/11/2023 |  |  | 03/31/2025 |  | 10/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 |  |  | ABIOMED, INC. |  |  | US |  | 2024326936 |  | 1000100 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/10/2023 |  |  |  |  | The user facility reported that a 59-year-old male patient implanted with the impella 5.5 for mechanical circulatory support developed thrombus. upon explant of the pump, a clot was observed inside inlet cage of the device. no further information was provided. there was no reported harm to the patient. | [5 empty] | The investigation into the reported thrombus was completed. the device was not returned for evaluation. based on the available information, the root cause of the thrombus in the pump on explant was most likely related to patient condition. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33750 | 1220648-2023-03993 | 17914571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17914571&pc=OZD | 09/12/2023 | 03/29/2024 |  | 10/11/2023 |  |  | 09/12/2023 | 10/11/2023 |  |  | 06/16/2023 |  | 10/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024351223 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Failure to Power Up | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/12/2023 |  |  |  |  | The user facility reported a 47-year-old male in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. the impella cp was backloaded over the .018 wire and the wire was then removed to initiate start. however, the motor would not turn on and failure alarm turned on. the physician requested another impella cp and that worked fine. there was no harm to the patient as a result of this incident. | [5 empty] | The investigation for the reported pump stop has been completed. the pump was not returned for investigation. data logs show that the pump was never started during implantation. according to clinical details, the pump was successfully replaced due to being unable to start. the cause of the pump did not start issue was not determined since product was not returned and sufficient clinical information was not provided. the instructions for use states ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
33751 | 1220648-2023-03994 | 17914631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17914631&pc=OZD | 09/13/2023 | 11/16/2023 |  | 10/11/2023 |  |  | 09/13/2023 | 10/11/2023 |  |  | 02/28/2025 |  | 10/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024341965 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 73-year-old male noted for high risk coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported the impella was prepped for placement but when it was plugged in, the console showed "defective catheter" alarm. the staff unplugged the impella and plugged it back in without resolution. a new impella was opened and prepped successfully. | [5 empty] | The investigation into the pump not detected issue has been completed since the original report was filed. the device was returned for investigation. the cause of the pump not detected issue was use related, as pump did not go through or complete the case start procedure on the console per clinical and field investigation. case start was successfully completed on the returned product and the issue was unable to be reproduced. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
33752 | 1220648-2023-03991 | 17914663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17914663&pc=OZD | 08/30/2023 | 01/17/2024 |  | 10/11/2023 |  |  | 08/30/2023 | 10/11/2023 |  |  | 04/30/2025 |  | 10/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024338712 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/13/2023 |  |  |  |  | The user facility reported a 77-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support . the team placed the 5.5 via the axillary and there was an access site bleed. the bleed was treated by three units of blood and surgical washout of the site. | [5 empty] | Since the medical center did not return the impella device, no benchtop analysis was possible. the cause of the access site bleeding issue was not determined since product was not returned and due to insufficient clinical information. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33753 | 1220648-2023-03995 | 17914723 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17914723&pc=OZD | 01/16/2023 | 10/11/2023 |  | 10/11/2023 |  |  | 01/16/2023 | 10/11/2023 |  |  | 09/30/2024 |  | 10/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2023267343 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Male |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/16/2023 |  |  |  |  | The user facility in the eu reported a 48-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. impella catheter shaft became broken. an observation was made of the red plug in the tapered end, has cracked and now are the wires inside the catheter shaft are visible. patient was noted to be stable. this occurred after just under 1 month of support. the pump remained on for 4 more days. | [5 empty] | No data logs were returned for this case. the product was returned and analyzed. a tear in the catheter jacket could be seen just after the kink protector. the root cause of the product damage was most likely use related as excessive/repetitive force would have to have been applied to the joint to result in a tear of the catheter jacket. this report is being filed as part of a retrospective review of historical records.
33754 | 1220648-2023-03996 | 17914776 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17914776&pc=OZD | 09/13/2023 | 02/26/2024 |  | 10/11/2023 |  |  | 09/13/2023 | 10/11/2023 |  |  | 06/30/2025 |  | 10/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024379701 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a 57-year-old female patient was implanted with an impella cp device for mechanical circulatory support for a left heart catheterization. on the same day, the patient did not have doppler pulse. therefore, decision was made to explant the impella. afterwards, the patient was noted to have had an arterial clot. the iliac was cleared and femoral and distal flow were achieved. the patient was stable. | [5 empty] | The investigation into the limb ischemia ¿ reversible issue has been completed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33755 | 1220648-2023-03997 | 17916050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17916050&pc=OZD | 05/18/2023 | 10/11/2023 |  | 10/11/2023 |  |  | 05/18/2023 | 10/11/2023 |  |  | 12/31/2024 |  | 10/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024299951 |  | 1000100 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient with a previous medical history of chronic heart failure and smoking presented to the md office complaining shortness of breath. the patient was diagnosed with triple vessel disease and reduced ef of 20%. the md determined that the patient was high risk cabg and planned the impella placement. once the device was implanted it was determined that the inlet of the impella was near the mitral valve. the md attempted to manipulate the pump and place it in the right position, however, he was unsuccessful. as a consequence the decision was made to remove impella. the patient was returned in the icu in stable condition. there was no harm to the patient. | [5 empty] | The investigation into the reported event has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that the most likely cause of the issue was use related. the failure modes will be monitored and trended.
33756 | 1220648-2023-03998 | 17916217 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17916217&pc=OZD | 09/15/2023 | 02/08/2024 |  | 10/11/2023 |  |  | 09/15/2023 | 10/11/2023 |  |  | 05/31/2025 |  | 10/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024366779 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 51-year-old male noted as high risk percutaneous cardiac intervention (hrpci) was implanted with an impella cp device for mechanical circulatory support. the complainant reported that patient's heparin was discontinued on (b)(6). patient was on cp since september 3rd to vent va ecmo and on p4. then, on (b)(6) the pump stopped, console was switched out but unable to restart pump. | [5 empty] | The investigation into the aic - pump stop issue has been completed. the device was not returned for investigation. data logs on the console show an rpm deviation pump stop alarm but are missing the end of the case right when this event occurs. the root cause of the pump stop issue was not determined since product was not returned and data logs are insufficient. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
33757 | 1220648-2023-03999 | 17916331 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17916331&pc=OZD | 09/15/2023 | 02/26/2024 |  | 10/11/2023 |  |  |  | 10/11/2023 |  |  | 05/31/2025 |  | 10/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024350118 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 62-year-old female patient received hemodynamic support from a femorally inserted impella cp smartassist heart pump due to acute myocardial infarction/cardiogenic shock. leg ischemia was noted after switching from the peel-away introducer sheath to the repositioning sheath. an external bypass was discussed. there was no intervention. it was later reported that there was no longer any lack of blood circulation. the team explanted the pump from the femoral after 5 days of support and did a new pump via the axillary access instead. | [5 empty] | The investigation into the limb ischemia ¿ reversible issue has been completed. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33758 | 1220648-2023-03915 | 17916471 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17916471&pc=OZD | 01/16/2023 | 10/10/2023 |  | 10/11/2023 |  |  | 01/16/2023 | 10/11/2023 |  |  | 06/30/2024 |  | 10/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023246669 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/16/2023 |  |  |  |  | The user facility reported a 68 year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient was returned to the operating room to reposition impella due to positioning towards mitral valve. unfortunately, it was unsuccessful. the impella was removed and another impella 5.5 was placed. there was no patient harm reported. | [5 empty] | The investigation for the reported positioning issues has been completed. the device was not returned for evaluation. however, after reviewing the clinical information, it was determined that the cause of the positioning issues was most likely use related, specifically improper repositioning technique. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33759 | 1220648-2023-04000 | 17916576 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17916576&pc=OZD | 08/19/2023 | 01/24/2024 |  | 10/11/2023 |  |  |  | 10/11/2023 |  |  | 01/31/2025 |  | 10/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP JP |  |  | OZD |  |  | IMPELLA CP JP |  |  | ABIOMED, INC. |  |  | US |  | 2024309523 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 88 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 88-year-old female patient noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the complainant had a cp-ecmo support ongoing for a surgical/transcatheter aortic valve replacement (tavr) patient. there was systemic bleeding, inclusive of the impella cp access site. there was noted rbcs given of an unknown amount. no vascular repair was needed. | [5 empty] | The investigation into the patient's access site bleeding has been completed. product was not returned for review. the root cause of access site bleeding was most likely due to patient condition. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33760 | 1220648-2023-04001 | 17916585 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17916585&pc=OZD | 09/08/2023 | 06/06/2024 |  | 10/11/2023 |  |  |  | 10/11/2023 |  |  | 04/30/2025 |  | 10/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2024338010 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/12/2023 |  |  |  |  | The user facility reported an 84-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. on the third day of impella support, the patient expired. the cause of death is multiple organ failure. the relationship between the cause of death and impella is unknown. the impella operated as intended. | [5 empty] | The investigation for the reported vascular damage has been completed. no product was returned for analysis. the root cause of the vascular damage - peripheral vessels could not be determined due to insufficient clinical information. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 added code 925 as it was inadvertently left off the previously submitted manufacture device report. h.6 code 4114 was reported incorrectly on the previously submitted report. new code has been added to type of investigation codes. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted report incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures. additional manufacturer narrative on the previously submitted report stated that the device was not returned but the device was discarded. type of investigation codes has been updated to reflect discarded. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33761 | 1220648-2023-04002 | 17917364 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17917364&pc=OZD | 05/18/2022 | 10/11/2023 |  | 10/11/2023 |  |  | 05/18/2022 | 10/11/2023 |  |  | 01/31/2024 |  | 10/11/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023184035 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 26 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/18/2022 |  |  |  |  | The user facility reported a 26 year old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there were frequent impella in aorta and position alarms. the position was confirmed under echo and support continued. there was no patient harm reported. | [5 empty] | The investigation for the reported positioning issues has been completed. the device was not returned for evaluation. however, the clinical report provided sufficient details to determine the root cause. the root cause of the placement signal alarm issue was most likely a result of patient condition. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33762 | 1220648-2023-03916 | 17904082 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17904082&pc=OZD | 08/31/2021 | 09/26/2024 |  | 10/10/2023 |  |  | 08/31/2021 | 10/10/2023 |  |  | 10/31/2022 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2021013141 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/31/2021 |  |  |  |  | The user facility reported a 57yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there were questionable flow and suspected hemolysis. there was no medical intervention mentioned. no patient harm reported. | [5 empty] | Impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr was provided in sections h1, and h10. | The investigation for the reported low flow and hemolysis issues has been completed. as no product or logs were returned, the root cause of the low flow and hemolysis could not be determined. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33763 | 1220648-2023-03917 | 17904111 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17904111&pc=OZD | 05/09/2023 | 10/10/2023 |  | 10/10/2023 |  |  | 05/09/2023 | 10/10/2023 |  |  | 03/31/2025 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024320130 |  | 1000100 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/09/2023 |  |  |  |  | The user facility reported 64yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the impella 5.5 had diminished flows following insertion in which significant clot burden was observed in the graft and aorta. the physician was not willing to replace the impella at the time. the pump was explanted eleven days after implant successfully. there was no patient harm. | [5 empty] | The investigation for the low flow and the positioning issues has been completed. no product was returned. as per clinical report, false impella in aorta alarms were observed and later echo was done to confirm position. the impella 5.5 had low flows following insertion in which significant clot burden was observed in the graft and aorta with patient having lv thrombus condition. data logs were reviewed, low flow with motor current spike and several suction alarms found during whole case. false impella in position alarm was observed after the motor current spike per log. the cause of the optical signal issue was most likely biomaterial with false 'impella in aorta' alarm observed after motor current spike per log and clinical review. the cause of the low pump flow issue was most likely biomaterial with motor current spike and suction alarms observed per log review and patient having lv thrombus and significant clot in graft and aorta post insertion per clinical. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33764 | 1220648-2023-03918 | 17904167 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17904167&pc=OZD | 02/15/2023 | 09/22/2024 |  | 10/10/2023 |  |  | 02/15/2023 | 10/10/2023 |  |  | 12/31/2024 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024291917 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/15/2023 |  |  |  |  | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | The user facility reported a 67yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that during repositioning, the impella was pulled out of the ventricle and unable to recross aortic valve. decision was made to swap out the pump. there was no patient harm reported. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03918 was provided in sections b1, b5, d6, h1, and h6. note: corrected information provided in h6 is an addition to previous coding. | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the positioning issue was improper technique as the pump was accidently pulled back across the aortic valve and the physician was unable to wirelessly re-cross the valve. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33765 | 1220648-2023-03919 | 17904191 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17904191&pc=OZD | 02/24/2023 | 10/10/2023 |  | 10/10/2023 |  |  | 02/24/2023 | 10/10/2023 |  |  | 07/31/2024 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023253793 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/24/2023 |  |  |  |  | The user facility reported a 68yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the impella was unable to run at p6 due to increased ectopy and then the patient subsequently decompensates. the positioning adjusted on echo but was unable to optimally seat completely in left ventricular free space. the medical team decided to reposition the device and run it at p5 and the alarms stopped. there was no patient harm reported. | [5 empty] | The investigation for the reported positioning issues has been completed. the device nor the data logs were returned for evaluation. additionally, the clinical information was insufficient to determine a root cause alone. therefore, the root cause of the reported positioning issues could not be determined. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33766 | 1220648-2023-03920 | 17904240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17904240&pc=OZD | 07/03/2023 | 09/19/2025 |  | 10/10/2023 |  |  | 07/03/2023 | 10/10/2023 |  |  | 04/30/2025 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024341021 |  | 1000112 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  | Hispanic | Asian | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/03/2023 |  |  |  |  | The user facility in japan reported a 81-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. there was left ventricle (lv) waveform dissociation. the lv waveform changed after imaging from about 5 days after using impella. it equaled the aortic pressure values, the position in aorta alarms sounded, so the echo confirmed the impella pump position multiple times. the position was confirmed by imaging, but it was in the same position and there was no problem with the positioning of the pump itself. additionally, the complainant reported upon removal of the impella device, there was a thrombus like substance near the impeller. | [5 empty] | Pump was returned for investigation. foreign material was observed on impeller. no other physical abnormalities were observed. data logs shows the drop in snr while pump was set to p-level 8. also, there were some improper positioning and the impella in aorta alarms reported. the cause of thrombus the issue was patient condition related. this report is being filed as part of a retrospective review of historical records. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. pump was returned for investigation. no burr/sharp edges were found on return product during visual inspection. no other physical abnormalities were observed. during test, 5s optical parameters were within specification. data logs shows improper positioning and the impella position in aorta alarms reported with active suction alarm. the root cause of the optical signal issue was patient condition related based on data logs and returned product review. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 updated with product problem. d6a/d6b updated with additional information. f6 and f8 should have been left blank. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures from the initial manufacturer device report number 1220648-2023-03920. h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2023-03920.
33767 | 1220648-2023-03921 | 17904278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17904278&pc=OZD | 12/04/2022 | 10/10/2023 |  | 10/10/2023 |  |  | 12/04/2022 | 10/10/2023 |  |  | 07/31/2024 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023251513 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Device Sensing Problem | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/04/2022 |  |  |  |  | The user facility reported a 60yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there were position/sensor alarms and saturated flows. the team made the choice to remove/replace the pump for va ecmo to be placed. there was no patient harm reported. | [5 empty] | The investigation for the reported saturated flows and optical issues has been completed. the data logs were reviewed and show a motor current spike followed by a lack of pulsatility, suction alarms, and an increase in purge pressure. once the pump was placed on p-9 saturated flow occurred. motor current reduced over the next couple of hours. the pump was returned and no abnormalities were found. the saturated flow was unable to be reproduced. the evidence found in the logs suggest that the root cause of the saturated flow was most likely biomaterial ingestions. ingested biomaterial was most likely the root cause of the false impella in aorta alarms. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33768 | 1220648-2023-03922 | 17904323 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17904323&pc=OZD | 06/16/2023 | 01/22/2024 |  | 10/10/2023 |  |  | 06/16/2023 | 10/10/2023 |  |  | 02/28/2025 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2024307316 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Male |  |  |  | Malfunction | No | Yes | Use of Device Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/16/2023 |  |  |  |  | The user facility reported a male patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the clinic reported that blood was leaking at the red plug from the impella. site was cleaned and further observed to see if blood was leaking again. no intervention required. pump is still running. | [5 empty] | The investigation for the reported hole in the catheter/leakage has been completed. the device and data logs were not returned. per clinical information, blood leak was observed from the red handle. the cause of the product damage issue was not determined since the product was not returned and sufficient clinical information was no provided. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
33769 | 1220648-2023-03923 | 17904484 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17904484&pc=OZD | 10/17/2022 | 03/10/2025 |  | 10/10/2023 |  |  | 10/17/2022 | 10/10/2023 |  |  | 07/31/2024 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023252095 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/17/2022 |  |  |  |  | The user facility reported a 50-year-old male patient was implanted with an impella 5.5 device for mechanical circulatory support. during support, it was noted that the pump was accidentally pulled into the aorta during repositioning attempts. reinsertion attempts were unsuccessful and the decision was made to take out the 5.5 and re-gain ventricular access with mpa catheter and wire. access gained again and the 5.5 pump was re-implanted successfully. as support continued, it was also documented that the patient experienced several ventricular tachycardia events. the medical team was concerned it was related to impella position. bedside echo was performed with ultrasound and 5.5 was measured to be 6.1 cm. the medical staff pulled back the impella 1 cm using us guidance. support was maintained throughout the event. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported vf/vt/af/at and positioning issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the positioning issue was most likely use related, as the impella was noted on ecmo to be flipped toward posterior wall and accidently pulled into the aorta with failed reinsertion efforts as per the clinical description provided. the root cause of the vf/vt/af/at issue was most likely patient condition related, as the vt/vf occurred during the placement/repositioning of the impella as per the clinical information provided. this report is being filed as part of a retrospective review of historical records. h4: device manufacture date corrected
33770 | 1220648-2023-03924 | 17904680 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17904680&pc=OZD | 11/10/2022 | 01/10/2024 |  | 10/10/2023 |  |  | 11/10/2022 | 10/10/2023 |  |  | 08/31/2024 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023256984 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/10/2022 |  |  |  |  | The user facility reported a 58-year-old male patient was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the impella 5.5 pump was pulled back out of position during patient support. due to the inadvertent displacement, the decision was made to explant the pump. there was no patient harm. | [5 empty] | The investigation into the positioning issues has been completed. the product was not returned. per the clinical information provided, the root cause of the positioning issue was use related. the impella was pulled back and out of position. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
33771 | 17904742 | 17904742 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17904742&pc=OZD | 08/14/2023 | 09/12/2023 | 09/12/2023 | 10/10/2023 | 09/12/2023 | 10/10/2023 |  | 10/10/2023 |  |  |  |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA |  |  | ABIOMED INC. |  |  | US |  | 1715449 |  | 0043-001 |  | Yes |  |  | Not Available |  |  | No |  |  |  | User Facility report | Yes | Other Health Care Professional | [6 empty] | 28470 DA |  | Male |  |  |  | Malfunction | No | Yes | [7 empty] | Insufficient Information | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | The rn went to change the impella purge cassette, upon changing it was noted there was a leak within the prior cassette.
33772 | 1220648-2023-03925 | 17904968 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17904968&pc=OZD | 02/05/2022 | 01/26/2024 |  | 10/10/2023 |  |  | 02/05/2022 | 10/10/2023 |  |  | 11/30/2023 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023162151 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/05/2022 |  |  |  |  | The user facility reported a 50-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was documented that the patient developed a hematoma at the pump access site resulting in a decrease in hemoglobin levels. as a result, the patient was given two units of packed red blood cells and systemic heparin was halted. support was maintained with the implanted pump. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the access site bleeding - major has been completed. the device was not returned for investigation. per the provided clinical information, the root cause of the access site bleeding was unable to be determined as limited clinical details and no product were returned for investigation.
33773 | 1220648-2023-03926 | 17905114 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17905114&pc=OZD | 07/06/2023 | 10/10/2023 |  | 10/10/2023 |  |  | 07/06/2023 | 10/10/2023 |  |  |  |  | 10/10/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1613084 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the aic showed an battery failure alarm. the user tried multiple times to reset it but pushing the button but it continued to alarm. there was no adverse event reported as a result of this incident. | [5 empty] | The investigation into the reported events has been completed. aic logs confirmed the reported failure. there was a battery failure alarm due to low pack voltage. the data logs from the complaint date showed cell imbalance with battery 1 which caused battery pack internal electronics to shut down that battery. batttest utility was run on the console, which confirmed that cell 1 of battery 1 was imbalanced with the other cells in battery. the evaluation revealed that the cause of the aic issue was a defective battery. the failure modes will be monitored and trended.
33774 | 1220648-2023-03928 | 17905154 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17905154&pc=OZD | 06/11/2023 | 09/19/2024 |  | 10/10/2023 |  |  | 06/11/2023 | 10/10/2023 |  |  | 04/30/2025 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2024334365 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | Yes | Restricted Flow rate | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/11/2023 |  |  |  |  | Additional information received that patient did not survive the procedure. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | The user facility in the eu reported a male patient admitted in cardiogenic shock was implanted with an impella device for mechanical circulatory support. the console alarmed during the 5.5 support with high purge pressure, low purge flow; they troublseshot with lysis which was unsuccessful, and so switched to new 5.5 pump.there was no harm done to the patient. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr was provided in sections b1, b5,d6, h1, and h6. | Pump was sent back for investigation with cut patient cable. original purge cassette was not returned. high purge pressure did not reproduce on pump. no other physical abnormalities were observed on the pump. the root cause of the high purge pressure issue was most likely kinked purge line based on provided clinical information. this report is being filed as part of a retrospective review of historical records.
33775 | 1220648-2023-03927 | 17905156 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17905156&pc=OZD | 10/16/2022 | 03/10/2025 |  | 10/10/2023 |  |  | 10/16/2022 | 10/10/2023 |  |  | 08/31/2024 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023252853 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Female |  |  | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an implanted impella 5.5 pump developed high purge pressure during patient support. lysis therapy was initiated and was found to be effective in reobtaining acceptable purge pressure levels. support continued with the implanted pump following the intervention. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported high purge pressure issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
33776 | 1220648-2023-03929 | 17905162 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17905162&pc=OZD | 11/09/2022 | 01/26/2024 |  | 10/10/2023 |  |  | 11/09/2022 | 10/10/2023 |  |  | 06/30/2024 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023229613 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/09/2022 |  |  |  |  | The user facility reported a 55-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that following implant of an impella 5.5 pump for patient support, the site continued to bleed. a jackson-pratt drain was already in place and the patient was given 2 units red blood cell (rbc) and 4 units of platelets to counteract the blood loss. support continued with the implanted pump. | [5 empty] | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for investigation. per the clinical information provided, the root cause of the access site bleed was unable to be determined as sufficient clinical information and product were not provided for the investigation. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
33777 | 1220648-2023-03930 | 17905180 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17905180&pc=OZD | 07/05/2023 | 10/10/2023 |  | 10/10/2023 |  |  | 07/05/2023 | 10/10/2023 |  |  |  |  | 10/10/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1250484 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The us complainant had an impella cp support ongoing when there were aic alarms. the yellow alarm prompted the aic to be replaced. no harm or injury occurred from the aic replacment and interrupted support. | [5 empty] | The investigation into the reported events has been completed. logs were unable to be accessed. the complaint description describes a controller error that could have been caused by a variety of factors. when the console was received for investigation, there was a system self-check console error on boot up. there was both a breakdown in communication between the eep and the 4-in-1, that could have caused the controller error seen in the field, and a corrupt flash card. swapping out the impellatronic board, resulted in the return of communication between the impellatronic and the 4-in-1. however, the system self-check error persisted until the compact flash cards were also replaced. the evaluation revealed that the cause of the system self-check error was a defective impellatronics board. the failure modes will be monitored and trended.
33778 | 1220648-2023-03931 | 17905229 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17905229&pc=OZD | 08/31/2021 | 03/10/2025 |  | 10/10/2023 |  |  | 08/31/2021 | 10/10/2023 |  |  | 03/31/2023 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2022059253 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Unknown |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/31/2021 |  |  |  |  | The user facility reported an 80-year-old patient was implanted with an impella cp device for mechanical circulatory support. during support, it was noted that a thrombus developed in the left ventricle prompting the pump to be explanted. the patient was later supported with an impella 5.5 pump roughly two weeks after the cp was explanted. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported thrombus issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
33779 | 1220648-2023-03932 | 17905233 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17905233&pc=OZD | 02/09/2023 | 10/10/2023 |  | 10/10/2023 |  |  | 02/09/2023 | 10/10/2023 |  |  |  |  | 10/10/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1108306 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The us complainant had to replace an aic due to priming/purge issues. the ifu was followed and a backup controller used for the cp support. there was no reported patient harm or injury. | [5 empty] | The investigation into the reported events has been completed. analysis of the console logs from the reported day of event are consistent with complaint and show that in the beginning of case start, the priming was stopped upon purge cassette removal (after prompt to insert purge disk), but not restarted after it was reinserted. the pump was disconnected and reconnected, presumably in an attempt to troubleshoot purge issue, but case start wizard had not been aborted or restarted - instead the aic was powered down and case successfully started on a backup console. console ic1738 and its purge system was extensively tested in the lab, but no irregularities were observed, and the device operated within specification. the evaluation revealed that the most likely cause of the aic issue was a software issue. the failure modes will be monitored and trended.
33780 | 1220648-2023-03933 | 17905260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17905260&pc=OZD | 11/04/2022 | 03/10/2025 |  | 10/10/2023 |  |  | 11/04/2022 | 10/10/2023 |  |  | 06/30/2023 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022104940 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Material Twisted/Bent | [6 empty] | Perforation of Vessels | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/04/2022 |  |  |  |  | The user facility reported a 42-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the us complainant had a cp that needed care escalation. the team attempted to place the 5.5 via the axillary graft with an rp introducer. placement failed due to anatomy and the catheter began to kink. the patient was reported to have stenosis with a vessel diameter of <7mm. resistance was felt at the past innominate artery. additionally, two vascular tears were found proximal to the anastomosis that prompted repair. the team aborted impella 5.5 insertion and placed an impella cp. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation for the reported vascular damage - peripheral vessel and delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the catheter kink was likely due to excessive force attempting to deliver. the cause of the delivery issue was patient condition related since the patient had stenosis with a vessel diameter of <7mm. this report is being filed as part of a retrospective review of historical records.
33781 | 1220648-2023-03934 | 17905276 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17905276&pc=OZD | 09/21/2023 | 03/13/2025 |  | 10/10/2023 |  |  | 09/21/2023 | 10/10/2023 |  |  | 08/31/2023 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022137347 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 18 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/21/2023 |  |  |  |  | The user facility reported a 18-year-old male was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that a patient on impella support had developed a hematoma at the access site for which they had to be taken to the operating room for evacuation. support with the implanted pump continued following the intervention. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. no product was returned. the root cause of the access site bleeding issue was not determined since product was not returned.
33782 | 1220648-2023-03935 | 17905406 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17905406&pc=OZD | 05/19/2021 | 05/27/2025 |  | 10/10/2023 |  |  | 05/19/2021 | 10/10/2023 |  |  | 03/31/2023 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022057826 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 65 YR |  | Male |  |  |  | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Unspecified Infection | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/15/2021 |  |  |  |  | The complainant reported a 65 year old male patient presenting with cardiomyopathy for impella support. the patient spiked up a fever on (b)(6) 2021 and blood cultures came up as positive. this prompted the medical team to remove all central lines (including the impella) and treat the patient with antibiotics. in addition, the patient was found to have hematoma, a washout had to be performed to remove it and the pump. a new impella was ultimately placed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿infusion filter the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The root cause of the access site bleeding cannot be determined since the product was not returned and enough clinical info was not provided. the root cause of the infection was most likely due to patient condition. the root cause of the positioning issue was most likely due to use issue. this report is being filed as part of a retrospective review of historical records.
33783 | 1220648-2023-03936 | 17905418 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17905418&pc=OZD | 05/11/2022 | 01/15/2024 |  | 10/10/2023 |  |  | 05/11/2022 | 10/10/2023 |  |  | 02/29/2024 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023193684 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/25/2023 |  |  |  |  | The user facility reported a 61-year-old male patient was implanted with an impella 5.5 device for mechanical circulatory support. during support, it was noted that the patient developed bleeding from multiple sites for which six units packed red blood cells, two units of fresh frozen plasma, and two units of platelets were given. support continued with the implanted pump following the intervention. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the reported access site bleeding has been completed since the original report was filed. no product was returned. the root cause of the access site bleeding issue was not determined since product was not returned and limited clinical information was provided.
33784 | 1220648-2023-03937 | 17905445 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17905445&pc=OZD | 07/21/2021 | 01/22/2024 |  | 10/10/2023 |  |  | 07/15/2021 | 10/10/2023 |  |  | 03/31/2023 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022067479 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/15/2021 |  |  |  |  | The user facility reported a 64-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient experienced an infection in the graft of the impella. the infection was controlled by the medical staff. | [5 empty] | The investigation for the reported infection has been completed. the device was not returned for evaluation. clinical information provided stated that the infection was in the shaft of the pump and was controlled. the pump continued to be in use till the end of the case and no concerns were reported that related to the pump. the root cause of the infection was patient condition related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿infusion filter: the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ this report is being filed as part of a retrospective review of historical records.
33785 | 1220648-2023-03939 | 17905590 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17905590&pc=OZD | 05/03/2021 | 11/19/2024 |  | 10/10/2023 |  |  | 05/03/2021 | 10/10/2023 |  |  | 02/28/2023 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 371 SET, US |  |  | ABIOMED, INC. |  |  | US |  | N/A |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a 59-year old male patient on impella cp and ecmo support developed ischemia in their leg. a fem-fem bypass was performed to treat the condition and the patient was escalated to an impella 5.5 pump for ongoing support. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | As the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
33786 | 1220648-2023-03940 | 17906719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17906719&pc=OZD | 07/03/2022 | 09/22/2025 |  | 10/10/2023 |  |  | 07/03/2022 | 10/10/2023 |  |  | 02/28/2025 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024341047 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 07/03/2022 |  |  |  |  | Additional information noted care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 56-year-old female was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced bleeding from an unknown source. the patient was taken to the or and the surgeon attempted to recannulate and ligate. pt had severe decompensation during the length of the intervention and more than 30 units of blood were given. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided. | The impella device was not returned by the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2023-03940. d6a and d6b revised from the initial manufacturer device report 1220648-2023-03940 in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2023-03940 in accordance with updated procedures. g1 reporting contact facility name was inadvertently omitted on the initial manufacturer device report number 1220648-2023-03940. g1 reporting contact email revised on manufacturer device report 1220648-2023-03940 in accordance with updated procedures.
33787 | 1220648-2023-03941 | 17906733 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17906733&pc=OZD | 09/13/2023 | 01/26/2024 |  | 10/10/2023 |  |  | 09/13/2023 | 10/10/2023 |  |  |  |  | 10/10/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1085518 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during the process of updating the software to version 12.2, the automated impella controller (aic) had a battery failure (red alarm). | [5 empty] | The investigation into the reported aic battery failure has been completed since the original report was filed. aic logs confirmed the reported failure. there was a battery failure alarm (transport connection blown/ missing ¿ event set #360) due to low pack voltage. the batttest utility was run on the aic via a telnet connection and provided detailed battery information that showed a cell imbalance in battery 2. cell 4 of battery 2 is unbalanced at 387 mv lower than the other cells in battery 2. the failure mode was reproduced on an engineering lab bench. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure, a known pattern failure. | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
33788 | 1220648-2023-03942 | 17906779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17906779&pc=OZD | 03/10/2022 | 10/10/2023 |  | 10/10/2023 |  |  | 03/10/2022 | 10/10/2023 |  |  | 10/31/2023 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2022149005 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 16 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/10/2022 |  |  |  |  | The user facility in hte eu reported a 16-year-old male with cardiomyopathy was implanted with an impella device for mechanical circulatory support. the complainant reported a drop in flows due to a thrombus. the pump was removed and replaced. there was no harm to the patient in relationship to the impella device. | [5 empty] | The impella was returned with only half the catheter intact. the other half that is attached to the pump itself was missing. data logs showed that the motor current was non-pulsatile and the flow rate decreased upon the sudden occurrence of continuous "suction" alarms. after reviewing the clinical information, it was determined that the cause of the low pump flow was most likely ingested biomaterial. the improper anticoagulation regimen very likely contributed to thrombus development.
33789 | 1220648-2023-03943 | 17906942 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17906942&pc=OZD | 10/01/2021 | 01/26/2024 |  | 10/10/2023 |  |  | 10/01/2021 | 10/10/2023 |  |  | 04/30/2023 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022070079 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/01/2021 |  |  |  |  | The user facility reported a 63-year-old male with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support inserted from the left axillary, experienced impella in the aorta alarms. the medical team attempted to reposition the pump, however, the md decided to replace the pump with a new one inserted from the right axillary artery. | [5 empty] | The investigation into the reported positioning issue has been completed since the original report was filed. no product was returned. the root cause of the positioning issue was most likely use related, as there were impella position in aorta alarms observed as per the clinical description provided. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
33790 | 1220648-2023-03944 | 17906950 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17906950&pc=OZD | 10/26/2022 | 03/13/2025 |  | 10/10/2023 |  |  | 10/26/2022 | 10/10/2023 |  |  | 03/31/2024 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023208498 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/26/2022 |  |  |  |  | The user facility reported a 77-year-old male was implanted with an impella 5.5 device for mechanical circulatory support. during support, it was noted that the pump was experiencing a low purge pressure and high purge flows. troubleshooting identified a fluid leak from the yellow luer on the purge sidearm. a purge bypass was performed to resolve the leak; however, the system then registered a high purge pressure. tissue plasminogen activator (tpa) was utilized for the purge line and support with the implanted pump continued. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ this report is being filed as part of a retrospective review of historical records. | The root cause of the purge cassette leak was unable to be determined, since the leak cannot be verified and the root cause cannot be determined as no product was returned for investigation. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
33791 | 1220648-2023-03945 | 17906988 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17906988&pc=OZD | 10/11/2021 | 09/26/2024 |  | 10/10/2023 |  |  | 10/11/2021 | 10/10/2023 |  |  | 07/31/2023 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022113564 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/11/2021 |  |  |  |  | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 48-year-old male admitted in cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported an encapsulated thrombus that became dislodged during impella support via a 5.5 device inserted surgically in the left axillary artery. the thrombus had been caught in the device's inflow cage, but ultimately passed into the aorta where it became lost under imagery. patient then endured ischemia in the right lower limb that required intervention via cutdown. the patient expired on support, however, the cause and/or relationship to the impella device is unknown. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03945 was provided in sections b5, h1 and h6. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
33792 | 1220648-2023-03946 | 17906990 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17906990&pc=OZD | 03/02/2022 | 01/26/2024 |  | 10/10/2023 |  |  | 03/02/2022 | 10/10/2023 |  |  | 11/30/2023 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023169684 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/02/2022 |  |  |  |  | The user facility reported a 55-year-old female presenting for mechanical circulatory support via the impella device. the complainant reported bleeding from the access site that prompted multiple units of blood products delivered. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
33793 | 1220648-2023-03948 | 17907055 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17907055&pc=OZD | 06/22/2022 | 03/28/2025 |  | 10/10/2023 |  |  | 06/22/2022 | 10/10/2023 |  |  | 02/29/2024 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023198265 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/22/2022 |  |  |  |  | The user facility reported a 58-year-old male admitted in cardiomyopathy and cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the medical staff was having difficulty delivering an impella pump due to the patient's anatomy. after several attempts the md decided to exchange the pump with a new one and retry delivery again. again with the new pump the md had difficulty with the delivery and proceeded to remove the second pump. the md was also having difficulty removing the 2nd pump due to to kink in the canula. finally they were able to remove the 2nd pump and the procedure was aborted. | [5 empty] | The root cause of the delivery issue anatomy was most likely due to the patient condition. the root cause of the product damage (cannula kink) could not be determined as insufficient clinical details were provided. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
33794 | 1220648-2023-03947 | 17907056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17907056&pc=OZD | 10/14/2022 | 01/12/2024 |  | 10/10/2023 |  |  | 10/14/2022 | 10/10/2023 |  |  | 05/31/2024 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023228262 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 22 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Advance | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/14/2022 |  |  |  |  | The user facility reported a 22-year-old female was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that difficulty was encountered while attempting to implant an impella 5.5 pump for patient support. due to the smaller anatomy of the patient, resistance was met when advancing the impella for delivery. this subsequently led to some bleeding for which the patient was given a unit of packed red blood cells (prbc). the pump was implanted successfully and support was initiated. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the reported delivery issue and bleeding has been completed since the original report was filed. the root cause of the the delivery issue and is because of patient condition. since the right axillary vessel deemed not optimal for graft, the doctor has to use the left axillary for insertion. even through the left axillary, because of smaller anatomy the ct surgeon had difficulty advancing the impella. hence, the root cause is patient condition. the root cause of access sit bleeding was determined to be patient condition because of diseases/small vessels.
33795 | 1220648-2023-03949 | 17907059 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17907059&pc=OZD | 03/23/2022 | 01/26/2024 |  | 10/10/2023 |  |  | 03/23/2022 | 10/10/2023 |  |  | 10/31/2023 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022157310 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella 5.5 support experienced a small amount of bleeding from the access site. the patient was taken to the or to have the source ligated and control the small bleeding. | [5 empty] | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for investigation. per the clinical information provided, the root cause of the access site bleeding was unable to be determined as limited clinical details and no product were returned for investigation. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
33796 | 1220648-2023-03950 | 17907087 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17907087&pc=OZD | 12/15/2022 | 01/26/2024 |  | 10/10/2023 |  |  | 12/15/2022 | 10/10/2023 |  |  | 10/31/2024 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023274392 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/15/2022 |  |  |  |  | The user facility reported a 67-year-old male suffering from cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support as the care needed escalation from the impella cp. the patient inadvertently pulled the impella device out of position when they became agitated during support. due to the forced removal out of position, the decision was made to replace the impella pump. there was no patient harm. | [5 empty] | The investigation into the positioning issue has been completed. the device was not returned for investigation. per the provided clinical information, the root cause of the positioning issue was patient movement, the patient pulled out the impella device and there were impella position in aorta alarms observed in the data logs. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
33797 | 1220648-2023-03951 | 17907130 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17907130&pc=OZD | 12/04/2022 | 03/13/2025 |  | 10/10/2023 |  |  | 12/04/2022 | 10/10/2023 |  |  | 10/31/2024 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023270762 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/04/2022 |  |  |  |  | The user facility reported a 57-year-old male admitted in cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 pump experienced persistent positioning issues throughout the course of patient supported. as a result of the noted complication, the decision was made to explant the pump. there was no patient harm. patient further supported on protech duo. | [5 empty] | The impella device was discarded by the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The root cause of the positioning issue was most likely due to use issue. this report is being filed as part of a retrospective review of historical records.
33798 | 1220648-2023-03952 | 17907136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17907136&pc=OZD | 10/17/2022 | 03/13/2025 |  | 10/10/2023 |  |  | 10/17/2022 | 10/10/2023 |  |  | 07/31/2024 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023248924 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/17/2022 |  |  |  |  | The user facility reported a 66-year-old male was implanted with an impella 5.5 device for mechanical circulatory support. the impella pump triggered a purge blocked alarm during patient support. the purge cassette was replaced in an attempt to troubleshoot the issue, but the failure persisted. tissue plasminogen activator was subsequently run through the purge to assist with lowering the purge pressure. upon reaching the end of support, the pump was placed off the aorta and tunneled out to the left supraclavicular. throughout and following the removal of the pump, the patient required seven units of prbc and three units of platelets. the patient was taken to the operating room for the removal and a washout was performed of the site. the patient was considered to be stable following the event. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
33799 | 1220648-2023-03953 | 17907140 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17907140&pc=OZD | 07/01/2022 | 01/12/2024 |  | 10/10/2023 |  |  | 07/01/2022 | 10/10/2023 |  |  | 03/31/2024 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023199544 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/01/2022 |  |  |  |  | The user facility reported a 63-year-old male implanted with an impella device for mechanical circulatory support. the patient on impella support experienced bleeding from the access site which required multiple blood products to be given. in addition, the patient had also a hematoma at the axillary would which required a washout. | [5 empty] | The investigation into the reported access site bleeding ¿ major has been completed since the original report was filed. the root cause of the access site bleeding was unable to be determined as limited clinical details and no product were returned for investigation. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
33800 | 1220648-2023-03954 | 17907145 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17907145&pc=OZD | 08/04/2022 | 09/20/2024 |  | 10/10/2023 |  |  | 08/04/2022 | 10/10/2023 |  |  | 05/31/2024 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023218333 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/04/2022 |  |  |  |  | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced a low purge flow and high purge pressures. the rn reviewed the potential causes but could not identify any kinks or other potential sources that would case the issues. |  |  |  |  | B2: date of patient death is unknown. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03954 was provided in sections b2, b5, h1, and h6. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
33801 | 1220648-2023-03955 | 17907151 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17907151&pc=OZD | 08/23/2023 | 09/19/2024 |  | 10/10/2023 |  |  | 08/23/2023 | 10/10/2023 |  |  | 06/30/2024 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023234709 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Life Threatening | Other | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/23/2023 |  |  |  |  | Additional information noted the patient's family decided to withdraw care and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 64-year-old black male admitted in cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced a large hemorrhagic stroke. the head ct scan also revealed brain tumor. the medical staff decided to withdraw care and the patient expired. |  |  |  |  | The investigation into the stroke/cva has been completed. the device was not returned for investigation. as the clinical information was not provided, the true root cause of the stroke/cva could not be determined. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b2, b5, h1, and h6. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
33802 | 1220648-2023-03956 | 17907160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17907160&pc=OZD | 08/21/2022 | 01/26/2024 |  | 10/10/2023 |  |  | 08/21/2022 | 10/10/2023 |  |  | 03/31/2024 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023206427 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient developed a hematoma around the impella access site during support. the patient was taken to the operating room where the hematoma was evacuated. a jp drain was placed for any residual fluid and support with the implanted pump continued. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major has been completed. the device was not returned for investigation. per the clinical information provided, the root cause of the access site bleeding was unable to be determined as limited clinical details and no product were returned for investigation.
33803 | 1220648-2023-03958 | 17907185 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17907185&pc=OZD | 08/24/2022 | 03/13/2025 |  | 10/10/2023 |  |  | 08/24/2022 | 10/10/2023 |  |  | 03/31/2024 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023208505 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hemolysis | No Clinical Signs, Symptoms or Conditions | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/24/2022 |  |  |  |  | The user facility reported a 81-year-old patient implanted with an impella device for mechanical circulatory support. the complainant reported concerns for hemolysis during support. the physician was aware, however, chose to continue using the device. the pump ultimately stopped and could not be restarted, therefore, the device was removed. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
33804 | 1220648-2023-03957 | 17907187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17907187&pc=OZD | 07/15/2022 | 01/26/2024 |  | 10/10/2023 |  |  | 07/15/2022 | 10/10/2023 |  |  | 09/30/2022 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2021002047 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/15/2022 |  |  |  |  | The user facility reported a 49-year-old male patient was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that a patient lost a liter of blood from their right axillary site during impella support. the patient's hemoglobin levels were down to 7 as a result of the blood loss. one unit of packed red blood cells (prbc) was given and a jackson-pratt drain (jp drain) was added to treat the bleed. the site was dressed with best practice and support continued with the pump. | [5 empty] | The investigation for the access site bleed has been completed. however, with limited clinical details nor the impella device, the root cause of the reported issue could not be determined. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
33805 | 1220648-2023-03959 | 17907220 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17907220&pc=OZD | 09/21/2022 | 03/13/2025 |  | 10/10/2023 |  |  | 09/21/2022 | 10/10/2023 |  |  | 04/30/2023 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023212678 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Death | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/21/2022 |  |  |  |  | The user facility reported a 57-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported a patient developing sepsis during support. it was alleged that the physician may have "missed the window" to replace lines in patient. the pump was ultimately removed and the patient expired shortly after. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ this report is being filed as part of a retrospective review of historical records. | An infection/sepsis can occur during or after impella support. when diagnosing the cause of an infection in a patient, the following clinical evidence would need to be considered holistically: patient symptoms and medical history obtained by the treating clinician, physical examination findings, results of special investigations: laboratory test results & imaging studies, microbiological culture and sensitivity results, response to previous treatments and therapies, any relevant epidemiological information or exposure history, concomitant devices in use, preexisting patient condition such as underlying infection, remote septic sources such as indwelling lines/catheters or other bioprosthetic material in situ. care of access site. because this evidence has not been provided, the root cause of the infection cannot be determined. this report is being filed as part of a retrospective review of historical records.
33806 | 1220648-2023-03960 | 17907232 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17907232&pc=OZD | 10/14/2022 | 01/12/2024 |  | 10/10/2023 |  |  | 10/14/2022 | 10/10/2023 |  |  | 06/30/2024 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023229609 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/14/2022 |  |  |  |  | Hold for dn 10/10 the user facility reported a 55-year-old male implanted with an impella device for mechanical circulatory support. the complainant reported difficulty delivering the impella device due to calcification. several attempts were made in both the right and left axillary arteries, to no avail. the pump was noted as "destroyed", therefore, insertion was aborted and a new impella was used. | [5 empty] | The investigation into the delivery issues- use has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, the root cause of the delivery issue was most likely due to the patient condition. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
33807 | 1220648-2023-03961 | 17907234 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17907234&pc=OZD | 06/24/2021 | 01/19/2024 |  | 10/10/2023 |  |  | 06/24/2021 | 10/10/2023 |  |  | 10/31/2021 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1464263 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/24/2021 |  |  |  |  | The user facility reported a 67-year-old male admitted in cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that impella was malpositioned at the septal wall and the medical team was unable to manipulate after multiple attempts, unable to torque to new position, or to go past p4 without suction. decision made to place back on central extra-corporeal membrane oxygenation (ecmo) and use impella to unload left ventricle (lv) with plan to wean off ecmo and back onto impella in future. impella might have been higher in ventricle, but hypertrophic ventricle did not allow the medical team to obtain flows if advanced deeper. all flows were normal, and labs were fine and there was no harm to the patient reported. | [5 empty] | The investigation into has been completed since the original report was filed. no product was returned. the root cause of the positioning issue was use related. | The impella device was discarded by the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
33808 | 1220648-2023-03963 | 17907239 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17907239&pc=OZD | 04/18/2021 | 01/23/2024 |  | 10/10/2023 |  |  | 04/18/2021 | 10/10/2023 |  |  | 08/31/2022 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 1496975 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/18/2021 |  |  |  |  | The user facility reported a 56-year-old male implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that a patient on impella support experienced oozing at impella site and 2 stitches were placed to make bleeding stop. the patient received 1.5 l of fluid overnight and 1 unit of packed red blood cells (prbc). | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Since the medical center did not return the impella device, no benchtop analysis was possible. the cause of access site bleeding was not determined since product was not returned and enough information was not provided.
33809 | 1220648-2023-03962 | 17907240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17907240&pc=OZD | 12/21/2022 | 01/24/2024 |  | 10/10/2023 |  |  | 12/21/2022 | 10/10/2023 |  |  | 07/31/2024 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023256377 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Tachycardia | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/21/2022 |  |  |  |  | In addition to the reported product malfunction, it was found that the patient endured vt during delivery as well as post placement of the impella device. the patient required 2 shocks as treatment. | The user facility reported a 749 year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported there was leaking from the yellow lumen on impella 5.5. no alarms were present. the purge was running sodium bicarbonate. the patient was to be explanted the next day so the leaking continued until then. there was no patient harm reported. |  |  |  |  | The investigation into the purge leak - pump and the ventricular tachycardia has been completed. the device was not returned for investigation. review of the clinical notes revealed the location of the hematoma to be the groin. based on the clinical information, the cause of the purge leak was most likely sidearm yellow luer damage due to interaction with bicarb. the root cause of the vt was determined to be the patient condition because the patient experienced prolonged runs of pulsatile ventricular tachycardia during this portion during the time of placement of impella. | B(5) updated to include vt endured during delivery as well as post placement of the impella device. the patient required 2 shocks as treatment. a supplemental mdr will be filed upon completion of our investigation. | The investigation into the purge sidearm leak has been completed. the device was not returned for evaluation. however, the clinical information was sufficient in determining the root cause, given the results of prior failure investigations. it was concluded that the cause of the purge leak was sidearm yellow luer damage due to sodium bicarbonate use. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33810 | 1220648-2023-03964 | 17907258 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17907258&pc=OZD | 09/09/2021 | 07/08/2024 |  | 10/10/2023 |  |  | 09/09/2021 | 10/10/2023 |  |  | 07/31/2023 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022112643 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | Device Handling Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/09/2021 |  |  |  |  | The user facility reported a 59-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the user dropped the console leading the white cord to break and leading to pump failure. multiple attempts were made to restart the pump, however these were unsuccessful. | [5 empty] | The investigation for the pump stop has been completed. no product was returned. however, the clinical details were sufficient to determine the root cause. the root cause of the electrical open issue was an open electrical line, the console was dropped resulting in damaged white cord as per the clinical description. the instructions for use for the impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock states the following: section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ d4-added model number, catalog number, serial number, lot number, expiration date, and udi number. added- h4 added-d6a/d6b. d1-corrected brand name. d2-corrected common device name. d4-corrected model number, catalog number, serial number, lot number, expiration date, and udi number. f6/f8 should have been left blank in the initial report. added- h4. h5 changed to yes. h8-changed to initial use of device. h6 impact code changed to 4629. h6 med dev prob code changed to 1503. h6 component code changed to 765. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
33811 | 1220648-2023-03965 | 17907296 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17907296&pc=OZD | 09/09/2021 | 03/13/2025 |  | 10/10/2023 |  |  | 09/09/2021 | 10/10/2023 |  |  |  |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/09/2021 |  |  |  |  | The user facility reported a 73-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support had the purge side arm break. the pump was exchanged for a new one. the serial number is now known for this device. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. the root cause of the issue was a leak from the sidearm but the exact location of the leak was unknown.
33812 | 1220648-2023-03966 | 17907299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17907299&pc=OZD | 09/19/2022 | 03/10/2025 |  | 10/10/2023 |  |  | 09/19/2022 | 10/10/2023 |  |  | 06/30/2024 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023237106 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/19/2022 |  |  |  |  | The user facility reported a 78-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the medical staff decided to implant an impella rp and also an impella 5.5 in a patient. after impella 5.5 placement the patient started bleeding from the access site and plates and factor 7 were given. the medical staff could not keep enough volume and the patient had to be taken off the pump. | [5 empty] | The impella device was discarded by the customer and therefore,can evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The cause of the access site bleeding cannot be determined since the product was not returned and enough clinical info was not provided.
33813 | 1220648-2023-03967 | 17907302 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17907302&pc=OZD | 06/07/2021 | 03/10/2025 |  | 10/10/2023 |  |  | 06/07/2021 | 10/10/2023 |  |  | 09/30/2022 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 1499125 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/07/2021 |  |  |  |  | The user facility reported a 65-year-old male implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the medical team who was explanting an impella pump noticed a clot which was removed from graft and graft was trimmed, clipped, and over-sewn. the chest washout was completed and chest was re-closed. sternum was broken due to soft closure so sternal plating was used to hold the sternum together. incision was closed. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | As the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined.
33814 | 1220648-2023-03968 | 17907307 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17907307&pc=OZD | 06/26/2021 | 10/10/2023 |  | 10/10/2023 |  |  | 06/26/2021 | 10/10/2023 |  |  | 01/31/2023 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2022041533 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/26/2021 |  |  |  |  | The user facility reported a 68-year-old male was implanted with an impella 5.5 device for mechanical circulatory support. during support, it was noted via echo that a clot was on the tip of the inflow. the clot coincided with a report of low flows and the decision was made to explant the pump. | [5 empty] | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
33815 | 1220648-2023-03969 | 17907636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17907636&pc=OZD | 09/23/2022 | 10/10/2023 |  | 10/10/2023 |  |  | 09/23/2022 | 10/10/2023 |  |  | 02/29/2024 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023190523 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hematoma | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/23/2022 |  |  |  |  | The user facility reported that a 65-year-old male patient implanted with the impella 5.5 for mechanical circulatory support experienced positioning issues. the physician attempted to perform slight manipulation of the device when crossing the valve to get it to flip away from the mitral valve upon initial implant. transesophageal echocardiogram measured 5.58cm. the following morning, the patient experienced a "purge flow decreased" alarm followed by an "impella in aorta" alarm. echo was performed and the device was noted to be right at the aortic valve but not across or into the left ventricle. the surgeon subsequently, attempted advancement with no success. additionally, a hematoma was noted at the insertion site that required evacuation. the patient was taken back to for pump exchange. no further information was provided. there were no reports of harm to the patient. | [5 empty] | The investigation into the reported positioning issues was completed. the device was returned for evaluation. visual inspection of the pump revealed that the catheter was cut. there were no cannula/catheter kinks observed and the touhy-borst was tight and functional. based on the available information, the root cause of the positioning issue was use-related. this report is being filed as part of a retrospective review of historical records. the failure modes will be monitored and trended.
33816 | 1220648-2023-03970 | 17908219 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17908219&pc=OZD | 02/07/2022 | 10/10/2023 |  | 10/10/2023 |  |  | 02/07/2022 | 10/10/2023 |  |  | 09/30/2023 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022143121 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The us complainant has an impella 5.5 supporting a patient for over a month. the pump support is ongoing while awaiting transplant. on day 37 of support there was a malfunctioning purge cassette that prompted purge cassette replacement. on a later date the patient experienced low pump pressure due to alleged positioning issues which required the pump to be repositioned to fix the issue. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿.
33817 | 1220648-2023-03971 | 17908232 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17908232&pc=OZD | 07/13/2023 | 10/10/2023 |  | 10/10/2023 |  |  | 07/13/2023 | 10/10/2023 |  |  | 01/31/2025 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024333694 |  | 0048-0045 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | [9 empty] | The complainant reported that the medical team had trouble inserting the impella cp pump across the aortic valve. the decision was made to remove that pump and insert a new one which was able to be inserted successfully. there was no harm to the patient as a result of this incident. | [5 empty] | The investigation into the reported events has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. data logs shows the first pump was run for 12 min and had an impella position in the aorta alarm. doctor had difficulty in repositioning the pump without a guidewire. the second cp pump was placed without issue. the evaluation revealed that the the cause of the issue was wireless re-access. per ifu 0048-9007, the impella device is not indicated for wireless re-access. the failure modes will be monitored and trended.
33818 | 1220648-2023-03972 | 17908249 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17908249&pc=PYX | 03/09/2023 | 10/10/2023 |  | 10/10/2023 |  |  | 03/09/2023 | 10/10/2023 |  |  | 12/31/2024 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024304801 |  | 0046-0035 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The us complainant reported during intensive care unit (icu) check-in, the nurse noticed placement signal numbers and waveforms represented more right ventricle (rv). chest x-ray confirmed that the impella rp had become dislodged into rv. therefore, a choice was made to explant the malpositioned rp. the impella rp was explanted within 3 hours of admission to unit. | [5 empty] | The investigation into the reported events has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the cause of the issue was not determined since product was not returned. the failure modes will be monitored and trended. instructions for use for the related event are as follows: ¿the impella rp flex with smartassist system is contraindicated for use with patients experiencing any of the following conditions: disorders of the pulmonary artery wall that would preclude placement or correct positioning of the impella rp flex with smartassist device; mechanical valves, severe valvular stenosis or valvular regurgitation of the tricuspid or pulmonary valve; mural thrombus of the right atrium or vena cava; anatomic conditions precluding insertion of the pump; presence of a vena cava filter or caval interruption device, unless there is clear access from the internal jugular vein to the right atrium that is large enough to accommodate a 22 fr catheter.¿
33819 | 1220648-2023-03974 | 17908296 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17908296&pc=OZD | 09/01/2022 | 10/01/2025 |  | 10/10/2023 |  |  | 09/01/2022 | 10/10/2023 |  |  | 10/31/2023 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022145177 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/01/2022 |  |  |  |  | The user facility in the eu reported a 60-year-old male admitted with cardiomyopathy and cardiogenic shock, was implanted with an impella device for mechanical circulatory support. purge leak between purge cassette and pump. purge cassette change did not resolve the issue. purge bypass placed until pump will be exchanged. patient survived and is still on bridge to transplant. | [5 empty] | The impella device was received from the customer and therefore,an evaluation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records. | D3 manufacture name address, country and fax number were erroneously entered on the initial manufacturer device report number 1220648-2023-03974. d6a and d6b revised from the initial manufacturer device report 1220648-2023-03974 in accordance with updated procedures. (d6b is unknown). f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2023-03974 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2023-03974 in accordance with updated procedures. g1 manufacturing site name was erroneously entered on the initial manufacturer device report number 1220648-2023-03974. g1 manufacturing site fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2023-03974.
33820 | 1220648-2023-03975 | 17908351 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17908351&pc=OZD | 04/14/2022 | 08/14/2025 |  | 10/10/2023 |  |  | 04/14/2022 | 10/10/2023 |  |  | 11/30/2023 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023169277 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Malfunction | Yes | Yes | Break | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 04/14/2022 |  |  |  |  | Additional information noted care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 65-year-old female patient was implanted with an impella 5.5 device for mechanical circulatory support. during support, it was noted that the pump had a fractured cable near the red impella plug. it was documented that the grey connector was broken and that wires were visible from the cable. the staff was advised to place tegaderm over the exposed area to prevent moisture and foreign material from entering the site. there was no patient harm and support was maintained with the implanted pump. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections b1, b2, b5, h1, and h6. | The impella device was not returned by the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Per the clinical information provided, the root cause of the product damage (bond failure) issue was not determined since product was not returned for investigation. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2023-03975. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2023-03975 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2023-03975 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2023-03975.
33821 | 1220648-2023-03973 | 17908391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17908391&pc=OZD | 09/14/2023 | 01/29/2024 |  | 10/10/2023 |  |  | 09/14/2023 | 10/10/2023 |  |  | 06/30/2025 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024376350 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 81-year-old male noted as high risk percutaneous cardiac intervention (hrpci) was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient had signs of hemolysis, which were amber urine and labs elevated. patient on continuous venovenous hemofiltration. the cause of acute hf remains unclear, possible delayed covid myocarditis. the impella device was positioned well and confirmed with echo. no blood products were given to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the hemolysis has been completed. the product and data were not returned for review. the root cause of the hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review.
33822 | 1220648-2023-03976 | 17908399 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17908399&pc=OZD | 10/07/2022 | 09/19/2024 |  | 10/10/2023 |  |  | 10/07/2022 | 10/10/2023 |  |  | 11/30/2023 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023169311 |  | 0550-0008 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/07/2022 |  |  |  |  | The user facility reported a 56-year-old male admitted in cardiogenic shock and suffering from cardiomyopathy was implanted with an impella device for mechanical circulatory support. the us complainant had a 5.5 support ongoing for just over 14 days (indicated use time) when the flows and mc were not optimal and the pump had "saturation" alarms. the patient was fully dependant on the pump while awaiting vad or transplant. the pump was explanted and the patient made a cmo/care and comfort only and expired. the outcome occurred after explant when not on impella. the team made choice not to replace the pump despite dependance. | The procedural outcome noted that the patient expired. medical safety review of the event for the patient's demise noted: there is not enough evidence to exclude impella as an associated factor in the patient's outcome. patient's peri-procedural condition was poor. |  |  |  |  | The pump was returned with the catheter off. a significant mass of biomaterial was found within the introducer graft. otherwise, no damage or abnormalities were found on the returned part of the impella. after reviewing the clinical information, it was determined that the cause of the saturated flow was most likely ingested biomaterial. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03976 was provided.
33823 | 1220648-2023-03977 | 17908518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17908518&pc=OZD | 05/07/2023 | 06/06/2024 |  | 10/10/2023 |  |  | 09/15/2023 | 10/10/2023 |  |  | 09/30/2024 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023276260 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient had a chronic total occlusion of the left anterior descending artery which was treated with good angiographic result. the impella was weaned and explanted. the site was perclosed twice. the site is intact. | The user facility reported a 60-year-old male noted as high risk coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant enrolled the patient in the protect iv registry. the team later noted in the crf that there was atrial fibrillation (af) with multiple etiologies: heart failure, pericarditis, and procedural complications. the cec determined the team to be a definite relationship to the impella despite the pi statement of it not being related. |  |  |  |  | The investigation into the atrial fibrillation and the vascular damage - great vessel issues has been completed since the original report was submitted. the product was not returned for investigation. according to the clinical details, the impella shifted in the aorta during removal of extracorporeal membrane oxygenation causing vascular complications. the root cause of the vascular damage - great vessel is most likely due to use. as the necessary information was not provided, the root cause of the atrial fibrillation was not determined. b.3 date of event revised as previously entered incorrectly. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted report incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33824 | 1220648-2023-03978 | 17908567 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17908567&pc=OZD | 06/16/2021 | 01/24/2024 |  | 10/10/2023 |  |  | 06/16/2021 | 10/10/2023 |  |  | 03/31/2023 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022071442 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Activation, Positioning or Separation Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/16/2021 |  |  |  |  | The user facility reported a 38-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that difficulty was encountered while attempting to implant an impella 5.5 pump for patient support. while advancing the device, the guidewire came out of the left ventricle and rested above the aortic valve. attempts to re-advance the guidewire into the ventricle were unsuccessful and the decision was made to remove the wire and pump. coinciding with the delivery issue, it was also noted that bleeding developed at the anastomosis. sutures were placed and blood was given to treat the condition. the patient remained supported by the already implanted cp pump. | [5 empty] | The investigation into the reported positioning issue has been completed since the original report was filed. the root cause of the access site bleeding was unable to be determined as limited clinical details an no product were returned for investigation. the root cause of the delivery issue anatomy was most likely due to the patient¿s condition. correction : during the initial report the delivery issue was not a reportable failure mode, as per new revised safety guideline and update it is now a reportable failure mode so its been added to this follow up report. correction section h6: the medical device problem code and component codes are added to reflect the new reportable guideline | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
33825 | 1220648-2023-03979 | 17908576 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17908576&pc=OZD | 06/09/2022 | 09/14/2024 |  | 10/10/2023 |  |  | 06/09/2022 | 10/10/2023 |  |  | 11/30/2023 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023169289 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Death | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 06/09/2022 |  |  |  |  | Additional information noted the patient expired on support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 76-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during impella support, the patient went into ventricular fibrillation in response to an increase in support levels from p4 to p9. the support level was reduced in response to the condition and it was documented that support levels had to remain at p4 or lower to avoid suction. the pump remained implanted after position was verified via transesophageal echocardiography (tee). |  |  |  |  | The investigation for the reported ventricular fibrillation has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the ventricular fibrillation could not be determined. added: h5-yes. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03979 was provided. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
33826 | 1220648-2023-03980 | 17908579 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17908579&pc=OZD | 01/26/2022 | 03/08/2024 |  | 10/10/2023 |  |  | 01/26/2022 | 10/10/2023 |  |  | 10/31/2023 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022157686 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/26/2022 |  |  |  |  | The user facility reported a 76-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that a patient in a serious medical condition had to be placed on impella support. a computed tomography (ct) of the head performed on monday did show that patient had a watershed stroke. it is unknown, however, if the stroke occurred prior to the impella implant or after. | [5 empty] | The investigation into the stroke/cva issue has been completed since the original report was filed. the device was not returned for investigation. as no clinical information was not provided, the root cause of the stroke/cva could not be determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
33827 | 1220648-2023-03981 | 17908650 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17908650&pc=OZD | 10/17/2022 | 10/10/2023 |  | 10/10/2023 |  |  | 10/17/2022 | 10/10/2023 |  |  | 04/30/2024 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023230541 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 43 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/17/2022 |  |  |  |  | The user facility in the eu reported a 43-year-old female admitted in cardiogenic shock was implanted with an impella 5.5, via the axillary artery, for mechanical circulatory support. due to visible clot at the inlet & left ventricle, a decision to remove it . no neurological complication but patient complaint about pain in the left arm. visible thrombus in the brachialis artery was successuflly aspirated. pump successfully weaned & explanted. | [5 empty] | Data logs were reviewed and did confirm that suction began shortly after activation and fluctuating pump flow and p-levels persisting throughout the case indicating that the biomaterial was likely ingested during support. there were no motor current spikes early in the case but on 10/16 motor current spikes did occur and correlated with the reported pump in aorta alarms when the position was verified indicating a likely persisting interaction with the pump impeller. the return pump was visually inspected and biomaterial was observed in the outflow via borescope view. the pump was run in a basin of water and suction with high pump flow occurred. the cause of the issue was biomaterial.
33828 | 1220648-2023-03982 | 17909506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17909506&pc=OZD | 09/11/2023 | 06/20/2024 |  | 10/10/2023 |  |  |  | 10/10/2023 |  |  | 07/31/2025 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024376657 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | Aortic Valve Insufficiency/ Regurgitation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | New information received via long-term outcome and quality indicator (loqi) impella registry notification that indicates this patient to have developed thrombocytopenia while on impella support, with a possible causal relationship alleged. in response, multiple platelet transfusions were required. | The user facility reported that in operating room (or), after explanting the impella 5.5 device, the patient was noted to have new moderate aortic insufficiency. |  |  |  |  | B(5): additional information was received indicating patient endured thrombocytopenia prompting platelet transfusion. the device was not returned. a supplemental will be filed upon completion of our evaluation. | The investigation for the reported thrombocytopenia and heart valve damage has been completed. product was not returned for investigation. due to limited clinical details provided, the root cause of the thrombocytopenia as well as the heart valve damaged cannot be determined. therefore, the root cause of the thrombocytopenia and heart valve damage could not be determined. the instructions for use for the impella 5.5 for use during cardiogenic shock in section: potential adverse events (united states) states ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- clinical code 4431. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33829 | 1220648-2023-03983 | 17909508 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17909508&pc=OZD | 09/11/2023 | 01/29/2024 |  | 10/10/2023 |  |  | 09/11/2023 | 10/10/2023 |  |  | 04/30/2025 |  | 10/10/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024351045 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Female |  |  |  | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 28-year-old female noted as high risk coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant had an impella cp support ongoing for a younger patient with non-hodgkins lympoma. the patient needed the support for a hrpci. the pump was placed but despite all re-positioning attempts the team found flows to be low. the pump remained on for the support of the hrpci and was then explanted. | [5 empty] | The investigation into the low pump flow has been completed. the pump was not returned for investigation. data logs show consistently low pump flow throughout the case without any motor current spikes. additionally, there were reported downward trends in the placement signal and suction condition from the beginning. as per the clinical, patient had short aortic root. the cause of the low pump flow issue was most likely patient condition related based on data log review and provided clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
33830 | 1220648-2023-03715 | 17894661 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17894661&pc=OZD | 06/28/2021 | 09/18/2024 |  | 10/09/2023 |  |  | 06/28/2021 | 10/09/2023 |  |  | 03/31/2023 |  | 10/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022065621 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/28/2021 |  |  |  |  | The user facility reported a 71-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced bleeding and required blood products. in addition, the following day, the patient also went to ventricular fibrillation and had to be shocked a couple of times. the pump supported for 8 days til family made choice to withdraw care. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided.
33831 | 1220648-2023-03774 | 17895382 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17895382&pc=OZD | 09/08/2023 | 07/30/2024 |  | 10/09/2023 |  |  | 09/08/2023 | 10/09/2023 |  |  | 04/30/2025 |  | 10/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024348512 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/08/2023 |  |  |  |  | The user facility reported an 82-year-old male was implanted with an impella 5.5 as a bridge to transplant of left ventricle assist device. the patient experienced a hematoma that quickly subsided with manual pressure. echo done showing pump angled toward mitral valve and the pump was repositioned to mid left ventricle. the patient had multiple episodes of vomiting causing bleeding at the access site. hemostasis was achieved with mattress suture and manual pressure. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the major access site bleed was patient movement as the patient had an episode of vomiting before the bleed occurred. it was resolved with manual pressure and mattress sutures. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- h6 component code 925 added- d6a/d6b f6/f8 should have been left blank in the initial report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33832 | 1220648-2023-03765 | 17895663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17895663&pc=OZD | 07/14/2021 | 03/10/2025 |  | 10/09/2023 |  |  | 07/14/2021 | 10/09/2023 |  |  | 12/31/2022 |  | 10/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022024332 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 26 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/14/2021 |  |  |  |  | The user facility reported a 26-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that a blood clot was noticed on the patient's left arm. patient survived on the 5.5 support for over 7 days and bridged to transplant. | [5 empty] | The impella device was discarded by the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported thrombus issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
33833 | 1220648-2023-03769 | 17895664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17895664&pc=OZD | 09/06/2023 | 01/25/2024 |  | 10/09/2023 |  |  | 09/06/2023 | 10/09/2023 |  |  |  |  | 10/09/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1590813 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during training, upon opening the purge chamber on the console the purge disk that was inserted appeared to be loose. upon further assessment and careful removal of demo purge cassette, the blue prongs were pulled forward and there appeared to be a crack in the blue purge disk stabilizer. no patient involvement. | [5 empty] | The investigation into the reported aic purge issue has been completed since the original report was filed. the console was tested after arriving in fs. the purge disc retainer was confirmed to be broken. the cause of the issue was a defective component. | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
33834 | 1220648-2023-03772 | 17895666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17895666&pc=OZD | 09/07/2023 | 02/06/2024 |  | 10/09/2023 |  |  | 09/07/2023 | 10/09/2023 |  |  | 06/30/2025 |  | 10/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024367249 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 22 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/07/2023 |  |  |  |  | The user facility reported a 29-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support as a bridge to transplant. the team delivered the pump and had high purge pressure alarms. the alarm also fired off for "purge system blocked" alarms. this prompted the pump to be explanted. the explanted pump was observed to have a large clot burden adhered to it. there was no noted embolization. the team implanted a new impella 5.5. | [5 empty] | The investigation into the reported high purge pressure has been completed since the original report was filed. data review: data logs confirmed the failure mode & clinical narrative. logs showed after pump started, mc spiked followed high pp & low flow that triggered purge pressure high, auto suction response & suction alarms. the high pp remained to the rest of the case. product was not returned for review. based on clinical description & data review, the root cause of high purge pressure was most likely due to biomaterial ingestion. the root cause of thrombus was due to patient condition. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
33835 | 1220648-2023-03768 | 17895668 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17895668&pc=OZD | 09/06/2023 | 01/22/2024 |  | 10/09/2023 |  |  | 09/06/2023 | 10/09/2023 |  |  | 02/29/2024 |  | 10/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 JP |  |  | OZD |  |  | IMPELLA 5.0 JP |  |  | ABIOMED, INC. |  |  | US |  | 2023190995 |  | 005066 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 48-year-old female noted as high risk percutaneous cardiac insertion was implanted with an impella 5.0 device for mechanical circulatory support. the complainant reported low purge pressure alarms that would occur then stop repeatedly. the staff replaced the purge cassette and purge pressure dropped further. the 5% glucose solution was changed to 10% glucose solution, and the purge pressure drop alarm occurred repeatedly at 260-330 mmhg and disappeared. physician replaced the impella pump successfully. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation into the purge leak ¿ pump issue has been completed since the original report was filed. the device was returned for investigation. it was reported that there was no problem found with the product. the root cause is most likely due to a use issue.
33836 | 1220648-2023-03672 | 17895669 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17895669&pc=OZD | 04/14/2023 | 10/04/2023 |  | 10/09/2023 |  |  | 04/14/2023 | 10/09/2023 |  |  | 12/31/2024 |  | 10/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET ¿ (LMR) |  |  | ABIOMED, INC. |  |  | US |  | 2024318856 |  | 1000100 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | Pressure Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/14/2023 |  |  |  |  | The user facility reported 69 yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that after placement of the 5.5 for the surgical patient, there was a high purge pressure alarm. upon evaluation, there was leaking noted inside the purge side arm cover. troubleshooting was performed by changing the purge cassette, which did not remedy the issue. the pump and console was replaced. there was no patient harm reported. | [5 empty] | The investigation for the reported purge side arm leak and the high pump pressure has been completed. the field data log was reviewed and upon activation, placement signal waveform was ventricular and high purge pressure occurred. no low purge pressure alarms. purge cassette replacement confirmed. the cause of the high purge pressure was not determined since product was not returned. the cause of the issue was a leak from the sidearm but the exact location of the leak was unknown. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33837 | 1220648-2023-03771 | 17895671 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17895671&pc=OZD | 09/07/2023 | 02/01/2024 |  | 10/09/2023 |  |  | 09/07/2023 | 10/09/2023 |  |  |  |  | 10/09/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1643465 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 09/07/2023 |  |  |  |  | The user facility reported a 70-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. while on support, the abiomed field service representative noticed a controller error alarm on the automated impella controller (aic). the patient was stable and did not require additional hemodynamic support. staff walked through aic to aic transfer with resulting resolution of alarm. aic in question tagged and sent to the user facility's biomed. | [5 empty] | The investigation into the aic - loss of opm data and the aic - controller error (yellow alarm) has been completed since the original report was filed. the device was returned for investigation. imc logs confirmed the reported failure mode given there was a controller error alarm triggered during the clinical procedure. real time logs confirmed that this was a pattern failure mode in which all signals received from optical bench were interpreted as -16384 values. the console was inspected and evaluated on an engineering lab bench. it was reported that the failure mode was not able to be reproduced while running with a known good pump. this known pattern failure, which is characterized by a temporary loss of optical data (placement, snr, contrast, lamp, gain) and triggering of a controller error alarm, has a low probability of reoccurrence. the cause of the loss of optical data issue was not determined because it could not be reproduced. | H.6 type of investigation code 10 was inadvertently left off the last manufacturer device report submitted and was added. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
33838 | 1220648-2023-03673 | 17895672 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17895672&pc=OZD | 09/07/2023 | 11/06/2023 |  | 10/09/2023 |  |  | 09/07/2023 | 10/09/2023 |  |  |  |  | 10/09/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1340487 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 YR |  | Unknown |  |  | Asian | Malfunction | No | Yes | Activation Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility in japan reported that the automated impella controller (aic) failed to start. the aic was not in use when the issue occurred. the aic was replaced successfully. no patient harm or injury. | [5 empty] | The investigation into the aic - system self-check error has been completed since the original report was filed. the console was returned for investigation. as per clinical, the cause of the aic issue was contamination. the failure mode will be monitored and trended. | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
33839 | 1220648-2023-03770 | 17895674 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17895674&pc=OZD | 08/05/2021 | 03/10/2025 |  | 10/09/2023 |  |  | 08/05/2021 | 10/09/2023 |  |  | 04/30/2023 |  | 10/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022032571 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/05/2021 |  |  |  |  | The user facility reported a 65-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support had the pump stop due to high motor current. there was no adverse event to the patient as a result of this incident. the pump support continued on a second 5.5 pump. | [5 empty] | The investigation for the reported pump stop issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records. | The impella device was discarded by the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
33840 | 1220648-2023-03674 | 17895675 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17895675&pc=OZD | 02/28/2023 | 10/04/2023 |  | 10/09/2023 |  |  | 02/28/2023 | 10/09/2023 |  |  | 04/30/2024 |  | 10/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023212021 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported swelling in the impella5.5/picc line inserted limb. no intervention was required. on the same day, a mural thrombus was discovered in the patient's left ventricle. while this was not attributed to the device and did not affect performance, the pump was removed out of caution. | [5 empty] | The investigation into the reported thrombus has been completed. the medical center did not return any impella products for benchtop analysis; only the clinical report was evaluated. the evaluation revealed that root cause of the patient injury/ swelling in the impella limb was unable to be determined due to limited clinical details. the thrombus was unrelated to the use of impella. instructions for use notes: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33841 | 1220648-2023-03675 | 17895676 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17895676&pc=OZD | 09/07/2023 | 07/31/2024 |  | 10/09/2023 |  |  | 09/07/2023 | 10/09/2023 |  |  | 06/30/2025 |  | 10/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024373026 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional treatment included 2 fresh frozen plasma and 1-unit platelets, and 3500 crystalloids | The user facility reported a 65-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient experienced bleeding from the right femoral impella cp access site. the bleeding did not resolve despite manual pressure. five units of blood were given, and the patient was taken to the operating room for vascular repair of femoral artery with purse-string suture around repositioning sheath. the access site was left open while on impella cp support. the bleeding was thought to be due to the size difference in peel-away sheath and repositioning sheath. |  |  |  |  | The investigation for access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was most likely lack of vessel recoil as it was reported that best practices were followed, and it was thought to be due to smaller size repositioning sheath from larger peel away. instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b5 additional treatment information provided. h6 component code added the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-03675: in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. g1 reporting contact fax number missing. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33842 | 1220648-2023-03767 | 17895677 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17895677&pc=OZD | 07/30/2021 | 03/01/2024 |  | 10/09/2023 |  |  | 07/30/2021 | 10/09/2023 |  |  | 05/31/2023 |  | 10/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022088848 |  | 0550-0008 |  | No |  |  | Yes |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/30/2021 |  |  |  |  | The user facility reported a 57-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced a stroke. | [5 empty] | The investigation into the stroke/cva event has been completed. the device was not returned for investigation. as the clinical information was not provided, the true root cause of the stroke/cva could not be determined. | The impella device was discarded by the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
33843 | 1220648-2023-03773 | 17895678 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17895678&pc=OZD | 08/20/2021 | 10/05/2023 |  | 10/09/2023 |  |  | 08/20/2021 | 10/09/2023 |  |  | 05/31/2023 |  | 10/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022091932 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 15 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/20/2021 |  |  |  |  | The user facility reported a 15-year-old male with cardiomyopathy and bridging to/awaiting transplant was implanted with an impella 5.5 device for mechanical circulatory support. the patient on impella support experienced sudden changes in purge flow/pressure. explained it was most likely due to kinking during patient movement. there was no harm to the patient reported as a result of this incident. support remained on for over 19 days til explanted and a heartmate 3 placed. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
33844 | 1220648-2023-03775 | 17895679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17895679&pc=OZD | 08/14/2021 | 03/10/2025 |  | 10/09/2023 |  |  | 08/14/2021 | 10/09/2023 |  |  | 05/31/2023 |  | 10/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022085290 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/14/2021 |  |  |  |  | The user facility reported a 67-year-old male with heart failure and cardiomyopathy awaiting heart transplant was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that a patient on impella support experienced hematuria. echocardiogram completed revealed that the device was deep with inflow by mitral structures, and it repositioned and pointed towards apex. four days later the patient experienced low purge pressure which was resolved by impella being repositioned at bedside. the pump reamined on for 12 days til removed and transplant performed. | [5 empty] | The investigation for the reported high purge pressure and positioning issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings, cautions and precautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ correct impella catheter position ¿for optimal positioning of the impella catheter, the inlet area of the catheter should be 3.5 cm below the aortic valve annulus and well away from papillary muscle and subannular structures. the outlet area should be well above the aortic valve.¿ this report is being filed as part of a retrospective review of historical records. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
33845 | 1220648-2023-03766 | 17895690 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17895690&pc=OZD | 07/29/2021 | 01/26/2024 |  | 10/09/2023 |  |  | 07/29/2021 | 10/09/2023 |  |  | 05/31/2023 |  | 10/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022091944 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  |  |  | Injury | Yes | No | Activation, Positioning or Separation Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/29/2021 |  |  |  |  | The user facility reported a 56-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient experienced bleeding from the access site which required the patient to go to the or and receive 12 units of blood. 11 days after this event the patient also experienced some positioning issues and the medical team discussed to reposition the pump. the pump remained on for support for 4 weeks and then bridged to a durable vad. | [5 empty] | The investigation for the reported positioning issue and the access site bleed has been completed. the device was not returned for evaluation. however, clinical information was provided is sufficient to determine the root cause of the positioning. the cause of the positioning issue was use-related due to impella being placed too deep per clinical description. the root cause of the access site bleeding could not be determined as there were limited clinical details. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
33846 | 1220648-2023-03893 | 17896013 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17896013&pc=OZD | 09/10/2023 | 11/30/2023 |  | 10/09/2023 |  |  | 09/10/2023 | 10/09/2023 |  |  | 06/30/2025 |  | 10/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024371214 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 65-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported while on impella support the patient had limb ischemia. a percutaneous fem-fem bypass was performed. no blood products were given to the patient. | [5 empty] | The investigation into the patient's limb ischemia has been completed. the impella cp was not returned for evaluation. as per clinical, patient lost pulses on impella leg with fem-fem bypass resolving ischemia. the cause of the limb ischemia issue was patient condition related with fem-fem bypass resolving ischemia on right leg per clinical review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33847 | 1220648-2023-03894 | 17896064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17896064&pc=OZD | 03/07/2023 | 10/09/2023 |  | 10/09/2023 |  |  | 03/07/2023 | 10/09/2023 |  |  | 09/30/2023 |  | 10/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2022145181 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/07/2023 |  |  |  |  | The user facility reported a 59-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. access site bleeding was one that only required manual compression and consult with surgery for redo to control the bleeding. | [5 empty] | The root cause of the access site bleeding issue was not determined since product was not returned and sufficient clinical information was not provided. the impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
33848 | 1220648-2023-03895 | 17896386 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17896386&pc=OZD | 09/10/2023 | 07/31/2024 |  | 10/09/2023 |  |  | 09/10/2023 | 10/09/2023 |  |  | 05/31/2025 |  | 10/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024361563 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/10/2023 |  |  |  |  | The user facility reported a 71-year-old female in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. when the impella cp was explanted, there was clot on the pigtail. the physician stated this was due to the pump being positioned too deep in the left ventricle. there were no signs of embolism and the clot was sent off to pathology. | [5 empty] | The investigation for thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings and cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." d6a and d6b added the following have been reported incorrectly or missing from manufacturer device report 1220648-2023-03895: a5 ethnicity. b2 outcomes attributed to adverse event missing. f6/f8-should have been left blank. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
33849 | 1220648-2023-03896 | 17896398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17896398&pc=OZD | 09/11/2023 | 01/16/2024 |  | 10/09/2023 |  |  | 09/11/2023 | 10/09/2023 |  |  | 12/31/2024 |  | 10/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024305762 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 51-year-old male patient noted as high risk percutaneous cardiac insertion was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient was experiencing hemolysis. the impella was removed and upon removal a clot was found in the inlet of the impella. hemolysis was resolved after pump removal. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into thrombus and hemolysis has been completed. the device was not returned for benchtop evaluation and investigation. based on clinical description and data analysis, a clot was found when the pump was removed. the root cause of the hemolysis was most likely was a biomaterial ingestion.
33850 | 1220648-2023-03897 | 17896420 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17896420&pc=OZD | 09/11/2023 | 01/26/2024 |  | 10/09/2023 |  |  | 09/11/2023 | 10/09/2023 |  |  |  |  | 10/09/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1687316 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Device Fell | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/11/2023 |  |  |  |  | The user facility reported a 71-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. during support, the user facility's automated impella controller (aic) feel over and had visible damage to the front right area where the impella white cable is located. the white cable loosely fit and the dust cover was broken. the impella 5.5 was operating normally and there was no patient harm. the aic will be returned for repair. | [5 empty] | The investigation into the aic ¿ damage has been completed since the original report was filed. the console was returned for investigation. the console was inspected and evaluated on an engineering lab bench. physical inspection of the front housing revealed damage of the blue optical receptacle. the root cause of the aic issue was a damaged blue receptacle. | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
33851 | 1220648-2023-03898 | 17896446 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17896446&pc=PYX | 09/10/2023 | 04/26/2024 |  | 10/09/2023 |  |  | 09/11/2023 | 10/09/2023 |  |  | 12/31/2024 |  | 10/09/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024296512 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  |  |  | Death | Yes | Yes | Improper Flow or Infusion | [6 empty] | Disseminated Intravascular Coagulation (DIC) | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/11/2023 |  |  |  |  | The user facility reported a 70-year-old female (race unknown) who had dilated right ventricle and increasing pulmonary hypertension was implanted with an impella rp device for mechanical circulatory support. the patient was in icu when the clinician pulled back the existing swan catheter that was inserted before the impella and the impella rp pump flows dropped and did not recover. the patient also had disseminated intravascular coagulation (dic) condition and was bleeding profusely from multiple areas not related to the impella. the patient deteriorated with persistent low flows and expired on impella support. it is unknown if the impella cause/contributed to the death. | [5 empty] | The impella device has been received from the customer however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the low pump flow has been completed. the impella device was not returned for evaluation. data logs shows the motor current spike followed by low pump flow for remaining of the case run. as per the given clinical details, when the swan was pulled back, flows dropped and did not recover; the patient deteriorated with persistent low flows and coded. the cause of the low pump flow issue was most likely biomaterial, based on data log review. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
33852 | 1220648-2023-03899 | 17896535 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17896535&pc=OZD | 09/11/2023 | 07/25/2024 |  | 10/09/2023 |  |  | 09/11/2023 | 10/09/2023 |  |  | 04/30/2025 |  | 10/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024351984 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was further reported that the patient expired after care was withdrawn. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there was not enough information to associate use of the device with the outcome. | The user facility reported the .018" guidewire (packaged with impella cp) was damaged beyond use when it was removed from the plastic ring. wire was described by physician and staff as being in a "corkscrew" once removed and prior to using with patient. wire was discarded and a .018" wire was removed from a second impella kit to complete insertion for the impella cp support. |  |  |  |  | The investigation for the product damage has been completed. according to the clinical, before beginning impella cp support, a 0.018 guidewire was damaged beyond use when it was removed from the plastic ring. the wire was discarded and another 0.018 wire was removed from a second impella kit to complete insertion. the root cause of the of the product damage (guidewire damage - peripheral vessels) cannot be determined due to lack of product returned. added- b5 mso review added-d6a/d6b. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
33853 | 1220648-2023-03900 | 17897494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17897494&pc=OZD | 01/29/2022 | 03/10/2025 |  | 10/09/2023 |  |  | 01/29/2022 | 10/09/2023 |  |  | 09/30/2023 |  | 10/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022145873 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Activation, Positioning or Separation Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/29/2022 |  |  |  |  | The user facility reported a 76-year-old male suffering from cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. the patient who was just transitioned from an impella cp to an impella 5.5 had experienced positioning issues confirmed by echocardiogram (echo). when the impella 5.5 pump was taken out the medical team noticed a clot at the tip of the pump. that prompted the medical team to remove the pump and replaced with the new one. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported thrombus and positioning issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the positioning issue was most likely a use issue due to poor placement. the root cause of thrombus was most likely due to the patient condition. this report is being filed as part of a retrospective review of historical records.
33854 | 1220648-2023-03901 | 17897643 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17897643&pc=OZD | 12/10/2021 | 10/28/2025 |  | 10/09/2023 |  |  | 12/10/2021 | 10/09/2023 |  |  | 03/31/2023 |  | 10/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022123479 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Device Sensing Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/10/2021 |  |  |  |  | The user facility reported a 57-year-old male undergoing aortic valve surgery was implanted with an impella 5.5 device for mechanical circulatory support. the medical staff was having difficulty delivering the device to a patient but after several attempts and 4 units of blood were given they were able to complete the delivery successfully. seven days later the patient experienced false positioning in the aorta alarms. the alarm was disabled and wavelengths were monitored more closely. | The complainant reported that the medical staff was having difficulty delivering the device to a patient but after several attempts and (b)(4) units of blood were given they were able to complete the delivery successfully. seven days later the patient experienced false positioning in the aorta alarms. the alarm was disabled and wavelengths were monitored more closely. |  |  |  |  | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Investigation was completed. the root cause of the delivery use issue was use issue due to loss of wire access and then wireless insertion. the impella device is not indicated for wireless insertion. also, use of non-abiomed product like wire v-18 was performed which is not per ifu protocols. optical issue- no product was returned. the root cause of the optical issue was most likely biomaterial, since there was a spike/trend in the motor current as per the data logs.
33855 | 1220648-2023-03902 | 17897760 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17897760&pc=OZD | 04/29/2022 | 03/10/2025 |  | 10/09/2023 |  |  | 04/29/2022 | 10/09/2023 |  |  | 12/31/2023 |  | 10/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023170360 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Handling Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/29/2022 |  |  |  |  | The user facility reported a 69-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device and ecmo for mechanical circulatory support. the patient had the impella placed direct via 10mm graft at the base of the innominate artery. the impella was kinked and unable to use. requiring a replacement. there was no harm to the patient as a result of this incident. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported product damage (cannula kink) issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
33856 | 1220648-2023-03903 | 17897832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17897832&pc=OZD | 04/05/2022 | 01/26/2024 |  | 10/09/2023 |  |  | 04/05/2022 | 10/09/2023 |  |  | 12/31/2023 |  | 10/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023172883 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/05/2022 |  |  |  |  | The user facility reported a 84-year-old male was implanted with an impella 5.5 device for mechanical circulatory support as support was needing escalation from the impella cp. the complainant reported that an impella 5.5 pump was unable to pass through the subclavian during attempted delivery. after multiple unsuccessful attempts, an angiogram was performed and revealed a dissection to the vasculature as well as the presence of heavy calcification. the implant was subsequently aborted as a result of the noted difficulties. an already implanted impella cp pump was left in place for ongoing support. | [5 empty] | The investigation into the reported vascular damage has been completed since the original report was filed. no product was returned. the root cause of the delivery issue - anatomy was most likely due to the patient condition. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records.
33857 | 1220648-2023-03904 | 17899949 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17899949&pc=OZD | 06/14/2022 | 01/26/2024 |  | 10/09/2023 |  |  | 06/14/2022 | 10/09/2023 |  |  | 10/31/2023 |  | 10/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022153060 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/14/2022 |  |  |  |  | The user facility reported a 64-year-old male was implanted with an impella 5.5 device for mechanical circulatory support as it needed escalation from the impella cp. when impella cp was removed and 5.5 placed, the patient had blood loss requiring blood products to be given and the site had to be surgically closed. | [5 empty] | The impella device was discarded by the customer and therefore, a benchtop evaluation of the device was not possible. the clinical report alone was analyzed and the root cause of the access site bleed was patient condition. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the reported access site bleeding has been completed since the original report was filed. the root cause of the access site bleeding was unable to be determined as limited clinical details an no product were returned for investigation.
33858 | 1220648-2023-03905 | 17899974 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17899974&pc=PYX | 09/12/2023 | 12/04/2023 |  | 10/09/2023 |  |  | 09/12/2023 | 10/09/2023 |  |  | 09/30/2023 |  | 10/09/2023 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023186652 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that after multiple attempts to place the impella rp, in conjunction with adequate wire purchase and buddy wire/catheter techniques, the initial rp could not be advanced and began to bow/kink in the ivc. this kinking became severe enough to restrict further use, and a second impella rp was opened for more attempts. | [5 empty] | The investigation into product malfunctions (cannula kink and delivery issues) has been completed. no product was returned. as per clinical, the impella rp pump catheter kinked during insertion in inferior vena cava and the pump was not able to advance past the right ventricle outflow tract (rvot) and pulmonic valve after several attempts. second rp pump was also unable to be advanced after multiple attempts. the cause of the delivery issue was patient condition related due to catheter kinked during insertion and the rp device unable to advance past the rvot and pulmonic valve per clinical review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿.
33859 | 1220648-2023-03906 | 17900093 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17900093&pc=OZD | 01/10/2023 | 10/09/2023 |  | 10/09/2023 |  |  | 01/10/2023 | 10/09/2023 |  |  |  |  | 10/09/2023 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1682677 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The us complainant had an aic removed and replaced due to fuzzy lines on the display. the ifu was followed and a backup was used for the patient support. no harm or injury noted from the interruption in the support. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. the evaluation revealed that the root cause of the display artifacts was a defective display. the failure modes will be monitored and trended.
33860 | 1220648-2023-03907 | 17900195 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17900195&pc=OZD | 06/23/2021 | 09/18/2024 |  | 10/09/2023 |  |  | 06/23/2021 | 10/09/2023 |  |  |  |  | 10/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | XXXXXX |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 06/23/2021 |  |  |  |  | The user facility reported a 68-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the medical staff experienced difficulty delivering the device in a patient and ultimately aborted the procedure opting to use a impella cp instead. as a result of this event, the patient experienced bleeding from the access site around anastomosis requiring blood products, multiple sutures, jackson-pratt drain and protamine. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the reported delivery issues- use and access site bleeding has been completed since the original report was filed. delivery issues - the cause for the delivery issue was patient condition as the pump could not pass beyond the anastomosis.access site bleeding (major) - the cause for the access site bleeding cannot be determined as enough clinical information is not available. correction: section d4: the model number of the pump was inadvertently left out in the initial report which is been added in this report. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided.
33861 | 1220648-2023-03908 | 17900279 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17900279&pc=OZD | 10/06/2022 | 10/09/2023 |  | 10/09/2023 |  |  | 10/06/2022 | 10/09/2023 |  |  | 02/28/2023 |  | 10/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022043492 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/06/2022 |  |  |  |  | The user facility reported a 67-year-old female suffering from cardiogenic shoc, was implanted with an impella 5.5 device for mechanical circulatory support. the patient was transferred to the tertiary center for care escalation with the 5.5 and a protekduo. during the impella 5.5 implant procedure, the introducer cracked and began leaking. the introducer was swapped and the subsequent implant was successful. there was no patient harm. | [5 empty] | The impella device was discarded by the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
33862 | 1220648-2023-03909 | 17900920 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17900920&pc=OZD | 08/24/2022 | 09/20/2024 |  | 10/09/2023 |  |  | 08/24/2022 | 10/09/2023 |  |  | 05/31/2024 |  | 10/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023222074 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/24/2022 |  |  |  |  | The user facility reported a 59-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the the medical staff was having difficulty delivering the impella past the aortic valve. on the first attempt the wire doubled back on itself and when the md pulled the impella back he lost wire access. the md had to start insertion all over again, but eventually got the impella across the valve. after the impella was in good working order, he reconfigured the ecmo so that the patient was on a protek duo set up. . | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The investigation for the delivery issue has been completed. the impella device was not returned for evaluation. however, clinical information provided was sufficient to determine the root cause of the reported delivery issues. according to clinical details, as the impella was being pulled back, wire access was lost. therefore, the root cause of the reported delivery issues was user-related. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2023-03909 was provided.
33863 | 1220648-2023-03910 | 17901042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17901042&pc=OZD | 09/12/2022 | 10/28/2025 |  | 10/09/2023 |  |  | 09/12/2022 | 10/09/2023 |  |  | 02/29/2024 |  | 10/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023188530 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/12/2022 |  |  |  |  | The user facility reported a 35-year-old female admitted in cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 experienced low purge pressure following a purge cassette exchange. the purge cassette was replaced, but the alarm and low purge pressure persisted. upon further troubleshooting, it was discovered that there was a fluid leak at the purge sidearm connection. a purge bypass was subsequently performed to address the leak. the following day, it was noted that flow rates had begun to drop as a result of a distended right ventricle. due to the purge bypass and drop in flow rates, the decision was made to place a new 5.5 pump, as well as an impella rp pump, for ongoing support. there was no lasting patient harm. | The complainant reported that the impella 5.5 experienced low purge pressure following a purge cassette exchange. the purge cassette was replaced, but the alarm and low purge pressure persisted. upon further troubleshooting, it was discovered that there was a fluid leak at the purge sidearm connection. a purge bypass was subsequently performed to address the leak. the following day, it was noted that flow rates had begun to drop as a result of a distended right ventricle. due to the purge bypass and drop in flow rates, the decision was made to place a new 5.5 pump, as well as an impella rp pump, for ongoing support. there was no lasting patient harm. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records. | Investigation was completed. purge leak-pump: no product was returned. the root cause of the purge leak-pump issue was not determined since product was not returned, as there was a leak observed but not enough information is provided in the clinical description provided. low pump flow- the root cause of low pump flow was not determined because was no determined and sufficient information was not provided to determine the root cause.
33864 | 1220648-2023-03911 | 17901301 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17901301&pc=OZD | 09/09/2022 | 01/26/2024 |  | 10/09/2023 |  |  | 09/09/2022 | 10/09/2023 |  |  | 05/31/2024 |  | 10/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023219346 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/09/2022 |  |  |  |  | The user facility reported a 68-year-old male needing escalation of support due to underlying disease, was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that a patient who transitioned from impella cp to impella 5.5 experienced oozing from the access site and as a result was required to receive blood products. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
33865 | 1220648-2023-03912 | 17901334 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17901334&pc=OZD | 10/16/2022 | 01/27/2024 |  | 10/09/2023 |  |  | 10/16/2022 | 10/09/2023 |  |  | 04/30/2024 |  | 10/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023211245 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Unspecified Infection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/16/2022 |  |  |  |  | The user facility reported a 48-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient developed a hematoma at their impella access site during pump support. it was documented that distention was noted at the patient's neck and, despite blood products being given, the patient's hemoglobin levels were dropping. further inspection identified a large infected hematoma at the impella pocket. the patient was taken to the operating room where the pump was removed and the hematoma was evacuated. vancomycin was given to the patient to treat the infection. the patient was stable following the event. | [5 empty] | The investigation into the reported access site bleeding and infection has been completed since the original report was filed. the root cause of the access site bleeding was unable to be determined as limited clinical details and no product were returned for investigation. no product was returned. the root cause of the infection/sepsis issue was patient condition related, as there was severe infection and hematoma observed and the relationship to impella is unknown. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ this report is being filed as part of a retrospective review of historical records.
33866 | 1220648-2023-03913 | 17901555 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17901555&pc=OZD | 07/05/2023 | 10/09/2023 |  | 10/09/2023 |  |  | 07/05/2023 | 10/09/2023 |  |  | 02/28/2025 |  | 10/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 SMARTASSIST SET, JP |  |  | ABIOMED, INC. |  |  | US |  | 2024311741 |  | 1000211 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  |  | Asian | Injury | Yes | No | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/05/2023 |  |  |  |  | The user facility in japan reported a 57-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the pump had a low flow and high purge pressure occur. the patient had the device for 51 days when the doctor decided that it was time to remove the device and replace it with another one on the opposite site. upon removal of the first impella the md noticed thrombus in the inflow and outflow cages of the pump. there was no serious adverse event to the patient as a result of these incidents. | [5 empty] | Data logs were not returned. pump was sent back for investigation. a significant amount of biomaterial found on impeller, purge gap and inflow cage. the root cause of the low pump flow issue was biomaterial. this report is being filed as part of a retrospective review of historical records.
33867 | 1220648-2023-03914 | 17901560 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17901560&pc=OZD | 09/12/2023 | 05/28/2024 |  | 10/09/2023 |  |  | 09/12/2023 | 10/09/2023 |  |  | 04/30/2025 |  | 10/09/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024352869 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 63-year-old female patient noted as high risk coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported that during the impella insertion a moderate amount of calcium was noted in the vessel and post impella placement the patient went into complete heart block requiring transvenous cardiac pacing (tvp). tvp placed and hr stabilized. after the pci procedure was completed the impella was weaned and removed without complication. | [5 empty] | The investigation into the vf/vt/af/at (complete heart block) issue has been completed since the original report was submitted. the pump was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined. section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
33868 | 1220648-2023-03885 | 17892597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17892597&pc=OZD | 09/07/2023 | 02/07/2024 |  | 10/07/2023 |  |  | 09/07/2023 | 10/07/2023 |  |  | 04/30/2025 |  | 10/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024376363 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/07/2023 |  |  |  |  | The user facility reported a 67-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. after the device was implanted, the perclose failed three times with angioseal closer on the right femoral artery. this created dissection and required a covered stent. this was not at the impella access site, but at the first stick site that was too high. the impella was removed to facilitate the stent and hemostasis. the impella was replaced in left femoral artery (lfa) and case continued with successful stenting of the ramus coronary artery. the impella was removed and attempted closure of the lfa with manata closure also failed twice. manual pressure was applied and access was gained in right renal artery (rra) for angiogram and possible balloon tamponade. hemostasis was achieved with just manual pressure. there was no issue related to the impella device. | [5 empty] | The investigation for the reported vascular dissection and access site bleed has been completed. the device was not returned for investigation. according to the clinical notes, the failure of the closure device three times led to the dissection at the first stick site. additionally, after removal of the impella from the second site where support was ultimately initiated, the site had bleeding that was resolved using manual pressure, a covered stent, and balloon tamponade the cause of the vascular damage and access site bleeding issue was most likely perclose device use-related, based on given clinical information. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
33869 | 1220648-2023-03886 | 17892598 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17892598&pc=OZD | 09/07/2023 | 01/31/2024 |  | 10/07/2023 |  |  | 09/07/2023 | 10/07/2023 |  |  | 02/28/2025 |  | 10/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024322714 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 72-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant reported after the impella implant the patient had very brisk oozing, so pressure held for 45 minutes with no resolution. fem angio showed brisk arterial bleeding at insertion site. figure eight sutures placed and tightened with continued bleeding and hematoma at site. site dressed with angle matching. fem stop applied and repeat contrast injection performed to confirm bleeding had stopped. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported access site bleeding has been completed since the original report was filed. product was not returned for investigation. as per the given clinical, patient had atrial bleeding at insertion site and figure 8 suture, fem stop and contrast injection was applied to stop the bleeding. the cause of the issue was not determined since product was not returned and sufficient clinical information was not provided.
33870 | 1220648-2023-03887 | 17892599 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17892599&pc=OZD | 09/01/2023 | 01/30/2024 |  | 10/07/2023 |  |  | 09/08/2023 | 10/07/2023 |  |  | 02/28/2025 |  | 10/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024332425 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 76-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the complainant had a cp support initiated when after one day of support (28hours) the team observed a purge leak. the leak prompted the team to explant the pump. there was no lasting harm or injury from the pump explant and purge leak. the pump was not replaced. | [5 empty] | The investigation into the purge leak ¿ pump issue has been completed. the product was not returned for review. the root cause of the pump leak was most likely a damaged yellow luer. the root cause of damaged yellow luer was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
33871 | 1220648-2023-03888 | 17892600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17892600&pc=OZD | 09/09/2023 | 05/22/2024 |  | 10/07/2023 |  |  | 09/09/2023 | 10/07/2023 |  |  | 03/31/2025 |  | 10/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024344889 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 17 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the purge system broke between the purge reservoir and filter impeding purge solution flow. it was believed it might have happened when the patient was transported from or to icu. therefore, the patient was taken back to the cath lab for impella exchange. | [5 empty] | The investigation for the reported purge leak has been completed. the product was not returned. however, clinical details was sufficient to determine the root cause. the root cause of the damaged sidearm is most likely due to use as the clinical stated that is occurred most likely due to transport to the icu. the instructions for use referenced in the initial report is from ifu for impella cp with smartassist for use during cardiogenic shock and high risk cpi, section: cleaning. a1-added patient identifier. d4-updated model number. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿.
33872 | 1220648-2023-03889 | 17892601 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17892601&pc=OZD | 09/09/2023 | 05/23/2024 |  | 10/07/2023 |  |  | 09/09/2023 | 10/07/2023 |  |  | 03/31/2025 |  | 10/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024344889 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 17 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the impella purge tubing completely broke off at the junction of purge reservoir and filter. therefore, a new impella cp was inserted through same site. | [5 empty] | The investigation into the purge leak ¿ pump issue has been completed since the original report was submitted. the pump was not returned for investigation. the root cause of the purge leak - pump is most likely due to a damaged sidearm, however it cannot be determined how/where the damaged occurred. impella cp with smartassist for use during cardiogenic shock and high risk cpi section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
33873 | 1220648-2023-03890 | 17892602 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17892602&pc=OZD | 09/09/2023 | 01/31/2024 |  | 10/07/2023 |  |  | 09/09/2023 | 10/07/2023 |  |  | 06/30/2025 |  | 10/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024379195 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/09/2023 |  |  |  |  | The user facility reported a 72-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. an access site complication occurred with a dissection of the femoral artery. staff managed the dissection with a balloon tamponade. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation has been completed since the original report was filed. the pump was not returned for investigation. according to clinical records, a balloon was placed following an access site complication. the cause of the issue was not determined since the product was not returned and sufficient clinical information was not provided.
33874 | 1220648-2023-03891 | 17892610 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17892610&pc=OZD | 09/17/2021 | 10/07/2023 |  | 10/07/2023 |  |  | 09/17/2021 | 10/07/2023 |  |  | 06/30/2023 |  | 10/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022097937 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/17/2021 |  |  |  |  | The user facility reported a 64yo male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the 5.5 kinked while positioning. the pump was replaced without harm or injury and support proceeded on second pump. there was no patient harm reported. | [5 empty] | The investigation for the kinked catheter has been completed. examination of the submitted photograph confirmed a catheter kink which reportedly resulted from attempted pump repositioning into the ventricle. based on the provided information that the patient was being moved at the time, the root cause of the reported positioning issues was most likely use-related. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
33875 | 1220648-2023-03892 | 17892611 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=17892611&pc=OZD | 11/15/2021 | 10/07/2023 |  | 10/07/2023 |  |  | 11/15/2021 | 10/07/2023 |  |  | 07/31/2023 |  | 10/07/2023 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022113030 |  | 0550-0008 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/15/2021 |  |  |  |  | The user facility reported a 69-year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the catheter was kinked and it was replaced with another pump. there was no patient harm reported. | [5 empty] | The investigation for the reported kinked catheter has been completed. the device was not returned for evaluation. the clinical report did not provide sufficient details to determine the root cause. therefore, the root cause of the reported issue could not be determined. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.