28983 | 1220648-2024-08827 | 18986819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18986819&pc=OZD | 08/08/2022 | 03/12/2025 |  | 03/27/2024 |  |  | 08/08/2022 | 03/27/2024 |  |  | 10/31/2022 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2021015168 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 08/08/2022 |  |  |  |  | The user facility reported a 83-year-old male patient needing mechanical circulatory support. it was reported that the patient was disoriented and continually tried to get up out of bed, which caused bleeding from groin site and development of small hematoma, which was resolved. unable to obtain pedal pulses for duration of support. echo was performed early in the morning. physician declined repositioning, decided to wean pump and explant due to patient's mental status. pump was successfully weaned and removed at bedside. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
28984 | 1220648-2024-08828 | 18986854 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18986854&pc=OZD | 08/10/2022 | 03/11/2025 |  | 03/27/2024 |  |  | 08/10/2022 | 03/27/2024 |  |  | 02/29/2024 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023201622 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/10/2022 |  |  |  |  | The complainant reported a 56 year old female patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported that there was access site bleeding at the groin site. due to the patient's body habitus and as there was extensive bleeding despite applying manual pressure, it was decided to explant the pump. the impella was successfully weaned and explanted. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | The investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
28985 | 1220648-2024-08829 | 18986860 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18986860&pc=OZD | 06/24/2022 | 05/20/2025 |  | 03/27/2024 |  |  | 06/24/2022 | 03/27/2024 |  |  | 01/31/2024 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023192797 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 92 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/24/2022 |  |  |  |  | The user facility reported an 92-year-old male was scheduled for high risk percutaneous coronary interventions with impella cp support. the physician had difficulty crossing the arteriovenous. a straight wire was used to successfully cross the valve and exchanged for a j wire. the impella was then inserted and started, and the patient then became severely hypotensive and went into tamponade. there was a left ventricular perforation which was identified by echo. transversus abdominis plane was performed and two units of blood were given. the decision was made to abort the procedure. the impella was weaned and removed per protocol. the patient remained stable then transferred to the intensive care unit. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28986 | 1220648-2024-08830 | 18986902 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18986902&pc=OZD | 08/12/2022 | 05/16/2025 |  | 03/27/2024 |  |  | 08/12/2022 | 03/27/2024 |  |  | 04/30/2024 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023211593 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/12/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while implanting they encountered delivery issues as they were unable to push the pump past the bifurcation of the patient's iliac artery. subsequently, placement of the pump was aborted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28987 | 1220648-2024-08831 | 18987009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18987009&pc=OZD | 07/16/2022 | 03/12/2025 |  | 03/27/2024 |  |  | 07/16/2022 | 03/27/2024 |  |  | 02/29/2024 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023197502 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 07/16/2022 |  |  |  |  | The user facility reported a 58-year-old male patient needing mechanical circulatory support. the complainant reported that the the patient experienced groin site bleeding/hematoma after the second peel away sheath was placed. the access site complication was managed utilizing suturing, protamine, and angle matching techniques. it was reported the surgical closure was successful. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
28988 | 1220648-2024-08832 | 18987090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18987090&pc=OZD | 06/24/2022 | 03/27/2024 |  | 03/27/2024 |  |  | 06/24/2022 | 03/27/2024 |  |  | 03/31/2024 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023211280 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/24/2022 |  |  |  |  | The user facility reported a 74-year-old female was on impella cp support due to ventricular fibrillation episode after admission for nstemi. while on support, a hematoma formed at the access site during preclosing which went away with manual pressure and placing the impella sheath. after the sheath removal, a second perclose did not take and there was a lot of bleeding. the team was unable to get control with a third perclose and still oozing around 12f sheath. vascular surgery came in for cutdown and repair. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records.
28989 | 1220648-2024-08833 | 18987136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18987136&pc=OZD | 03/12/2024 | 04/12/2024 |  | 03/27/2024 |  |  | 03/12/2024 | 03/27/2024 |  |  | 07/31/2025 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024431933 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/12/2024 |  |  |  |  | The user facility reported a 71-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported the patient had impella cp support due to needing continued support and to wean off pressor support. while on support, the impella site started to bleed profusely and the staff applied a fem stop. one unit of packed red blood cells were provided. the patient had a large hematoma and the staff was monitoring. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. per clinical details provided, the root cause of the access site bleeding was most likely due to anticoagulation management since the act at the time of bleed was 304 which above the limit range. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
28990 | 1220648-2024-08834 | 18987295 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18987295&pc=OZD | 03/14/2024 | 04/03/2024 |  | 03/27/2024 |  |  | 03/14/2024 | 03/27/2024 |  |  |  |  | 03/27/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1401075 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Charge | Battery Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/14/2024 |  |  |  |  | The user facility reported a 65-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that an automated impella controller (aic) displayed a battery failure alarm. the plug port was switched and the plug indicator was on, but the battery still failed to charge. the aic was swapped as a result of the failure. there was no patient harm. | [5 empty] | The investigation into the aic - battery charging/other issues has been completed since the initial report was submitted. the console was returned, visually inspected, and tested. battery failure alarm issue was able to be reproduced. the root cause of this issue was internal battery cell imbalance. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed.
28991 | 1220648-2024-08835 | 18987332 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18987332&pc=OZD | 07/08/2022 | 03/12/2025 |  | 03/27/2024 |  |  | 07/08/2022 | 03/27/2024 |  |  | 03/31/2024 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204060 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/08/2022 |  |  |  |  | The user facility reported a 57-year-old male patient needing mechanical circulatory support. it was reported after being placed on impella cp support the patient lost pulses after cp removal. thrombectomy was performed. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28992 | 1220648-2024-08836 | 18987342 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18987342&pc=OZD | 06/24/2022 | 03/27/2024 |  | 03/27/2024 |  |  | 06/24/2022 | 03/27/2024 |  |  | 03/31/2024 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204856 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/24/2022 |  |  |  |  | The user facility reported a 77-year-old male was implanted with an impella device for mechanical circulatory support. during support the p levels were increased and they had suction alarms, so they reduced the p levels. the central venous pressure was also low, so fluids were given. the patient was also having hemolysis. the impella position was checked based on waveforms and the physician was happy with the impella position. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ . ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿. this report is being filed as part of a retrospective review of historical records.
28993 | 1220648-2024-08837 | 18987519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18987519&pc=OZD | 06/28/2022 | 03/07/2025 |  | 03/27/2024 |  |  | 06/28/2022 | 03/27/2024 |  |  | 03/31/2024 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023211191 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | Activation, Positioning or Separation Problem | [5 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/28/2022 |  |  |  |  | The user facility reported a 59-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported during impella cp support, the patient experienced arrhythmia which required defibrillation intervention. the physician ordered the pump removed due to patient anatomy being incompatible with the impella. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28994 | 1220648-2024-08838 | 18987579 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18987579&pc=OZD | 06/24/2022 | 03/27/2024 |  | 03/27/2024 |  |  | 06/24/2022 | 03/27/2024 |  |  | 03/31/2024 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023211490 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 27 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/24/2022 |  |  |  |  | The user facility reported a 27-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient developed bleeding/oozing at the impella cp insertion site. a femstop was placed on the groin due to the oozing. the patient received one unit of blood. the patient was not on any anticoagulation and bicarb was in the purge. the patient was withdrawn from care due to condition. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
28995 | 1220648-2024-08839 | 18987629 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18987629&pc=OZD | 03/14/2024 | 05/02/2024 |  | 03/27/2024 |  |  | 03/14/2024 | 03/27/2024 |  |  | 12/31/2024 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024318730 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/14/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. during support, there was a prolonged purge system blocked alarm and tissue plasminogen activator protocol was unsuccessful. purge flow less than two cubic centimeters per hour for over eight hours. the impella 5.5 was replaced with a new impella 5.5. there was no harm to the patient. | [5 empty] | The investigation into the high purge pressure issue has been completed since the original report was filed. the pump was returned and evaluated. upon visual inspection, no biomaterial was found in purge gap and no kinks were found in purge line. the failure was not reproducible during testing. however, data logs show a build-up in purge pressure over approximately a 2-hour time period and purge flow gradually dropped as well as impella stopped. the root cause of the high purge pressure was most likely due to a gradual buildup on biomaterial in the purge gap. the root cause of the temporary pump stop was most likely due to a use issue as pump was plugged back in shortly after and did not cause any prolonged loss of support. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
28996 | 1220648-2024-08840 | 18987791 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18987791&pc=OZD | 06/28/2022 | 05/19/2025 |  | 03/27/2024 |  |  | 06/28/2022 | 03/27/2024 |  |  | 03/31/2024 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023205487 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | Asian | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/28/2022 |  |  |  |  | The user facility reported a male patient in acute myocardial infarction/cardiogenic shock was placed on impella cp for mechanical circulatory support. it was reported while the patient was being implanted impella cp the guidewire became loose from the left ventricle. the guidewire was removed from the body, cleaned with distilled water and inserted again successfully. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
28997 | 1220648-2024-08841 | 18987805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18987805&pc=PYX | 03/14/2024 | 04/26/2024 |  | 03/27/2024 |  |  | 03/14/2024 | 03/27/2024 |  |  | 08/31/2025 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024404276 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/14/2024 |  |  |  |  | The user facility reported a 72-year-old male undergoing cardiac surgery was implanted with an impella rp device for mechanical circulatory support. it was reported that the purge pressure was rising, and the purge flow was dropping. the cassette was changed and there were no visible kinks. tissue plasminogen activator protocol appeared to have cleared high purge pressure and low purge flow alarms. there was no patient harm. | [5 empty] | The investigation for the reported high purge pressure has been completed. no product was returned for evaluation. data logs were reviewed and lt logs show a gradual increase in purge pressure over a period of approximately 2 hours before high purge pressure alarms were triggered. a decrease in purge flow was also noted over that time. no change to motor current or pump flow. clinical details noted that tpa protocol was initiated and resolved high purge pressure alarms. the root cause of the high purge pressure was most likely due to a gradual buildup of biomaterial in the purge gap. the instructions for use for impella rp with smartassist system states in section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
28998 | 1220648-2024-08842 | 18987823 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18987823&pc=OZD | 06/25/2022 | 09/12/2024 |  | 03/27/2024 |  |  | 06/25/2022 | 03/27/2024 |  |  | 03/31/2024 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | ABIOMED, INC. |  |  | ABIOMED, INC. |  |  | US |  | 2023206832 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  |  |  | Injury | Yes | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | Thrombosis/Thrombus | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [10 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/25/2022 |  |  |  |  | Survival outcome at explant noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | The user facility reported a 50-year-old female in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump experienced low pump flow. the physician removed the impella cp and clots were found in the inlet and outlet of the device. a new impella cp device was placed with no issues. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08842 was provided in sections b2, b5, h1, and h6. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
28999 | 1220648-2024-08843 | 18987971 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18987971&pc=OZD | 06/29/2022 | 05/19/2025 |  | 03/27/2024 |  |  | 06/29/2022 | 03/27/2024 |  |  | 02/29/2024 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023197906 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/29/2022 |  |  |  |  | The user facility reported a 81-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. it was reported the device was weaned and removed without any issues; however, during the closure process complete hemostasis was not achieved. pressure was held at the site for 15 minutes until there was no more oozing and two perclose proglide devices were used to complete the closure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29000 | 1220648-2024-08844 | 18988063 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18988063&pc=OZD | 03/14/2024 | 05/02/2024 |  | 03/27/2024 |  |  | 03/14/2024 | 03/27/2024 |  |  |  |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025457831 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Integrity Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/14/2024 |  |  |  |  | The user facility reported a 75-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that when trying to insert the impella 5.5, the medical doctor tried pushing the pump in and damaged the impella catheter near the motor housing. a new impella 5.5 was opened and used successfully. the patient is stable. | [5 empty] | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned by the user facility for investigation. the root cause of the delivery issue was most likely patient condition as the patient had a tight turn from the axillary artery. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29001 | 1220648-2024-08845 | 18988073 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18988073&pc=OZD | 03/17/2024 | 05/06/2024 |  | 03/27/2024 |  |  | 03/17/2024 | 03/27/2024 |  |  | 06/30/2025 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024427298 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/17/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after the impella implant, the patient had some bleeding. the impella was angled, but a hematoma formed and a fem stop was applied for light pressure. additionally, the patient had a gastrointestinal (gi) bleed on support and blood products were provided. upon explanting the impella, the patient had a worsening hematoma. it was noted that the patient had bleeding from the large side branch of the profunda artery. vascular surgery was performed and a coil was placed to control the bleeding. | [5 empty] | The investigation into the vascular damage - peripheral vessel and the access site bleeding - minor issues has been completed since the original report was submitted. the device was not returned by the user facility for investigation. the cause of the vascular damage - peripheral vessel and the access site bleeding - minor issues was not determined due to insufficient clinical information and unknown relation to impella. d6a, d6b. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
29002 | 1220648-2024-08846 | 18988131 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18988131&pc=OZD | 06/25/2022 | 05/19/2025 |  | 03/27/2024 |  |  | 06/25/2022 | 03/27/2024 |  |  | 03/31/2024 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023205213 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/25/2022 |  |  |  |  | The user facility reported a 72-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. upon explant of the impella cp device, a visible clot was identified on the end of pigtail. the staff allowed adequate bleed back before applying manual pressure. manual pressure was applied for 50 minutes with good hemostasis and the patient remained stable. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Thrombus. analysis summary: the failure mode of ¿thrombus¿ is to be used when a thrombus is noted but cannot be connected to any other adverse events or decrease in pump performance. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions description of the location and characteristics of the thrombus (e.g., size, shape, consistency). relevant laboratory test results, such as coagulation profiles and platelet counts imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). medications or treatments that the patient was receiving at the time of thrombus formation. surgical or procedural details if applicable (e.g, cardiac catheterization) any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined.
29003 | 1220648-2024-08847 | 18988159 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18988159&pc=OZD | 03/18/2024 | 05/04/2024 |  | 03/27/2024 |  |  | 03/18/2024 | 03/27/2024 |  |  |  |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025447956 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/18/2024 |  |  |  |  | The user facility reported a 69-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support the patient had hemolyzed labs. the pump position was confirmed. | [5 empty] | The investigation for the reported hemolysis has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hemolysis could not be determined. potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." "hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis." "patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis." | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.
29004 | 1220648-2024-08848 | 18988259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18988259&pc=OZD | 03/02/2024 | 05/07/2024 |  | 03/27/2024 |  |  | 03/18/2024 | 03/27/2024 |  |  |  |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025442043 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/18/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support as a bridge to transplant. while on support, the impella site was red, swollen and tender. the patient complained about pain at the impella site. the patient also had an infection and fever. a washout was performed and tested and proved to be negative for infection. additionally, the access site was oozing. the staff planned for another washout. the patient was undergoing heart transplant at the end of case details. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
29005 | 1220648-2024-08849 | 18988277 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18988277&pc=OZD | 06/25/2022 | 04/07/2025 |  | 03/27/2024 |  |  | 06/25/2022 | 03/27/2024 |  |  | 01/31/2024 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191122 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/25/2022 |  |  |  |  | The user facility reported a 57-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support it was reported that the partial thromboplastin time goals remained higher due to left ventricle thrombus. the extracorporeal membrane oxygenation (ecmo) and impella cp device were removed together without complications. | [5 empty] | Section d6a / d6b: implant and explant dates added. when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions. ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency). ¿relevant laboratory test results, such as coagulation profiles and platelet counts. ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). ¿medications or treatments that the patient was receiving at the time of thrombus formation. ¿surgical or procedural details if applicable (e.g, cardiac catheterization). ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29006 | 1220648-2024-08859 | 18988591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18988591&pc=OZD | 03/20/2024 | 05/07/2024 |  | 03/27/2024 |  |  | 03/20/2024 | 03/27/2024 |  |  | 07/31/2025 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, IINC. |  |  | US |  | 2025442536 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | Improper Flow or Infusion | Material Twisted/Bent |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 75-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. the admitting diagnosis was congestive heart failure and non-st-segment elevation myocardial infarction. the impella device was placed; however, the physician noted the positioning was "a little deep." the pump was pulled back, and low flow was observed. fluoroscopy was checked, and it was noted the cannula was kinked. the physician attempted to advance the catheter and pull back to reduce the kinked area, which was unsuccessful. therefore, the physician elected to remove the pump and place a new one. there was no report of adverse effect to the patient. | [5 empty] | The investigation into the low pump flow issue has been completed since the original report was filed. the pump was returned and evaluated. visual inspection found a kink on the cannula. data logs confirmed the low flow when pump started. the root cause of low flow was most likely a kinked cannula. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer at the time of this mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29007 | 1220648-2024-08860 | 18988651 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18988651&pc=OZD | 06/29/2022 | 05/16/2025 |  | 03/27/2024 |  |  | 06/29/2022 | 03/27/2024 |  |  | 06/30/2023 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMART ASSIST SET APAC |  |  | ABIOMED, INC. |  |  | US |  | 2022093250 |  | 0048-0047 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/29/2022 |  |  |  |  | The user facility reported a 48-year-old female was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient while on impella cp support had bleeding at the impella access site. the patient received blood transfusions and he impella cp was removed for artery repair by a vascular surgeon. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. corrected b1 to serious injury.
29008 | 1220648-2024-08861 | 18988680 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18988680&pc=OZD | 03/21/2024 | 05/06/2024 |  | 03/27/2024 |  |  | 03/21/2024 | 03/27/2024 |  |  |  |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025443043 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | Patient-Device Incompatibility | Material Twisted/Bent |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/21/2024 |  |  |  |  | The user facility reported a 66-year-old female with cardiomyopathy was to be implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the physician was unable to pass pump through patient¿s native vessel anatomy, which resulted in cable kinking. no patient harm was reported. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely due to patient condition since patient's vessel diameter < 7mm.. added- d6a/d6b. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29009 | 1220648-2024-08862 | 18988720 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18988720&pc=OZD | 06/23/2022 | 05/19/2025 |  | 03/27/2024 |  |  | 06/23/2022 | 03/27/2024 |  |  | 01/31/2024 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183106 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/23/2022 |  |  |  |  | The user facility reported a 76-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. during support, the pt was turned and the impella site started bleeding. femostop was place, but bleeding worsened. decision was made to explant the device. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29010 | 1220648-2024-08863 | 18988791 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18988791&pc=OZD | 03/21/2024 | 05/13/2024 |  | 03/27/2024 |  |  | 03/21/2024 | 03/27/2024 |  |  |  |  | 03/27/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1548488 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | Computer Operating System Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/21/2024 |  |  |  |  | No patient involvement: the us complainant had an aic (automated impella controller) that was removed from service with a note for malfunction attached to it. the team later determined the pump-aic had a pump stop that prompted the removal from service the month prior. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the initial report was submitted. the console was returned, tested, and visually inspected. the failure was unreproducible. the cause of the reset during patient support was due to the software partially resetting the console after one of the processes were unresponsive. the cause for unresponsiveness of the process could not be determined since the event could not be reproduced. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29011 | 1220648-2024-08864 | 18988823 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18988823&pc=OZD | 06/26/2022 | 03/27/2024 |  | 03/27/2024 |  |  | 06/26/2022 | 03/27/2024 |  |  | 12/31/2023 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023180128 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/26/2022 |  |  |  |  | The user facility reported a male in acute anterior myocardial infarction was implanted with an impella cp device for mechanical circulatory support. while on support, femstop on the impella site noted to be pushing the impella down. the impella was not secured with forward tension sutures; therefore, the sutures were redone and pressure dressing was applied. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
29012 | 1220648-2024-08865 | 18988885 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18988885&pc=OZD | 03/16/2024 | 05/17/2024 |  | 03/27/2024 |  |  | 03/16/2024 | 03/27/2024 |  |  | 10/31/2025 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025440168 |  | 0048-0047 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/16/2024 |  |  |  |  | The user facility reported a 61-year-old male of unknown race and ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced access site bleeding and four units of red blood cells were given as treatment and the impella was removed. the patient was reported as stable. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use referenced impella cp with smart assist system section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury".
29013 | 1220648-2024-08866 | 18989142 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18989142&pc=OZD | 06/26/2022 | 03/27/2024 |  | 03/27/2024 |  |  | 06/26/2022 | 03/27/2024 |  |  | 03/31/2024 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023211666 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/29/2022 |  |  |  |  | The user facility reported a 63-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella support, echocardiography was performed to verify impella cp placement. it was observed that the pigtail was around a papillary muscle. the surgeon attempted to reposition the impella without success. the impella was removed and a new sheath kit was used to regain access to the left ventricle, and the impella was successfully placed. it was also reported during impella cp support, the patient become septic and the impella weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿infusion filter the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ this report is being filed as part of a retrospective review of historical records.
29014 | 1220648-2024-08867 | 18989170 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18989170&pc=OZD | 08/20/2022 | 03/27/2024 |  | 03/27/2024 |  |  | 08/20/2022 | 03/27/2024 |  |  | 05/31/2024 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2023230687 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/20/2022 |  |  |  |  | The user facility reported a 65-year-old male patient was implanted with an impella cp for mechanical circulatory support. the patient experienced a moderate hematoma. the doctor elected to hold off on heparin. a thrombectomy was performed followed by x1 drug with good results. the patient had a successful impella cp implant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29015 | 1220648-2024-08868 | 18989291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18989291&pc=OZD | 03/17/2024 | 05/04/2024 |  | 03/27/2024 |  |  | 03/17/2024 | 03/27/2024 |  |  | 11/30/2025 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025438224 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/17/2024 |  |  |  |  | The user facility reported a 83-year-old female with a high risk of percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the doctor was unable to obtain pedal pulses. the impella cp was replaced with an intra-aortic balloon pump. | [5 empty] | The investigation into the limb ischemia ¿ reversible has been completed since the original report was submitted. the device was not returned by the user facility for investigation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation." | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
29016 | 1220648-2024-08869 | 18989444 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18989444&pc=OZD | 03/17/2024 | 05/13/2024 |  | 03/27/2024 |  |  | 03/17/2024 | 03/27/2024 |  |  | 11/30/2025 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024431397 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/17/2024 |  |  |  |  | The user facility reported an 81-year-old male with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that during insertion, the left ventricle was perforated. the team performed a sternotomy and repaired the perforation and transfused blood. | [5 empty] | The investigation into the ventricle perforation issue has been completed since the original report was submitted. the device was not returned by the user facility for investigation. data logs were reviewed and several improper positioning in aorta, ventricle alarms were observed. surgical mode was enabled post these events. the cause of the ventricle perforation issue was most likely due to guidewire puncture. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿
29017 | 1220648-2024-08870 | 18989516 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18989516&pc=OZD | 06/30/2022 | 05/19/2025 |  | 03/27/2024 |  |  | 06/30/2022 | 03/27/2024 |  |  | 03/31/2024 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023205220 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/30/2022 |  |  |  |  | The user facility reported a 77-year-old male patient experiencing heart transplant rejection and subsequent cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported during impella cp support that the patient experienced significant bleeding at the impella access site. as a result of the bleeding heparin was discontinued. it was also reported after being placed on impella cp support the patient experienced episodes of ventricular tachycardia due to impella being placed too deeply into the left ventricle. the impella was repositioned and the patient was placed on veno-arterial extracorporeal membrane oxygenation (va-ecmo), but due to poor right heart function the patient experienced continuous ventricular tachycardia rhythm. the impella remained implanted for continued mechanical support to unload the left ventricle. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29018 | 1220648-2024-08871 | 18989568 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18989568&pc=OZD | 06/30/2022 | 05/19/2025 |  | 03/27/2024 |  |  | 06/30/2022 | 03/27/2024 |  |  | 02/29/2024 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023194879 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/30/2022 |  |  |  |  | The user facility reported a 67-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed access site bleeding. as a result, the impella cp was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29019 | 1220648-2024-08872 | 18989609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18989609&pc=OZD | 06/26/2022 | 05/16/2025 |  | 03/27/2024 |  |  | 06/26/2022 | 03/27/2024 |  |  | 03/31/2024 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023210435 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/26/2022 |  |  |  |  | The user facility reported a 61-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient presented no pedal pulses bilaterally. the patient was taken to or for escalation to extracorporeal membrane oxygenation (ecmo), while in or the patient decompensated and expired on support. although the patient expired while on impella support, no allegations were made towards the impella for cause of death. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Limb ischemia. analysis summary: limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition. any concomitant medication. vascular size or variations thereof. detailed description of the symptoms experienced by the patient. physical examination findings, including pulse assessment and visual examination of the affected limb. diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. laboratory test results, including blood tests for markers of inflammation or clotting disorders. any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
29020 | 1220648-2024-08873 | 18989801 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18989801&pc=OZD | 09/08/2022 | 06/02/2025 |  | 03/27/2024 |  |  | 09/08/2022 | 03/27/2024 |  |  | 06/30/2023 |  | 03/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022099317 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/08/2022 |  |  |  |  | The user facility reported an 84-year-old patient in non-st-elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that during impella support the patient experienced oozing at the impella access site. the patient was non-compliant with maintaining a straight leg to help resolve bleeding. angle matching and redressing of the site improved the bleed, but the patient insisted on removal of the device. impella cp was weaned and removed, and the access site was closed with a perclose device and manual pressure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29021 | 1220648-2024-08714 | 18976494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18976494&pc=OZD | 06/28/2022 | 05/16/2025 |  | 03/26/2024 |  |  | 06/28/2022 | 03/26/2024 |  |  | 03/31/2024 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023209737 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/28/2022 |  |  |  |  | The user facility reported an impella cp in 79-year-old female patient for mechanical circulatory support. it was reported that the impella cp pump was unable to pass through the aortic valve with pigtail after a few attempts the decision was made to switch to al-2 to gain ventricular access. impella was inserted with slight difficulty as patient had very tortuous aorta. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29022 | 1220648-2024-08715 | 18976501 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18976501&pc=OZD | 06/29/2022 | 05/16/2025 |  | 03/26/2024 |  |  | 06/29/2022 | 03/26/2024 |  |  | 03/31/2024 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023211198 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/29/2022 |  |  |  |  | The user facility reported a 77-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding that was resolved with cessation of anticoagulants. additionally, the patient experienced atrial fibrillation, later to be reported as stable with no intervention specified. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29023 | 1220648-2024-08716 | 18976504 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18976504&pc=OZD | 06/29/2022 | 05/16/2025 |  | 03/26/2024 |  |  | 06/29/2022 | 03/26/2024 |  |  | 03/31/2024 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023213038 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/29/2022 |  |  |  |  | The user facility reported a 59-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the device was flipped towards the mitral valve and was stuck. repositioning was attempted by aggressive maneuvers to set the device free. unfortunately, there was no success freeing the pump with maneuver tactics. subsequently, the impella was inadvertently pulled out of the left ventricle.. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29024 | 1220648-2024-08717 | 18976510 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18976510&pc=OZD | 07/01/2022 | 05/16/2025 |  | 03/26/2024 |  |  | 07/01/2022 | 03/26/2024 |  |  | 01/31/2024 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023189368 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/01/2022 |  |  |  |  | The user facility reported a 47-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient lost pulses in her right leg and heparin-induced thrombocytopenia was suspected which was treated by switching to dextrose 5 percent in water. one unit of blood was also administered to the patient. eventually the device was weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Thrombocytopenia: the cause of the issue was not determined due to insufficient clinical details. limb ischemia: when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition ¿ any concomitant medication ¿ vascular size or variations thereof ¿ detailed description of the symptoms experienced by the patient ¿ physical examination findings, including pulse assessment and visual examination of the affected limb ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension ¿ patient's response to previous treatments or interventions for vascular conditions with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
29025 | 1220648-2024-08718 | 18976515 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18976515&pc=OZD | 07/01/2022 | 05/16/2025 |  | 03/26/2024 |  |  | 07/01/2022 | 03/26/2024 |  |  | 03/31/2024 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023210544 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/01/2022 |  |  |  |  | The user facility reported a 81-year-old male in acute myocardial infarction/cardiogenic shock. while the device was being implanted, there was significant resistance in the left femoral artery (lfa), due to extreme tortuosity of the iliac artery. the pump was withdrawn and multiple reattempts were performed with no success. the right femoral artery was evaluated for the impella post-percutaneous coronary intervention (pci); however, the guide catheters were difficult to advance as well. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29026 | 1220648-2024-08719 | 18976519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18976519&pc=OZD | 07/05/2022 | 05/16/2025 |  | 03/26/2024 |  |  | 07/05/2022 | 03/26/2024 |  |  | 03/31/2024 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023206501 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/05/2022 |  |  |  |  | The user facility reported a 50-year-old male on impella cp for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient had ventricular fibrillation arrest. therefore, the impella cp was removed. the patient was upgraded to the impella 5.5. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. timing of the arrythmia event in relation to impella support (during or post-implant). electrocardiogram (ekg) recordings of the arrhythmia episodes. evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. response to any antiarrhythmic treatments or interventions during the event. any documented device-related issues or complications during impella support. evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined.
29027 | 1220648-2024-08720 | 18976525 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18976525&pc=OZD | 07/06/2022 | 05/16/2025 |  | 03/26/2024 |  |  | 07/06/2022 | 03/26/2024 |  |  | 07/31/2023 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022108937 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/06/2022 |  |  |  |  | The user facility reported a patient on impella cp for mechanical circulatory support. it was reported that the patient experienced access site bleeding. the patient required a blood transfusion to resolve the bleeding. the impella cp was weaned off and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29028 | 1220648-2024-08721 | 18976527 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18976527&pc=OZD | 07/05/2022 | 05/16/2025 |  | 03/26/2024 |  |  | 07/05/2022 | 03/26/2024 |  |  | 01/31/2024 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183799 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/05/2022 |  |  |  |  | The user facility reported a 70-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump came out of position while repositioning patient. surgeon was unable to advance under echo. p2 position alarm was disabled. impella was unable to be repositioned back across the valve. decision was made to explant impella and place new impella cp. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29029 | 1220648-2024-08722 | 18977133 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18977133&pc=OZD | 03/12/2024 | 03/26/2024 |  | 03/26/2024 |  |  | 03/12/2024 | 03/26/2024 |  |  |  |  | 03/26/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1251624 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 03/12/2024 |  |  |  |  | The complainant reported that an automated impella controller (aic) experienced a controller failure alarm. the device was removed from service as a result of the failure. there was no patient involvement. | [5 empty] | The investigation has been completed. console logs from the reported day of event were consistent with the complaint and confirmed the reported failure mode given. there was a controller error alarms triggered during the clinical procedure. real time (rt) logs confirmed that all signals received from optical bench were for approximately four minutes, interpreted as -16384 values. the console was investigated. the failure was unable to be reproduced throughout testing and was characterized by a temporary loss of optical data. the cause of the issue was not determined since the issue could not be reproduced
29030 | 1220648-2024-08723 | 18977235 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18977235&pc=OZD | 07/29/2022 | 03/26/2024 |  | 03/26/2024 |  |  | 07/29/2022 | 03/26/2024 |  |  | 04/30/2024 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023223675 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/29/2022 |  |  |  |  | The user facility reported a 61-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that during impella support, high purge pressure (pp)/low pf alarms were present. doctor repositioned the pump and purge pressure increased. no patient harm reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29031 | 1220648-2024-08725 | 18977428 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18977428&pc=OZD | 08/08/2022 | 05/16/2025 |  | 03/26/2024 |  |  | 08/08/2022 | 03/26/2024 |  |  | 04/30/2024 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2023226088 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/08/2022 |  |  |  |  | The user facility reported a 76-year-old female patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported during impella support the flows from the pump were lower than expected. the impella was attempted to be repositioned, but flows did not improve. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29032 | 1220648-2024-08726 | 18977772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18977772&pc=OZD | 06/27/2023 | 03/26/2024 |  | 03/26/2024 |  |  | 06/27/2023 | 03/26/2024 |  |  | 07/31/2024 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMART ASSIST SET APAC |  |  | ABIOMED, INC. |  |  | US |  | 2024285406 |  | 0048-0047 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 34 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Material Twisted/Bent | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/27/2023 |  |  |  |  | The user facility reported a 34-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. complainant from taiwan reported the impella pump kinked during implantation. repositioning was not possible. pump was removed and a new one inserted, without problems. | [5 empty] | The impella device was received from the customer on 21-aug-2023. product was returned and investigated. a kink in the cannula was observed. per the clinical information provided, the root cause of the positioning issue was use related as when the repo sheath was advanced the pump came out of position causing the cannula to kink as a result. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿
29033 | 1220648-2024-08727 | 18977817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18977817&pc=OZD | 03/06/2024 | 04/22/2024 |  | 03/26/2024 |  |  | 03/06/2024 | 03/26/2024 |  |  |  |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025447447 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/05/2024 |  |  |  |  | The user facility reported a 56-year-old male with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed access site bleeding, bivalrudin was paused and femstop was placed for 3 hours. the site was re-dressed in acute angle matching technique and one unit of packed red blood cells and albumin x1 was administered. the patient was reported as stable. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding issue was most likely use related as the impella site was redressed with acute angle matching done as per the clinical description. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
29034 | 1220648-2024-08728 | 18977822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18977822&pc=OZD | 06/24/2022 | 05/21/2025 |  | 03/26/2024 |  |  | 06/24/2022 | 03/26/2024 |  |  | 02/29/2024 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023201435 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/24/2022 |  |  |  |  | The user facility reported experiencing difficulty while implanting impella cp in 68-year-old patient for mechanical circulatory support due to aortic dissection. ultimately the medical team aborted the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29035 | 1220648-2024-08729 | 18977871 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18977871&pc=OZD | 03/06/2024 | 04/17/2024 |  | 03/26/2024 |  |  | 03/06/2024 | 03/26/2024 |  |  |  |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024435519 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had impella 5.5 pump placed to support 68-year-old male patient admitted in cardiogenic shock (cgs). the device was able to be delivered but experienced immediate alarms for motor current and saturated flows. the pump was therefore explanted. there was no reported patient harm. | [5 empty] | The investigation into the saturated flow issue has been completed since the original report was submitted. the device was not returned for investigation. clinical details noted that lv thrombus was present and when pump was removed, clot was observed in motor housing. the root cause of the saturated flow was most likely due to the ingested biomaterial in the motor housing. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29036 | 1220648-2024-08730 | 18977913 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18977913&pc=OZD | 03/07/2024 | 05/08/2024 |  | 03/26/2024 |  |  | 03/07/2024 | 03/26/2024 |  |  |  |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 EU |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2025454409 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported tat during insertion of impella 5.5 for 70-year-old male patient, the device was unable to pass through the anastomosis/graft and the catheter kinked behind the motor housing. therefore, another device was successfully used. there was no reported patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the pump was returned for investigation and kink in the catheter behind the motor housing was noticed. the root cause of the delivery is determined to be use issue because the graft was small and they were able to pass the pump through the graft and resulted in kinking of the catheter behind the motor housing. hence, the root cause is determined to be use issue. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
29037 | 1220648-2024-08731 | 18977967 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18977967&pc=OZD | 03/07/2024 | 03/26/2024 |  | 03/26/2024 |  |  | 03/07/2024 | 03/26/2024 |  |  |  |  | 03/26/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1122786 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Material Too Rigid or Stiff | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported there was an automated impella controller's (aic) front door would not close. | [5 empty] | (h3) the investigation into the reported 'aic - damage" has been completed since the original report was filed. product was returned for investigation. the inspection of the console confirmed inability to keep the purge door closed, as the latch for the door was found to be ¿sticking¿, due to insufficient lubrication (lubricant most likely dried up or migrated since console¿s manufacturing in 2012). during inspection it¿s been also discovered that the optical pump connector was contaminated, and could not be fully cleaned, this is unrelated to the reported complaint. the console logs do not contain any entries on the reported date of 2024-03-07. on the last day the console was used ((b)(6) 2024), although there are no direct indications of purge door being open or closed, during case start same purge cassette was repeatedly removed and reinserted, and pump was ultimately not connected to the console which, possibly indicates that the reported issue was observed.
29038 | 1220648-2024-08732 | 18977985 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18977985&pc=OZD | 06/24/2022 | 03/26/2024 |  | 03/26/2024 |  |  | 06/24/2022 | 03/26/2024 |  |  | 01/31/2024 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023182770 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/24/2022 |  |  |  |  | The user facility reported a 55-year-old male was implanted with an impella device for mechanical circulatory support and that the patient was experiencing a hematoma at the access site and lost pulses in the lower extremities which became mottled. the impella cp device was removed. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29039 | 1220648-2024-08733 | 18978006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18978006&pc=OZD | 03/07/2024 | 04/24/2024 |  | 03/26/2024 |  |  | 03/07/2024 | 03/26/2024 |  |  | 11/30/2025 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2024434630 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Difficult to Insert | Patient-Device Incompatibility |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp pump met resistance during attempted delivery. it was noted that the pump was unable to advance into the patient's ventricle as a result of their aortic anatomy. the implant was subsequently aborted. there was no patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery is determined to be patient condition because they were unable to be inserted into ventricle due to aortic anatomy. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
29040 | 1220648-2024-08734 | 18978033 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18978033&pc=OZD | 03/07/2024 | 04/24/2024 |  | 03/26/2024 |  |  | 03/07/2024 | 03/26/2024 |  |  | 09/30/2025 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024407128 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the 63-year-old male patient was on impella cp support due to having acute hemodynamic decompensation. while on support, the patient had limb ischemia. the impella cp was removed and replaced with the impella 5.5. | [5 empty] | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿.
29041 | 1220648-2024-08735 | 18978083 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18978083&pc=OZD | 02/26/2024 | 09/26/2025 |  | 03/26/2024 |  |  | 03/08/2024 | 03/26/2024 |  |  |  |  | 03/26/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1138486 |  | 0042-0000-EU |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that shortly after impella implantation for patient of unknown age and gender, the ps and left ventricle (lv) disappeared with message saying controller was not reliable. pump was noted to be running on auto mode at 3.2l/min. however, patient became unstable during main stem intervention, and decision was made to exchange the pump and the automated impella controller (aic). there was no reported patient harm. | The complainant reported that shortly after impella implantation for patient of unknown age and gender, the placement signal (ps) disappeared with message saying controller was not reliable. there was no reported patient harm. |  |  |  |  | The investigation into the aic - display issue has been completed since the initial report was submitted. the console was returned and tested. the console operated as designed, and optical bench parameters were within specification. further investigation confirmed that placement signal loss was pump related. the issue was unrelated to this device. the device functioned/operated to specification. | The automated impella controller (aic) was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. console logs from the reported day of event are consistent with complaint and show that shortly after pump start, console started alarming ¿placement signal not reliable¿. rtlog data shows that snr abruptly dropped to near 0, and raw placement signal was clipped at -32768, while contrast was reduced and gain values became erratic. this data pattern is considered consistent with pump sensor or fiber damage. testing of the console at aachen fs did not reveal any anomalies - the console operated as designed, and optical bench 5s parameters were within specification. separate investigation 24-23454-1 into the pump 482792 confirmed that placement signal loss was pump-related. repair: perform functional check of the console. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issue "2917" as part of a retrospective review of optical signal issues in accordance with fda recommendation.
29042 | 1220648-2024-08736 | 18978177 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18978177&pc=OZD | 07/18/2022 | 05/16/2025 |  | 03/26/2024 |  |  | 07/18/2022 | 03/26/2024 |  |  | 12/31/2023 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023165867 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/15/2022 |  |  |  |  | The user facility reported a 56-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient developed heparin induced thrombocytopenia. heparin was discontinued in the purge system. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29043 | 1220648-2024-08737 | 18978179 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18978179&pc=OZD | 03/08/2024 | 05/31/2024 |  | 03/26/2024 |  |  | 03/08/2024 | 03/26/2024 |  |  |  |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP EU |  |  | OZD |  |  | IMPELLA CP EU |  |  | ABIOMED, INC. |  |  | US |  | 2025449167 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 37 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | There is not enough information to exclude outcome from use of device. | The complainant reported that during impella cp support for 37-year-old male patient, the automated impella controller (aic) displayed "motor current high" and "controller failure". the aic was switched to the backup. however, the same issue and messages "pump stop" and "motor current high" were displayed. the second aic was shut down. the pump was noted to have been running for 15 days. the patient's ptt at the time of the pump failures was 49. there was no reported patient harm and is on ecmo support. |  |  |  |  | The investigation into the pump stop issue has been completed since the original report was filed. the device was returned for investigation. upon visual inspection, a large purge gap was noted with biomaterial within the purge gap. no data logs were returned for evaluation. the root cause of the pump stop is due to bearing wear beyond indicated design life of 8 days per design. b5, d6a, d6b. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
29044 | 1220648-2024-08738 | 18978185 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18978185&pc=OZD | 06/26/2022 | 05/16/2025 |  | 03/26/2024 |  |  | 06/26/2022 | 03/26/2024 |  |  | 12/31/2023 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023181069 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Injury | Yes | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/26/2022 |  |  |  |  | The user facility reported experiencing difficulty while implanting impella cp in a patient of undetermined gender or age, for mechanical circulatory support due to patient going into ventricular tachycardia and ventricular fibrillation every time placement was attempted. ultimately the medical team aborted the procedure. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29045 | 1220648-2024-08739 | 18978217 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18978217&pc=OZD | 03/09/2024 | 04/29/2024 |  | 03/26/2024 |  |  | 03/09/2024 | 03/26/2024 |  |  | 11/30/2025 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024432675 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the 61-years-old male patient was on impella cp support due to having 90% distal left main artery. while on support, the patient's left leg with the impella cp became cold and had no doppler pulses. the impella was removed. | [5 empty] | The investigation into the reported limb ischemia ¿ reversible issues has been completed since the original report was filed. the pump was returned and evaluated. however, all the necessary clinical information was not provided. the root cause of the limb ischemia was not determined. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
29046 | 1220648-2024-08740 | 18978259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18978259&pc=OZD | 03/09/2024 | 03/26/2024 |  | 03/26/2024 |  |  | 03/09/2024 | 03/26/2024 |  |  |  |  | 03/26/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1226867 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during impella cp support for patient of unknown age and gender, the automated impella controller (aic) displayed 'impella defective" alarm. the aic was disconnected and connected without success. therefore, another device was used. there was no reported patient harm. | [5 empty] | (h3) the investigation into the reported "aic - controller error (yellow alarm)" has been completed since the original report was filed. product was returned for investigation. the console was tested at aachen fs. the failure mode was not reproduced while running an optical pump simulator. there were no issues with the console hardware. the console logs from the reported day of event are mostly consistent with complaint and show that during case start pump 494459 was disconnected upon prompt, due to unreliable placement signal (most likely due to air gap caused by pump connector not being fully inserted) indicated by ¿placement signal unreliable¿ alarm (#1036, due to opm signal invalid). after reconnecting, the issue was briefly resolved, but soon after, console presented with controller error alarm (#1043, opm comm stopped). the logs also showed processsamplingdatafromopticalbench: sentstopacquisitioncmd ack, indicating ossiopsens receive thread gets a data sampling packet with an unexpected length. real time (rt) logs confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) were lastly reported as +16384 values and then no more samples were recorded. the absence of snr samples impacted the console¿s ability to recognize the pump disconnection and required the user to perform an emergency shutdown. the log entry processsamplingdatafromopticalbench: sentstopacquisitioncmd ack indicates the sentstopacquisitioncmdflag was set when entering ossi_sampling_stopped_state which eventually led to a false ¿communication stopped¿ condition. the root cause of the controller error was due to an automated impella controller (aic) software issue.
29047 | 1220648-2024-08742 | 18978351 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18978351&pc=OZD | 06/27/2022 | 05/16/2025 |  | 03/26/2024 |  |  | 06/27/2022 | 03/26/2024 |  |  | 03/31/2024 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204094 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/27/2022 |  |  |  |  | The user facility reported an impella cp in 55-year-old male patient for mechanical circulatory support. during the insertion, the patient had multiple episodes of ventricular tachycardia, which required defibrillation use. while on impella cp support, the impella had ¿placement signal low" alarms which was resolved by adjusting the levels and the pump was repositioned under echocardiogram. additionally, the patient developed some oozing at the impella site. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29048 | 1220648-2024-08741 | 18978361 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18978361&pc=OZD | 03/09/2024 | 05/20/2024 |  | 03/26/2024 |  |  | 03/10/2024 | 03/26/2024 |  |  |  |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 EU |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2025454459 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella 5.5 support for patient of unknown age and gender, there was pump stop. the team was able to restart the pump. however, the pump was exchanged for a new device. the patient's act was between 160/180. there was no reported patient harm. | [5 empty] | The investigation into the temporary pump stop issue has been completed since the original report was submitted. the device was not returned for investigation. no data logs were returned for evaluation. clinical details noted that high motor current occurred at time of pump stop. the root cause of the temporary pump stop could not be determined as no product or data logs were returned. instructions for use for the related event are as follows: impella 5.5 with smart assist for use during cardiogenic shock entering cardiac output (impella cp with smart assist) section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ d.6a and d.6b revised from the initial submission in accordance with updated procedures. g.1 revised manufacturing site name and address section as previously entered incorrectly. h.6 code 2199 was reported incorrectly on the previously submitted report. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
29049 | 1220648-2024-08743 | 18978410 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18978410&pc=OZD | 03/10/2024 | 04/30/2024 |  | 03/26/2024 |  |  | 03/10/2024 | 03/26/2024 |  |  | 11/30/2025 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024425144 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella 5.5 pump to support a 56-year-old male patient who arrested and was taken to the er and found to have cad. the team placed an iabp and awaited cabg. when in the or the cabg was initiated with 5.5 pump support. the impella 5.5 pump was positioned deep and a perforation of the left ventricle (lv) occurred. the team repaired the perforation and was sent to the icu on support. | [5 empty] | The investigation into the ventricle perforation issue has been completed since the original report was submitted. the device was not returned for investigation. per the clinical information mentioned, the pump was seated deep while torquing the pump for optimal position. imaging confirmed deep positioning. this potentially led to the ventricle perforation. the root cause of the ventricle perforation issue was most likely improper positioning of the device. potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. to conform with updated procedures, sections d6 and h10 were revised with updated information. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
29050 | 1220648-2024-08744 | 18978539 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18978539&pc=OZD | 03/11/2024 | 04/26/2024 |  | 03/26/2024 |  |  | 03/11/2024 | 03/26/2024 |  |  | 11/30/2025 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024426679 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/11/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was to be implanted with an impella 5.5 for mechanical circulatory support during an endovascular aneurysm repair surgery. due to anatomic limitations the impella 5.5 failed to deliver. the placement was aborted. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely due to patient condition since multiple attempts to advance past subclavian artery were made but were unsuccessful. potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29051 | 1220648-2024-08745 | 18978575 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18978575&pc=OZD | 06/28/2022 | 03/26/2024 |  | 03/26/2024 |  |  | 06/28/2022 | 03/26/2024 |  |  | 02/29/2024 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023193219 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 22 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/28/2022 |  |  |  |  | The user facility reported a 22-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and that the patient experienced tissue oozing from all three sites and numerous blood products were given throughout the procedure. eventually, the device was removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records.
29052 | 1220648-2024-08746 | 18978614 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18978614&pc=OZD | 03/07/2024 | 03/26/2024 |  | 03/26/2024 |  |  | 03/12/2024 | 03/26/2024 |  |  |  |  | 03/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1107573 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during prep, the automated impella controller (aic) failed auto test and displayed "system self-check error". there was no reported patient harm. | [5 empty] | (h3) the investigation into the reported "aic - system self-check error" has been completed since the original report was filed. product was returned for investigation. the console was tested after arrival and the failure mode was reproduced so when the console was turned on, it immediately output a system self check fail. the super cap on the pbm was confirmed to have corroded the pbm communication trace next to it. during data analysis. the automated impella controller (aic) logs for (b)(6) show pbm1 communication errors. the cause of the aic issue was contamination of a communication line on the pbm.
29053 | 1220648-2024-08747 | 18978665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18978665&pc=OZD | 03/08/2024 | 03/26/2024 |  | 03/26/2024 |  |  | 03/12/2024 | 03/26/2024 |  |  |  |  | 03/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | ABIOMED, INC. |  |  | US |  | 1412639 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported during prep, the automated impella controller (aic) did not detect purge cassette/disc. the cassette was exchanged, but the issue did not resolve. therefore, the aic was exchanged and the issue resolved. there was no reported patient harm. | [5 empty] | (h3) the investigation into the reported low pump flow has been completed since the original report was filed. product was returned for investigation. the console was inspected and missing/broken purge disk detect flag was observed. the console logs from the reported day of event are consistent with complaint and show intermittent purge disk not detected alarm (#402) during case start, prior to pump connection. the cause of the issue was a defective component.
29054 | 1220648-2024-08748 | 18978754 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18978754&pc=OZD | 03/13/2024 | 04/23/2024 |  | 03/26/2024 |  |  | 03/13/2024 | 03/26/2024 |  |  | 10/31/2025 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024414257 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Insufficient Flow or Under Infusion | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | ECMO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/13/2024 |  |  |  |  | The user facility reported a 59-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support, there were suction alarms. additionally, a distal perfusion catheter was placed on the impella leg to help with diminished flows. the patient continued to decline and the impella was removed. there were still no flows on the impella leg post removal. the staff did a surgical repair to establish flows. | [5 empty] | The investigation into the low pump flow and the limb ischemia irreversible issue has been completed since the original report was submitted. the device was not returned for investigation. based on clinical description and data review, the root cause of low flow was most likely due to patient condition. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.
29055 | 1220648-2024-08749 | 18978776 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18978776&pc=OZD | 03/14/2024 | 05/02/2024 |  | 03/26/2024 |  |  | 03/14/2024 | 03/26/2024 |  |  | 11/30/2025 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024435570 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/14/2024 |  |  |  |  | The user facility reported a 64-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had placed a cp to support a patient admitted with a stemi (elevation myocardial infarction) and mvd (mitral valve disease) cad (coronary artery disease). the cp had low flows which prompted pump troubleshooting/flushing but when the flow remained suboptimal, the pump was replaced. the team replaced it after just 28 minutes with a new successful pump. | [5 empty] | The investigation for the reported low pump flow has been completed. data logs confirmed the low flow when pump started and remained to the rest of the case that triggered auto suction response. product was not returned for review. the root cause of low flow was unable to be determined as limited clinical details were provided and no product was returned for review. added- h6 impact code 4629. d4-updated model number. added- d6a/d6b. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
29056 | 1220648-2024-08750 | 18978798 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18978798&pc=OZD | 06/28/2022 | 03/26/2024 |  | 03/26/2024 |  |  | 06/28/2022 | 03/26/2024 |  |  | 01/31/2024 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023184904 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/28/2022 |  |  |  |  | The user facility reported a 70-year-old female was implanted with an impella device for mechanical circulatory support. during support, hematuria and frank bleeding from the impella site was noted requiring multiple transfusions. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
29057 | 1220648-2024-08751 | 18978828 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18978828&pc=OZD | 03/14/2024 | 05/24/2024 |  | 03/26/2024 |  |  | 03/14/2024 | 03/26/2024 |  |  |  |  | 03/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1210877 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/14/2024 |  |  |  |  | The patient's expiration while in the procedure area has been reviewed by abiomed's medical safety officer. it was determined that there was not enough information to associate outcome with the use of the device. | The user facility reported a 61-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the aic was powered on and impella catheter was hooked up to the aic before advancing to start the case and an "impella defective¿ yellow alarm appeared. the staff found another aic and plugged the catheter in, and the case continued. |  |  |  |  | The investigation for the reported console (aic) yellow alarm has been completed. the console was returned for evaluation. functional tested found that the console was unable to reproduce the reported failure mode. the console was tested with three different uninitialized impella cp test pumps. upon connecting each pump without the case start wizard, the pump was detected, and the aic navigated through the case start wizard automatically. there was no pump detection issue or impella defective alarm. data logs were reviewed which showed that the console received the ¿impella defective (error reading eeprom) ¿ alarm, event #3.the pump was unplugged and re-plugged into the aic 3 more times, but event #3 was persistent every time when the pump was re-plugged, which confirms the reported failure mode. the root cause of the impella defective alarm was not determined due to the inability to reproduce. b5-updated with medical safety officer review d2-updated common device name. d4-updated model number. | The impella device was received from the customer on (b)(6)2024 and therefore,an evaluation of the device is underway. upon investigation closure, a supplemental mdr will be filed.
29058 | 1220648-2024-08752 | 18978900 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18978900&pc=OZD | 03/14/2024 | 04/30/2024 |  | 03/26/2024 |  |  | 03/14/2024 | 03/26/2024 |  |  | 10/31/2025 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024418179 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | Material Twisted/Bent | Optical Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/14/2024 |  |  |  |  | The user facility reported a 77-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that upon removal of the peel away sheath, the cannula folded in the left ventricle. despite the cannula unfolding slowly, the displayed waveforms became inaccurate and the motor current began to increase. as a result of the noted issue, the decision was made to remove the device. there was no patient harm. | [5 empty] | The investigation into the optical signal issue and the low pump flow issues has been completed since the original report was filed. the pump was returned and evaluated. upon visual inspection, a biomaterial was found to be wrapped around the impeller blade. a slight catheter kink was found on the cannula. pump flow was found to be saturated/clipping for the rest of the case. the root cause of the positioning issues was a use issue due to pump folding into the left ventricle when peel-away sheath was removed. the root cause of the saturated flow was due to biomaterial present on the impeller blade. the root cause of the low pump flow was most likely due to improper positioning of the pump when it was pushed into the ventricle and kinked. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29059 | 1220648-2024-08753 | 18978965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18978965&pc=OZD | 03/14/2024 | 05/01/2024 |  | 03/26/2024 |  |  | 03/14/2024 | 03/26/2024 |  |  | 09/30/2025 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024400710 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/14/2024 |  |  |  |  | The user facility reported an 82-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support, a doppler was performed and was unable to locate pulse on the patient's leg. the impella was repositioned under echocardiogram and a fem-fem bypass was performed to perfuse the leg. following the intervention pulse and color returned and patient was reported as stable. | [5 empty] | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
29060 | 1220648-2024-08754 | 18979032 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18979032&pc=OZD | 03/15/2024 | 05/02/2024 |  | 03/26/2024 |  |  | 03/15/2024 | 03/26/2024 |  |  |  |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025449395 |  | 100080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 19 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/15/2024 |  |  |  |  | The user facility reported a 19-year-old female with cardiomyopathy underwent an attempted impella cp device implant for mechanical circulatory support. the implant procedure was aborted due to the patient¿s cardiac anatomy. the patient was transferred to the cardiothoracic intensive care unit and has been reported to be stable. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition. potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ d4 updated model number. to conform to updated procedures, sections d6 & h10 were revised with updated information. | The impella device has not been returned by the customer, therefore, a device evaluation has not been possible. upon device return and/or investigation closure, a supplemental mdr will be filed.
29061 | 1220648-2024-08755 | 18979092 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18979092&pc=OZD | 03/16/2024 | 05/02/2024 |  | 03/26/2024 |  |  | 03/16/2024 | 03/26/2024 |  |  | 08/31/2025 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024398755 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/16/2024 |  |  |  |  | The user facility reported a 74-year-old male undergoing percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that after insertion of the impella device the patient developed ischemia of the right lower extremity as evidenced by mottling and loss of doppler pulses. as a result, the impella device was explanted. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ to conform to updated procedures, section d6 & h10 were revised with updated information. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
29062 | 1220648-2024-08757 | 18979168 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18979168&pc=OZD | 03/17/2024 | 05/24/2024 |  | 03/26/2024 |  |  | 03/17/2024 | 03/26/2024 |  |  | 08/31/2025 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024386695 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The expiration of the patient after the patient was withdrawn from care was reviewed by abiomed's medical safety officer. it was determined that there was not enough information to associate use of the device with patient outcome. | The user facility reported an (b)(6) male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient developed a significant access site hematoma. femstop was placed and heparin was held until groin was stable. |  |  |  |  | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use referenced in the initial report is from ifu for the impella cp with smart assist system section: general patient care considerations, section: entering cardiac output, and section: potential adverse events (united states) which now includes the statement "cardiac or vascular injury (including ventricular perforation).¿ b5-added medical safety officer review. d4-updated model number. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was discarded by the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
29063 | 1220648-2024-08759 | 18979218 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18979218&pc=OZD | 03/07/2024 | 04/05/2024 |  | 03/26/2024 |  |  | 03/07/2024 | 03/26/2024 |  |  |  |  | 03/26/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1635044 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility's biomedical engineer reported that the automated impella controller (aic) would not go pass the self-check test. | [5 empty] | The automated impella controller (aic) controller was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation into the aic - system self-check error has been completed since the original report was filed. the pump was returned, tested, and evaluated. the reported complaint was confirmed; the console was unable to fully boot due to a system self-check failure. the cause of the aic issue was a defective impellatronics board. due to the defect, adequate voltage was not being delivered to the motor driving circuitry on the impellatronics board.
29064 | 1220648-2024-08758 | 18979221 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18979221&pc=OZD | 03/17/2024 | 04/19/2024 |  | 03/26/2024 |  |  | 03/17/2024 | 03/26/2024 |  |  | 11/30/2025 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024432637 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | IMPELLA RP | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/17/2024 |  |  |  |  | The user facility reported a 58-year-old male with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported that patient experienced access site bleeding and treatment included femstop and received 2 units of blood products. the patient was reported as stable. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.
29065 | 1220648-2024-08760 | 18979225 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18979225&pc=OZD | 06/28/2022 | 03/07/2025 |  | 03/26/2024 |  |  | 06/28/2022 | 03/26/2024 |  |  | 03/31/2024 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023211175 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/28/2022 |  |  |  |  | The user facility reported a male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. while implanting the device, ventricular fibrillation arrhythmia occurred during insertion which required defibrillation one time to terminate; leading to a successful return of sinus rhythm. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | Analysis summary: cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease ¿timing of the arrythmia event in relation to impella support (during or post-implant) ¿electrocardiogram (ekg) recordings of the arrhythmia episodes ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities ¿response to any antiarrhythmic treatments or interventions during the event ¿any documented device-related issues or complications during impella support ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand ¿review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined.
29066 | 1220648-2024-08761 | 18979272 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18979272&pc=OZD | 08/22/2022 | 05/16/2025 |  | 03/26/2024 |  |  | 08/22/2022 | 03/26/2024 |  |  | 01/31/2024 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023189624 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/22/2022 |  |  |  |  | The user facility reported a 73-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient experienced bloody urine. the patient was started on heparin purge. additionally, the patient had oozing at the impella access site upon removal of the impella cp. the physician was unable to achieve hemostasis with perclose sutures. a surgical wound closure was performed, and hemostasis was achieved. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29067 | 1220648-2024-08763 | 18979343 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18979343&pc=OZD | 03/19/2024 | 05/04/2024 |  | 03/26/2024 |  |  | 03/19/2024 | 03/26/2024 |  |  | 10/31/2025 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025439315 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Disability | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 30-year-old female with severe preeclampsia and concern for peripartum cardiomyopathy was implanted with an impella device for mechanical circulatory support. the team evaluated and found bi-ventricular failure. the patient had an emergent cesarean section. upon extubation the patient went pulseless electrical activity, cardiopulmonary resuscitation was performed, placed upon extracorporeal membrane oxygenation, and an impella device via the right leg/femoral. the limb became ischemic and decision made to explant the impella device after 88 hours. amputation was required of the left leg. | [5 empty] | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned by the user facility for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined. potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation." | The impella device was discarded by the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/pos, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
29068 | 1220648-2024-08764 | 18979346 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18979346&pc=OZD | 10/01/2022 | 05/16/2025 |  | 03/26/2024 |  |  | 10/01/2022 | 03/26/2024 |  |  | 03/31/2024 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213059 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/01/2022 |  |  |  |  | The user facility reported a 67-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. during insertion, resistance was felt upon removal of the .018 wire. the pump was pulled and the wire was retrieved. the valve was recrossed with a new .035 wire. a second impella was placed. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29069 | 1220648-2024-08765 | 18979381 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18979381&pc=OZD | 03/20/2024 | 05/07/2024 |  | 03/26/2024 |  |  | 03/20/2024 | 03/26/2024 |  |  | 09/30/2025 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024411321 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO). | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 54-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient also had extracorporeal membrane oxygenation (ecmo) ongoing and there were limb issues. the team observed thrombus at the access site. the impella device was explanted after 213 hours, and the patient survived. | [5 empty] | The investigation into the thrombus issue has been completed since the original report was submitted. the device was not returned by the user facility for investigation. necessary clinical information was not provided and the device was not returned. therefore, the root cause of the thrombus was not determined. impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." d.4 revised model number as previously entered incorrectly. d.6a and d.6b revised from the initial submission in accordance with updated procedures. e.1 first name was revised as previously entered incorrectly. d.10 updated as was inadvertently omitted from the previous submitted report. h.6 code 4641 was reported incorrectly on the previously submitted report. new code has been added to health effect - impact code. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | The impella device was discarded by the customer, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
29070 | 1220648-2024-08766 | 18979483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18979483&pc=OZD | 03/20/2024 | 05/10/2024 |  | 03/26/2024 |  |  | 03/20/2024 | 03/26/2024 |  |  |  |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025448787 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 68-year-old male with cardiomyopathy was implanted with an impella device for mechanical circulatory support. the oozing was observed at the access site. after removing the dressing and adjusting the angle match, the oozing resolved. both forward pulling sutures in place and blue hub to skin noted. one unit of red blood cells was given. the patient was noted to be in stable condition. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned by the user facility for investigation. the root cause of the access site bleeding issue was most likely use related. the oozing was resolved by adjusting the angle match and both forward pulling sutures were in place as per the clinical description. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b.7 revised as information was inadvertently omitted from the previously submitted report. revised model number was revised from the initial submission in accordance with updated procedures. revised from the initial submission in accordance with updated procedures. reporting contact facility name was revised as information was inadvertently omitted from the previously submitted report. additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer at the time of this mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
29071 | 1220648-2024-08767 | 18979517 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18979517&pc=OZD | 03/07/2024 | 05/08/2024 |  | 03/26/2024 |  |  | 03/07/2024 | 03/26/2024 |  |  | 01/31/2025 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024313094 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support for the patient, a purge filter leakage was identified leading to pump explant. there was no reported patient harm. | [5 empty] | The investigation into the purge leak ¿ pump issue has been completed since the original report was filed. the pump was returned for analysis. upon visual inspection, no cracks were present on purge sidearm. pump was plugged into aic and pump was purged. while purging through the sidearm filter, a leak was observed from the filter holes. clinical details also observed a leak from the sidearm filter. the root cause of the purge leak is due to damaged sidearm filter. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
29072 | 1220648-2024-08768 | 18979572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18979572&pc=OZD | 03/21/2024 | 05/10/2024 |  | 03/26/2024 |  |  | 03/21/2024 | 03/26/2024 |  |  | 08/31/2025 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024394323 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 75-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. while on patient support, the impella device stopped. restart was unsuccessful. the pump was disconnected from the console and left in the left ventricle as the team was concerned of clot formation in the descending aorta due to pump being completely off. inotropes and vasopressors were increased. several hours later the patient was taken to the cath lab and a new impella device was inserted. the patient was noted to be in stable condition. | [5 empty] | The investigation for pump stop has been completed. the pump and purge cassette were returned for evaluation. electrical testing was performed and no abnormalities were found. pump was plugged into aic and primed using purge cassette that was send back with the pump. pump was plugged in for approximately 15 minutes before high purge pressure alarms were triggered. a window in the catheter was cut to check for blood ingress, purge line was clear. catheter was cut and purge pressure was released. pump teardown was performed and dried blood was found in the motor and in the bearing. data logs were reviewed and an increase in motor current and purge pressure over a period of 30 minutes was observed before a pump stop occurred. pump was able to be restarted and run for approximately 15 minutes with high purge pressure before another pump stop occurred. the root cause of the pump stop was due to biomaterial ingress into the motor. | The impella device was not received from the customer at the time of this mdr writing, and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
29073 | 1220648-2024-08771 | 18979861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18979861&pc=OZD | 08/09/2022 | 03/26/2024 |  | 03/26/2024 |  |  | 08/09/2022 | 03/26/2024 |  |  | 04/30/2024 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023224772 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/09/2022 |  |  |  |  | The user facility reported a 65-year-old female patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. as the patient continued to have significant bleeding, it was decided to remove the pump. an angio-seal closure was used and no more bleeding occurred. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
29074 | 1220648-2024-08772 | 18980045 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18980045&pc=OZD | 08/11/2022 | 03/26/2024 |  | 03/26/2024 |  |  | 08/11/2022 | 03/26/2024 |  |  | 03/31/2024 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023199540 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/11/2022 |  |  |  |  | The user facility reported a 52-year-old male patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the decision was made to continue venoarterial extracorporeal membrane oxygenation (va ecmo) support with upgrade to impella 5.5 on a patient that was seen for heart transplant evaluation and possible left ventricular assist device (lvad) or transplant. during the right axillary 5.5 insertion site noticed jackson pratt-drain bleeding overnight, multiple blood products given, and bedside washout was performed. the bleeding was stopped. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
29075 | 1220648-2024-08773 | 18980369 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18980369&pc=OZD | 03/08/2024 | 03/26/2024 |  | 03/26/2024 |  |  | 03/08/2024 | 03/26/2024 |  |  |  |  | 03/26/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1137666 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The hospital had a console with a controller error observed by the hospital biomedical department. there was no patient use and no harm or injury was possible. it was found and removed. the automated impella controller (aic) will be replaced by a backup console. | [5 empty] | The investigation has been completed. the product was returned for investigation. the console logs did not show any entries on the reported day of event but several weeks before there were multiple and persistent error messages indicating communication loss and the console was unable to fully boot up. when the console was inspected, board corrosion was confirmed. there were no defects or anomalies discovered. the cause of the automated impella controller (aic) issue was corrosion on the board.
29076 | 1220648-2024-08774 | 18980539 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18980539&pc=OZD | 10/22/2022 | 05/16/2025 |  | 03/26/2024 |  |  | 10/22/2022 | 03/26/2024 |  |  | 07/31/2024 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023254668 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/22/2022 |  |  |  |  | The user facility reported a 32-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. while on support, the patient experienced access site bleeding. due to worsening organ function, the clinical team connected the patient to venoarterial extracorporeal membrane oxygenation. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
29077 | 1220648-2024-08775 | 18981216 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18981216&pc=OZD | 10/13/2022 | 05/16/2025 |  | 03/26/2024 |  |  | 10/13/2022 | 03/26/2024 |  |  | 06/30/2024 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023237489 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/13/2022 |  |  |  |  | The user facility reported a 62-year-old male patient was on impella cp support due to post cardiotomy cardiogenic shock on veno-arterial extracorporeal membrane oxygenation emco with left ventricle distension. while on support, the patient experienced access site bleeding that required a stitch to be placed. the patient also received one unit of packed red blood cells, and the pump was repositioned via trans-esophageal echo. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29078 | 1220648-2024-08776 | 18981256 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18981256&pc=OZD | 10/16/2022 | 05/16/2025 |  | 03/26/2024 |  |  | 10/16/2022 | 03/26/2024 |  |  | 07/31/2024 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023252366 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/16/2022 |  |  |  |  | The user facility reported a 59-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. us complainant reported that while the patient was on impella cp support, patient had significant amount of oozing from the groin site. the patient also had hemolysis. therefore, the impella was removed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿assess access site for bleeding and hematoma.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29079 | 1220648-2024-08777 | 18981271 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18981271&pc=OZD | 06/28/2022 | 05/16/2025 |  | 03/26/2024 |  |  | 06/28/2022 | 03/26/2024 |  |  | 03/31/2024 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023205501 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/28/2022 |  |  |  |  | The user facility reported a 75-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and that the patient¿s bilateral lower extremities were "dusky". however, the patient expired while on support. no allegations were made towards the device for cause of death. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Limb ischemia. analysis summary: limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition. any concomitant medication. vascular size or variations thereof. detailed description of the symptoms experienced by the patient. physical examination findings, including pulse assessment and visual examination of the affected limb. diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. laboratory test results, including blood tests for markers of inflammation or clotting disorders. any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
29080 | 1220648-2024-08778 | 18981713 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18981713&pc=OZD | 11/15/2022 | 05/19/2025 |  | 03/26/2024 |  |  | 11/15/2022 | 03/26/2024 |  |  | 05/31/2024 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023222813 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 22 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/15/2022 |  |  |  |  | The user facility reported a 22-year-old female patient needing mechanical circulatory support. it was reported that the patient was on impella cp support and developed bleeding and a hematoma. the patient required blood transfusions. additionally, it was reported the patient experienced episodes of ventricular tachycardia. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29081 | 1220648-2024-08779 | 18981807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18981807&pc=OZD | 09/02/2022 | 09/11/2024 |  | 03/26/2024 |  |  | 09/02/2022 | 03/26/2024 |  |  | 05/31/2024 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231354 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/02/2022 |  |  |  |  | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 64-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on support, the patient had a bleed; therefore, systemic heparin was stopped, the purge was cut to 25u/ml, and the patient was transfused with 5 units of blood. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08779 was provided in sections b2, b5, h1, and h6. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29082 | 1220648-2024-08780 | 18982318 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18982318&pc=OZD | 10/27/2022 | 05/19/2025 |  | 03/26/2024 |  |  | 10/27/2022 | 03/26/2024 |  |  | 07/31/2024 |  | 03/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023255877 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/27/2022 |  |  |  |  | The user facility reported a 68-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. during impella cp support, there was a saturated flow event, and the patient experienced ventricular tachycardia for one hour. the device was then exchanged for a different impella cp. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Vf/vt/af/at: cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. timing of the arrythmia event in relation to impella support (during or post-implant). electrocardiogram (ekg) recordings of the arrhythmia episodes. evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. response to any antiarrhythmic treatments or interventions during the event. any documented device-related issues or complications during impella support. evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. saturated flow: the cause of the issue was not determined due to insufficient clinical details.
29083 | 1220648-2024-08656 | 18967215 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18967215&pc=OZD | 07/04/2022 | 04/07/2025 |  | 03/25/2024 |  |  | 07/04/2022 | 03/25/2024 |  |  | 03/31/2024 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023209384 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/04/2022 |  |  |  |  | The user facility reported a 77-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced blood in urine. impella p level was reduced, and the patient was transfused with one unit of packed red blood cells. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29084 | 1220648-2024-08657 | 18967253 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18967253&pc=OZD | 07/05/2022 | 04/07/2025 |  | 03/25/2024 |  |  | 07/05/2022 | 03/25/2024 |  |  | 03/31/2024 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023210429 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 07/05/2022 |  |  |  |  | The user facility reported a 47-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced bleeding at the groin area. systemic heparin was discontinued, and the patient received several units of packed red blood cells. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29085 | 1220648-2024-08658 | 18967291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18967291&pc=OZD | 07/06/2022 | 04/08/2025 |  | 03/25/2024 |  |  | 07/06/2022 | 03/25/2024 |  |  | 04/30/2024 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023215944 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  |  | White | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/06/2022 |  |  |  |  | The user facility reported a 70-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported during insertion of the impella the canula prolapsed. the physician removed the impella through the sheath and was about to successfully re-inserted after washing device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
29086 | 1220648-2024-08659 | 18967371 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18967371&pc=OZD | 07/08/2022 | 04/08/2025 |  | 03/25/2024 |  |  | 07/08/2022 | 03/25/2024 |  |  | 03/31/2024 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023192043 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 54 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Deliver | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/08/2022 |  |  |  |  | The user facility reported a 54-year-old male patient in acute myocardial infarction/cardiogenic shock was was being placed on an impella cp device for mechanical circulatory support. it was reported that during insertion of the impella, the healthcare staff was unable to advance the device through the patient¿s vasculature due to a kink in the cannula. the procedure was aborted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29087 | 1220648-2024-08660 | 18967391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18967391&pc=OZD | 07/01/2022 | 03/07/2025 |  | 03/25/2024 |  |  | 07/01/2022 | 03/25/2024 |  |  | 12/31/2024 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023176005 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/01/2022 |  |  |  |  | The user facility reported a 62-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while the patient was being transferred to a stretcher the "pump in aorta" alarm and the device was pulled back to 82cm within the aorta. due to the patient being unstable the physicians continued with the transfer and reposition the device upon arrival. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29088 | 1220648-2024-08661 | 18967404 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18967404&pc=OZD | 07/02/2022 | 03/25/2024 |  | 03/25/2024 |  |  | 07/02/2022 | 03/25/2024 |  |  | 03/31/2024 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023212183 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 07/02/2022 |  |  |  |  | The user facility reported a 69-year-old patient needing mechanical circulatory support. it was reported that the patient had an attempted impella cp implant but aborted related to patient anatomy. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29089 | 1220648-2024-08662 | 18967435 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18967435&pc=OZD | 07/04/2022 | 03/25/2024 |  | 03/25/2024 |  |  | 07/04/2022 | 03/25/2024 |  |  | 03/31/2024 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023206803 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 21 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 07/04/2022 |  |  |  |  | The user facility reported a 21-year-old patient needing mechanical support. it was reported that while on support, the patient continued to have persistent vt/ventricular fibrillation (vf) after multiple attempts to reposition the catheter. impella was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29090 | 1220648-2024-08663 | 18967493 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18967493&pc=OZD | 07/04/2022 | 03/25/2024 |  | 03/25/2024 |  |  | 07/04/2022 | 03/25/2024 |  |  | 02/29/2024 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202229 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 07/04/2022 |  |  |  |  | The user facility reported a 68-year-old patient needing mechanical circulatory support. it was reported that while on support, the patient experienced access site bleeding that was resolved with sutures and femstop. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records.
29091 | 1220648-2024-08664 | 18967526 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18967526&pc=PYX | 02/29/2024 | 04/17/2024 |  | 03/25/2024 |  |  | 02/29/2024 | 03/25/2024 |  |  | 03/31/2025 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024390038 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Increase in Pressure | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant reported the 65-year-old male patient was on impella rp flex support and there were low purge flows. tissue plasminogen activator was given, however the motor current spiked. the impella rp was therefore replaced. at the time of the explant, there was a clot noted on the catheter. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the high purge pressure and thrombus issues has been completed since the original report was submitted. the device was not returned for investigation. based on the data log provided, there were a few instances of purge pressure spikes with hpp alarms and purge system blocked alarms sustained for about 3 minutes before returning to normal ranges. tpa was run due to suspected clot. the root cause of the low pump flow was most likely biomaterial.
29092 | 1220648-2024-08665 | 18967540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18967540&pc=OZD | 03/06/2024 | 04/18/2024 |  | 03/25/2024 |  |  | 03/06/2024 | 03/25/2024 |  |  | 11/30/2025 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024424877 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/06/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the impella was in the aorta the patient was taken back to lab for repositioning. it was unclear how the impella moved. after the repositioning, the pump appeared to be close to the inlet cage however, the doctor declined the repositioning. the patient remained on impella support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the positioning issue was use related.
29093 | 1220648-2024-08666 | 18967618 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18967618&pc=OZD | 03/07/2024 | 04/23/2024 |  | 03/25/2024 |  |  | 03/07/2024 | 03/25/2024 |  |  |  |  | 03/25/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1146151 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/07/2024 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) the console displayed "optical sensor error" when initially setting up. there was no patient involvement. the device was swapped out with no further issues. | [5 empty] | The investigation into the optical signal issue has been completed since the original report was filed. the pump was returned, tested, and evaluated. the failure mode was able to be reproduced by not completely plugging in the pump. fully plugging in the pump resolved the controller error. the cause of the optical signal failure was most likely an air gap from between the optical cable and the pump. | The impella controller was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
29094 | 1220648-2024-08667 | 18967730 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18967730&pc=OZD | 02/21/2024 | 05/02/2024 |  | 03/25/2024 |  |  | 02/27/2024 | 03/25/2024 |  |  | 10/31/2025 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024416074 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported 67-year-old female patient on impella cp support had confluent hematomas from the right groin into the right retroperitoneum. hematomas were detected by computed tomography (ct) after fall with dislocated arterial access. hemoglobin dropped but could be explained by the blood loss from the dislocated radial pressure measurement. the extended focus search showed the small retroperitoneal hematoma on the ct. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical details provided was sufficient to determine the root cause. the root cause of the access site bleeding was most likely due to use issue of patient movement, the patient had fallen/mobilized as per the clinical description. added-d6a/d6b
29095 | 1220648-2024-08668 | 18967758 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18967758&pc=OZD | 03/08/2024 | 04/24/2024 |  | 03/25/2024 |  |  | 03/08/2024 | 03/25/2024 |  |  | 06/30/2025 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024375541 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/08/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while swapping the peel away sheath with the repositioning sheath, the user accidentally advanced the entire pump into the left ventricle. the device kinked and repositioning was completed successfully. however, shortly after, the placement signal began to malfunction and pump flows dropped. due to the noted failures, the decision was made to replace the pump. there was no patient harm. | [5 empty] | The investigation into the low pump flow and the optical signal issue issues has been completed since the original report was submitted. the device was not returned for investigation. the cause of the low pump flow issue was most likely due to kinked cannula per clinical and log review. the cause of the temporary pump stop issue was most likely biomaterial. the cause of the positioning issue was most likely use related due to improper technique. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29096 | 1220648-2024-08669 | 18967794 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18967794&pc=OZD | 07/06/2022 | 03/07/2025 |  | 03/25/2024 |  |  | 07/06/2022 | 03/25/2024 |  |  | 09/30/2023 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2022144078 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/06/2022 |  |  |  |  | The user facility reported a 82-year-old female in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella cp. it was reported that during insertion, the pump was unable to be advanced to the left ventricle due to the patient's advanced aortic stenosis. after multiple attempts, the impella cp could not be inserted successfully and the procedure was aborted. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29097 | 1220648-2024-08670 | 18967849 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18967849&pc=OZD | 06/28/2022 | 04/08/2025 |  | 03/25/2024 |  |  | 06/28/2022 | 03/25/2024 |  |  | 03/31/2024 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023211239 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/28/2022 |  |  |  |  | The user facility reported a 80-year-old male noted as high risk percutaneous coronary intervention was being implanted with an impella cp device for mechanical circulatory support. the patient¿s anatomy created such difficulty in delivering the impella that the physician decided to abort the procedure. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29098 | 1220648-2024-08671 | 18968021 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18968021&pc=OZD | 06/28/2022 | 05/16/2025 |  | 03/25/2024 |  |  | 06/28/2022 | 03/25/2024 |  |  | 03/31/2024 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191974 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Female |  |  |  | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/28/2022 |  |  |  |  | The user facility reported that 84-year-old female patient was scheduled for high risk percutaneous coronary insertion and was being implanted with an impella cp device for mechanical circulatory support, however the procedure was aborted due to very calcified iliaca and the impella cp could not pass the iliaca. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29099 | 1220648-2024-08672 | 18968076 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18968076&pc=PYX | 03/09/2024 | 04/18/2024 |  | 03/25/2024 |  |  | 03/09/2024 | 03/25/2024 |  |  | 08/31/2025 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024397320 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/09/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella rp for mechanical circulatory support. during support, the impella rp had evidence of clot ingestion with significant decrease in flows. the impella was removed. the patient is stable. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the reported low pump flow has been completed. the impella device was evaluated finding a large clot wrapped around the impeller. the optical parameters were within specification. data logs showed a motor current spike and upwards placement signal shift during a decrease in p-level. flows were slightly decreased. at the same time, there is a psnr alarm for 4 seconds indicating that the clot is on the optical sensor. later in the case there is a secondary ingestion or clot migration while at p-6 with low flows following. the root cause of the low pump flow was determined to be biomaterial. the instructions for use ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿
29100 | 1220648-2024-08673 | 18968087 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18968087&pc=OZD | 02/28/2024 | 04/17/2024 |  | 03/25/2024 |  |  | 02/28/2024 | 03/25/2024 |  |  |  |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025440646 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the 72-year-old male patient developed a radial hematoma and experienced acute anemia during impella cp support. the hematoma was stable throughout support and the acute anemia was treated via blood transfusions. the impella procedure was deemed possibly related to the acute anemia, while the impella device was deemed possibly related to the hematoma. | [5 empty] | The investigation into the access site bleeding ¿ minor and the hemolysis issues has been completed since the original report was submitted. the device was not returned for investigation. it was reported that best practices were not followed as a 6fr medtronic sheath was used. therefore, the root cause of the access site bleeding issue was use related to improper technique. the root cause of the hemolysis issue was not determined since product was not returned and insufficient clinical information provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
29101 | 1220648-2024-08674 | 18968111 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18968111&pc=OZD | 03/10/2024 | 04/11/2024 |  | 03/25/2024 |  |  | 03/10/2024 | 03/25/2024 |  |  | 11/30/2025 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025454025 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/10/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient had moderate to large retroperitoneal hematoma. ct scan showed no active extravasation. the patient was treated with two units of packed red blood cells. the patient remained on impella support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding cannot be determined since the product was not returned and limited clinical details were provided.
29102 | 1220648-2024-08675 | 18968121 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18968121&pc=OZD | 07/07/2022 | 05/16/2025 |  | 03/25/2024 |  |  | 07/07/2022 | 03/25/2024 |  |  | 03/31/2024 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023214412 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/07/2022 |  |  |  |  | The user facility reported a 59-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support the patient developed a hematoma on their right side and also experienced pulselessness bilaterally in their legs. exploratory surgery was performed to investigate the hematoma and a external femoral-femoral bypass was performed, but due to lack of pulses in the legs, the pump was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29103 | 1220648-2024-08676 | 18968142 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18968142&pc=OZD | 03/11/2024 | 05/02/2024 |  | 03/25/2024 |  |  | 03/11/2024 | 03/25/2024 |  |  | 11/30/2025 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024424245 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/11/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella cp pump experienced persistent positioning issues during patient support. after attempts to reposition the device were unsuccessful, the decision was made to replace the pump. there was no patient harm. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the cause of the positioning issue was use related due to patient movement as the pump was confirmed to be in the incorrect position following patient transfer to another hospital. b.7 was revised. the patient's past medical history was inadvertently omitted from the previously reported report. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 health effect - impact code was revised as code was inadvertently omitted from the previously reported report. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29104 | 1220648-2024-08677 | 18968160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18968160&pc=OZD | 10/17/2022 | 05/16/2025 |  | 03/25/2024 |  |  | 10/17/2022 | 03/25/2024 |  |  | 06/30/2024 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023237649 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/17/2022 |  |  |  |  | The user facility reported a 66-year-old male was implanted with an impella cp device for mechanical circulatory support due to society for cardiovascular angiography & interventions stage b. while on support, oozing with hematoma was identified around the site. the impella cp was removed and escalated to an impella 5.5. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29105 | 1220648-2024-08678 | 18968272 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18968272&pc=OZD | 07/08/2022 | 05/16/2025 |  | 03/25/2024 |  |  | 07/08/2022 | 03/25/2024 |  |  | 02/29/2024 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023199825 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/08/2022 |  |  |  |  | The user facility reported a 87-year-old female patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, patient experienced access site bleeding that was resolved with two perclose devices and light pressure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29106 | 1220648-2024-08679 | 18968387 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18968387&pc=OZD | 03/14/2024 | 05/02/2024 |  | 03/25/2024 |  |  | 03/14/2024 | 03/25/2024 |  |  | 11/30/2025 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025439359 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/14/2024 |  |  |  |  | The user facility reported an 80-year-old male with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported that the implanted impella cp pump was found to be underneath the papillary muscle following implant. repositioning was attempted, but the pump was pulled out of the left ventricle. the staff was unable to re-advance the pump across the aortic valve and the decision was made to remove the device. there was no patient harm. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the positioning issue was use related. added-h6 impact code 4628 added- d6a/d6b. g1-corrected spelling for reporting contact last name. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29107 | 1220648-2024-08680 | 18968415 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18968415&pc=OZD | 10/24/2022 | 03/25/2024 |  | 03/25/2024 |  |  | 10/24/2022 | 03/25/2024 |  |  | 07/31/2024 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023250549 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/24/2022 |  |  |  |  | The user facility reported an 87-year-old female on impella cp for mechanical circulatory support to keep the left ventricle from distending, and to help with hemostasis in the left ventricle. the patient experienced profuse bleeding and will be weaned off the impella cp over next several days. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29108 | 1220648-2024-08681 | 18968422 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18968422&pc=OZD | 03/15/2024 | 05/29/2024 |  | 03/25/2024 |  |  | 03/15/2024 | 03/25/2024 |  |  | 10/31/2025 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024411530 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/15/2024 |  |  |  |  | The patient expired in the procedure area. upon review by abiomed's medical safety officer, there was not enough information to exclude the patient's outcome from the use of the device. | The user facility reported a 60-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient was diagnosed with sepsis prior to placement which created communication and conducting issues prior to placing leads and the impella pump. the impella was inserted, but it did not direct into the left ventricle, echocardiogram showed it across the aortic-left atrium (la) communication. the doctor noted there was blood coming into the endotracheal tube. the doctor removed the impella and tried to repair the aorta-la. perfusion struggled with flows during this time. the patient expired in the procedural room. the doctor stated it was not the impella's fault. no allegation was made towards the impella for the cause of death. |  |  |  |  | The investigation for the reported vascular damage has been completed. the impella device was not returned for evaluation. the cause of the vascular damage-great vessel issue was most likely patient condition related per clinical review. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during section: warnings & cautions: cautions. ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ h.1 was revised due to medical safety officer review. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 code 4114 was reported incorrectly on the previously submitted report. new code has been added to type of investigation code. h.10 additional manufacturer narrative on the previously submitted report stated that the device was not returned but the device was discarded. type of investigation code has been updated to reflect discarded. additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury
29109 | 1220648-2024-08682 | 18968447 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18968447&pc=PYX | 03/15/2024 | 05/01/2024 |  | 03/25/2024 |  |  | 03/15/2024 | 03/25/2024 |  |  |  |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC |  |  | US |  | 2025453786 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Insufficient Flow or Under Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/15/2024 |  |  |  |  | The user facility reported a 39-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. it was reported flows for the device would not exceed 2.5 l/min on power level p9. troubleshooting was unable to resolve the issue and the pump was removed. upon explant, it was noted that a thrombus had been ingested and was wrapped around the proximal impeller shaft. a second rp flex device was subsequently implanted for support. | [5 empty] | The investigation into the low pump flow issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of low flow was most likely biomaterial ingestion as an ingestion signature was identified in the logs and the patient's act was reported to be below the recommended range. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29110 | 1220648-2024-08683 | 18968478 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18968478&pc=OZD | 03/15/2024 | 05/02/2024 |  | 03/25/2024 |  |  | 03/15/2024 | 03/25/2024 |  |  | 07/31/2025 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025443283 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 19 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/15/2024 |  |  |  |  | The user facility reported a 19-year-old female with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported that after 15 hours of support the impella cp was being repositioned and came out of the aorta. the decision was made to explant the pump. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed since the original report was submitted. the device was not returned by the user facility for investigation. data logs confirmed pump was in aortic area for about 9.75 hours into the case. pump then was stopped manually & disconnected from the console. based on clinical description & data analysis, the root cause of positioning issue was most likely use related. potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ to conform to updated procedures, section d6 & h10 were revised with updated information.
29111 | 1220648-2024-08684 | 18968513 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18968513&pc=OZD | 03/16/2024 | 05/08/2024 |  | 03/25/2024 |  |  | 03/16/2024 | 03/25/2024 |  |  | 05/31/2025 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024369990 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/16/2024 |  |  |  |  | The user facility reported a 51-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that after delivery of the pump in the left ventricle the patient went into ventricular tachycardia and was shocked once, which restored sinus rhythm. the patient additionally had limb ischemia and persistent groin site bleeding. blood products were transfused and a femostop was placed. the patient's anticoagulation was adjusted. the patient remained on support. | [5 empty] | The investigation for the reported access site bleed, limb ischemia, and ventricular tachycardia/ventricular fibrillation (vt/vf) has been completed. the impella device was not returned for evaluation. clinical details provided were sufficient to determine the root cause of the access site bleed. the root cause of access site bleeding was determined to be use issue because the bleeding was significantly reduced after femostop placed to belt pressure and angle of entry re-matched. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia and vt/vf was not determined. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d.6a and d.6b revised from the initial submission in accordance with updated procedures. g.1 last name was revised as previously entered incorrectly. added code 4641 as it was inadvertently left off the previously submitted manufacture device report. codes 3038 were reported incorrectly on the previously submitted report. additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and 4643apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. ischemia potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.
29112 | 1220648-2024-08685 | 18968565 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18968565&pc=OZD | 07/08/2022 | 09/11/2024 |  | 03/25/2024 |  |  | 07/08/2022 | 03/25/2024 |  |  | 03/31/2024 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023214683 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/08/2022 |  |  |  |  | Information noted the patient care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 60-year-old noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient developed ventricular tachycardia while in care. the decision was made to leave the impella device in and to later reassess. it was reported that as the patient was not a surgical candidate, the team elected for high-risk percutaneous coronary intervention with impella support. it was reported that the patient received a unit of blood overnight. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr was provided in sections b2, b5, h1, and h6. | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records.
29113 | 1220648-2024-08687 | 18968617 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18968617&pc=OZD | 03/17/2024 | 05/07/2024 |  | 03/25/2024 |  |  | 03/17/2024 | 03/25/2024 |  |  |  |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025445908 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 21 YR |  | Female |  |  |  | Injury | Yes | Yes | Insufficient Flow or Under Infusion | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/17/2024 |  |  |  |  | The user facility reported a 21-year-old female of unknown race and ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support, was unable to get the pump above p2 and the patient had poor distal perfusion to the lower extremities. the decision was made to remove the impella. | [5 empty] | The investigation has been completed. the impella device was not returned for investigation. data logs were not available to review. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. the root cause of the low pump flow was unable to be determined as no product or data was returned for analysis. impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 codes 4641, 2993 and 3038 were reported on the previously submitted report inadvertently and should not have been. additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
29114 | 1220648-2024-08688 | 18968650 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18968650&pc=PYX | 03/17/2024 | 01/13/2025 |  | 03/25/2024 |  |  | 03/18/2024 | 03/25/2024 |  |  | 09/30/2025 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC |  |  | US |  | 2024408346 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 03/18/2024 |  |  |  |  | Additional information was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry stating that the patient endured worsening thrombocytopenia with a "possible" relationship to the impella device used: thrombocytopenia likely consumptive. heparin was held until heparin induced thrombocytopenia results showed negative. argatroban was restarted and then replaced with a heparin drip. the patient received two units of platelets transfused. the patient recovered with no sequelae. | The user facility reported a 69-year-old female post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella rp for mechanical circulatory support. it was reported that while on support, the patient developed hemolysis and experienced a drop in platelets. one unit of packed red blood cells was given as treatment and the patient was reported as stable. |  |  |  |  | The investigation into the hemolysis issue has been completed since the original report was submitted. the device was not returned by the user facility for investigation. there was a motor current spike, placement signal shift, and speed deviation immediately prior to a high motor current pump stop. the pump restarted about two seconds later and continued to run with lower flows for the remainder of the case. there were a few more potential ingestion signatures. it is unknown if the activated clotting time was maintained throughout support. the pump was confirmed to be in good position. hemolysis was reported two days later. blood products were given and it is unknown if that helped resolve the issue. the cause of the hemolysis issue most likely was biomaterial ingestion. b.7 revised as this information was inadvertently omitted from the previously submitted report. d.6a and d.6b revised from the initial submission in accordance with updated procedures. | B.5 additional information was received and was added. b.6 relevant labs were added. h.6 codes were added due to new information that was received. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. b.3 revised date of event due to new information received after the initial manufacturer device report number 1220648-2024-08688 was submitted b.7 information was added as it was inadvertently omitted from the previously submitted manufacturer device report number 1220648-2024-08688. d.6a and d.6b revised from the initial submission of manufacturer device report number 1220648-2024-08688 in accordance with updated procedures. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-08688 and should not have been. g.1 revised mdr reporting contact email in accordance with updated procedures since the initial manufacturer device report number 1220648-2024-08688 was submitted. h.6 codes were added that were inadvertently omitted from the initial manufacturer device report number 1220648-2024-08688 in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of thrombocytopenia and hemolysis has been completed. the impella device was not returned for evaluation. pump stop failure mode added since it was reported for the patient update. the cause of the temporary pump stop issue and hemolysis is most likely was biomaterial ingestion. the cause of the thrombocytopenia issue cannot be determined due to insufficient clinical details.
29115 | 1220648-2024-08689 | 18968662 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18968662&pc=OZD | 03/19/2024 | 05/04/2024 |  | 03/25/2024 |  |  | 03/19/2024 | 03/25/2024 |  |  | 11/30/2025 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024434349 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/19/2024 |  |  |  |  | The user facility reported a 64-year-old female was escalated from and impella cp to impella 5.5 device for mechanical circulatory support. it was reported that while on impella 5.5 support there were high purge pressures noted. tissue plasminogen activator was added to the purge. no patient harm was reported. | [5 empty] | The investigation for the reported high purge pressure has been completed. product was not returned for investigation. the data logs confirmed a gradual rise in purge pressure about 12 hours into support with multiple "high purge pressure" and "purge blocked" alarms. tpa was administered after which the purge pressure stabilized. no product was returned. the root cause of high purge pressure was most likely biomaterial build up. the instructions for use for the impella 5.5 with smartassist in section: using the automated impella controller with the impella catheter states "unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the "purge pressure high" alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen." | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29116 | 1220648-2024-08690 | 18968680 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18968680&pc=OZD | 10/25/2022 | 02/18/2025 |  | 03/25/2024 |  |  | 10/25/2022 | 03/25/2024 |  |  | 07/31/2024 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023256460 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 20 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Fluid/Blood Leak | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/25/2022 |  |  |  |  | The user facility reported a 20-year-old female on impella cp for mechanical circulatory support that the patient required two units of packed red blood cells due to bleeding. the patient experienced blood pooling on top of the pressure side arm disc. there was some concern for effusion, and the patient became hypoxemic. the physician attempted to place the swan but the patient was decompensating with worsening hypoxemia and the patient was crashed onto veno-arterio-venous (triple cannulation). it was decided to remove the impella cp. the impella cp removal was successful with no issues. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the reported issue has been completed. the device history record review confirmed that the pump passed all post-sterile inspection checks. no device was received for evaluation. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29117 | 1220648-2024-08691 | 18968736 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18968736&pc=OZD | 03/19/2024 | 04/05/2024 |  | 03/25/2024 |  |  | 03/19/2024 | 03/25/2024 |  |  |  |  | 03/25/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1122630 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/19/2024 |  |  |  |  | The user facility biomed department reported that an automated impella controller (aic) experienced a controller error during preparation for support. the aic was removed from service as a result of the noted issue. there was no patient harm. | [5 empty] | The investigation into the aic - controller error (yellow alarm) issue has been completed since the initial report was submitted. the console was returned, visually inspected, and tested. the reported failure mode was reproduced. communication issues causing by ribbon cable defect was discovered. the root cause of the issue was a defective component. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
29118 | 1220648-2024-08692 | 18968791 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18968791&pc=OZD | 03/10/2024 | 05/01/2024 |  | 03/25/2024 |  |  | 03/11/2024 | 03/25/2024 |  |  |  |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025438881 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Required Intervention | [5 empty] | IMPELLA RP | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/11/2024 |  |  |  |  | The user facility reported a 53-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed compartment syndrome in the right leg after being on bipella support for multiple days. there had been prior perfusion issues with this patient due to low perfusion bilaterally but was caused due to patient hemodynamics and clinical state. once compartment syndrome was diagnosed, the patient was brought to the operating room for bipella explant and a fasciotomy of the right lower extremity. after the procedure, the leg was found to be mobile and obtaining doppler pulses appropriately. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
29119 | 1220648-2024-08693 | 18968975 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18968975&pc=OZD | 02/29/2024 | 04/17/2024 |  | 03/25/2024 |  |  | 02/29/2024 | 03/25/2024 |  |  | 05/31/2025 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024352359 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 13 YR |  | Male |  | Non Hispanic |  | Injury | Yes | Yes | Infusion or Flow Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Stroke/CVA | Pericardial Effusion | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/29/2024 |  |  |  |  | The user facility reported a 13-year-old male patient of an unknown race was implanted with an impella 5.5 as a bridge to transplant. it was reported that while on support, the patient had multiple occurrences of changes in waveforms and impella flows. the left ventricle would uncouple, and flows would drop. position was verified. it was reported that during the heart transplant, the physicians stated when the chest was opened, there was a large pericardial effusion in the pericardial sac. there was also particulate matter on the outflow cage that was visualized. the patient suffered a stroke, the physician stated they are unsure if the impella was related. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve. | The investigation for the reported low pump flow, ventricle perforation, and stroke has been completed. the product and data logs were not returned for investigation. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. the cause of the low pump flow issue was not determined since product and data logs were not returned. the cause of the ventricle perforation issue was not determined since product and logs were not returned and due to insufficient clinical information.
29120 | 1220648-2024-08695 | 18969039 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18969039&pc=OZD | 03/11/2024 | 04/29/2024 |  | 03/25/2024 |  |  | 03/11/2024 | 03/25/2024 |  |  |  |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025446461 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/11/2024 |  |  |  |  | The user facility reported a 72-year-old male in st- segment elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that there was no flow in the right lower extremity. the impella cp was removed, and the patient was escalated to the impella 5.5. after removal, a one centimeter long vegetation was found on the impella cp tip. the patient was reported as stable. | [5 empty] | The investigation into the reported issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia and the thrombus was not determined. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
29121 | 1220648-2024-08696 | 18969054 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18969054&pc=OZD | 02/27/2024 | 04/23/2024 |  | 03/25/2024 |  |  | 03/04/2024 | 03/25/2024 |  |  | 09/30/2025 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024398184 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/04/2024 |  |  |  |  | The user facility reported a 64-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient was experiencing re-occurring arrhythmias. tte (transesophageal echocardiography) performed to assess placement; cannula body not lying on septum while the patient is still. inflow still measuring 5.1cm below the aortic valve. the patient is stable. | [5 empty] | The investigation for the reported ventricular fibrillation/ventricular tachycardia (vf/vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf/vt could not be determined. the instructions for use states ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ d4-updated model number. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29122 | 1220648-2024-08697 | 18969123 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18969123&pc=OZD | 03/09/2024 | 05/06/2024 |  | 03/25/2024 |  |  | 03/18/2024 | 03/25/2024 |  |  | 11/30/2025 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024431396 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/18/2024 |  |  |  |  | The user facility reported a 54-year-old male with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that on day 2 of support the patient developed a hematoma at the access site. the hematoma was evacuated, and pressure dressing was applied. the patient remains on support and is reported to be in stable condition. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned by the user facility for investigation. the root cause of the access site bleeding issue was not determined since product was not returned and not much information is provided in the clinical description. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
29123 | 1220648-2024-08698 | 18969187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18969187&pc=OZD | 11/02/2023 | 10/03/2025 |  | 03/25/2024 |  |  | 11/02/2023 | 03/25/2024 |  |  | 08/31/2025 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024382376 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Stroke/CVA | [12 empty] | Required Intervention | Death |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The complainant reported that 49-year-old male patient on impella support experienced a placement signal not reliable. there was no harm to the patient as a result of this incident. nine days after this event, the patient who had a known 70% occlusion to his right middle cerebral artery prior, experienced a right middle cerebral artery cerebrovascular accident (cva). as a result the patient was taken to the icu where it was determined that the cerebral artery was fully occluded. evacuation was attempted but they were unable to extract all of the embolus. the patient's act was between 160 and 180 seconds. d5w + 25 (meq/l) sodium bicarbonate was being used as the purge solution. the patient was not considered a transplant candidate. family decision was to have the impella device removed and patient to go home on comfort care. biomaterial was observed in the outflow cage upon explant. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08698 was provided in sections b2, b5, h1, and h6. | The investigation into the reported ¿stroke/cva¿ and the ¿placement signal issue¿ has been completed. no products were returned for investigation. data logs show an abrupt drop in snr, contrast, gain, and light. there is an abrupt increase in ps and lamp. there is also ¿placement signal not reliable¿ alarms at this time. these 5s parameters recover 26 hours later and return to normal values. the root cause of the optical signal failure was most likely an air gap, but the location of the gap is unable to be determined since no product was returned. without imaging evidence to confirm the source of the stroke was due to a clot that passed through the impella device, the relation to impella is unknown. without a pre-existing ischemic stroke, if heparin was in the purge, it is unknown if the cause of the stroke was due to heparin. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. h6 updated to include placement signal coding. b7 other relevant history, including preexisting medical conditions updated. d6a / d6b updated. d10 concomitant medical products and therapy dates updated. g1 reporting contact email updated.
29124 | 1220648-2024-08699 | 18969214 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18969214&pc=OZD | 03/11/2024 | 04/25/2024 |  | 03/25/2024 |  |  | 03/11/2024 | 03/25/2024 |  |  |  |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025446058 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Advance | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/11/2024 |  |  |  |  | The user facility reported a 79-year-old male undergoing cardiac surgery was to be implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that complications were encountered during attempted delivery of an impella 5.5 pump. multiple attempts were made to insert the pump into the graft, but the device kept bending behind the motor housing. due to the severity of the bend, the decision was made to remove the device and a new impella 5.5 pump was placed for patient support. there was no patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely due to patient condition since the patient had a 90 degree innominate to aorta. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29125 | 1220648-2024-08700 | 18969377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18969377&pc=PYX | 03/10/2024 | 10/02/2024 |  | 03/25/2024 |  |  | 03/11/2024 | 03/25/2024 |  |  | 06/30/2025 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024371151 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Disability | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/11/2024 |  |  |  |  | The user facility reported a 53-year-old male undergoing cardiac surgery was implanted with impella devices for mechanical circulatory support. the us complainant reported that the patient was on bipella (impella rp-cp) support ongoing for over 44 hours when the impella rp and impella cp were weaned and explanted. the impella rp accessed limb had compartment syndrome diagnosed and treated by fasciotomy. after the surgical treatment, limb flow returned. | [5 empty] | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-08700. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact fax number missing. h6 component code revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿the impella rp flex with smartassist system is contraindicated for use with patients experiencing any of the following conditions: disorders of the pulmonary artery wall that would preclude placement or correct positioning of the impella rp flex with smartassist device; mechanical valves, severe valvular stenosis or valvular regurgitation of the tricuspid or pulmonary valve; mural thrombus of the right atrium or vena cava; anatomic conditions precluding insertion of the pump; presence of a vena cava filter or caval interruption device, unless there is clear access from the internal jugular vein to the right atrium that is large enough to accommodate a 22 fr catheter.¿
29126 | 1220648-2024-08701 | 18969970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18969970&pc=OZD | 03/11/2024 | 04/29/2024 |  | 03/25/2024 |  |  | 03/11/2024 | 03/25/2024 |  |  |  |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025446649 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Insufficient Flow or Under Infusion | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/11/2024 |  |  |  |  | The user facility reported a 55-year-old female noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that immediately following implant, the impella cp pump began to experience suction issues at performance level p-9. an x-ray was performed, and a kink was noted under the outlet. attempts to reposition the pump were unsuccessful in straightening the kink. despite the noted damage, the physician opted to continue support at performance level p4 with suboptimal flows. following a scheduled procedure, one last attempt was made to reposition straighten out the pump, however the patient's short ascending aorta made the task difficult, and the decision was made to wean and remove the pump. there was no patient harm resulting from the issue. | [5 empty] | The investigation for the reported low pump flow has been completed. data logs were reviewed which confirmed the low flow when pump started and remained to the rest of the case that triggered auto suction response and suction alarms. the impella was returned for evaluation. visual inspection found a kink on the cannula about 15 mm from inflow cage and the cannula bonding site. no other issues were found on the rest of the pump. the root cause of low flow was most likely a kinked cannula. added- d9 device return date, h6 impact code 4628 to conform to updated procedures, section d6 has revised with updated information. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29127 | 1220648-2024-08702 | 18970254 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18970254&pc=OZD | 10/12/2022 | 05/16/2025 |  | 03/25/2024 |  |  | 10/12/2022 | 03/25/2024 |  |  | 03/31/2024 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204819 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Deliver | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/12/2022 |  |  |  |  | The user facility reported a 80-year-old male patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during insertion the surgeon tried to place the perclose, however, when it was deployed it would not retract and stuck in the left femoral artery (lfa). the perclose device was surgically removed and the lfa was closed. the decision was made to place the impella via the right femoral artery (rfa) and proceed with percutaneous coronary intervention (pci) via a single access. during support, the patient decompensated and vasopressin, epinephrine and neo-synephrine drips were added, the drips were able to be turned off after the impella was inserted. the patient was also having some oozing at the site which resolved. the impella was weaned down and removed. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined.
29128 | 1220648-2024-08703 | 18970506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18970506&pc=OZD | 07/21/2022 | 05/23/2025 |  | 03/25/2024 |  |  | 07/21/2022 | 03/25/2024 |  |  | 01/31/2024 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 2023207455 |  | 005062-RO |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/21/2022 |  |  |  |  | The complainant reported a patient, race and ethnicity unknown, with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the complainant reported that the patient suffered from an infected peripherally inserted central catheter (picc) line. the impella cp was placed in the cardiac catheterization laboratory (ccl). the health care professional (hcp) communicated that high purge pressure and purge system blocked alarms were received. purge pressure increased with a pump flow of 2.3 milliliters per hour. per instructions for use (ifu), the catheter was checked for kinks when a rise in motor current was noted. no kinks were reported to be observed in impella cp catheter tubing. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29129 | 1220648-2024-08704 | 18970655 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18970655&pc=OZD | 08/01/2022 | 05/16/2025 |  | 03/25/2024 |  |  | 08/01/2022 | 03/25/2024 |  |  | 04/30/2024 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023217741 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/01/2022 |  |  |  |  | The complainant reported a patient, race and ethnicity unknown, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient was on impella cp support due to ventricular fibrillation (vf) arrest. while on support, the patient experienced distal limb perfusion issues. the interventional cardiologist placed a jump catheter in the intensive care unit (icu). at first, he used the sidearm of the impella catheter, then he placed a right radial arterial line and connected the jump catheter to the right superficial femoral artery (sfa) distal to the impella. discussed typical presentation of bilateral fem-fem distal perfusion catheters and the need to flush/inability to assess clot within tubing. the patient was escalated to the impella 5.5 | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition. any concomitant medication. vascular size or variations thereof. detailed description of the symptoms experienced by the patient. physical examination findings, including pulse assessment and visual examination of the affected limb. diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. laboratory test results, including blood tests for markers of inflammation or clotting disorders . any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension . patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
29130 | 1220648-2024-08705 | 18970951 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18970951&pc=OZD | 08/09/2022 | 05/14/2025 |  | 03/25/2024 |  |  | 08/09/2022 | 03/25/2024 |  |  | 12/31/2023 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2023170397 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/09/2022 |  |  |  |  | The user facility reported a 60-year-old male patient, with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the complainant reported the patient experienced hemolysis. treatment information for hemolysis is unknown, however, the issue was resolved. the complainant reported that the patient had access site bleeding with a hematoma. the access site bleeding was treated with a blood transfusion. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a/d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29131 | 1220648-2024-08706 | 18972169 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18972169&pc=OZD | 03/12/2024 | 04/26/2024 |  | 03/25/2024 |  |  |  | 03/25/2024 |  |  | 08/31/2024 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023262919 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/12/2024 |  |  |  |  | The user facility reported a 77-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported after the implant, a hematoma started to form around the access site. pressure was applied but the patient's mean arterial pressure slowly dropped and the patient coded. return of spontaneous circulation was achieved. fluids and two units of packed red blood cells were provided. the doctor believes there was a perforation of the femoral artery that occurred while attempting to place the perclose sutures. the peel away sheath was removed and a balloon tamponade was performed. vascular surgery was consulted for surgical closure. | [5 empty] | The investigation for the reported vascular damage and subsequent bleeding has been completed. no product was returned for evaluation. clinical details noted that physician attempted to place preclose sutures twice but was unsuccessful, then proceeded to insert impella dilators and peel-away introducer sheath when bleed was observed. patient coded and upon spontaneous recirculation returning, 2 units of blood were given. patient was taken to or to repair femoral bleed and surgical closure. the root cause of the vascular damage was most likely due to use during the impella access procedure. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿
29132 | 1220648-2024-08707 | 18972790 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18972790&pc=OZD | 03/07/2024 | 05/23/2024 |  | 03/25/2024 |  |  | 03/12/2024 | 03/25/2024 |  |  |  |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025437483 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Death | Required Intervention |  |  |  |  | EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/12/2024 |  |  |  |  | The user facility reported a 60-year-old male patient that underwent ascending aortic replacement for sepsis, aortic dissection, and the left and right coronary arteries also had false lumen occlusion. coronary artery bypass graft was performed on the right and percutaneous coronary intervention was performed on the left, and the patient was managed with extracorporeal membrane oxygenation - impella (ecpella). surgery was performed and impella cp was placed. postoperatively, the patient was managed with an open chest. eleven days later, the coronary artery bypass graft was repaired, and impella cp was exchanged for impella 5.5, and the chest was closed. immediately thereafter, signs of infection were observed, and hemodynamic failure occurred. the infection was attributed to the prolonged thoracotomy procedure and was not related to the impella support. it is unclear whether the change to impella 5.5 is related to the procedure. it was further reported that that when patient was escalated from impella cp to 5.5 the patient's heart failure worsened after the device was replaced, and the patient expired of severe heart failure. the relationship between death and impella is unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ | The investigation has been completed since the original report was submitted. no product or data logs were returned for evaluation. clinical details noted that the patient had replacement of artificial ascending blood vessel due to sepsis and was being managed on ecpella support with chest open when the decision was made to escalate to impella 5.5 for more flows. after placement of 5.5 via direct aortic access, the patient developed an infection and passed away. the clinical team noted that there were no issues with the function of the impella and the infection was attributed to previous prolonged thoracotomy. the cause of death was reported to be due to the patient's worsening heart failure and there was no implication of a relation to impella. evidence has not been provided, therefore, the root cause of the infection could not be determined. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock & cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 code 4114 was reported incorrectly on the previously submitted report. new code has been added to type of investigation codes. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. additional manufacturer narrative on the previously submitted report stated that the device was not returned but the device was discarded.
29133 | 1220648-2024-08708 | 18973258 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18973258&pc=OZD | 09/14/2022 | 03/25/2024 |  | 03/25/2024 |  |  |  | 03/25/2024 |  |  | 10/31/2023 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022154757 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/14/2022 |  |  |  |  | The user facility reported a 70-year-old male patient needing an impella cp for mechanical circulatory support, due to triple vessel disease. while on support, the patient had a hematoma. blood products were provided and heparin was withheld to manage the bleeding. additionally, the patient had a drop in platelets. heparin drip was discontinued due to the drop in platelets. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29134 | 1220648-2024-08709 | 18973879 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18973879&pc=OZD | 03/12/2024 | 04/29/2024 |  | 03/25/2024 |  |  |  | 03/25/2024 |  |  | 11/30/2025 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024428672 |  | 1000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/12/2024 |  |  |  |  | The user facility reported a 53-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. over a month later, the pump reportedly stopped working. the console was swapped, but the pump remained inactive. the pump was subsequently removed, and evidence of a clot was observed. a new impella 5.5 pump was placed for ongoing support. there was no patient harm. | [5 empty] | The investigation for the reported pump stop has been completed. pump was not returned for investigation. data logs were returned and investigated. alarm 119 was seen in the logs along with a motor current spike of 30000. no purge pressure alarms were seen. the seen characteristics are similar to that of electrical open. since product was not returned, the root cause of the pump stop issue was most likely an open electrical line. added- d6a/d6b in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
29135 | 1220648-2024-08710 | 18974184 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18974184&pc=OZD | 06/27/2022 | 03/25/2024 |  | 03/25/2024 |  |  |  | 03/25/2024 |  |  | 02/29/2024 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023201596 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/27/2022 |  |  |  |  | The user facility reported a 48-year-old male patient in an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced ventricular tachycardia (vt) that resolved with no intervention specified. additionally, the patient experienced access site bleeding/hematoma that was resolved with femostop. the patient remained on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
29136 | 1220648-2024-08711 | 18974293 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18974293&pc=PYX | 03/12/2024 | 05/01/2024 |  | 03/25/2024 |  |  |  | 03/25/2024 |  |  |  |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025450362 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Insufficient Flow or Under Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/12/2024 |  |  |  |  | The user facility reported a 78-year-old female post cardiothoracic surgery was implanted with an impella rp flex device for mechanical circulatory support. it was reported that after swan manipulation without reducing the impella p level, flows dropped significantly with concern for clot ingestion. the impella rp flex was explanted. there was no patient harm reported. | [5 empty] | The investigation into the low pump flow issue has been completed since the original report was filed. the pump was returned, tested, and evaluated. upon investigation of the returned pump, a clot was observed around the impeller. when the pump was run in the lab, low flows were reproduced. the root cause of the low pump flow was determined to be biomaterial. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 codes 4114 and 3233 were inadvertently left off the previously submitted report. the device has since been returned and this investigation has been completed. codes to reflect this information can be seen in section h.6. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29137 | 1220648-2024-08712 | 18975150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18975150&pc=OZD | 06/27/2022 | 05/16/2025 |  | 03/25/2024 |  |  |  | 03/25/2024 |  |  | 02/29/2024 |  | 03/25/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195177 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient Device Interaction Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/27/2022 |  |  |  |  | The user facility reported a 58-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the impella experienced delivery issues. it would not cross the iliac due to tortuous anatomy. the impella was removed, sheath exchanged for long sheath, and a new impella inserted. the patient remained on support. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29138 | 1220648-2024-08647 | 18965220 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18965220&pc=OZD | 07/17/2022 | 05/27/2025 |  | 03/23/2024 |  |  | 07/17/2022 | 03/23/2024 |  |  | 03/31/2024 |  | 03/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023209700 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Insufficient Flow or Under Infusion | Difficult to Advance | [5 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 07/17/2022 |  |  |  |  | The user facility reported a 46-year-old patient needing mechanical support. it was reported that the patient was on impella cp support due to cardiomyopathy, congestive heart failure, atrial fibrillation, and decompensated heart failure. while on support, the patient experienced a positioning issue, continued atrial fibrillation that did not resolve while on support, suction alarms indicating low pump flow, access site weeping resulting in heparin being withheld, and optical sensor issue for ps low alarm occurring. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29139 | 1220648-2024-08648 | 18965319 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18965319&pc=OZD | 07/18/2022 | 05/14/2025 |  | 03/23/2024 |  |  | 07/18/2022 | 03/23/2024 |  |  | 04/30/2024 |  | 03/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023215085 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/18/2022 |  |  |  |  | The user facility reported a 63-year-old female was implanted with an impella cp device for mechanical circulatory support due to multi vessel disease. during support, post implant bleeding occurred. in addition, the patient was not able to obtain full hemostasis using best practices even after re-suturing with forward tension and angle matching. a large hematoma was observed, and the doctor explanted the impella for patient safety. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29140 | 1220648-2024-08649 | 18965596 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18965596&pc=OZD | 12/05/2022 | 05/13/2025 |  | 03/23/2024 |  |  | 12/05/2022 | 03/23/2024 |  |  | 08/31/2024 |  | 03/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023270477 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Patient-Device Incompatibility | Difficult to Advance | [5 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/05/2022 |  |  |  |  | The user facility reported an impella cp in a 74-year-old male patient for mechanical circulatory support. during insertion, it was reported that once guide was engaged there was a severe dampening of pressure. the right coronary artery (rca) was ballooned several times. a reperfusion arrhythmia occurred and a guideliner was utilized, and the stent would not cross. the medical team decided to change the slender to a 6fr destination but was having difficulty advancing it into the impella sheath. another guide was introduced, and rca was wired but was found to be in a dissection plane. the patient decompensated, and levo was started, balloon unable to cross and kicked the guide pack, but patient began to stabilize. after several attempts, the impella was left overnight. sheath was used and access site was closed with mynx control. the patient was presented with good hemostasis. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29141 | 1220648-2024-08650 | 18965679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18965679&pc=OZD | 10/06/2022 | 05/14/2025 |  | 03/23/2024 |  |  | 10/06/2022 | 03/23/2024 |  |  | 01/31/2024 |  | 03/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023173685 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | Activation, Positioning or Separation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/06/2022 |  |  |  |  | The user facility reported a 64-year-old male patient on an impella cp for mechanical circulatory support. it was reported that the patient was placed on impella cp support, but the device fell out into the aorta, so it was removed. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29142 | 1220648-2024-08651 | 18965853 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18965853&pc=OZD | 10/11/2022 | 02/21/2025 |  | 03/23/2024 |  |  | 10/11/2022 | 03/23/2024 |  |  | 04/30/2024 |  | 03/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023215498 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Difficult to Insert | Failure to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/11/2022 |  |  |  |  | The user facility reported an 83-year-old male on impella cp for mechanical circulatory support. it was reported the patient presented with chest pain, sob, and elevated troponins. the patient was taken to cath lab to place impella cp. it was reported that they were unable to cross lesion in the left anterior descending artery (lad). at the conclusion of procedure, the patient was transferred to ormc and taken to cath lab for high risk percutaneous coronary intervention (hrpci). the patient expired while on support. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. initial mfr 1220648-2024-08651 was sent erroneously. upon review, the complaint was not related to the impella cp. the reported incident was related to the percutaneous insertion, and unrelated to the impella device.
29143 | 1220648-2024-08652 | 18965890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18965890&pc=OZD | 10/11/2022 | 05/21/2025 |  | 03/23/2024 |  |  | 10/11/2022 | 03/23/2024 |  |  | 07/31/2024 |  | 03/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023250111 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/11/2022 |  |  |  |  | The user facility reported a 46-year-old male patient in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during repositioning attempt on the impella cp, it began suctioning and left ventricle/thoracic aorta waveforms became erratic and inconsistent. they attempted to stop the pump and restart, which was successful for only a few minutes. despite repositioning attempts, the pump would not resume normal flows or waveforms, so decision was made to replace the pump. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29144 | 1220648-2024-08654 | 18965945 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18965945&pc=OZD | 10/14/2022 | 03/22/2024 |  | 03/23/2024 |  |  | 10/14/2022 | 03/23/2024 |  |  | 06/30/2024 |  | 03/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023249018 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/14/2022 |  |  |  |  | The user facility reported a 45-year-old male patient needing an impella cp for mechanical circulatory support. the decision was made for high risk percutaneous coronary interventions (hrpci) with impella cp support. post procedure, an echocardiogram was performed at bedside and showed the impella was too deep. dressing was clean, dry, and intact. patient was resting comfortably; however, it was reported the impella was removed due to a hematoma. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
29145 | 1220648-2024-08655 | 18966163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18966163&pc=OZD | 10/22/2022 | 04/04/2025 |  | 03/23/2024 |  |  | 10/22/2022 | 03/23/2024 |  |  | 08/31/2024 |  | 03/23/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023248109 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/22/2022 |  |  |  |  | The user facility reported a 76-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support it was reported that the patient experienced pain to left lower extremity (lle), concerns for ischemia with mottling to leg. the pump was pulled back from left ventricular outflow tract (lvot) after cardiac arrest. cardiac arrest related to severity of acute myocardial infarction (ami). vascular fellow at bedside reports dopplering pulse. after arresting multiple times, team spoke with family and made patient do-not-resuscitate order (dnr). patient arrested shortly after. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
29146 | 1220648-2024-08546 | 18957969 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18957969&pc=OZD | 08/26/2022 | 03/19/2025 |  | 03/22/2024 |  |  | 08/26/2022 | 03/22/2024 |  |  | 04/30/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023214031 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/26/2022 |  |  |  |  | The user facility reported a 60-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced a hematoma at the groin access site and systemic heparin was decreased as a result. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | The investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29147 | 1220648-2024-08547 | 18957984 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18957984&pc=OZD | 12/15/2022 | 03/11/2025 |  | 03/22/2024 |  |  | 12/15/2022 | 03/22/2024 |  |  | 07/31/2023 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023249543 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/15/2022 |  |  |  |  | The user facility reported a 69-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while the patient was on impella cp support, the patient had a period of ventricular fibrillation (vf) requiring multiple shocks. additionally, femstop was in place due to oozing from the site (access site bleeding). bicarbonate was ordered for purge. the patient was monitored and there was blood in the urine (gross hematuria). lab results were analyzed, which had indicated that everything was hemolysed, gross hematuria also remained. central venous pressure was 15. it was advised that an echo be performed for position even if the position appeared to be perfect. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29148 | 1220648-2024-08548 | 18958029 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18958029&pc=OZD | 07/13/2022 | 03/19/2025 |  | 03/22/2024 |  |  | 07/13/2022 | 03/22/2024 |  |  | 03/31/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023209959 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/13/2022 |  |  |  |  | The user facility reported a 88-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that a hematoma was at the insertion site. the heparin drip was stopped, due to the hematoma. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in d6, h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29149 | 1220648-2024-08549 | 18958053 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18958053&pc=OZD | 07/25/2022 | 05/23/2025 |  | 03/22/2024 |  |  | 07/25/2022 | 03/22/2024 |  |  | 02/29/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023201615 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic |  | Death | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Positioning Problem | Pumping Problem |  |  |  |  | Hemolysis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 07/25/2022 |  |  |  |  | The complainant reported that after the pump was implanted the patient had hemolysis. this was treated with repositioning, however, hemolysis remained present. after the patient was placed on va-ecmo, the hemolysis issue improved dramatically. additionally, the complainant reported that there were also frequent suction alarms at p2. in response to the suction alarms, the patient was given fluid bolus and two units of blood, however, suction alarms persisted with low pump flows. subsequently, the proper steps were demonstrated on how to disable suction alarms. however, the pump continued to have positioning issues as it kept interacting with the sub valvular apparatus and lateral and inferior wall. as a result, it was decided to explant the pump. | Procedural outcome was that patient expired and medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08549 was provided in sections b2, b5,d6, h1, and h6. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29150 | 1220648-2024-08550 | 18958071 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18958071&pc=OZD | 02/26/2024 | 04/17/2024 |  | 03/22/2024 |  |  | 02/26/2024 | 03/22/2024 |  |  |  |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025437258 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/26/2024 |  |  |  |  | The user facility reported a 65-year-old male patient in st-segment elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it reported that during impella cp removal, a large piece of a clot came out with the device. the patient was reported stable. | [5 empty] | The investigation for the reported thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombus could not be determined. the instructions for use states ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. it also states ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ h6 impact code-corrected 4614 should be 2199. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29151 | 1220648-2024-08551 | 18958076 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18958076&pc=OZD | 07/14/2022 | 03/20/2025 |  | 03/22/2024 |  |  | 07/14/2022 | 03/22/2024 |  |  | 03/31/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023211183 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/14/2022 |  |  |  |  | The user facility reported a 75-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported after the impella was successfully implanted, an impella in aorta alarm occurred. staff reduced the impella to p2. imaging showed impella pulled back, but pigtail was still in left ventricle. it was reported the patient experienced a hematoma at the access site. the medical team decided to explant the impella cp. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. section d6a / d6b: added implant and explant dates.
29152 | 1220648-2024-08552 | 18958099 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18958099&pc=OZD | 03/05/2024 | 04/24/2024 |  | 03/22/2024 |  |  | 03/05/2024 | 03/22/2024 |  |  |  |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025446077 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | American Indian Or Alaskan Native | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/05/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was to be implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 pump was unable to advance through the patient's vasculature during attempted delivery. as a result, the impella 5.5 pump was removed and an impella cp pump was placed for planned support. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, the impella was unable to pass through the patient's vasculature as per the clinical description provided.
29153 | 1220648-2024-08553 | 18958166 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18958166&pc=OZD | 03/05/2024 | 04/24/2024 |  | 03/22/2024 |  |  | 03/05/2024 | 03/22/2024 |  |  | 07/31/2025 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025443417 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 76-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella cp was inserted without complications and percutaneous coronary intervention procedure was completed. the impella cp was explanted in the procedural area due to concerns for ischemia. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
29154 | 1220648-2024-08555 | 18958229 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18958229&pc=OZD | 03/05/2024 | 05/04/2024 |  | 03/22/2024 |  |  | 03/05/2024 | 03/22/2024 |  |  | 06/30/2025 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024378558 |  | 0048-0044 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/05/2024 |  |  |  |  | It was reported that a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support, the patient's foot became cool to the touch following implant of an impella cp device. it was noted that the patient's vessel was small prior to implant. a perfusion catheter was inserted into the vessel to treat the noted condition. it was then observed that the patient had developed some oozing at their access site. one unit of blood was given to manage the bleed and support was continued with the implanted pump. | [5 empty] | The investigation for the reported limb ischemia and access site bleed has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia and access site bleed could not be determined as the device was not returned nor sufficient clinical details. potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation." | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
29155 | 1220648-2024-08556 | 18958244 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18958244&pc=OZD | 07/17/2022 | 03/19/2025 |  | 03/22/2024 |  |  | 07/17/2022 | 03/22/2024 |  |  | 01/31/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183672 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/17/2022 |  |  |  |  | The user facility reported that a 45-year old male patient was placed on impella cp support after presenting with chest pains, st-segment elevation myocardial infarction and developed 2 episodes of ventricular fibrillation requiring defibrillation. the patient developed a hematoma at the impella cp insertion site that was noted upon removal of the peel-away introducer and insertion of the repositioning sheath. hemostasis was maintained with a femstop, and due to the bleeding the device was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29156 | 1220648-2024-08557 | 18958330 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18958330&pc=OZD | 03/05/2024 | 05/10/2024 |  | 03/22/2024 |  |  | 03/06/2024 | 03/22/2024 |  |  | 02/29/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 2023191009 |  | 005066-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 50-year-old female with cardiomyopathy was implanted with an impella 5.0 device for mechanical circulatory support. it was reported that from around the seventh day of impella support, the purge flow rate decreased. despite measures such as increasing the heparin concentration in the purge solution and replacing the purge housing, there was no improvement. the flow rate remained at around 2.0 milliliters. on the tenth day of impella support, the impella suddenly stopped. the impella was explanted, and no additional treatment was required as the patient¿s cardiac function was improving. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported low pump flow and the pump stop has been completed. pump was not returned. data logs were returned and investigated. data logs show purge pressure gradual increase leading high purge pressure and purge system blocked alarms. the purge pressure alarms were there for 3 days until pump stop 119 alarm occurred. there were no suction alarms. heparin was increased but no result. the root cause of the pump stop issue was not determined since product was not returned. data log review and clinical information is inconclusive. the root cause of the high purge pressure issue was not determined since product was not returned. data log review and clinical information is inconclusive.
29157 | 1220648-2024-08561 | 18958550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18958550&pc=OZD | 08/02/2022 | 03/21/2025 |  | 03/22/2024 |  |  | 08/02/2022 | 03/22/2024 |  |  | 04/30/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023215494 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/02/2022 |  |  |  |  | The user facility reported a 73-year-old male on impella cp for mechanical circulatory support. it was reported that the patient was placed on impella cp support for a pci after being a surgical turndown. the patient was going to have a pci with shockwave to the left main, left circumflex artery and proximal lad artery. the patient developed bleeding at the impella cp access site due to the perclose failing twice. the physician placed a 14x30 cook sheath and manual pressure held once activated clotting time was sub therapeutic. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29158 | 1220648-2024-08562 | 18958571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18958571&pc=OZD | 02/28/2024 | 05/31/2024 |  | 03/22/2024 |  |  | 03/06/2024 | 03/22/2024 |  |  |  |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025447950 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/06/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. a left ventricular thrombus was confirmed after impella cp insertion. a left ventricular thrombectomy was performed at the same time as mitral valve procedure. the impella cp was exchanged for an impella 5.5. | The patient's outcome was reviewed by abiomed's medical safety officer (mso). it was determined that there was not enough information to associate use of device with outcome. |  |  |  |  | The impella device was discarded and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombus could not be determined. the instructions for use for the impella cp with smart assist system in section: potential adverse events (united states) states, "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)".
29159 | 1220648-2024-08563 | 18958622 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18958622&pc=OZD | 07/10/2022 | 03/22/2024 |  | 03/22/2024 |  |  | 07/10/2022 | 03/22/2024 |  |  | 02/29/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023193222 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | Yes | Initial submission |  |  |  |  | 07/10/2022 |  |  |  |  | The user facility reported a 51-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that while the patient was on impella cp support, there was some existing blood with clot noted at the site. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29160 | 1220648-2024-08564 | 18958682 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18958682&pc=OZD | 07/05/2022 | 03/20/2025 |  | 03/22/2024 |  |  | 07/05/2022 | 03/22/2024 |  |  | 03/31/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023214258 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/05/2022 |  |  |  |  | The user facility reported a 64-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. while on support, the groin was oozing, and leg ischemia was reported. impella was explanted. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. g1 manufacturing site address line 1.
29161 | 1220648-2024-08565 | 18958694 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18958694&pc=OZD | 07/06/2022 | 05/02/2025 |  | 03/22/2024 |  |  | 07/06/2022 | 03/22/2024 |  |  | 02/29/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023197029 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  |  |  | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/06/2022 |  |  |  |  | The user facility reported a 46-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during impella support, the pump malpositioned causing heavy hemolysis. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29162 | 1220648-2024-08566 | 18958702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18958702&pc=OZD | 10/11/2022 | 03/24/2025 |  | 03/22/2024 |  |  | 10/11/2022 | 03/22/2024 |  |  | 06/30/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023246513 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/11/2022 |  |  |  |  | The user facility reported an 81-year-old male on impella cp for mechanical circulatory support. it was reported while the patient was on impella cp support, a hematoma was noted at the impella access site. the impella cp was explanted and two percloses were used with 5 mins of manual pressure until hemostasis was achieved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in d6, h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29163 | 1220648-2024-08567 | 18958761 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18958761&pc=OZD | 07/05/2022 | 03/25/2025 |  | 03/22/2024 |  |  | 07/05/2022 | 03/22/2024 |  |  | 01/31/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191235 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/05/2022 |  |  |  |  | The user facility reported a 75-year-old female patient experienced oozing at the groin site after being placed on impella cp for mechanical circulatory support. additionally, there was hematuria due to malposition of the pump. heparin was discontinued and femostop was placed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29164 | 1220648-2024-08568 | 18958821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18958821&pc=OZD | 10/13/2022 | 03/21/2025 |  | 03/22/2024 |  |  | 10/13/2022 | 03/22/2024 |  |  | 07/31/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023253553 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/13/2022 |  |  |  |  | The user facility reported a 67-year-old female on impella cp for mechanical circulatory support. it was reported that while the patient was on impella cp support, the right lower leg was ischemic due to a smaller vessel size and the medical team decided to explant the impella cp as a result. it was reported the patient experienced bleeding at the impella access site. the preclose that had been deployed at the access site became synched and an additional preclose was deployed successfully. after successfully explanting the impella device, hemostasis was achieved and pulses returned to the right foot. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
29165 | 1220648-2024-08569 | 18958848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18958848&pc=OZD | 07/04/2022 | 03/24/2025 |  | 03/22/2024 |  |  | 07/04/2022 | 03/22/2024 |  |  | 09/30/2023 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022143713 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/04/2022 |  |  |  |  | The user facility reported a 66-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that under echocardiogram it was found that the impella cp was out of position, hugging the posterior wall and possibly intermingled with the mitral valve. the impella cp device was explanted. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29166 | 1220648-2024-08570 | 18958863 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18958863&pc=OZD | 07/06/2022 | 03/22/2024 |  | 03/22/2024 |  |  | 07/06/2022 | 03/22/2024 |  |  | 03/31/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023206483 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/06/2022 |  |  |  |  | The user facility reported a 65-year-old female was implanted with an impella cp device for mechanical circulatory support as part of high risk percutaneous coronary intervention initiated due to elevated pro bnp. the patient experienced bleeding at the access site. pressure was held to the site and heparin was stopped. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29167 | 1220648-2024-08571 | 18958925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18958925&pc=OZD | 07/04/2022 | 03/25/2025 |  | 03/22/2024 |  |  | 07/04/2022 | 03/22/2024 |  |  | 01/31/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023188168 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/04/2022 |  |  |  |  | The user facility reported an impella cp in 66-year-old male patient for mechanical circulatory support. it was reported during impella support, the impella cp migrated out of ventricle. the impella cp device was explanted. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29168 | 1220648-2024-08572 | 18958991 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18958991&pc=OZD | 03/05/2024 | 04/22/2024 |  | 03/22/2024 |  |  | 03/05/2024 | 03/22/2024 |  |  | 11/30/2025 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024433360 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that impella cp support was ongoing for a 58-year-old female patient admitted in acute myocardial infarction (ami). after six days of support, the patient had heparin induced thrombocytopenia (hit) diagnosed and the team removed the heparin and started argatroban. the pump remains on for support. | [5 empty] | The investigation into the thrombocytopenia issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the thrombocytopenia cannot be determined due to insufficient clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29169 | 1220648-2024-08573 | 18958992 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18958992&pc=OZD | 07/09/2022 | 03/22/2024 |  | 03/22/2024 |  |  | 07/09/2022 | 03/22/2024 |  |  | 02/29/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202055 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 07/09/2022 |  |  |  |  | The user facility reported a 50-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. while on support, the patient experienced limb ischemia in one of their legs. it was reported that it was improving; however, no medical or surgical intervention was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
29170 | 1220648-2024-08574 | 18959038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18959038&pc=OZD | 07/03/2022 | 03/25/2025 |  | 03/22/2024 |  |  | 07/03/2022 | 03/22/2024 |  |  | 01/31/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023189166 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/03/2022 |  |  |  |  | The user facility reported a 59-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported during impella support, the patient¿s lost pulses in right lower leg. the right leg became cold and mottled. the impella cp was removed with 2 perclose and the patient regained pulses. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | When making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition. any concomitant medication. vascular size or variations thereof. detailed description of the symptoms experienced by the patient. physical examination findings, including pulse assessment and visual examination of the affected limb. diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. laboratory test results, including blood tests for markers of inflammation or clotting disorders. any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
29171 | 1220648-2024-08575 | 18959042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18959042&pc=OZD | 03/12/2024 | 04/25/2024 |  | 03/22/2024 |  |  | 03/12/2024 | 03/22/2024 |  |  | 11/30/2025 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025454023 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/12/2024 |  |  |  |  | The user facility reported a 64-year-old male was implanted with an impella cp device for mechanical circulatory support. it reported the patient was support developed limb ischemia. the team placed 2nd bypass to the cp limb in a y configuration. the pulses restored and support continued until care was escalated to the 5.5 pump. | [5 empty] | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.
29172 | 1220648-2024-08576 | 18959077 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18959077&pc=OZD | 11/10/2022 | 03/22/2024 |  | 03/22/2024 |  |  | 11/10/2022 | 03/22/2024 |  |  | 08/31/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023263786 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/10/2022 |  |  |  |  | The user facility reported a 59-year-old female patient on an impella cp for mechanical circulatory support. while on support the patient required a femorofemoral bypass due to limb ischemia. a post procedure echocardiogram was completed and revealed good device placement. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29173 | 1220648-2024-08577 | 18959146 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18959146&pc=OZD | 10/23/2022 | 03/31/2025 |  | 03/22/2024 |  |  | 10/23/2022 | 03/22/2024 |  |  | 10/31/2023 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022154191 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/23/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that there was an aorta alarm as the patient was being prepared for transfer to intensive care unit. an echocardiogram then showed that the impella was shallow and in the aorta, at p2 location, so it was advanced at bedside under echocardiogram and placed in p0. a line and knee immobilizer were placed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29174 | 1220648-2024-08578 | 18959166 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18959166&pc=OZD | 03/13/2024 | 04/23/2024 |  | 03/22/2024 |  |  | 03/13/2024 | 03/22/2024 |  |  |  |  | 03/22/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1574757 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/13/2024 |  |  |  |  | No patient involvement: the complainant reported that during an in-service, the aic displayed a ¿self-test failed¿ screen and is now inoperable. there was no patient involvement. | [5 empty] | The investigation for the reported console (aic) display failure has been completed. the console was returned for evaluation. the console was physically inspected and a firmware check was performed, which revealed that there was a lack of communication between the 4-in-1 and the power battery manager channel 1 on the pbm. this lack of communication triggered the self check failure. the pbm pca was physically inspected, revealing corrosion to the circuit traces that carry the communication signals between the 4-in-1 and the battery 1 circuitry on the pbm. this type of corrosion is known to be caused by leakage of the electrolytic supercapacitors c69 and c70 on the pbm. the data logs revealed that preceding the self check failure, there were multiple instances of disruptions in communication between the epc and the power battery manager channel 1. the cause of the aic issue was contamination of the pbm pca due to leakage of the supercapacitors. d4-updated model number. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29175 | 1220648-2024-08580 | 18959237 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18959237&pc=OZD | 10/17/2022 | 03/25/2025 |  | 03/22/2024 |  |  | 10/17/2022 | 03/22/2024 |  |  | 06/30/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023240390 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/17/2022 |  |  |  |  | The user facility reported a 50-year-old male on impella cp for mechanical circulatory support. it was reported that the patient was placed on impella cp support after the patient arrested. the patient developed bleeding at the impella cp insertion site. systemic heparin was withheld due to oozing from multiple sites. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Additional information provided in d6, h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29176 | 1220648-2024-08579 | 18959240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18959240&pc=OZD | 03/13/2024 | 03/03/2025 |  | 03/22/2024 |  |  | 03/13/2024 | 03/22/2024 |  |  |  |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1587053 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/13/2024 |  |  |  |  | Biomed reported; i would like to send in aic (b)(6) for service. the controller can¿t be turned on and is emitting constant muffled beeping noise. | [5 empty] | The impella device was received from the customer on 14-mar-2024 and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - a/c power issues has been completed. the console was for investigation. the reported complaint (unexpected shutdown) was confirmed by review of the data logs but could not be reproduced through testing. the root cause was not conclusively determined. d2 common device name was revised on manufacturer device report 1220648-2024-08579 in accordance with updated procedures. d4 model number was corrected from initial manufacturer device report 1220648-2024-08579 e2 health professional was corrected from initial manufacturer device report 1220648-2024-08579 e3 occupation was corrected from initial manufacturer device report 1220648-2024-08579 f6/f8-should have been left blank on manufacturer device report 1220648-2024-08579 g1 reporting contact email revised on manufacturer device report 1220648-2024-08579 in accordance with updated procedures.
29177 | 1220648-2024-08581 | 18959256 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18959256&pc=OZD | 07/15/2022 | 05/13/2025 |  | 03/22/2024 |  |  | 07/15/2022 | 03/22/2024 |  |  | 03/31/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023211631 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/15/2022 |  |  |  |  | The user facility reported a 44-year-old male was implanted with an impella cp device for mechanical circulatory support due to experiencing chest pain, two episodes of ventricular tachycardia with return of spontaneous circulation, and had cpr administered at home. while on support, the patient experienced repositioning issues two times due to pink urine. with repositioning the impella cp the hematuria cleared up. there were attempts to wean dobutamine; however, dobutamine had to restart due to pressure drop. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29178 | 1220648-2024-08582 | 18959298 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18959298&pc=OZD | 02/27/2024 | 04/08/2024 |  | 03/22/2024 |  |  | 03/06/2024 | 03/22/2024 |  |  | 09/30/2025 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024432483 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/06/2024 |  |  |  |  | The user facility reported a 55-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient had an anticoagulant held due to previous impella access site bleeding which has been resolved. the patient has been reported to be stable. | [5 empty] | The impella device was discarded and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instruction for use for related events states as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was most likely due to anticoagulation management since the bleeding started to resolve once the anticoagulation was held.
29179 | 1220648-2024-08583 | 18959351 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18959351&pc=OZD | 03/13/2024 | 05/01/2024 |  | 03/22/2024 |  |  | 03/13/2024 | 03/22/2024 |  |  |  |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025439555 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Optical Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/13/2024 |  |  |  |  | The user facility reported a 75-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and there were placement signal not reliable alarms. the impella and automated impella controller (aic) were replaced. reportable due to potential patient harm. | [5 empty] | The investigation for the optical issue has been completed. product was not returned for investigation. data logs were returned and reviewed. the 5s parameters of the pump were showing symptoms identical to that of an air gap and a fiber break. the placement signal not reliable (psnr) alarm was seen as soon as the pump was connected. the root cause of the optical signal issue was not determined since product was not returned for investigation and data logs are inconclusive. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk cpi under alarm message summary, table 8.2 automated impella controller alarm messages states the following: action-placement and suction monitoring are suspended 1. monitor patient hemodynamics and impella position with imaging. 2. check the patient cable for kinks. cause-there is a problem with the impella catheter sensor signal. added- d6a/d6b. d4-updated model number. g1-updated mfr site name and address. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29180 | 1220648-2024-08585 | 18959454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18959454&pc=OZD | 07/15/2022 | 03/22/2024 |  | 03/22/2024 |  |  | 07/15/2022 | 03/22/2024 |  |  | 03/31/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023205211 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/15/2022 |  |  |  |  | The user facility reported a 66-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding that was resolved by applying pressure on the hematoma. the patient also had a thrombectomy (thrombus) of the proximal left anterior descending and drug eluding stent was performed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records.
29181 | 1220648-2024-08584 | 18959456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18959456&pc=OZD | 03/13/2024 | 05/01/2024 |  | 03/22/2024 |  |  | 03/13/2024 | 03/22/2024 |  |  |  |  | 03/22/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1150571 |  | 0048-0045-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Optical Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/13/2024 |  |  |  |  | The user facility reported a 75-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and there were placement signal not reliable alarms. the impella and automated impella controller (aic) were replaced. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console (aic) loss of opm data has been completed. the console was returned for evaluation. the console was visually inspected and the failure mode could not be reproduced in the lab without manually inducing an air gap. the front optical connector was inspected and cleaned showing a clean contact surface. a known good pump and purge line was run for about 2 hours without any abnormalities. the root cause for the placement signal not reliable alarm could not be determined. the failure mode could not be reproduced in the lab without manually inducing an airgap. to conform with updated procedures, section d6 & has been revised with updated information.
29182 | 1220648-2024-08586 | 18959513 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18959513&pc=OZD | 03/06/2024 | 11/18/2025 |  | 03/22/2024 |  |  | 03/06/2024 | 03/22/2024 |  |  |  |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1462815 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Female |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller (aic) displayed 'purge cassette not identified" alarm. the team attempted to restart the aic but was unsuccessful. the grey button was unable to be pressed and the purge liquid was visible the whole time at the end of the purge line/motor gap. a second aic was used with the same purge cassette and everything went well, and the impella was successfully implanted. | [5 empty] | The investigation has been completed. data analysis: on 3/6, the purge cassette and disc is inserted before case started is begun. therefore, step 1 is followed by step 3 and proceeds to step 7 at which point the purge cassette is removed. the purge cassette is reinserted but then case start wizard is exited and purge cassette is removed. purge cassette is inserted again but then pump and purge are removed and console is shutdown. purge cassette not identified could not be seen in the logs as purge cassette was read every time as seen in the aic logs. the start button being grayed out could be neither confirmed nor denied but the "imcgui: casestart: msg c:90 e:3011 0x8047bf8 screen curr:8 prev:8 (3009)" should have allowed the user to start the pump. device analysis: the failure mode was not reproduced during testing and console was able to successfully start after case start. the purge cassette was a user misunderstanding and the console successfully recognized purge cassettes. root cause: there was no issue with purge disc/flag found during the case. d2 common device name updated. d10 concomitant medical products and therapy dates updated. g1 reporting contact email updated. | The automated impella controller (aic) was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
29183 | 1220648-2024-08587 | 18959538 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18959538&pc=OZD | 03/13/2024 | 05/01/2024 |  | 03/22/2024 |  |  | 03/13/2024 | 03/22/2024 |  |  | 10/31/2025 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024422165 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 14 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/13/2024 |  |  |  |  | The user facility reported a 14-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient had an attempted impella 5.5 implant via right axillary approach. the impella would not advance and the pump scraped some of the epithelium of the medial aspect of the vessel. a repair patch was applied. no blood products were needed. the pump was going to be rinsed and attempted again however, a clot was observed in the outflow. a new pump was used. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ | The investigation into the delivery issues, the vascular damage - peripheral vessel, and the thrombus issue has been completed since the original report was submitted. the device was not returned by the user facility for investigation. the root cause of the delivery issue was most likely patient condition related to small patient vessels. the recommended vessel size for impella pump is minimum 7mm and patient¿s vessel size was measured 6mm. the root cause of the vascular damage - peripheral vessel issue was most likely patient condition related to small patient vessels. as the necessary information was not provided, the root cause of the thrombus was not determined. potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance". to conform with updated procedures, sections d6 & h10 have been revised with updated information.
29184 | 1220648-2024-08588 | 18959606 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18959606&pc=OZD | 10/18/2022 | 03/10/2025 |  | 03/22/2024 |  |  | 10/18/2022 | 03/22/2024 |  |  | 03/31/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204001 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/18/2022 |  |  |  |  | The user facility reported a 86-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported during the initial access of the right femoral artery, a spiral dissection occurred and was noted on angiogram. the physician stated that the dissection appeared to have self resolved due to natural retrograde flow down the right iliac. patient continued to decline, and bleeding was found from the right iliac vessel. the physician did a balloon tamponade of the right iliac once the bleed was confirmed, and the patient was transfused with 2 units of packed red blood cells (prbc). two covered stents were also placed to the right iliac. the decision was made to discontinue impella support and the patient was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported vascular damage - peripheral vessel issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
29185 | 1220648-2024-08590 | 18959664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18959664&pc=OZD | 11/24/2022 | 03/22/2024 |  | 03/22/2024 |  |  | 11/24/2022 | 03/22/2024 |  |  | 06/30/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023248175 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Pain | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/24/2022 |  |  |  |  | The user facility reported a 41-year-old male patient was implanted with an impella cp for mechanical circulatory support. the patient developed oozing (bleeding), a hematoma and pain at the impella cp insertion site. a bicarb purge was started and the systemic heparin was switched to bival. in addition, it was suspected that the patient was experiencing hemolysis. the impella was repositioned when the echocardiogram indicated the device was more posterior. it was later decided that the patient would need to be escalated to the impella 5.5 for more support. the impella cp was removed and impella 5.5 was placed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29186 | 1220648-2024-08591 | 18959825 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18959825&pc=OZD | 03/13/2024 | 04/23/2024 |  | 03/22/2024 |  |  | 03/13/2024 | 03/22/2024 |  |  |  |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1670339 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/13/2024 |  |  |  |  | The user facility reported a 57-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the automated impella controller (aic) stated that the device was defective and to replace the impella cp. a new impella pump was dropped and prepped. once it was connected to the aic, it was unable to recognize the new impella. the decision was made to replace the aic and once the new aic was connected to the pump, both devices functioned normally. | [5 empty] | The investigation into the aic - pump issue has been completed since the initial report was submitted. the console was returned and tested. the pump receptacle was physically examined, and a piece of plastic debris was located within the pump receptacle. the contamination was physically removed, and the pump was then able to be fully inserted. the root cause was determined to be the contamination within the pump receptacle. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29187 | 1220648-2024-08592 | 18959864 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18959864&pc=OZD | 03/06/2024 | 05/10/2024 |  | 03/22/2024 |  |  | 03/06/2024 | 03/22/2024 |  |  |  |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025449796 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported during impella cp support for 71-year-old female patient, the patient underwent complex ventricular tachycardia ablation. there was challenging access and the activated clotting time (act) during the procedure was around 600 seconds. prior to the repo sheath exchange, act was down to 350 seconds. fifteen minutes post sheath change, there was bleeding noted from the access site. bleeding was attempted to be managed with manual compression and pressure bandage. however, the pump was ultimately explanted since the bleeding continued. the access site required surgical closure to control bleeding. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
29188 | 1220648-2024-08593 | 18960002 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18960002&pc=OZD | 07/18/2022 | 03/22/2024 |  | 03/22/2024 |  |  | 07/18/2022 | 03/22/2024 |  |  | 02/29/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195418 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/18/2022 |  |  |  |  | The user facility reported a 72-year-old male was implanted with an impella cp device for mechanical circulatory support and that the patient developed bleeding and hematoma at the insertion site. the patient had a blood transfusion and surgical wound closure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29189 | 1220648-2024-08594 | 18960009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18960009&pc=OZD | 03/06/2024 | 04/23/2024 |  | 03/22/2024 |  |  | 03/06/2024 | 03/22/2024 |  |  | 01/31/2025 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024326488 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/06/2024 |  |  |  |  | Additional information states that on day 2 of support, patient experienced a graft anastomosis bleed at the access site and was taken back to or for repair. | The user facility reported a 68-year-old female with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. after 18 days of being admitted, the pump stopped for a total of three times on that day. the pump was then explanted and an intra-aortic balloon pump was implanted. |  |  |  |  | The investigation for the reported access site bleed and the pump stop has been completed. impella cp was returned for evaluation. upon visual inspection, a large purge gap was found. data logs were reviewed and a high motor current pump stop with alarm #13 occurred. multiple attempts were made to restart pump without success. clinical details noted that clinical team attempted to restart the pump on new console without success. the root cause of the pump stop was due to bearing wear as pump was run past indicated design life. clinical details noted that cp pump was inserted via axillary graft from failed 5.5 insertion attempt and on day 2 of support graft anastomosis bleed was occurring at access site. patient was taken to or to repair bleed. the root cause of the access site bleeding - major was a use issue as patient had to be taking to the or to repair anastomosis. added-b1 selected adverse event; b2 selected required intervention to prevent damage; b5 addl event narrative; h6 clinical code 1888; h6 impact code 4625 | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. the instructions for use state the following concerning pump stopping: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
29190 | 1220648-2024-08595 | 18960076 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18960076&pc=OZD | 03/13/2024 | 05/02/2024 |  | 03/22/2024 |  |  | 03/13/2024 | 03/22/2024 |  |  | 11/30/2025 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024433112 |  | 0042-0040-US |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation Problem | [6 empty] | Hemolysis | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/13/2024 |  |  |  |  | The user facility reported a 57-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the automated impella controller (aic) stated that the device was defective and to replace the impella cp. a new impella pump was dropped and prepped. once it was connected to the aic, it was unable to recognize the new impella. the decision was made to replaced the aic and once the new aic was connected to the pump, both devices functioned normally. the pump is being sent back for investigation. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported pump not starting has been completed. the pump was not returned for evaluation. however, data logs were reviewed which showed the left logs from the case do not show any positioning issues of the pump. ¿impella position in aorta¿ alarms were triggered at the beginning of the case but were resolved before hemolysis occurred. no suction alarms were triggered, no motor current spikes were observed. clinical details note that the impella pump but plugged into the aic but was unable to detect pump. aic was switched out and new aic was able to detect pump without issue. a separate investigation was conducted for the aic and it was found that there was contamination of the pump receptacle that caused the "impella defective" alarms and pump not detected alarms. no problem was found with the pump as the pump not detected was due to an issue with the aic pump receptacle . clinical details noted that patients free hemoglobin (pfhg) was elevated on day 2 of support, pump was repositioned and correct positioning was checked with imaging. on day 4 of support, pfhg was noted to be trending down. no indications of any hemolysis troubleshooting techniques were performed. the root cause of the hemolysis cannot be determined as insufficient clinical information was provided and data logs did not indicate any positioning issues. added- h6 clinical code 1886. d4-updated model number. added-d6a/d6b.
29191 | 1220648-2024-08596 | 18960125 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18960125&pc=OZD | 03/13/2024 | 04/12/2024 |  | 03/22/2024 |  |  | 03/13/2024 | 03/22/2024 |  |  |  |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025449575 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/13/2024 |  |  |  |  | The user facility reported a 64-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support due to having electrocardiogram changes. while on support, the patient had access site bleeding and a hematoma. the physician re-sutured with forward tension, tightened the perclose a little, and put a light pressure dressing. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical details provided was sufficient to determine the root cause. the root cause of the access site bleeding was most likely due to anticoagulation management since the act at the time of bleed was 311 which above the limit range.
29192 | 1220648-2024-08597 | 18960190 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18960190&pc=OZD | 03/13/2024 | 04/18/2024 |  | 03/22/2024 |  |  | 03/13/2024 | 03/22/2024 |  |  |  |  | 03/22/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1211523 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Electrical /Electronic Property Problem | Computer Operating System Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/13/2024 |  |  |  |  | After impella insertion and successfully started support, yellow alarm for console failure occurred. staff changed the controller. | [5 empty] | The automated impella controller device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - controller error (yellow alarm) issue has been completed since the initial report was submitted. the console was returned and tested but was unable to boot up. after further testing, it was discovered that the root cause for the controller error was a defective impellatronic board drawing current beyond the normal threshold.
29193 | 1220648-2024-08598 | 18960209 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18960209&pc=OZD | 07/24/2022 | 03/27/2025 |  | 03/22/2024 |  |  | 07/24/2022 | 03/22/2024 |  |  | 03/31/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023205058 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/24/2022 |  |  |  |  | The user facility reported a 50-year-old male patient was on impella cp support due to cardiac arrest and ventricular fibrillation. while on impella cp support, the patient experienced increased suction alarms. the patient was placed at p2 level and received additional volume. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Additional information provided in d6, h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29194 | 1220648-2024-08599 | 18960227 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18960227&pc=OZD | 12/03/2022 | 05/14/2025 |  | 03/22/2024 |  |  | 12/03/2022 | 03/22/2024 |  |  | 04/30/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023221032 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/03/2022 |  |  |  |  | The user facility reported that the decision for impella cp implant was made when a 79-year-old female lhc showed severe microvascular disease. during transport and on impella support, the patient became agitated requiring intubation. it was reported access site bleeding after agitation. femstop was removed and the nurse redressed and performed angle matching, which alleviating bleeding and heparin was discontinued. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29195 | 1220648-2024-08600 | 18960232 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18960232&pc=OZD | 03/13/2024 | 03/10/2025 |  | 03/22/2024 |  |  | 03/13/2024 | 03/22/2024 |  |  | 10/31/2025 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024418156 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | Thrombosis/Thrombus | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [10 empty] | Required Intervention | Death | Other | [22 empty] | A |  |  |  | Initial submission | Followup | Followup | Followup |  | 03/13/2024 |  |  |  |  | It was noted the patient care was withdrawn and the patient expired while on impella cp support. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the death outcome. the patient's peri-procedural condition was critical. | The user facility reported a 67-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient had impella cp support due to having a papillary muscle rupture. while on support there were persistent suction alarms. the pump was unable to flow higher than p-1. the pump was explanted and it was noted that biomaterial was present on the inlet cage. the pump was discarded. reportable due to low pump flows. | Additional information received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured worsening of cardiogenic shock with a "remote (unlikely) related" relationship to the impella device used: ¿after impella not flowing, map trending down, pressors requirement increasing. decision to remove impella cp and place iabp. after returning from cath lab patient with poor prognosis, as per cts and hf cardiac support should be withdrawn, patient's son agreed. ecmo and iabp turned off by perfusion. extubated by rt. patient declared deceased at 11:23 pm.¿. the patient outcome was fatal. |  |  |  | The investigation for the reported low pump flow has been completed. the impella device was not returned for investigation. data log were reviewed finding motor current spikes and many suction alarms occurred during impella support and the time of flow issue occurred. the cause of low pump flow most likely was biomaterial ingestion. added- h6 impact code 4627 d4-updated model number g1-updated mfr site name and address. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08600 was provided in sections b1, b2, b5, h1, and h6. | The impella device was reportedly discarded by the customer and therefore,¿an evaluation of the device will not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | Complaint device not returned for investigation. data log reviewed: suction alarms occurred during support. no mc spike observed at the time low flow issue reported. a few quick resolved impella position alarms in ventricle observed. the root cause of the low pump flow issue cannot be determined since the complaint device not available for analysis and insufficient data provided. as the necessary information was not provided, the root cause of the vt/vf was not determined. the following information was updated on manufacturer device report 1220648-2024-08600 to reflect the additional information that was received. b2: updated to death and date of death to reflect additional information received. b5: updated with additional information received h1: type of report updated to death. h6: health effect - impact code 1802 added.
29196 | 1220648-2024-08601 | 18960382 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18960382&pc=OZD | 10/21/2022 | 05/16/2025 |  | 03/22/2024 |  |  | 10/21/2022 | 03/22/2024 |  |  | 07/31/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023255929 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Improper Flow or Infusion | Infusion or Flow Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/24/2022 |  |  |  |  | The user facility reported a 44-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced bleeding at the access site. the patient was transfused with blood products due to the ongoing bleeding at the access site. it was reported the impella cp had alarms for suction, which were resolved after the patient began extracorporeal membrane oxygenation. it was reported a moderate thrombus was found in multiple coronary vessels and the patient was administered systemic integrilin. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29197 | 1220648-2024-08602 | 18960410 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18960410&pc=OZD | 03/06/2024 | 09/29/2025 |  | 03/22/2024 |  |  | 03/06/2024 | 03/22/2024 |  |  |  |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025442096 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/06/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. during support, a sudden change in placement signal with the impella 5.5. placement signal began rapidly increasing and becoming sporadic. additional information received states that no intervention was needed and that the impella was still active and working. | [5 empty] | The investigation for the reported optical issues has been completed. the impella 5.5 was returned for investigation. the pump was visually inspected finding no broken blades or cannula kinks. borescope view obtained through cannula found no biomaterial from the top view of impeller. bottom view of impeller with borescope confirmed biomaterial wrapped beneath impeller wings at outflow. the optical bench 5s parameters were verified and observed to be in normal ranges. the light continuity test passed, and no optical signal issues found. no other abnormalities relative to the pump seen. data logs were reviewed and upward shift in motor current and placement signal with p-level changing per auto suction response and low flows observed. there was no optical signal issue found per logs and therefore an optical issue could not be confirmed. the cause of the low pump flow issue (as seen on data logs) was biomaterial per return product investigation and log review. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Corrections that were made from the initial manufacturer device report number 1220648-2024-08602. f6 and f8 should not have had entries in the initial report. g1 reporting contact email updated.
29198 | 1220648-2024-08603 | 18960537 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18960537&pc=OZD | 03/06/2024 | 04/16/2024 |  | 03/22/2024 |  |  | 03/06/2024 | 03/22/2024 |  |  | 11/30/2025 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024435069 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Improper Flow or Infusion | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/06/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. following implant of the impella 5.5 pump, the device immediately experienced a spike in motor current coinciding with flow saturation. a thrombus was noted in the left ventricle and a clot was subsequently seen entrapped in the device. the pump was removed and replaced as a result of the failure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the saturated flow and the thrombus issues has been completed since the original report was submitted. the device was not returned for investigation. clinical details noted that clot was observed on pump upon explant and lv thrombus was found on echo. the root cause of the saturated flow was due to biomaterial ingestion.
29199 | 1220648-2024-08604 | 18960540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18960540&pc=OZD | 07/15/2022 | 09/13/2024 |  | 03/22/2024 |  |  | 07/15/2022 | 03/22/2024 |  |  | 04/30/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023214977 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 07/15/2022 |  |  |  |  | The user facility reported a 75-year-old patient needing mechanical circulatory support. it was reported that the patient was placed on impella cp support and it was noted that the patient was bleeding at all sites and the bleeding began to slow when heparin was stopped. patient was found to have a brain bleed. family decided to withdraw care. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08604 was provided.
29200 | 1220648-2024-08605 | 18960638 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18960638&pc=OZD | 08/03/2022 | 04/17/2025 |  | 03/22/2024 |  |  | 08/03/2022 | 03/22/2024 |  |  | 03/31/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023214636 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/03/2022 |  |  |  |  | The user facility reported a 83-year-old female noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. when gaining access for single access, the sheath cracked, so it was removed and a new sheath was used. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29201 | 1220648-2024-08606 | 18960672 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18960672&pc=OZD | 10/17/2022 | 03/31/2025 |  | 03/22/2024 |  |  | 10/17/2022 | 03/22/2024 |  |  | 07/31/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023253802 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Difficult to Insert | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/17/2022 |  |  |  |  | The user facility reported a 60-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, single access was unsuccessful due to not being able to advance wire up femoral artery and the surgeon decided to take the patient to or for a coronary artery bypass graft (cabg). the impella cp remained in and running during cabg and an aortic repair for a known an ascending thoracic aneurysm.. it was also reported of hemolysis present. echocardiography was performed at bedside and the surgeon repositioned the impella cp. the impella cp device was weaned and removed prior to clearing of hemolysis. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29202 | 1220648-2024-08607 | 18960723 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18960723&pc=PYX | 03/06/2024 | 03/22/2024 |  | 03/22/2024 |  |  | 03/06/2024 | 03/22/2024 |  |  |  |  | 03/22/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1143007 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 03/06/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. while the patient was on support, the customer was removing the purge cassette, purge disc, and the blue retainer that contains the purge disc popped off the automated impella controller. the device was swapped out. there was no patient harm. | [5 empty] | The investigation into the automated impella controller (aic) broken purge disc retainer has been completed. the product was returned. logs were irrelevant to this complaint as the clinical staff only reported that the consoles purge retainer broke off during removal of purge cassette and disc. the console was tested. upon examining for any visible defects, it was evident that the blue purge disc retainer was broken. the root cause of this issue was a broken blue purge disc retainer.
29203 | 1220648-2024-08608 | 18960732 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18960732&pc=OZD | 03/13/2024 | 04/26/2024 |  | 03/22/2024 |  |  | 03/13/2024 | 03/22/2024 |  |  | 10/31/2025 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024420573 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/13/2024 |  |  |  |  | The user facility reported a 84-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp being placed to support a patient coding in cath lab. the team found the pump could not deliver across aorta and to the lv (left ventricle) . the pump implant was aborted. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, the impella was unable to pass through as aorta was extremely tortuous and aneurysmal as per the clinical description. potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ d4-updated model number in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29204 | 1220648-2024-08609 | 18960762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18960762&pc=OZD | 03/13/2024 | 04/30/2024 |  | 03/22/2024 |  |  | 03/13/2024 | 03/22/2024 |  |  |  |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2025436750 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/13/2024 |  |  |  |  | The user facility reported a 77-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that impella 5.5 insertion was attempted but the outflow would not pass area of suspected innominate artery calcification despite significant push force and wire tension. the case was aborted. | [5 empty] | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue was most likely due to patient condition since attempted insertion and outflow would not pass area of suspected innominate artery calcification/narrowing despite significant push force and wire tension. potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29205 | 1220648-2024-08610 | 18960787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18960787&pc=OZD | 07/05/2022 | 03/26/2025 |  | 03/22/2024 |  |  | 07/05/2022 | 03/22/2024 |  |  | 12/31/2023 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023179466 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/05/2022 |  |  |  |  | The user facility reported a 54-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the pump experienced delivery issues while insertion was attempted into the left groin. this was resolved by insertion into the right groin. additionally, while on support, the patient experienced access site bleeding which required an additional suture. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29206 | 1220648-2024-08611 | 18960825 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18960825&pc=OZD | 03/13/2024 | 04/24/2024 |  | 03/22/2024 |  |  | 03/13/2024 | 03/22/2024 |  |  | 10/31/2025 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024415807 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 13 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/13/2024 |  |  |  |  | The user facility reported a 13-year-old male patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it reported that while on support the patient developed a thrombus in the left ventricle. the staff switched the systematic anticoagulation from heparin to angiomax. small clots were still present on echocardiogram. additionally, the patient tested positive for heparin induced thrombocytopenia. the patient had bival in the systemic and bicarb in the purge. the patient continued with support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the thrombocytopenia and the thrombus issues has been completed since the original report was submitted. the device was not returned for investigation. the cause of the thrombocytopenia issue was not determined since product was not returned and due to insufficient clinical information. as the necessary information was not provided, the root cause of the thrombus was not determined.
29207 | 1220648-2024-08622 | 18961292 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18961292&pc=OZD | 03/03/2024 | 04/22/2024 |  | 03/22/2024 |  |  | 03/06/2024 | 03/22/2024 |  |  | 09/30/2025 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024411380 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  | White | Malfunction | No | Yes | Fluid/Blood Leak | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support for a male patient, the automated impella controller (aic) had low purge flow alarms, and on inspection the purge side arm was leaking. the impella cp was explanted. there was no patient harm. | [5 empty] | The investigation for the reported purge leak and high purge pressure has been completed. no product was returned for evaluation so the cause of the damage to the sidearm cannot be determined. data logs were not returned for evaluation. clinical details noted that high purge flow alarms were triggered on aic and upon inspection, sidearm was found to be damaged and leaking. the root cause of the purge leak was most likely due to the damaged sidearm that was the source of the leak. the cause of the high purge pressure was likely due to damaged side arm. however, the cause of the damage is unknown as the product was not returned. the instructions for use states ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
29208 | 1220648-2024-08623 | 18961352 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18961352&pc=OZD | 08/28/2022 | 03/10/2025 |  | 03/22/2024 |  |  | 08/28/2022 | 03/22/2024 |  |  | 01/31/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023192912 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/28/2022 |  |  |  |  | The user facility reported a 84-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced vascular damage that necessitated the impella cp to be repositioned. heparin was discontinued due to the patient still experiencing blood loss, the patient continued to have a drop in pressure, small boluses of epi were administered in addition, the patient also received blood. however, the patient continue to bleed and heparin dosages were discontinued. therefore, the decision was made to have the pump removed. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported vascular damage - peripheral vessel and positioning issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
29209 | 1220648-2024-08624 | 18961368 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18961368&pc=OZD | 03/14/2024 | 03/22/2024 |  | 03/22/2024 |  |  | 03/14/2024 | 03/22/2024 |  |  |  |  | 03/22/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1133147 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | Detachment of Device or Device Component | No Apparent Adverse Event | Unexpected Shutdown |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 03/14/2024 |  |  |  |  | The user facility biomed department reported the automated impella controller (aic) had a controller error. the aic was removed from service. no patient involvement or injury was reported. | [5 empty] | Data analysis: console logs from the reported day of event are consistent with complaint and show that console presents with a controller error (#1035, due to opm comm stopped) immediately after boot-up, upon each boot-up cycle. device analysis: upon visual inspection of the console revealed disconnected power cable from optical bench, which was caused by compromised cable tie wrap for that cable (material degraded and fell apart). when correctly installed, the cable tie wrap prevents the attached ferrite from pulling the cable out of a socket, under its weight. repair: install cable tie wrap on cable (or replace cable assembly if needed) and perform full functional check of the console. the root cause of the automated impella controller (aic) issue was a loose connector.
29210 | 1220648-2024-08625 | 18961392 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18961392&pc=OZD | 03/15/2024 | 05/02/2024 |  | 03/22/2024 |  |  | 03/15/2024 | 03/22/2024 |  |  |  |  | 03/22/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1422863 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/15/2024 |  |  |  |  | The user facility reported a 58-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. complainant in japan reported the automated impella controller (aic) had a purge pressure transmitter installation failure alarm and priming was not possible. the aic was replaced. | [5 empty] | The impella device was received from the customer on (b)(6) 2024 and therefore,¿an evaluation of the device was is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console purge issues has been completed. the console was returned for evaluation. the logs were reviewed which confirmed the inability to detect the purge disc. the console was evaluated and visual inspection of the console reveals a missing purge flag and a crack in the purge retainer. the root cause for the purge disc not detected alarms was a defective purge pressure sensor due to missing purge flag. the retainer also had a crack. d4-updated model number. to conform to updated procedures, section d6 was revised with updated information.
29211 | 1220648-2024-08626 | 18961633 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18961633&pc=PYX | 03/06/2024 | 04/24/2024 |  | 03/22/2024 |  |  | 03/06/2024 | 03/22/2024 |  |  | 08/31/2025 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024404274 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Inaccurate Delivery | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient on bipella support with an impella 5.5 and was delivering the impella rp for support when the delivery failed. the impella rp was replaced by an impella rp flex which was able to be delivered. the patient had known comorbidities and was turned down for surgery. the rp flex is supporting the patient successfully to date. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation into the reported issue has been completed since the original report was filed. the pump was returned, tested, and evaluated. the root cause of the positioning issue was most likely due to use issue since despite multiple attempts, the device could not be implanted due to malpositioning. no issues were found with the returned product.
29212 | 1220648-2024-08627 | 18961817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18961817&pc=OZD | 03/06/2024 | 03/22/2024 |  | 03/22/2024 |  |  | 03/06/2024 | 03/22/2024 |  |  |  |  | 03/22/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1345073 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during prep, the pressure transducer of the automated impella controller (aic) did not detect the purge disc. there were no alarms. therefore, the aic was exchanged and the same purge disc was inserted into another aic without issue and worked. there was no reported patient harm. | [5 empty] | The investigation has been completed. data analysis showed the console repeatedly was unable to detect purge discs after a cassette change. the complaint description was accurate in the purge disc not detected but the "no alarms" described, was inaccurate as otherwise described in the complaint and seen in the logs. the console was tested and was missing the purge flag. the cause of the issue was a missing purge flag.
29213 | 1220648-2024-08628 | 18961909 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18961909&pc=OZD | 03/06/2024 | 04/18/2024 |  | 03/22/2024 |  |  | 03/06/2024 | 03/22/2024 |  |  | 11/30/2025 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024424829 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the impella cp was having continuous suction at p2 and would not flow more than 0.6 liters per minute. the impella was therefore explanted. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow issue has been completed since the original report was submitted. the device was not returned for investigation. per the data log, the root cause was most likely patient condition related as continuous suction was observed after patient was escalated to ecmo.
29214 | 1220648-2024-08629 | 18962006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18962006&pc=OZD | 09/08/2022 | 09/13/2024 |  | 03/22/2024 |  |  | 09/08/2022 | 03/22/2024 |  |  | 01/31/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023183172 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Death | No | Yes | Activation, Positioning or Separation Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/08/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. several attempts to place the impella 5.5 pump were made however, the device would not pass through axillary artery. the decision was made to abort 5.5 due to anatomy and to place impella cp. | Procedural outcome noted that the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | B2: date of death is unknown. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08629 was provided in sections b2, b5, h1, and h6.
29215 | 1220648-2024-08630 | 18962016 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18962016&pc=OZD | 03/11/2024 | 04/18/2024 |  | 03/22/2024 |  |  | 03/11/2024 | 03/22/2024 |  |  | 11/30/2025 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024428169 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/11/2024 |  |  |  |  | The user facility reported a 59-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed hemolysis, after difficult positioning. the impella cp was replaced. it was reported that during the explant the pigtail was possibly seen under the pap muscle. no patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the reported hemolysis has been completed. the impella device was not returned for evaluation. data logs confirmed pump was in improper position corresponding with clinical description that triggered alarms impella position in ventricle/aorta. no other issues were found on the logs. based on clinical description and data analysis, the root cause of hemolysis was most likely due to pump was not in the proper position.
29216 | 1220648-2024-08631 | 18962035 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18962035&pc=OZD | 03/15/2024 | 05/02/2024 |  | 03/22/2024 |  |  | 03/15/2024 | 03/22/2024 |  |  | 07/31/2025 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024377027 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/15/2024 |  |  |  |  | The user facility reported a 62-year-old female undergoing coronary artery bypass grafting was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the impella had a sudden motor spike that led to saturated flows of 4.4l/min at p4. the left ventricle signal rose and is above the placement signal. purge flows and purge pressures are normal and unchanged. it was further reported that the patient was weaned to p2 in preparation for explant due to the saturated pump flows and rising motor current. the pump motor stopped eventually but restarted at p2 until surgeon arrived for emergent explant at bedside. the pump was removed without issues and the patient is stable. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use for the related event are as follows: maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop and the saturated flow issues has been completed since the original report was filed. the pump was returned and evaluated. upon visual inspection, large clots of biomaterial were found in outflow cage on impeller blade, within inflow cage, and within the pump cannula. the root cause of the saturated flow and the pump stop was due to biomaterial build-up on the impeller blade due to insufficient anticoagulation while on support. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ to conform to updated procedures, sections d6 & h10 were revised with updated information.
29217 | 1220648-2024-08632 | 18962046 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18962046&pc=OZD | 09/13/2022 | 04/07/2025 |  | 03/22/2024 |  |  | 09/13/2022 | 03/22/2024 |  |  | 06/30/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235910 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/13/2022 |  |  |  |  | The user facility reported a 74-year-old female on impella cp for mechanical circulatory support. while on support, the patient experienced no pedals or doppler pulses and urine started turning pink. femoral a-line placed. the staff ultimately removed the impella cp successfully and without complication. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Section d6a / d6b: added implant and explant dates the cause of the limb ischemia was not determined since product was not returned and all the necessary information was not provided.
29218 | 1220648-2024-08633 | 18962114 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18962114&pc=PYX | 03/15/2024 | 05/03/2024 |  | 03/22/2024 |  |  | 03/15/2024 | 03/22/2024 |  |  | 02/28/2025 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP FLEX W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2024333411 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/15/2024 |  |  |  |  | The user facility reported a 39-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the us complainant had a bipella support ongoing with a cp-rp flex for the postop patient. the coronary artery bypass graft was done surgically. the rp flex was supporting for 23 hours when explanted for signs of hemolysis. the hemolysis was observed by urine color change and ldh labs. | [5 empty] | The investigation into the hemolysis has been completed since the original report was submitted. the device was not returned by the user facility for investigation. data logs were reviewed and motor current spike with an upward shift in placement signal and speed deviation while in p-9 was discovered. purge pressure also increases, and flow dropped slightly. patient was also concern for hit and a hypercoagulable state, thus act not maintained throughout support. the cause of the hemolysis issue most likely was biomaterial ingestion. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29219 | 1220648-2024-08634 | 18962384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18962384&pc=OZD | 09/14/2022 | 04/01/2025 |  | 03/22/2024 |  |  | 09/14/2022 | 03/22/2024 |  |  | 01/31/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023182477 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/14/2022 |  |  |  |  | The user facility reported an 83-year-old male on impella cp for mechanical circulatory support. during insertion they were unable to advance wire, dissection noted. access aborted, perclose deployed, and pressure held. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29220 | 1220648-2024-08635 | 18962643 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18962643&pc=OZD | 03/15/2024 | 05/02/2024 |  | 03/22/2024 |  |  | 03/15/2024 | 03/22/2024 |  |  | 10/31/2025 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024409328 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | No | Other | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/15/2024 |  |  |  |  | The user facility reported a 80-year-old male in acute myocardial infarction cardiogenic shock was implanted with an impella device for mechanical circulatory support. the patient had an attempted impella cp implant due to the patient going into cardiac arrest. the doctor attempted to cross the atrioventricular valve, but was unable due to aortic stenosis. the doctor decided to fix 100% proximal left anterior descending artery and 99% proximal circ first and attempt the impella insertion after they opened the arteries. on the second attempt, they were able to cross the valve and gain access with 0.018 wire. impella was inserted over the 0.018 wire, but it was unable to be advanced past prior abdominal aortic aneurysm (aaa) repair. the doctor decided to abort the impella insertion at the risk of damaging the aaa or harming the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, as the impella was unable to pass through due to aortic stenosis as per the clinical description provided. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ d4-updated model number. to conform to updated procedures, sections d6 & h10 were revised with updated information.
29221 | 1220648-2024-08636 | 18962791 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18962791&pc=PYX | 03/16/2024 | 05/03/2024 |  | 03/22/2024 |  |  | 03/16/2024 | 03/22/2024 |  |  | 06/30/2025 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC |  |  | US |  | 2024369391 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Positioning Problem | Material Split, Cut or Torn |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/16/2024 |  |  |  |  | The user facility reported a 76-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. it was reported that the physician dilated the patient using all supplied dilators but was unable to advance the sheath over the wire. the wire and the sheath were exchanged. the doctor noted that the tip of the dilator was damaged. the tip was split. additionally, the doctor attempted to remove the introducer and pump migrated into the right ventricle during each attempt. the introducer was left on, and the pump position became questionable. | [5 empty] | The investigation into the reported issues has been completed since the original report was submitted. the device was not returned by the user facility for investigation. as per the clinical, the doctor had to use excessive force to advance the sheath which led to the tip split. the root cause of the introducer damage issue was a sheath tip split caused most likely due to excessive force. the root cause of the positioning issue was use related to improper technique. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29222 | 1220648-2024-08637 | 18963062 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18963062&pc=OZD | 03/17/2024 | 04/19/2024 |  | 03/22/2024 |  |  | 03/17/2024 | 03/22/2024 |  |  |  |  | 03/22/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1200770 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/17/2024 |  |  |  |  | The user facility reported a 48-year-old male with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that during the support there was an issue with the automated impella controller (aic). the aic was not recognizing the purge cassette. the aic was exchanged and no patient harm was reported and continued on support. | [5 empty] | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were reviewed which confirmed the inability to detect the purge disc. the console was evaluated and it was confirmed that the purge disc retainer had broken off on the console. the root cause of this issue was a broken purge disc retainer. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
29223 | 1220648-2024-08638 | 18963484 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18963484&pc=PYX | 03/18/2024 | 05/04/2024 |  | 03/22/2024 |  |  | 03/18/2024 | 03/22/2024 |  |  |  |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP FLEX W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2025437984 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/18/2024 |  |  |  |  | The user facility reported a 70-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a 23 fr sheath being placed for a rp delivery for support of a patient admitted preop for a mitral valve repair . the sheath became kinked. the sheath was replaced and the rp was delivered successfully for the surgical case. no patient harm was reported. | [5 empty] | The investigation into the introducer damage ¿ mechanical issue has been completed since the original report was submitted. the device was not returned by the user facility for investigation. clinical details noted that physician had difficulty with dilation at 23fr and the sheath became kinked and had to be removed. sheath replaced with 24fr dry seal dilator and new sheath was placed successfully. additionally, it was noted that the patient had a short neck and was obese. the root cause of the introducer damage was most likely a sheath kink due to patient's anatomy. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29224 | 1220648-2024-08639 | 18963595 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18963595&pc=OZD | 12/29/2022 | 03/22/2024 |  | 03/22/2024 |  |  | 01/12/2023 | 03/22/2024 |  |  | 04/30/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 - JP |  |  | ABIOMED, INC. |  |  | US |  | 2023251517 |  | 1000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/12/2023 |  |  |  |  | The complainant reported a patient (unknown age and gender) was implanted with an impella 5.5 device for mechanical circulatory support for unknown primary indication. the pump implant was aborted after a kink was noted following insertion and an impella cp was implanted in its stead. there was no harm to the patient. | [5 empty] | The investigation into the delivery issues has been completed. no impella products were returned for evaluation. the cause of the delivery issue was most likely a catheter kink due to improper delivery technique.
29225 | 1220648-2024-08640 | 18963662 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18963662&pc=OZD | 04/13/2023 | 03/22/2024 |  | 03/22/2024 |  |  | 04/14/2023 | 03/22/2024 |  |  | 08/31/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2023256201 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/14/2023 |  |  |  |  | The complainant reported a patient (unknown age and gender) undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. upon the initial delivery attempt, it was noted that the pump failed to advance in the iliac due to stenosis. a second attempt was made from below the collarbone and it was observed that the shaft of the catheter had kinked. the pump was replaced as a result of the noted damage. the device was discarded. there was no patient harm | [5 empty] | The investigation into the delivery issues has been completed. the impella product was not returned for review. based on clinical description, the root cause of delivery issue was most likely due to patient condition of diseased vessels.
29226 | 1220648-2024-08641 | 18963689 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18963689&pc=OZD | 12/02/2022 | 06/24/2024 |  | 03/22/2024 |  |  | 12/03/2022 | 03/22/2024 |  |  | 04/30/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 SMARTASSIST SET, JP |  |  | ABIOMED, INC |  |  | US |  | 2023255863 |  | 1000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/03/2022 |  |  |  |  | The user facility reported a 65-year-old male with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient suffered a cerebral hemorrhage during impella 5.5 support. anticoagulation was discontinued in response to the condition. as the patient's hemodynamics had improved, the decision was subsequently made remove the pump. per the doctor, as heparin had been utilized at the time of the noted condition, a causal relationship between pump support and the cerebral hemorrhage could not be ruled out. | [5 empty] | The investigation into the stroke/cva issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided, the root cause of the stroke/cva was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
29227 | 1220648-2024-08642 | 18963768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18963768&pc=OZD | 06/28/2022 | 03/22/2024 |  | 03/22/2024 |  |  | 06/28/2022 | 03/22/2024 |  |  | 03/31/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023212190 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Female |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/28/2022 |  |  |  |  | The user facility reported a female of unknown age was implanted with an impella cp device for mechanical circulatory support. while on support, the impella alarmed continuous suction alarms. as a result of the noted issue, p-levels were reduced and administration of one bolus to resolve the alarm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29228 | 1220648-2024-08643 | 18964166 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18964166&pc=OZD | 10/02/2023 | 05/21/2024 |  | 03/22/2024 |  |  | 02/21/2024 | 03/22/2024 |  |  | 05/31/2025 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024369481 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Pseudoaneurysm | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 02/21/2024 |  |  |  |  | Additional information received a notification from our medical safety team via loqi (abiomed¿s long-term outcome and quality indicator impella registry), pertaining to the above case. it was noted that recurrent non-sustained ventricular tachycardia with a definite relationship to the impella cp device used on this patient. patient was hospitalized to address the issue. | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry regarding a 51 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, this patient endured a pseudoaneurysm of right groin involving right superficial femoral artery after reporting pain at the impella insertion site. intervention via an ultrasound-guided thrombin injection of the right groin was required. |  |  |  |  | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined. the instructions for use referenced in the initial report is from ifu for impella cp with smart assist system sections: general patient care considerations, entering cardiac output, and potential adverse events (united states), which now includes "cardiac or vascular injury (including ventricular perforation.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The investigation for the reported vascular damage has been completed. the impella device was not returned for evaluation. the root cause of the vascular damage - peripheral vessel issue was not determined since product was not returned for investigation and insufficient clinical details were provided related to the insertion and the failure observed. the instructions for use states ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ d4-updated model number.
29229 | 1220648-2024-08461 | 18964177 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18964177&pc=OZD | 10/21/2022 | 04/02/2025 |  | 03/22/2024 |  |  | 10/21/2022 | 03/22/2024 |  |  | 06/30/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023238092 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/21/2022 |  |  |  |  | The user facility reported a 52-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding at the groin insertion site. the patients heparin was held to resolve the bleeding. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
29230 | 1220648-2024-08621 | 18964208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18964208&pc=OZD | 03/14/2024 | 04/29/2024 |  | 03/22/2024 |  |  | 03/14/2024 | 03/22/2024 |  |  |  |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1456617 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/14/2024 |  |  |  |  | The user facility reported a 70-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the automated impella controller (aic) would not detect the purge discs. staff said this occurred a few times and they switched the purge cassette out to no avail. the aic was replaced successfully. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the initial report was submitted. the console was returned and tested. the purge disc not detected alarm issue was able to be reproduced, and purge disc retainer was found to be broken. the root cause of this issue was a broken purge disc retainer. portion of the purge retainer that is supposed to be fastened to the purge assembly was no longer intact. | The impella device was received from the customer and upon completion of an evaluation of the device, a supplemental mdr will be filed.
29231 | 1220648-2024-08617 | 18964209 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18964209&pc=OZD | 08/06/2022 | 03/22/2024 |  | 03/22/2024 |  |  | 08/06/2022 | 03/22/2024 |  |  | 04/30/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023221613 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/06/2022 |  |  |  |  | The user facility reported a 72-year-old male was implanted with an impella cp device for mechanical circulatory support. access site bleeding was noted upon pump implantation, which was treated by giving the patient platelets. additionally, the patient was found to have heparin-induced thrombocytopenia. this was managed by switching and staying on purge solution with 5 percent dextrose water. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29232 | 1220648-2024-08620 | 18964213 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18964213&pc=OZD | 08/11/2022 | 03/10/2025 |  | 03/22/2024 |  |  | 08/11/2022 | 03/22/2024 |  |  | 05/31/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023217271 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Female |  |  |  | Injury | Yes | No | Positioning Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/11/2022 |  |  |  |  | The user facility reported a 56-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that there was access site bleeding at the groin. this was treated by reducing heparin. additionally, the complainant reported that the pump slipped into the aorta. consequently, the device was explanted. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29233 | 1220648-2024-08614 | 18964214 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18964214&pc=PYX | 03/06/2024 | 04/23/2024 |  | 03/22/2024 |  |  | 03/06/2024 | 03/22/2024 |  |  | 11/30/2025 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024435273 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/06/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp flex device for mechanical circulatory support. on day two of impella rp flex support, the impella access site was oozing so additional sutures were placed around repositioning sheath. the oozing resolved. the patient remained on support. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the cause of the issue was use related because the additional sutures around the access site resolved the bleeding issue. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.
29234 | 1220648-2024-08460 | 18964217 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18964217&pc=OZD | 10/20/2022 | 03/21/2024 |  | 03/22/2024 |  |  | 10/20/2022 | 03/22/2024 |  |  | 06/30/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023240343 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 10/20/2022 |  |  |  |  | The user facility reported a 64-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient experienced right lower leg ischemia while on support. no medical or surgical intervention was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
29235 | 1220648-2024-08612 | 18964222 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18964222&pc=OZD | 03/13/2024 | 04/24/2024 |  | 03/22/2024 |  |  | 03/13/2024 | 03/22/2024 |  |  |  |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025457805 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Cardiac Perforation | Intraoperative Cardiac Injury | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/13/2024 |  |  |  |  | The user facility reported an (b)(6) female post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. the impella was implanted, and the device was noted to be deep. the clinical representative verbalized their concerns for the positioning. at the time, a small bleed was noticed of the posterior graft. while manipulating heart to access bleeding site, device perforated left ventricle (lv). device was immediately pulled back and attention was given to repairing perforation. repairing perforation led to additional impairments of lv wall ultimately causing the abandonment of impella device altogether. decision was made to explant the impella 5.5, and switch therapy switched to intra-aortic balloon pump. | [5 empty] | The investigation for the reported ventricle perforation and the positioning issues has been completed. the 5.5 pump was not returned for investigation. data logs were reviewed and suction alarms and impella position unknown events were observed. the cause of the ventricle perforation issue was most likely patient condition related due to weaker ventricle wall prior to impella placement and pump puncturing left ventricle during support. the cause of the positioning issue was most likely use related per clinical review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.
29236 | 1220648-2024-08459 | 18964223 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18964223&pc=OZD | 11/24/2022 | 03/21/2024 |  | 03/22/2024 |  |  | 11/24/2022 | 03/22/2024 |  |  | 08/31/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023267917 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/24/2022 |  |  |  |  | The user facility reported a 77-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after the procedure was completed, it was noted that the patient had a small hematoma at the groin site. protamine was ordered by the physician and pressure was held until the site was soft. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records.
29237 | 1220648-2024-08444 | 18964226 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18964226&pc=OZD | 10/08/2022 | 04/02/2025 |  | 03/22/2024 |  |  | 10/08/2022 | 03/22/2024 |  |  | 07/31/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023253298 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Life Threatening | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/08/2022 |  |  |  |  | The user facility reported a 53-year-old male was implanted with an impella device for mechanical circulatory support. the patient was successfully implanted with the impella cp with systemic heparin administered after the patient was found to have a 90% occlusion of the proximal right coronary artery and significant lesions in the left anterior descending artery and circumflex artery. while the patient was waiting for a transfer to an outside institution for surgery, it was reported that the patient expired while on impella cp support after experiencing a stroke. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the cause of the stroke/ cva was not determined since product was not returned and all the necessary information was not provided. this report is being filed as part of a retrospective review of historical records.
29238 | 1220648-2024-08644 | 18964227 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18964227&pc=OZD | 06/28/2022 | 03/22/2024 |  | 03/22/2024 |  |  | 06/28/2022 | 03/22/2024 |  |  | 01/31/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023189625 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/28/2022 |  |  |  |  | The user facility reported a 66-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced a hematoma at the access site. the purge was switched to sodium bicarbonate and heparin was discontinued. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29239 | 1220648-2024-08560 | 18964230 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18964230&pc=OZD | 03/06/2024 | 05/06/2024 |  | 03/22/2024 |  |  | 03/06/2024 | 03/22/2024 |  |  |  |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025438388 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during impella cp placement procedure, the physician attempted to start the device. the motor was increased, but the pump stopped. therefore, the device was explanted and a new impella device was placed. there was no reported patient harm. | [5 empty] | The investigation for the reported pump stop has been completed. the pump was returned for evaluation and upon visual inspection, a crack was found along the purge filter. the pump was run and the pump stop was not reproduced in the lab. clinical details state that the physician tried to start the impella cp, but the motor current increased and the impella stopped. it is unknown if there were any access complications. the data logs confirmed multiple motor current spikes in succession and low purge pressure alarms prior to attempts to start the pump. the root cause of the pump stop was determined to be a damaged sidearm filter. added- h6 impact code 4629 added-d6a/d6b | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
29240 | 1220648-2024-08559 | 18964232 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18964232&pc=OZD | 07/25/2022 | 04/02/2025 |  | 03/22/2024 |  |  | 07/25/2022 | 03/22/2024 |  |  | 01/31/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023175017 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  |  | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/25/2022 |  |  |  |  | The user facility reported a 69-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported there was a thrombus in the left ventricle coinciding with an impella in aorta alarm and the impella cp pump was explanted as a result. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29241 | 1220648-2024-08613 | 18964237 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18964237&pc=OZD | 08/07/2022 | 03/07/2025 |  | 03/22/2024 |  |  | 08/07/2022 | 03/22/2024 |  |  | 04/30/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023225160 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/07/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the pump accidently slipped into the patient's aorta. despite best practice and recommendation, as doctor felt that it would be too hard to cross it back into the left ventricle, the pump was left in the aorta and placement alarms were disabled. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29242 | 1220648-2024-08618 | 18964241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18964241&pc=OZD | 08/08/2022 | 03/10/2025 |  | 03/22/2024 |  |  | 08/08/2022 | 03/22/2024 |  |  | 04/30/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023223934 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/08/2022 |  |  |  |  | The user facility reported a 77-year-old male in post cardiotomy cardiogenic shock /low cardiac output syndrome was implanted with an impella cp device for mechanical circulatory support. while on support the patient had frequent abnormal heart rhythms and a lower gi bleed. 2 units of blood was given and heparin concentration was decreased for treatment. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events: acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleeding and vascular damage - peripheral vessel issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
29243 | 1220648-2024-08558 | 18964269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18964269&pc=OZD | 07/17/2022 | 04/03/2025 |  | 03/22/2024 |  |  | 07/17/2022 | 03/22/2024 |  |  | 03/31/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023212680 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/17/2022 |  |  |  |  | The user facility reported a 47-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support while undergoing cardiopulmonary resuscitation. after implantation, the motor spiked with no flow. the pump stopped and was subsequently removed and replaced with a second device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ instructions for use for the related event are as follows: | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29244 | 1220648-2024-08616 | 18964275 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18964275&pc=OZD | 03/14/2024 | 05/24/2024 |  | 03/22/2024 |  |  | 03/14/2024 | 03/22/2024 |  |  | 07/31/2025 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025441398 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/14/2024 |  |  |  |  | There is not enough information to associate use of device with outcome. | The user facility reported a 62-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support due to having right heart failure and sustained vt (ventricular tachycardia) for 12 hours. while on support, the impella had high motor current and the pump stopped. the patient may have been having cardiopulmonary resuscitation at the time of the failure. |  |  |  |  | The investigation into the saturated flow and the pump stop issues has been completed since the original report was submitted. the pump was not returned for investigation. per data logs provided, the pump was turned down and an increase in placement signal pulsatility due to cpr being performed. additionally, pump in aorta alarms occurred during cpr. multiple motor current spikes are observed in that time, and motor current becomes saturated after second motor current spike. the root cause of the pump stop was most likely due to biomaterial ingestion when cpr was being performed. b5, d6a, d6b. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
29245 | 1220648-2024-08462 | 18964308 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18964308&pc=OZD | 03/07/2024 | 04/24/2024 |  | 03/22/2024 |  |  | 03/07/2024 | 03/22/2024 |  |  | 11/30/2025 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024430516 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/07/2024 |  |  |  |  | The user facility reported a 80-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient on bipella support experienced hemolysis. the impella rp was in good position with no alarms, so the impella cp was likely the culprit of the hemolysis due to positioning issue. the medical doctor chose to try fluids first instead of repositioning due to the patient being on bipella support. the impella cp was also dislodged upon transferring to the intensive care unit (icu) bed from the table which could have also caused hemolysis prior to foley insertion and suction alarms in the icu. the patient also coded with cardiopulmonary resuscitation for 45 minutes with the impella cp in place prior to the impella rp being inserted which could also contributed to hemolysis. | [5 empty] | The investigation for the reported hemolysis has been completed. according to the clinical, during the first day of impella cp support the patient's urine was extremely hemolyzed with a pink/red color. following suction alarms, it was then discovered that the cp inlet cage was against the mitral valve. upon transfer to the icu, the impella was found to be unlocked and was accidentally pulled back into the aorta. following a successful repositioning, the patients urinary output was still hemolyzed and the family of the patient decided to withdraw care. no product was returned for a visual inspection. data log analysis confirmed positioning issues with suction alarms. the most likely cause of the hemolysis was improper positioning. the most likely cause of the positioning issue was use related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
29246 | 1220648-2024-08619 | 18964315 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18964315&pc=OZD | 03/06/2024 | 04/22/2024 |  | 03/22/2024 |  |  | 03/06/2024 | 03/22/2024 |  |  |  |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024436405 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Insufficient Flow or Under Infusion | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella 5.5 pump was placed to support a patient post surgery. the patient had suffered shock post surgery and the intra-aortic balloon pump (iabp) placed was not sufficient support. the pump experienced flow alarms and all lines were thrombosing so after 121 hours the pump was removed and replaced. the team placed a new impella 5.5 after bleeding. no lasting harm or injury. the second impella 5.5 pump supports to date. | [5 empty] | The investigation into the low pump flow and the thrombus issues has been completed since the original report was filed. the pump was returned, tested, and evaluated. visual inspection found biomaterial inside the pump housing & the purge gap. the root cause of the low flow was biomaterial ingestion. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed.
29247 | 1220648-2024-08645 | 18964805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18964805&pc=OZD | 06/28/2022 | 03/22/2024 |  | 03/22/2024 |  |  | 06/28/2022 | 03/22/2024 |  |  | 02/29/2024 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023193220 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/28/2022 |  |  |  |  | The user facility reported a 67-year-old male was implanted with an impella cp device for mechanical circulatory support. the medical team experienced difficulty implanting the impella 5.5 due to the patient's anatomy. therefore, the impella 5.5 implant was aborted, and the medical team decided to use an impella cp. the impella cp pump was also unable to pass through the anastomoses as well. the graft was subsequently clamped and placed more proximal, following which the impella cp insertion was successful. additionally systemic heparin was discontinued due to bleeding at access site. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
29248 | 1220648-2024-08646 | 18965147 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18965147&pc=OZD | 06/28/2022 | 09/11/2024 |  | 03/22/2024 |  |  | 06/28/2022 | 03/22/2024 |  |  | 11/30/2023 |  | 03/22/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023166356 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  |  |  | Death | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/28/2022 |  |  |  |  | It was reported by medical staff that the patient expired while on impella cp support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 79-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while the patient was on impella cp support, the patient had multiple rhythm disturbances including ventricular fibrillation. cardiopulmonary resuscitation was done, but it caused the impella to be dislodge and multiple attempts to reposition the impella were made. after the procedure, patient continued to have multiple arrests and coded five separate times and advanced cardiac life support protocol was initiated. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08646 was provided in sections b2, b5, h1, and h6. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29249 | 1220648-2024-08442 | 18948614 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18948614&pc=OZD | 03/04/2024 | 04/22/2024 |  | 03/21/2024 |  |  | 03/04/2024 | 03/21/2024 |  |  | 05/31/2025 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024376309 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 78-year-old male undergoing cardiac surgery was to be implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp pump was unable to be advanced through the patient's vasculature during attempted delivery. it was noted that the patient had severe peripheral vascular disease, which created resistance during the implant. the delivery was subsequently aborted as a result of the complication. there was no patient harm. | [5 empty] | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue was most likely due to patient condition since impella couldn't be placed through vascular anatomy due to severe pvd. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29250 | 1220648-2024-08443 | 18948702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18948702&pc=OZD | 10/17/2022 | 04/09/2025 |  | 03/21/2024 |  |  | 10/17/2022 | 03/21/2024 |  |  | 06/30/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023247423 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 10/17/2022 |  |  |  |  | The user facility reported a 68-year-old patient needing mechanical circulatory support. it was reported that while the patient was on impella cp support, there was some bleeding at the site. systemic heparin was withheld to resolve bleeding. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
29251 | 1220648-2024-08445 | 18948756 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18948756&pc=OZD | 12/03/2022 | 05/28/2025 |  | 03/21/2024 |  |  | 12/03/2022 | 03/21/2024 |  |  | 08/31/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023270099 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Suction Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/03/2022 |  |  |  |  | The user facility reported a 74-year-old male patient on an impella cp for mechanical circulatory support. during support, there were suction events reported despite repositioning. in addition, there was oozing from the access site and hemoglobin dropped. the patient was transfused with 3 units of packed red blood cell (prbcs). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29252 | 1220648-2024-08446 | 18948786 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18948786&pc=OZD | 07/15/2022 | 03/21/2024 |  | 03/21/2024 |  |  | 07/15/2022 | 03/21/2024 |  |  | 11/30/2023 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023166881 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 07/15/2022 |  |  |  |  | The user facility reported a 84-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. while on support, the patient experienced access site bleeding that required blood transfusion and cessation of heparin. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
29253 | 1220648-2024-08447 | 18948843 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18948843&pc=OZD | 01/10/2024 | 05/29/2025 |  | 03/21/2024 |  |  | 03/05/2024 | 03/21/2024 |  |  | 11/30/2025 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024426029 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Renal Impairment | [11 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/05/2024 |  |  |  |  | The user facility reported a 76-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 pump was placed in (b)(6) 2024 for support of a patient. the patient was enrolled in the impact protocol. in 3/2024, a document noted a "possible" relationship of the impella causing anemia. the anemia was observed post coronary artery bypass graft and the patient was given one unit of red blood cells with dropping hemoglobin levels. | Additional information received that the the study further reported upon acute kidney injury in which the labs showed rising creatinine. |  |  |  |  | The investigation into the access site bleeding ¿ major issue has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the access site bleeding cannot be determined | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The following fields were updated to reflect the additional information received from the clinical site: b2 updated to include other serious, b5, h6 health effect - clinical code and health effect - impact code d6a and d6b revised from the initial manufacturer device report number 1220648-2024-08447 in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2024-08447. g1 reporting contact email revised on manufacturer device report 1220648-2024-08447 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2024-08447.
29254 | 1220648-2024-08448 | 18948895 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18948895&pc=OZD | 07/15/2022 | 03/21/2024 |  | 03/21/2024 |  |  | 07/15/2022 | 03/21/2024 |  |  | 10/31/2023 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022149151 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Impairment | [11 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/15/2022 |  |  |  |  | The user facility reported a 57-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that hemolysis began during impella support. heparin was held. the patients care team believed the patient was in liver failure. subsequently, the patient expired. no allegations were made towards the impella for cause of death. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
29255 | 1220648-2024-08449 | 18948923 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18948923&pc=OZD | 03/06/2024 | 04/24/2024 |  | 03/21/2024 |  |  | 03/06/2024 | 03/21/2024 |  |  | 08/31/2025 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024386807 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  |  |  | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/06/2024 |  |  |  |  | The user facility reported a 47-year-old female of unknown race and ethnicity with cardiomyopathy was to be implanted with an impella 5.5 for mechanical circulatory support. it was reported that while attempting to advance through the anastomosis into the artery, the inlet advanced into the artery but was unable to advance further and cannula collapsed on itself. wire access was lost, and the devices were removed. the wire was swapped, and the pump was reinserted, but resistance was encountered, and the result was the same. bleeding was observed at the anastomosis site. four units of packed red blood cells were administered. the implant was aborted, and the artery repaired. the patient was reported as stable. | [5 empty] | The investigation into the reported issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the vascular damage was most likely due to use issue since s the artery was damaged during the repeated insertion attempts. the root cause of the delivery issue was most likely due patient condition since the pump was unable to pass through anastomosis/graft. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury"
29256 | 1220648-2024-08450 | 18949003 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18949003&pc=OZD | 03/07/2024 | 04/24/2024 |  | 03/21/2024 |  |  | 03/07/2024 | 03/21/2024 |  |  |  |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA |  |  | ABIOMED, INC |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Company representation | [5 empty] | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Material Separation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/07/2024 |  |  |  |  | It was reported that a demo 7 french low profile companion sheath separated while the device was being manipulating. further inspection noted that the luer had a defect which affected the retention of the components. there was no patient involvement at the time of the noted issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. catalog number, serial number, lot number, expiration date, unique identifier (udi) number, initial reporter and device manufacture date are unknown. | The investigation into the introducer damage ¿ mechanical issue has been completed since the original report was filed. the pump was returned, tested, and evaluated. all three sheath luers were confirmed to be crushed. the stack-up of the damaged parts could not retain the necessary components and resulted in the reported separation issue. the cause of introducer damage ¿ mechanical was a damaged luer lock connection from a vendor manufacturing machine misalignment.
29257 | 1220648-2024-08451 | 18949040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18949040&pc=OZD | 10/11/2022 | 03/13/2025 |  | 03/21/2024 |  |  | 10/11/2022 | 03/21/2024 |  |  | 07/31/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023251540 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Suction Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/11/2022 |  |  |  |  | The user facility reported a 47-year-old male on impella cp for mechanical circulatory support. during support, the patient experienced frank hemolysis in urine and ongoing suction alarms occurred. additionally, the patient did not have pulses on the impella leg and was sent to the operating room for emergency fasciotomy. it was also reported that there was no systemic heparin due to the patient experiencing bleeding from multiple sites. the patient was administered volume and blood products. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29258 | 1220648-2024-08452 | 18949113 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18949113&pc=OZD | 07/17/2022 | 03/13/2025 |  | 03/21/2024 |  |  | 07/17/2022 | 03/21/2024 |  |  | 04/30/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218672 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/17/2022 |  |  |  |  | The user facility reported a 43-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while under impella cp support, the patient's leg became ischemic. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29259 | 1220648-2024-08453 | 18949117 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18949117&pc=OZD | 03/07/2024 | 04/18/2024 |  | 03/21/2024 |  |  | 03/07/2024 | 03/21/2024 |  |  |  |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025438964 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/07/2024 |  |  |  |  | The user facility reported an 82-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was that that the implanted impella cp pump became entangled in the patient's papillary muscle during support. after multiple attempts, the pump was eventually freed from the muscle, but was incidentally pulled out of the left ventricle. after attempts to re-advance the pump were unsuccessful, the decision was made to replace the pump with a new impella cp device. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. according to the clinical, on the second day of impella cp support the team discovered that the pump was tangled in the pap muscle, shallow, and turned. multiple attempts were made to untangle the pump from the pap muscle with no success. another attempt was made the following day to reposition the pump, but instead the pump was pulled out of the lv and was unable to be advanced back over the valve. the decision was then made to explant the pump and replace it with another impella cp device. the root cause of the positioning issue was use related. added- h6 impact code 4628, g1 reporting contact last name updated.
29260 | 1220648-2024-08454 | 18949121 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18949121&pc=OZD | 10/20/2022 | 03/13/2025 |  | 03/21/2024 |  |  | 10/20/2022 | 03/21/2024 |  |  | 06/30/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023240868 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/20/2022 |  |  |  |  | The user facility reported a 76-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. a hematoma was noted on the right femoral artery impella site and manually compressed in the intensive critical unit. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29261 | 1220648-2024-08455 | 18949245 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18949245&pc=OZD | 10/13/2022 | 02/18/2025 |  | 03/21/2024 |  |  | 10/13/2022 | 03/21/2024 |  |  | 02/29/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195250 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/13/2022 |  |  |  |  | The user facility reported an 82-year-old female on impella cp for mechanical circulatory support. it was reported that the patient's right leg was absent pulses and was cold. the patient had an external bypass. care was withdrawn. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29262 | 1220648-2024-08456 | 18949296 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18949296&pc=OZD | 07/18/2022 | 09/12/2024 |  | 03/21/2024 |  |  | 07/18/2022 | 03/21/2024 |  |  | 03/31/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023209715 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/18/2022 |  |  |  |  | The user facility reported a 68-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. due to worsening shortness of breath, the patient was intubated. after intubation, the left anterior descending artery (lad) and left circumflex artery (lca) were opened up and the lad was filled with thrombus and 100% circumflex. both vessels were opened and the patient taken to icu. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08456 was provided.
29263 | 1220648-2024-08457 | 18949339 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18949339&pc=OZD | 03/07/2024 | 04/24/2024 |  | 03/21/2024 |  |  | 03/07/2024 | 03/21/2024 |  |  |  |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025438086 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/07/2024 |  |  |  |  | The user facility reported a 59-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that soon after impella 5.5 placement, the patient developed a large hematoma in the upper right subclavian area. heparin was held and one unit of red blood cells was given. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was unable to be determined as no product and limited clinical details were provided. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
29264 | 1220648-2024-08458 | 18949483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18949483&pc=OZD | 10/15/2022 | 04/07/2025 |  | 03/21/2024 |  |  | 10/15/2022 | 03/21/2024 |  |  | 06/30/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023235261 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/15/2022 |  |  |  |  | The user facility reported a 51-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during impella cp support the patient was in sustained ventricular tachycardia, was shocked 5 times in route, and was electively intubated, then successfully defibrillated. however, the impella cp was explanted and the impella 5.5 was implanted with direct access in place with no complications. upon implantation of the device the patient encountered transient ischemic attack (tia) with some speech deficits, and neuro was working to find a anticoagulate for the patient. unfortunately, the patient complication was noted to be cerebral vascular accident (cva). the patient's health progressively improved and the device was eventually explanted. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. to make a determination as to the cause of the stroke, the following clinical information must be considered: -patient's medical history, including any previous strokes, cardiovascular disease, or risk factors such as hypertension, diabetes, or atrial fibrillation -timing of the stroke in relation to impella support (during or post-implant) -detailed description of the symptoms experienced by the patient -neurological examination findings and any focal deficits observed -diagnostic imaging results, such as ct scan or mri of the brain, to identify the type and location of the stroke (ischemic or hemorrhagic) -laboratory test results, including coagulation profiles and cardiac markers -any relevant procedural or device-related factors, such as duration and settings of impella support, presence of embolic protection devices, or any documented procedural complications -response to any previous anticoagulation or antiplatelet therapies -evaluation of potential alternative causes of stroke, such as arrhythmias, embolic sources, or underlying vascular pathology. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. this report is being filed as part of a retrospective review of historical records.
29265 | 1220648-2024-08463 | 18949734 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18949734&pc=OZD | 07/26/2022 | 03/11/2025 |  | 03/21/2024 |  |  | 07/26/2022 | 03/21/2024 |  |  | 03/31/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023225928 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/26/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient pulses were unable to be dopplered in the catheterization laboratory. however, retrograde flow was seen on fluoroscopy. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29266 | 1220648-2024-08464 | 18949785 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18949785&pc=OZD | 07/18/2022 | 03/13/2025 |  | 03/21/2024 |  |  | 07/18/2022 | 03/21/2024 |  |  | 04/30/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213905 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 91 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | Tachycardia | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/18/2022 |  |  |  |  | The user facility reported a (b)(6) year-old male was being implanted with an impella cp device for mechanical circulatory support. the physician obtained access on the right femoral but was too low. access was then obtained on the left femoral artery. when the physician attempted to preclose, it was unsuccessful due to calcified artery. the physician decided to close the right femoral sheath with proglide and attempted to obtain access above the right femoral access. the second attempt was also unsuccessful. on the third attempt, the physician preclosed and predilated with the 10 and 12 fr dilator. the physician then advanced the impella 14 fr sheath. when trying to advance pigtail, the physician struggled to advance to the femoral anatomy. after crossing aortic valve into the left ventricle, the patient went into ventricular tachycardia twice and had to be shocked. physician then attempted to advance the impella over 0.18 wire but was unsuccessful due to patient anatomy. the procedure was aborted due to the patient¿s anatomy. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29267 | 1220648-2024-08465 | 18949821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18949821&pc=OZD | 03/08/2024 | 04/29/2024 |  | 03/21/2024 |  |  | 03/08/2024 | 03/21/2024 |  |  | 08/31/2025 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025441092 |  | 1000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/08/2024 |  |  |  |  | The user facility reported a 69-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient was on impella 5.5 support due to needing a bridge to a transplant. while on support, the patient had pain and soreness from the insertion site. there was also a hematoma noted and pressure dressing was applied. additionally, the patient tested positive for heparin induced thrombocytopenia. heparin was paused and bival was added. | [5 empty] | The investigation into the reported issues has been completed since the original report was submitted. the device was not returned for investigation. according to the clinical, patient had hematoma, pain in the arm, low platelet count, and tested positive for hit. the most likely cause of the minor access site bleeding was the patient's hit condition. the cause of the pain at the insertion site and the thrombocytopenia was not determined due to insufficient clinical details. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29268 | 1220648-2024-08467 | 18949930 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18949930&pc=OZD | 10/01/2022 | 09/12/2024 |  | 03/21/2024 |  |  | 10/01/2022 | 03/21/2024 |  |  | 06/30/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023236013 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Hispanic |  | Death | Yes | No | Activation, Positioning or Separation Problem | [6 empty] | Stroke/CVA | Hemolysis | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/01/2022 |  |  |  |  | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during impella support, the patient was rolled to be placed on bedpan and the impella pulled back into aorta. echocardiogram (echo) confirmed impella catheter in aorta with pigtail across aortic valve. the catheter was advanced, under echo guidance. the physician removed slack and tightened tuohy burst valve. the measure then verified by echo at 3.5 setting med-ventricular space clear of any structures. the patient returned to cardiac catheterization lab (ccl) for placement of pa catheter and a-line. the position was verified, and slack removed from impella catheter, tuohy burst valve locked. there were also suction alarms, that were resolved by a decrease in performance level, the pt finally regained adequate fluid status. in addition, a ct head scan revealed hemorrhagic stroke with midline shift. | The procedural outcome is patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | B2: date of death is unknown. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08467 was provided.
29269 | 1220648-2024-08468 | 18949942 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18949942&pc=OZD | 07/18/2022 | 03/13/2025 |  | 03/21/2024 |  |  | 07/18/2022 | 03/21/2024 |  |  | 04/30/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213921 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 07/18/2022 |  |  |  |  | The user facility reported a 58-year-old patient was implanted with an impella cp for mechanical circulatory support. it was reported that the patient was observed to have site bleeding at the right femoral artery. the impella cp was removed and replaced with an intra-aortic balloon pump (iabp) due to excessive bleeding. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29270 | 1220648-2024-08469 | 18949969 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18949969&pc=OZD | 10/04/2022 | 09/11/2024 |  | 03/21/2024 |  |  | 10/04/2022 | 03/21/2024 |  |  | 05/31/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231246 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hematoma | Ischemia | [10 empty] | Life Threatening | Required Intervention | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/04/2022 |  |  |  |  | It was reported patient coded and expired while on impella cp support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported at the end of the procedure, the impella leg had dye shot down through and the re-access port showed very slow flow. fem-fem bypass inserted, good blood flow after 2nd shot down leg. during support, a hematoma was noted, 2u packed red blood cells (prbc) were administered. a computer tomography (ct) scan confirmed an embolic stroke. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08469 was provided in sections b2, b5, h1, and h6. | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿.
29271 | 1220648-2024-08471 | 18950088 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18950088&pc=OZD | 03/09/2024 | 05/01/2024 |  | 03/21/2024 |  |  | 03/09/2024 | 03/21/2024 |  |  | 07/31/2025 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024388788 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Life Threatening | Required Intervention | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/09/2024 |  |  |  |  | The user facility reported a 73-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was on impella cp support due to severe cardiomyopathy and cardiogenic shock. while on support, the patient experienced hemolyzed labs. the patient had dark urine with clots. the pump was repositioned, and the plan was to have labs tested again, however the patient expired. no allegations were made towards the impella for cause of death. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿. | The investigation into the hemolysis and the thrombus issues has been completed since the original report was submitted. the device was not returned for investigation. the cause of the hemolysis issue was improper positioning related due to improper technique and the clinical data confirmed that the position was 3cm shallow. as the necessary information was not provided, the root cause of the thrombus was not determined. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ in order to conform to updated procedures, sections d6 & h10 have been revised with updated information.
29272 | 1220648-2024-08470 | 18950100 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18950100&pc=PYX | 10/07/2022 | 03/13/2025 |  | 03/21/2024 |  |  | 10/07/2022 | 03/21/2024 |  |  | 02/29/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023182487 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/07/2022 |  |  |  |  | The user facility reported a 70-year-old patient needing mechanical circulatory support was attempted to be implanted with an impella rp. it was reported that during insertion, the right femoral vein was used for access and an angiogram was performed where may-thurner disease was revealed in right iliac vein. the decision was made to abort rp placement due to stenosis. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29273 | 1220648-2024-08472 | 18950103 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18950103&pc=OZD | 12/06/2022 | 03/13/2025 |  | 03/21/2024 |  |  | 12/06/2022 | 03/21/2024 |  |  | 07/31/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023247159 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  |  | Asian | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 12/06/2022 |  |  |  |  | The user facility reported a 82-year-old male was implanted with an impella cp device for mechanical circulatory support. the physician expressed concern that the position of the impella may lead to acute aortic regurgitation, hemodynamic deterioration and left anterior descending nonflow. patient experienced aortic regurgitation and thrombus. the impella device was removed. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29274 | 1220648-2024-08473 | 18950223 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18950223&pc=OZD | 10/24/2022 | 03/13/2025 |  | 03/21/2024 |  |  | 10/24/2022 | 03/21/2024 |  |  | 07/31/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023253768 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Insufficient Flow or Under Infusion | Positioning Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/24/2022 |  |  |  |  | The user facility reported a 46-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. impella cp reposition was attempted at bedside under echo due to hematuria/hemolysis; however, the repositioning was unsuccessful. the patient was taken to cardiac cath laboratory to reposition under fluoroscopy. the device was pulled back as pigtail seemed to be stuck in pap muscle/chordate. the device jumped back and re-advancement of the device was attempted, but flows of 2.0 liters was achieved. the physician made the decision to replace the device with a new impella cp due to low flow. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29275 | 1220648-2024-08474 | 18950277 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18950277&pc=OZD | 03/10/2024 | 04/19/2024 |  | 03/21/2024 |  |  | 03/10/2024 | 03/21/2024 |  |  | 07/31/2025 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025441406 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/10/2024 |  |  |  |  | The user facility reported a 47-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that after the impella implant, the doctor found a perclose suture knot outside the body. the repositioning sheath was completely pulled back and the suture knot was advanced inside. there was bleeding noted around the site after the repositioning sheath was reinserted. perclose sutures were tightened, pressure was held, pressure dressing was applied, sandbag was placed, and a safeguard was applied. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical details provided was sufficient to determine the root cause. the root cause of the access site bleeding was most likely due to use issue because of the improper suturing that lead to reinsertion of the repo sheath that caused the bleed. the issue resolved the later day after perclose suture tightened, held pressure, pressure dressing applied, and sandbag placed at the bleeding site. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
29276 | 1220648-2024-08475 | 18950335 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18950335&pc=OZD | 08/03/2022 | 03/13/2025 |  | 03/21/2024 |  |  | 08/03/2022 | 03/21/2024 |  |  | 01/31/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183510 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/03/2022 |  |  |  |  | The user facility reported a 60-year-old male on impella cp for mechanical circulatory support. it was reported after a successful impella implant procedure, ct revealed extravasation in the vicinity of the descending aorta. the patient was sent to or for impella cp explant, aortic graft placement, pericardial window, and chest tube placement. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29277 | 1220648-2024-08476 | 18950443 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18950443&pc=OZD | 12/12/2022 | 03/13/2025 |  | 03/21/2024 |  |  | 12/12/2022 | 03/21/2024 |  |  | 07/31/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023246651 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 12/12/2022 |  |  |  |  | The user facility reported a 71-year-old male was implanted with an impella cp device for mechanical circulatory support due to heart failure and ejection fraction was in the 20 percent. when the impella was placed, it was difficult to position it in the direction of the rear wall, and it was removed once and then reinserted. the echo showed it was slightly facing the rear wall, the pump was turned on, and confirmed with p-8 that it was not suctioned, and it was fixed. | [5 empty] | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29278 | 1220648-2024-08477 | 18950543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18950543&pc=OZD | 03/11/2024 | 04/29/2024 |  | 03/21/2024 |  |  | 03/11/2024 | 03/21/2024 |  |  | 11/30/2025 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024431014 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/11/2024 |  |  |  |  | The user facility reported a 59-year-old male with multiple cardiac comorbidities was escalated from an intra-aortic balloon pump (iabp) to an impella 5.5 device for mechanical circulatory support. after 11 days of support, the patient was diagnosed with hemolysis demonstrated by an elevated plasma-free hemoglobin. the pump was replaced with a new impella 5.5, and support was continued. | [5 empty] | The investigation into the reported hemolysis issue has been completed since the original report was submitted. the device was not returned for investigation. per the data log, many suction alarms and placement signal downtrend were observed. it is unknown if the clot observed on outflow cage occurred during explant or during pump support. the cause of the hemolysis cannot be determined since the complaint device not returned for investigation and insufficient clinical data provided. adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ to conform to updated procedures, sections d6 and d10 have been revised with updated information. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
29279 | 1220648-2024-08478 | 18950564 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18950564&pc=OZD | 08/12/2022 | 02/18/2025 |  | 03/21/2024 |  |  | 08/12/2022 | 03/21/2024 |  |  | 04/30/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023226489 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Fluid/Blood Leak | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/12/2022 |  |  |  |  | The user facility reported a 44-year-old male in cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. during support, there was a leak in the purge line and the pump begin to have saturation. the decision was made for the device to be removed and a new impella cp to be placed within the patient. however, the patient had been bleeding through the hemostatic valve and 1 unit of blood was given for low hemoglobin levels. there was mild to moderate oozing from the groin site and 4 units of packed red blood cells (prbcs) were distributed for treatment. eventually the impella cp was explanted and the impella 5.5 was implanted. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. ¿assess access site for bleeding and hematoma.¿. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir.". this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. tthe investigation into the reported issue has been completed. no device waws receiced for evaluation. the device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29280 | 1220648-2024-08479 | 18950645 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18950645&pc=OZD | 10/27/2022 | 03/13/2025 |  | 03/21/2024 |  |  | 10/27/2022 | 03/21/2024 |  |  | 06/30/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023237428 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/27/2022 |  |  |  |  | The user facility reported a male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the pt was heparin-induced thrombocytopenia positiv. bicarb purge was administered. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29281 | 1220648-2024-08480 | 18950737 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18950737&pc=OZD | 06/29/2022 | 03/28/2025 |  | 03/21/2024 |  |  | 06/29/2022 | 03/21/2024 |  |  | 03/31/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023214399 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/29/2022 |  |  |  |  | The user facility reported a 80-year-old male was being implanted with an impella cp device for mechanical circulatory support for high risk percutaneous cardiac insertion. the impella cp was unable to be placed due to the patient¿s anatomy and aortic graft repair. the patient was subsequently brought to the cath lab later for axillary insertion of the impella cp. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The initial mfr 1220648-2024-08480 is a duplicate report and should be disregarded. please refer to mfr 1220648-2024-08670 for this complaint.
29282 | 1220648-2024-08481 | 18950749 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18950749&pc=OZD | 12/05/2022 | 03/13/2025 |  | 03/21/2024 |  |  | 12/05/2022 | 03/21/2024 |  |  | 08/31/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023266561 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/05/2022 |  |  |  |  | The user facility reported a 77-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient underwent ct scan of the abdomen & pelvis to rule out bleeding. hemaglobin/hematocrit was stable, and complete blood count was retested. the patient received two units of blood. only one of the perclose's was successful, so an 8fr angio seal was used with good result. manual pressure was held for 15 minutes. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29283 | 1220648-2024-08482 | 18950960 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18950960&pc=OZD | 07/08/2022 | 05/28/2025 |  | 03/21/2024 |  |  | 07/08/2022 | 03/21/2024 |  |  | 03/31/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023211465 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 07/08/2022 |  |  |  |  | The user facility reported a 74-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced excessive bleeding at the impella site and ventricular fibrillation arrest. all systemic heparin was stopped. | Patient had a ventricle fibrillation arrest. patient unfortunately expired on support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08482 was provided. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29284 | 1220648-2024-08483 | 18950980 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18950980&pc=OZD | 07/07/2022 | 04/16/2025 |  | 03/21/2024 |  |  | 07/07/2022 | 03/21/2024 |  |  | 04/30/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213862 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/07/2022 |  |  |  |  | The user facility reported a 50-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced positioning issues with the pump. the patient was noted to be hemodynamically stable, so no action was taken to address the issue. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29285 | 1220648-2024-08484 | 18951008 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18951008&pc=OZD | 07/11/2022 | 05/14/2025 |  | 03/21/2024 |  |  | 07/11/2022 | 03/21/2024 |  |  | 04/30/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023216020 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/11/2022 |  |  |  |  | The user facility reported a 48-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced bleeding at the impella site and hemolyzed labs indicating that the patient had developed hemolysis. as a result of the noted issue, two units of packed red blood cells (prbcs) and albumin were given, manual pressure was held, and hemostasis was achieved. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29286 | 1220648-2024-08485 | 18951014 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18951014&pc=OZD | 08/14/2022 | 03/21/2024 |  | 03/21/2024 |  |  | 08/14/2022 | 03/21/2024 |  |  | 05/31/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023227768 |  | CP, 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/14/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the central venous pressure (cvp) port was clotted and aspirated. the access site was flattened with small amount of ooze. the site was redressed, and angle matched. there was also slight oozing at groin site, which was redressed. the angle of entry was supported with 4x4 and one-unit packed red blood cells (prbcs) was administered. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29287 | 1220648-2024-08486 | 18951060 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18951060&pc=OZD | 08/11/2022 | 03/13/2025 |  | 03/21/2024 |  |  | 08/11/2022 | 03/21/2024 |  |  | 03/31/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023212045 |  | 0048-0047 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/11/2022 |  |  |  |  | The user facility reported a 78-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that hemolysis was found. this was treated by explanting the pump and inserting an intra-aortic balloon pump (iabp) instead for support. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29288 | 1220648-2024-08487 | 18951097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18951097&pc=OZD | 07/18/2022 | 03/13/2025 |  | 03/21/2024 |  |  | 07/18/2022 | 03/21/2024 |  |  | 04/30/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023217165 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/18/2022 |  |  |  |  | The user facility reported a 73-year-old female was implanted with an impella cp device for mechanical circulatory support. a large hematoma at the right femoral access impella site developed. the patient required two (2) units of blood and a higher dose of vasopressor support was required. the surgeon opted to explant the impella cp and replace it with a impella 5.5 at the right groin. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29289 | 1220648-2024-08488 | 18951141 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18951141&pc=OZD | 08/14/2022 | 03/13/2025 |  | 03/21/2024 |  |  | 08/14/2022 | 03/21/2024 |  |  | 04/30/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023224428 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/14/2022 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that after pump implant, the patient showed no distal flow to the left leg and their pulses were not audible. consequently, it was decided to escalate care by explanting the impella cp and changing to impella 5.5. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29290 | 1220648-2024-08489 | 18951213 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18951213&pc=OZD | 07/07/2022 | 03/13/2025 |  | 03/21/2024 |  |  | 07/07/2022 | 03/21/2024 |  |  | 03/31/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023205223 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/07/2022 |  |  |  |  | The user facility reported a 59-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the patient experienced access site bleeding that was resolved with pressure. additionally, the patient experienced sepsis that resolved with administration of antibiotics. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29291 | 1220648-2024-08490 | 18951276 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18951276&pc=OZD | 08/04/2022 | 04/08/2025 |  | 03/21/2024 |  |  | 08/04/2022 | 03/21/2024 |  |  | 03/31/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023205011 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup |  |  | 08/04/2022 |  |  |  |  | The user facility reported an 84-year-old patient needing mechanical circulatory support. it was reported while on impella support, the patient had a brain bleed and was unable to maintain pressures. family decided to withdraw support. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08490 was provided. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29292 | 1220648-2024-08491 | 18951335 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18951335&pc=OZD | 08/05/2022 | 03/21/2024 |  | 03/21/2024 |  |  | 08/05/2022 | 03/21/2024 |  |  | 04/30/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023216637 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 08/05/2022 |  |  |  |  | The user facility reported a 61-year-old patient needing mechanical circulatory support. it was reported that the patient was placed on impella cp support the patient developed bleeding when they came out of the operating room, so heparin was off and the patient was not given anticoagulant treatment until resolved. the bleeding was resolved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records.
29293 | 1220648-2024-08492 | 18951392 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18951392&pc=OZD | 07/07/2022 | 03/13/2025 |  | 03/21/2024 |  |  | 07/07/2022 | 03/21/2024 |  |  | 10/31/2023 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022156149 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/07/2022 |  |  |  |  | The user facility reported a 79-year-old male was implanted with an impella cp device for mechanical circulatory support due to percutaneous coronary intervention to address mitral valve disease. while on support, the patient experienced low pump flow. the pump was left in place due to the patient¿s aortic stenosis. the patient remained stable throughout the procedure. | [5 empty] | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29294 | 1220648-2024-08493 | 18951482 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18951482&pc=OZD | 08/14/2022 | 03/21/2024 |  | 03/21/2024 |  |  | 08/14/2022 | 03/21/2024 |  |  | 04/30/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023225120 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/14/2022 |  |  |  |  | The user facility reported a 64-year-old female patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. proglide was used for closure on the left groin during the impella placement. in addition, a hematoma formed after closure but was able to be expressed and femstop was placed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
29295 | 1220648-2024-08494 | 18951507 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18951507&pc=OZD | 08/06/2022 | 04/15/2025 |  | 03/21/2024 |  |  | 08/06/2022 | 03/21/2024 |  |  | 04/30/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023215981 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 08/06/2022 |  |  |  |  | The user facility reported a 68-year-old patient needing mechanical circulatory support. it was reported that the patient was placed on impella cp support . the patient developed a small hematoma at the groin impella site. there was no oozing present. staff redressed the site and planned on continuing to monitor site as there had been no hematoma growth over several hours. systemic anticoagulation on hold. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29296 | 1220648-2024-08495 | 18951536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18951536&pc=OZD | 07/18/2022 | 03/13/2025 |  | 03/21/2024 |  |  | 07/18/2022 | 03/21/2024 |  |  | 04/30/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023217689 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Deliver | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/18/2022 |  |  |  |  | The user facility reported a 86-year-old female was being placed on an impella cp for mechanical circulatory support. the complainant reported that an impella cp pump implant attempt had to be aborted aborted due to stenosis in right iliac bifurcation.. | [5 empty] | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29297 | 1220648-2024-08496 | 18951682 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18951682&pc=OZD | 08/22/2022 | 11/19/2024 |  | 03/21/2024 |  |  | 08/22/2022 | 03/21/2024 |  |  | 05/31/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023227193 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/22/2022 |  |  |  |  | The user facility reported a 56-year-old female patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was placed on an impella cp support after the patient started to decompensate while being treated for st-elevation myocardial infarction. during the insertion of the impella, the patient went into ventricular fibrillation. the patient was shocked and then went into a 2nd degree heart block. a temporary pacemaker was placed. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. a product history review was performed and confirmed that the pump passed all post sterile inspections checks. any device contacting the heart wall in conjunction with a poor patient condition could result in ventricular tachycardia/ventricular fibrillation (vt/vf), but since fluoroscopic imaging and/or sufficient clinical details were not provided, a relation between impella and the vt/vf is unable to be determined. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported cardiac arrhythmia has been completed. as the necessary clinical information was not provided and the device was not returned, the root cause of the cardiac arrhythmia was not determined. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29298 | 1220648-2024-08497 | 18951705 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18951705&pc=OZD | 07/07/2022 | 03/21/2024 |  | 03/21/2024 |  |  | 07/07/2022 | 03/21/2024 |  |  | 01/31/2023 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022032165 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/07/2022 |  |  |  |  | The user facility reported a 80-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while the patient was on impella cp support, there was bleeding at the site, therefore heparin was not initiated. the plan was to leave the impella in place overnight to recover and explant the following day. the patient remained on impella cp support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29299 | 1220648-2024-08498 | 18951728 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18951728&pc=OZD | 08/11/2022 | 03/07/2025 |  | 03/21/2024 |  |  | 08/11/2022 | 03/21/2024 |  |  | 04/30/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023221073 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/11/2022 |  |  |  |  | The user facility reported a 55-year-old female noted as high risk percutaneous cardiac insertion. during implantation of the device, there was difficulties placing the oscor 14fr x 25cm sheath after sequential dilation using the impella insertion kit. the sheath tip was bent on the dilator and a 16fr cook dilator as well as a short oscor 14fr x 13cm sheath were both utilized to manage the complication. eventually, the pump was explanted. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
29300 | 1220648-2024-08499 | 18951945 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18951945&pc=OZD | 10/07/2022 | 10/01/2025 |  | 03/21/2024 |  |  | 10/07/2022 | 03/21/2024 |  |  | 07/31/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023253304 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 95 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Patient-Device Incompatibility | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 10/07/2022 |  |  |  |  | The user facility reported a (b)(6) male patient needing an impella cp for mechanical circulatory support. during insertion, a 8f sheath access on the left femoral artery wire advanced up to right axillary artery. ultrasound guided access on right axillary artery. a 6f shores sheath revealed tortuous anatomy. a long 14f sheath was first used but kinked halfway and was exchanged to a short 14f sheath. a stent was implanted. patient remained hemodynamically stable throughout procedure. cp successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. b1 adverse event in initial mdr 1220648-2024-08499 should have been left unselected h1 type of reportable event has been corrected to product malfunction
29301 | 1220648-2024-08500 | 18951947 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18951947&pc=OZD | 07/07/2022 | 03/28/2025 |  | 03/21/2024 |  |  | 07/09/2022 | 03/21/2024 |  |  | 03/31/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023210476 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Detachment of Device or Device Component | Adverse Event Without Identified Device or Use Problem | Separation Problem |  |  |  |  | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/09/2022 |  |  |  |  | The user facility reported a 77-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported that while the patient was on impella cp support, the peel away sheath was occlusive, therefore, the plan was to explant the impella device. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿avoid manual compression of the inlet and outlet areas of the cannula assembly.¿ ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. the following was corrected from initial mfr 1220648-2024-08500. b1 type of reportable event was corrected to adverse event. b3 date of event was corrected. h1 type of reportable event was corrected to serious injury. section h6 health effect - clinical code was corrected to 4440 thrombus. medical device problem code was corrected to 2993 adverse event without identified.device or use problem.
29302 | 1220648-2024-08501 | 18951968 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18951968&pc=OZD | 07/19/2022 | 03/13/2025 |  | 03/21/2024 |  |  | 07/19/2022 | 03/21/2024 |  |  | 03/31/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023206590 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/19/2022 |  |  |  |  | The user facility reported a 83-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced bleeding that was resolved with surgical wound care. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29303 | 1220648-2024-08502 | 18952074 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18952074&pc=OZD | 07/28/2022 | 03/13/2025 |  | 03/21/2024 |  |  | 07/28/2022 | 03/21/2024 |  |  | 09/30/2023 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022143718 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/18/2022 |  |  |  |  | The user facility reported a male patient of unknown age was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced a gastrointestinal bleed and heparin was discontinued. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29304 | 1220648-2024-08503 | 18952139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18952139&pc=OZD | 10/07/2022 | 03/10/2025 |  | 03/21/2024 |  |  | 10/07/2022 | 03/21/2024 |  |  | 06/30/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023235253 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/07/2022 |  |  |  |  | The user facility reported a 75-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had hematuria with probable hemolysis. there was also some postoperative bleeding from the chest tubes and the patient was changed to heparin free purge, due to chest tube output. however, there was an increase in bleeding from the chest tubes and no bleeders were found fair to generalize the oozing. eventually, the device was explanted . | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis" this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29305 | 1220648-2024-08504 | 18952226 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18952226&pc=OZD | 10/10/2022 | 03/07/2025 |  | 03/21/2024 |  |  | 10/10/2022 | 03/21/2024 |  |  | 07/31/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023231447 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Premature Separation | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/10/2022 |  |  |  |  | The user facility reported a 70-year-old female noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. upon removal of the pump the peel away sheath was accidentally pulled back which resulted in loss of vessel access. hematoma at site and manual pressure was held for 1hr with hemostasis. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29306 | 1220648-2024-08505 | 18952235 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18952235&pc=OZD | 09/13/2022 | 09/11/2024 |  | 03/21/2024 |  |  | 09/13/2022 | 03/21/2024 |  |  | 04/30/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023223718 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Life Threatening | Other | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/13/2022 |  |  |  |  | It was reported the patient expired while on impella cp support after experiencing a massive infarction and intercranial bleed. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 67-year-old female was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient had a possible brain death. the patient was posturing, and the brain computed tomography (ct) indicated massive infarct and intercranial bleed. due to the possible brain death, the decision was made to not escalate. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08505 was provided in sections b2, b5, h1, and h6. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29307 | 1220648-2024-08506 | 18952307 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18952307&pc=OZD | 10/10/2022 | 03/10/2025 |  | 03/21/2024 |  |  | 10/10/2022 | 03/21/2024 |  |  | 04/30/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023215157 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/10/2022 |  |  |  |  | The user facility reported a 53-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient encountered distal limb ischemia of the impella extremity. an external perfusion bypass was placed from the re-access port to the antegrade sheath. there was a filling defect noticed on post closure angiogram with concerns for a dissection and noticeably reduced flow down the right common femoral. a covered stent was placed in the right common femoral artery to cover the filling defect. great angiographic result post stent deployment with brisk flow down common femoral and superficial femoral artery (sfa). mild-moderate disease still present in the right profunda. eventually, the device was successfully weaned. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported limb ischemia issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
29308 | 1220648-2024-08507 | 18952310 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18952310&pc=OZD | 07/10/2022 | 03/21/2024 |  | 03/21/2024 |  |  | 07/10/2022 | 03/21/2024 |  |  | 10/31/2023 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022149319 |  | 0048-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  |  | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/10/2022 |  |  |  |  | The user facility reported a 61-year-old male was being implanted with an impella cp device via femoral access location for mechanical circulatory support. due to vascular anatomy, there were access complications and an inability to implant the device. therefore, the implantation was aborted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29309 | 1220648-2024-08508 | 18952363 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18952363&pc=OZD | 07/19/2022 | 03/13/2025 |  | 03/21/2024 |  |  | 07/19/2022 | 03/21/2024 |  |  | 04/30/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023215595 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/19/2022 |  |  |  |  | The user facility reported a 74-year-old male was implanted with an impella cp device for mechanical circulatory support due to segment elevated myocardial infarction and occluded left anterior descending coronary artery branch. while on support, the patient experienced access site bleeding that was resolved with manual pressure, site redressing, and re-suturing. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29310 | 1220648-2024-08509 | 18952391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18952391&pc=OZD | 10/11/2022 | 04/16/2025 |  | 03/21/2024 |  |  | 10/11/2022 | 03/21/2024 |  |  | 07/31/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023253052 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 10/11/2022 |  |  |  |  | The user facility reported a 50-year-old patient needing mechanical circulatory support. it was reported that the patient was taken to the cath lab with a lesion on the left main, left anterior descending artery (lad), and huge thrombus ostial at the right coronary artery (rca). the impella cp was prepped, but unable to be implanted due to patient's anatomy. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29311 | 1220648-2024-08510 | 18952488 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18952488&pc=OZD | 07/20/2022 | 03/13/2025 |  | 03/21/2024 |  |  | 07/20/2022 | 03/21/2024 |  |  | 04/30/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023219489 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/20/2022 |  |  |  |  | The user facility reported a 70-year-old male was implanted with an impella cp device for mechanical circulatory support. the patient had some oozing in the groin area and received some angle matching and tightening of the sutures and was resolved with the re-dressing. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29312 | 1220648-2024-08511 | 18952494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18952494&pc=OZD | 07/11/2022 | 03/13/2025 |  | 03/21/2024 |  |  | 07/11/2022 | 03/21/2024 |  |  | 04/30/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023216038 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/11/2022 |  |  |  |  | The user facility reported a 71-year-old female was implanted with an impella cp device for mechanical circulatory support. the patient had decreased and loss of pulsatility with balloon inflations. patient with numbness secondary to ischemia throughout procedure as well as following removal of peel away and site closure. the device was explanted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29313 | 1220648-2024-08512 | 18953325 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18953325&pc=OZD | 12/02/2022 | 03/21/2024 |  | 03/21/2024 |  |  | 12/02/2022 | 03/21/2024 |  |  | 08/31/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023269967 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/02/2022 |  |  |  |  | The user facility reported a 59-year-old female was implanted with an impella cp device for mechanical circulatory support and developed bleeding at the impella cp insertion site when the patient was turned. the dressing was changed and pressure was held. bleeding was controlled after a purse string suture was placed and the patient was propped up. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29314 | 1220648-2024-08513 | 18953588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18953588&pc=OZD | 03/11/2024 | 04/29/2024 |  | 03/21/2024 |  |  | 03/11/2024 | 03/21/2024 |  |  | 11/30/2025 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024434988 |  | 1000100 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/11/2024 |  |  |  |  | The user facility reported a 81-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that an impella 5.5 pump was inserted in the operating room for support of a patient previously supported by impella cp-ecmo (extracorporeal membrane oxygenation). the impella 5.5 was placed but the pump had purge alarms. the team attempted to change the cassette without resolution of alarms. the pump was removed and replaced after just under six hours. | [5 empty] | The investigation into the high purge pressure has been completed since the original report was submitted. the device was not returned for investigation. it was reported that upon running the pump, there was a motor current spike followed by gradual buildup in the purge pressure for next 10 minutes. saturated flows were also seen during this time. therefore, per the data log review, the root cause of the high purge pressure issue was most likely biomaterial ingestion. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
29315 | 1220648-2024-08515 | 18953838 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18953838&pc=OZD | 12/04/2022 | 03/13/2025 |  | 03/21/2024 |  |  | 12/04/2022 | 03/21/2024 |  |  | 07/31/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023259174 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/04/2022 |  |  |  |  | The user facility reported a 85-year-old male was being implanted with an impella cp device for mechanical circulatory support and the impella cp pump was unable to advance due to the patient's vasculature. therefore, the impella cp implant was aborted. | [5 empty] | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29316 | 1220648-2024-08516 | 18953888 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18953888&pc=OZD | 03/11/2024 | 04/18/2024 |  | 03/21/2024 |  |  | 03/11/2024 | 03/21/2024 |  |  |  |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025438146 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/11/2024 |  |  |  |  | The user facility reported a 69-year-old male with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the impella surgical site was red and inflamed causing pain and swelling at the site and down the patient¿s arm. after investigating the problem, a small hematoma was evacuated from the impella pocket and the symptoms improved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use states ¿assess access site for bleeding and hematoma.¿
29317 | 1220648-2024-08518 | 18953953 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18953953&pc=OZD | 08/01/2022 | 03/13/2025 |  | 03/21/2024 |  |  | 08/01/2022 | 03/21/2024 |  |  | 12/31/2023 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023165859 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/01/2022 |  |  |  |  | The user facility reported a male patient noted as high risk percutaneous coronary intervention needing an impella cp device for mechanical circulatory support. it was reported that the medical team experienced difficulty implanting the impella due to the patient's vascular anatomy and to calcified lesions. | [5 empty] | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records
29318 | 1220648-2024-08519 | 18953987 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18953987&pc=OZD | 07/05/2022 | 03/13/2025 |  | 03/21/2024 |  |  | 07/05/2022 | 03/21/2024 |  |  | 11/30/2023 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023161905 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/05/2022 |  |  |  |  | The user facility reported a 50-year-old male was implanted with an impella cp device for mechanical circulatory support. there was no left femoral access and the right iliac was adequate at access, but it was narrowed and calcified up near the aortic bifurcation, therefore, shockwave was used to open vessels and the impella cp was placed successfully. | [5 empty] | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29319 | 1220648-2024-08517 | 18953997 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18953997&pc=OZD | 11/02/2022 | 05/13/2025 |  | 03/21/2024 |  |  | 11/02/2022 | 03/21/2024 |  |  | 07/31/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023251657 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Adverse Event Without Identified Device or Use Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 11/02/2022 |  |  |  |  | The user facility reported a 35-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, echocardiography was performed and it revealed the impella device was too far into the left ventricle by 2cm and a thrombus was visualized. the patient was escalated to an impella 5.5 prior to repositioning the device. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. correction section h6: the medical device problem code was inadvertently left out in the initial report which has been added to this report.
29320 | 1220648-2024-08514 | 18954024 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18954024&pc=OZD | 03/11/2024 | 04/29/2024 |  | 03/21/2024 |  |  | 03/11/2024 | 03/21/2024 |  |  | 11/30/2025 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP) |  |  | ABIOMED, INC. |  |  | US |  | 2025438300 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | No | Other | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/11/2024 |  |  |  |  | The user facility reported a 73-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient had an attempted impella cp implant and there was a catheter kink. there were multiple interactions trying to pass bilateral iliac stents at the bifurcation. the impella flows would not exceed 1.2 liters per minute. the impella sheath and pump were replaced. | [5 empty] | The investigation for the reported low pump flow and delivery issues has been completed. neither the impella device nor the data logs were returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the cause of the low pump flow issue most likely was cannula kink due to patient condition as the patient had stents at bifurcation. the cause of the delivery issue was patient condition related due patient vascular condition. potential adverse events (united states). "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ to conform to updated procedures, section d6 has revised with updated information. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29321 | 1220648-2024-08520 | 18954029 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18954029&pc=OZD | 11/23/2022 | 03/13/2025 |  | 03/21/2024 |  |  | 11/23/2022 | 03/21/2024 |  |  | 07/31/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023253391 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/23/2022 |  |  |  |  | The user facility reported a 63-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the access site oozing was noticed during a dressing. the physician placed a suture around skin near reposition sheath and oozing slowed down. the site was dressed using best practices and angle match was performed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29322 | 1220648-2024-08522 | 18954126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18954126&pc=OZD | 03/11/2024 | 04/29/2024 |  | 03/21/2024 |  |  | 03/11/2024 | 03/21/2024 |  |  | 11/30/2025 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024430032 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Fracture | Difficult to Remove | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/11/2024 |  |  |  |  | The user facility reported a 86-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp supported a percutaneous coronary intervention procedure successfully. upon removal of the impella cp, there was limited ability to maintain hemostasis after the 14 french sheath fractured, leading to significant bleeding in the access site. after wire access was restored, the 14 french gore sheath was placed to obtain hemostasis. fluids and two units of blood were given. the patient expired in the procedural area. | [5 empty] | The investigation for the reported introducer damage, access site bleeding, and removal issues has been completed. the impella device was not returned for evaluation. clinical details were not sufficient to determine the root cause. the root cause of the access site bleeding issue was not determined since product was not returned and the clinical information was enough to be conclusive of the root cause of the bleeding. the root cause of the introducer damage issue was not determined since product was not returned and insufficient clinical information regarding the location and reason of the damage was given. potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ to conform to updated procedures, sections d6 and h10 have been revised with updated information. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
29323 | 1220648-2024-08521 | 18954129 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18954129&pc=OZD | 11/30/2022 | 03/13/2025 |  | 03/21/2024 |  |  | 11/30/2022 | 03/21/2024 |  |  | 08/31/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023270038 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/30/2022 |  |  |  |  | The user facility reported a 80-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the impella sheath was found to be totally occluded distal perfusion. the physician made the decision to explant the impella to minimize complications | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition. any concomitant medication. vascular size or variations thereof. detailed description of the symptoms experienced by the patient. physical examination findings, including pulse assessment and visual examination of the affected limb. diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. laboratory test results, including blood tests for markers of inflammation or clotting disorders. any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
29324 | 1220648-2024-08523 | 18954160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18954160&pc=OZD | 07/05/2022 | 05/21/2025 |  | 03/21/2024 |  |  | 07/05/2022 | 03/21/2024 |  |  | 12/31/2023 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023178579 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/05/2022 |  |  |  |  | The user facility reported a 51-year-old male was implanted with an impella cp device for mechanical circulatory support and that the pump stopped overnight, was restarted, and stopped again after several hours and was subsequently removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29325 | 1220648-2024-08524 | 18954333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18954333&pc=OZD | 12/04/2022 | 03/13/2025 |  | 03/21/2024 |  |  | 12/04/2022 | 03/21/2024 |  |  | 02/29/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023194512 |  | 0048-0044 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/04/2022 |  |  |  |  | The user facility reported a 67-year-old female was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced oozing around the reposition sheath. the oozing was treated with pressure and temporary discontinuation of heparin. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29326 | 1220648-2024-08525 | 18954648 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18954648&pc=OZD | 07/11/2022 | 03/13/2025 |  | 03/21/2024 |  |  |  | 03/21/2024 |  |  | 02/29/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023203091 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/11/2022 |  |  |  |  | The user facility reported a 58-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported while being placed on impella cp support, the pump flows were unable to titrate beyond p6 without suction. the impella was found to be too deep into the left ventricle with transesophageal echocardiography and the cannula was drawn back. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29327 | 1220648-2024-08526 | 18955019 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18955019&pc=OZD | 03/11/2024 | 05/02/2024 |  | 03/21/2024 |  |  |  | 03/21/2024 |  |  | 09/30/2025 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025444342 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Perforation of Vessels | [10 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/11/2024 |  |  |  |  | The user facility reported a 77-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was on impella cp support due to being hypotensive and unstable. during the impella implant, the patient went into ventricular tachycardia (vt) and required multiple shocks. after the impella implant, prior to transport, the patient had oozing from the access site. purse string sutures and manual pressure were applied. after the transport, there was severe bleeding from the left femoral artery. the patient was given eight units of blood products. the left femoral artery was repaired by the vascular team and the impella was removed. reportable due to patient's vt, access site bleeding and vascular damage. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ | The investigation has been completed. the impella cp pump was not returned. as per clinical, bleeding did not resolve with purse string suture and vascular surgeon was consulted for left femoral artery bleeding. hemostasis achieved by surgical intervention. the cause of the vascular damage issue was most likely use related per clinical. as the necessary information was not provided, the root cause of the vt/vf was not determined. to determine the true root cause of the ventricular tachycardia/ventricular fibrillation (vt/vf), the clinical information would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d.4 and h.4 lot number, expiration date, unique identifier (udi) # and device manufacture date were revised as previously entered incorrectly. d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 health effect - clinical code and health effect - impact code were revised as some codes were inadvertently left off the previously submitted manufacture device report. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures.
29328 | 1220648-2024-08527 | 18955035 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18955035&pc=OZD | 07/11/2022 | 03/13/2025 |  | 03/21/2024 |  |  |  | 03/21/2024 |  |  | 03/31/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023214394 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/11/2022 |  |  |  |  | The user facility reported a 74-year-old female was implanted with an impella cp device for mechanical circulatory support. it was reported after impella cp was inserted, a clot was found in the impella cp. the impella cp pump was explanted and replaced. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. the failure mode of ¿thrombus¿ is to be used when a thrombus is noted but cannot be connected to any other adverse events or decrease in pump performance. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions. ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency). ¿relevant laboratory test results, such as coagulation profiles and platelet counts. ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). ¿medications or treatments that the patient was receiving at the time of thrombus formation. ¿surgical or procedural details if applicable (e.g, cardiac catheterization). ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined.
29329 | 1220648-2024-08528 | 18955441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18955441&pc=OZD | 07/06/2022 | 10/01/2025 |  | 03/21/2024 |  |  |  | 03/21/2024 |  |  | 02/29/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200023 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | Positioning Problem | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/06/2022 |  |  |  |  | The user facility reported a 69-year-old male was implanted with an impella cp device for mechanical circulatory support. the impella cp was pulled back to descending aorta during grafts being sutured on, and a large hematoma was noticed. additionally, there was oozing at the groin access site. the impella cp device was explanted. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. d6a and d6b revised from the initial manufacturer device report 1220648-2024-08528 in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2024-08528 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-08528 in accordance with updated procedures. g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2024-08528. h6 medical device problem code 2906 was erroneously entered on the initial manufacturer device report number 1220648-2024-08528.
29330 | 1220648-2024-08529 | 18955566 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18955566&pc=OZD | 09/17/2022 | 03/17/2025 |  | 03/21/2024 |  |  |  | 03/21/2024 |  |  | 06/30/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023236828 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/17/2022 |  |  |  |  | The user facility reported a 42-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the physician repositioned the sheath with echo guidance and the pump's insertion site was oozing profusely, but significantly improved with angle matching and additional forward tension with manual pressure. the device was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29331 | 1220648-2024-08530 | 18955656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18955656&pc=OZD | 10/17/2022 | 03/17/2025 |  | 03/21/2024 |  |  |  | 03/21/2024 |  |  | 07/31/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023251699 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/17/2022 |  |  |  |  | The user facility reported a 77-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. a foley was placed and the patient immediately showed signs of hematuria. there was a slow wean of pressors to decrease systemic vascular resistance. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29332 | 1220648-2024-08531 | 18955741 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18955741&pc=OZD | 10/19/2022 | 06/03/2025 |  | 03/21/2024 |  |  |  | 03/21/2024 |  |  | 07/31/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023255936 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/19/2022 |  |  |  |  | The user facility reported a 68-year-old female was implanted with an impella cp device for mechanical circulatory support. the patient experienced limb ischemia. an external femorofemoral bypass was performed. the patient was successfully weaned off the impella cp. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition. any concomitant medication. vascular size or variations thereof. detailed description of the symptoms experienced by the patient. physical examination findings, including pulse assessment and visual examination of the affected limb. diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. laboratory test results, including blood tests for markers of inflammation or clotting disorders. any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
29333 | 1220648-2024-08532 | 18955783 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18955783&pc=OZD | 11/11/2022 | 09/26/2024 |  | 03/21/2024 |  |  |  | 03/21/2024 |  |  | 07/31/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023256315 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/11/2022 |  |  |  |  | The user facility reported a 55-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient experienced a stroke while on support. the patient later expired while on the impella cp support. no allegations were made towards the impella cp for cause of death. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08532 was provided in sections b5, h1 and h6.
29334 | 1220648-2024-08533 | 18955878 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18955878&pc=OZD | 10/21/2022 | 03/17/2025 |  | 03/21/2024 |  |  |  | 03/21/2024 |  |  | 04/30/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023223891 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/21/2022 |  |  |  |  | The user facility reported a 81-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. during support, there were no pulses dopplerable and the patient's leg was mottled. contralateral access was obtained and balloon tamponade was performed, along with covered stent placement. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | When making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
29335 | 1220648-2024-08534 | 18955953 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18955953&pc=OZD | 08/05/2022 | 03/17/2025 |  | 03/21/2024 |  |  |  | 03/21/2024 |  |  | 04/30/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023221608 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/05/2022 |  |  |  |  | The user facility reported a 70-year-old female patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that after the pump implant through the right femoral artery, the patient's right leg became cold and had no detectable pulse. consequently, the pump was inserted via an external bypass through the left femoral artery into the right superficial femoral artery, an antegrade sheath was put in place and the refusion sheath was upsized. the pump was explanted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29336 | 1220648-2024-08535 | 18956100 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18956100&pc=OZD | 08/06/2022 | 03/21/2024 |  | 03/21/2024 |  |  |  | 03/21/2024 |  |  | 03/31/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023209262 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/06/2022 |  |  |  |  | The user facility reported a 75-year-old female patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that there was bleeding/oozing noted at the impella access site. the patient was given one unit of blood and a femostop was applied to control the bleeding. the impella cp was removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records.
29337 | 1220648-2024-08536 | 18956149 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18956149&pc=OZD | 07/06/2022 | 03/17/2025 |  | 03/21/2024 |  |  |  | 03/21/2024 |  |  | 08/31/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023193209 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/06/2022 |  |  |  |  | The user facility reported a 52-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding which was resolved by balloon tamponade and wiring of the leg. additionally, the pump experienced low pump flow which was resolved by removal of the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29338 | 1220648-2024-08538 | 18956183 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18956183&pc=OZD | 07/07/2022 | 03/07/2025 |  | 03/21/2024 |  |  | 07/07/2022 | 03/21/2024 |  |  | 03/31/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023211210 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/07/2022 |  |  |  |  | The user facility reported a 68-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the echo at bedside measured inlet at 3.7cm and it appeared oriented towards the mitral valve, but physician opted to not reposition. however, the pump was pulled back in the aorta and the pigtail was still across on echo while the physician advanced with the mid left ventricle position. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29339 | 1220648-2024-08539 | 18956187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18956187&pc=OZD | 08/10/2022 | 03/21/2024 |  | 03/21/2024 |  |  |  | 03/21/2024 |  |  | 05/31/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2023227236 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Patient-Device Incompatibility | Difficult to Advance | Positioning Problem |  |  |  |  | Hematoma | Ischemia | [36 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 08/10/2022 |  |  |  |  | The user facility reported a 53-year-old female patient needing an impella cp for mechanical circulatory support. the complainant reported that the pump kept alarming a placement signal low alarm. in addition, there was limb ischemia as the sheath became occlusive in the patient's iliac. attempts to position the device were unsuccessful as this proved to be volatile and unstable. to prevent any further complications the decision was made to explant the pump. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
29340 | 1220648-2024-08540 | 18956195 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18956195&pc=OZD | 07/08/2022 | 03/10/2025 |  | 03/21/2024 |  |  |  | 03/21/2024 |  |  | 03/31/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023204522 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/08/2022 |  |  |  |  | The user facility reported a 65-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the patient decompensated into extracorporeal membrane oxygenation (ecmo) and is in active diabetic ketoacidosis (dka). persistent oozing was noted at the cannulation access sites. one (1) unit of packed red blood cells (prbc) was required. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleeding issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
29341 | 1220648-2024-08542 | 18956202 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18956202&pc=OZD | 08/10/2022 | 03/21/2024 |  | 03/21/2024 |  |  | 08/10/2022 | 03/21/2024 |  |  | 02/29/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200319 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/10/2022 |  |  |  |  | The complainant reported a patient undergoing coronary artery bypass was planned to be implanted with an impella cp for mechanical circulatory support. the surgeon was planning to implant the impella cp via direct approach after suturing the graft onto the aorta. the delivery of the impella cp was unsuccessful due to the anatomy of the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29342 | 1220648-2024-08543 | 18956214 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18956214&pc=OZD | 08/12/2022 | 05/19/2025 |  | 03/21/2024 |  |  |  | 03/21/2024 |  |  | 04/30/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218681 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/12/2022 |  |  |  |  | The user facility reported a 66-year-old male patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient had hemolysis and thrombocytopenia during pump support. treatment information is unknown, however, the issues persisted until the device was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29343 | 1220648-2024-08544 | 18956233 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18956233&pc=OZD | 07/11/2022 | 09/30/2025 |  | 03/21/2024 |  |  |  | 03/21/2024 |  |  | 12/31/2022 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022023649 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 07/11/2022 |  |  |  |  | The user facility reported a 63-year-old patient needing mechanical circulatory support. it was reported that the impella cp was repositioned due to the patient having hemolysis. the impella cp was repositioned, which alleviated the patient¿s bleeding at the access site. no blood products were required. the patient was weaned and the impella cp was explanted. | It was reported that shortly after the impella cp was explanted the patient decompensated and expired. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined. b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2024-08544. b2 death and date of death was inadvertently omitted on the initial manufacturer device report number 1220648-2024-08544. b5 patient outcome was inadvertently omitted on the initial manufacturer device report number 1220648-2024-08544. d6a and d6b revised from the initial manufacturer device report 1220648-2024-08544 in accordance with updated procedures. e4 should have been left blank on the initial manufacturer device report number 1220648-2024-08544 as it is unknown if the initial reporter also sent the report to the food and drug administration. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2024-08544 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-08544 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2024-08544. h1 type of report was erroneously entered on the initial manufacturer device report number 1220648-2024-08544. h6 health effect - clinical code 1888, health effect - impact code 1802, and medical device problem code 3009 was inadvertently omitted on the initial manufacturer device report number 1220648-2024-08544. h6 medical device problem code 2993 was erroneously entered on the initial manufacturer device report number 1220648-2024-08544. h6 component code revised from the initial submission in accordance with updated procedures.
29344 | 1220648-2024-08545 | 18956241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18956241&pc=OZD | 08/13/2022 | 09/13/2024 |  | 03/21/2024 |  |  |  | 03/21/2024 |  |  | 05/31/2024 |  | 03/21/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA |  |  | ABIOMED, INC. |  |  | US |  | 2023228825 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/13/2022 |  |  |  |  | The user facility reported a 62-year-old male patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. there was access site bleeding. this was treated through applying manual pressure at the site, forward suturing and angle match dressing. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08532 was provided.
29345 | 1220648-2024-08322 | 18939191 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18939191&pc=OZD | 10/20/2022 | 03/07/2025 |  | 03/20/2024 |  |  | 10/20/2022 | 03/20/2024 |  |  | 07/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023256244 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/20/2022 |  |  |  |  | The user facility reported a 51-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that there were a couple instances of the impella diving too deep into the left ventricle, but were able to remove slack and pull back a couple of centimeters. proper position was achieved and single access was gained. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29346 | 1220648-2024-08323 | 18939296 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18939296&pc=OZD | 10/21/2022 | 03/20/2024 |  | 03/20/2024 |  |  | 10/21/2022 | 03/20/2024 |  |  | 07/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023256439 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 10/21/2022 |  |  |  |  | The user facility reported a 55-year-old patient needing mechanical circulatory support. it was reported that the patient was found to be experiencing minor hemolysis. due to the hemolysis, the impella was explanted a day early. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
29347 | 1220648-2024-08324 | 18939313 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18939313&pc=OZD | 10/21/2022 | 03/20/2024 |  | 03/20/2024 |  |  | 10/21/2022 | 03/20/2024 |  |  | 06/30/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023237210 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 10/21/2022 |  |  |  |  | The user facility reported an 86-year-old patient needing mechanical circulatory support. it was reported that the patient on impella cp had a small groin hematoma during implantation. repositioning sheet, manual pressure and angle matching, advancement of repositioning of the sheet did not help the hematoma which was larger at the insertion site. pump was weaned off and was explanted with balloon tamponade. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records.
29348 | 1220648-2024-08325 | 18939350 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18939350&pc=OZD | 10/21/2022 | 03/20/2024 |  | 03/20/2024 |  |  | 10/21/2022 | 03/20/2024 |  |  | 07/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023256462 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 10/21/2022 |  |  |  |  | The user facility reported a 65-year-old patient needing mechanical circulatory support. it was reported that the patient was on impella cp support due to non-disclosed cardiovascular medical history. while on support, the patient developed a hematoma at the access site. the patient was weaned from the impella followed by the impella being explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29349 | 1220648-2024-08326 | 18939375 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18939375&pc=OZD | 09/13/2022 | 09/26/2025 |  | 03/20/2024 |  |  | 09/13/2022 | 03/20/2024 |  |  | 04/30/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023215519 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/13/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 54-year-old female patient for mechanical circulatory support. it was reported while on support, the patient required repositioning of the pump under echocardiography. procedure was completed without further issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29350 | 1220648-2024-08328 | 18939653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18939653&pc=OZD | 10/06/2022 | 03/20/2024 |  | 03/20/2024 |  |  | 10/06/2022 | 03/20/2024 |  |  | 03/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023225724 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/06/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 71-year-old female patient for mechanical circulatory support. it was reported that after the impella cp was weaned and removed, a post angiogram revealed no distal flow. the patient was referred for cutdown procedure with a vascular surgeon. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records.
29351 | 1220648-2024-08329 | 18939681 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18939681&pc=OZD | 03/04/2024 | 03/20/2024 |  | 03/20/2024 |  |  | 03/04/2024 | 03/20/2024 |  |  |  |  | 03/20/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1505149 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 03/04/2024 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) displayed a controller error (yellow alarm). the aic was returned for evaluation. there was no patient involvement. | [5 empty] | The investigation has been completed. the console was tested. the failure mode was not reproduced while running an optical pump simulator. there were no issues with the console hardware. console logs from the reported day of event did not show any placement signal issues, as the pump was not connected. logs were consistent with the complaint and showed that during case start, the pump was momentarily disconnected, and after reconnecting, the console presented with controller error alarm. the logs also reported processsamplingdatafromopticalbench: sentstopacquisitioncmd ack, indicating ossiopsens received thread got a data sampling packet with an unexpected length. real time (rt) logs confirmed that all signals received from the optical bench were lastly reported as +16384 values and then no more samples were recorded. the absence of signal to noise ratio (snr) samples impacted the console¿s ability to recognize the pump disconnection and required the user to perform a hard shutdown. the log entry processsamplingdatafromopticalbench: sentstopacquisitioncmd ack indicates the sentstopacquisitioncmdflag was set when entering ossi_sampling_stopped_state which eventually led to a false ¿communication stopped¿ condition. the root cause of the controller error was due to an aic software issue.
29352 | 1220648-2024-08330 | 18939825 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18939825&pc=OZD | 11/03/2022 | 03/12/2025 |  | 03/20/2024 |  |  | 11/03/2022 | 03/20/2024 |  |  | 04/30/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023224024 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/03/2022 |  |  |  |  | The user facility reported a 76-year-old male in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. upon receiving the pump, the patient began moving on the stretcher and the groin site began to ooze. manual pressure was held, femstop applied and bleeding had subsided. the patient was later transferred for coronary artery bypass graft evaluation and no oozing was noted. the device was weaned from the patient. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29353 | 1220648-2024-08331 | 18939832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18939832&pc=OZD | 03/04/2024 | 04/23/2024 |  | 03/20/2024 |  |  | 03/04/2024 | 03/20/2024 |  |  |  |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025443042 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/04/2024 |  |  |  |  | The user facility reported a 70-year-old female in acute myocardial infarction cardiogenic shock was to be implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the impella 5.5 pump was unable to advance past the subclavian innominate angle during attempted delivery. after multiple failed attempts, the impella 5.5 implant was aborted and an impella cp was placed for patient support. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because of resistance at innominate artery in spite of best practices being followed.
29354 | 1220648-2024-08332 | 18939962 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18939962&pc=OZD | 03/04/2024 | 04/22/2024 |  | 03/20/2024 |  |  | 03/04/2024 | 03/20/2024 |  |  |  |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025443142 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | American Indian Or Alaskan Native | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/04/2024 |  |  |  |  | The user facility reported a 61-year-old female in acute myocardial infarction cardiogenic shock was to be implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the impella 5.5 pump was unable to cross the aortic valve during attempted delivery. after multiple failed attempts, the decision was made to abort the implant. venous-arterial extracorporeal membrane oxygenation (va ecmo) was subsequently placed for patient support. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue was most likely due to patient condition since patient had small femoral vessels and since the md was unable to cross aortic valve with catheter and wire.
29355 | 1220648-2024-08333 | 18940017 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18940017&pc=OZD | 10/06/2022 | 03/20/2024 |  | 03/20/2024 |  |  | 10/06/2022 | 03/20/2024 |  |  | 03/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022057010 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Atrial Fibrillation | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/06/2022 |  |  |  |  | The user facility reported a 69-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported an echocardiography was performed and revealed the impella cp catheter was wrapped around a papillary muscle. the physician attempted to reposition the impella bedside, but was unsuccessful in freeing the pigtail from the papillary muscle, although was able to improve the pump flows. it was reported while patient was on impella cp support, the patient experienced atrial fibrillation with rapid ventricular response and hemolysis. the pump was repositioned again, with echocardiography and the device was successfully positioned away from the papillary muscle. the patient received a unit of blood and began continuous renal replacement therapy as a result of the hemolysis and the patient received pressor support after the episode of atrial fibrillation. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
29356 | 1220648-2024-08334 | 18940059 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18940059&pc=OZD | 11/03/2022 | 03/12/2025 |  | 03/20/2024 |  |  | 11/03/2022 | 03/20/2024 |  |  | 12/31/2023 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023178882 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/03/2022 |  |  |  |  | The user facility reported a 65-year-old male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the patient had some bleeding (oozing) at the impella cp insertion site. in the operating room, the groin bleeding could not be controlled, even with manual pressure applied. the device was left in place as a result of the bleeding and the patient was transferred to the icu. femstop was placed above the repositioning sheath and no bleeding at site was noted. while in the icu, the impella has been explanted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29357 | 1220648-2024-08335 | 18940207 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18940207&pc=OZD | 10/21/2022 | 03/12/2025 |  | 03/20/2024 |  |  | 10/21/2022 | 03/20/2024 |  |  | 07/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023251059 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/21/2022 |  |  |  |  | The user facility reported a 55-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. patient had oozing around the sheath. angle matching was done, and purse string suture was utilized around the sheath to control the bleeding due to the skin nick. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29358 | 1220648-2024-08336 | 18940232 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18940232&pc=OZD | 11/04/2022 | 03/20/2024 |  | 03/20/2024 |  |  | 11/04/2022 | 03/20/2024 |  |  | 07/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023255505 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/04/2022 |  |  |  |  | The user facility reported a 72-year-old male on impella cp support for the treatment of acute myocardial infarction/ cardiogenic shock. the patient developed bleeding at the impella cp insertion site. the bleeding was treated with the application of pressure dressing and switching the patient to a bicarbonate purge. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
29359 | 1220648-2024-08337 | 18940267 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18940267&pc=OZD | 10/21/2022 | 03/10/2025 |  | 03/20/2024 |  |  | 10/21/2022 | 03/20/2024 |  |  | 06/30/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023237418 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 25 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/21/2022 |  |  |  |  | The user facility reported a 25-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed oozing at the pump site. the oozing was initially resolved after the dressing was changed. however, the groin site started to ooze again heavily and surgicel, gauze and pressure was applied to the site to gain control and the dressing was reapplied. the patient was also provided with 2 units of packed red blood cells. after the bleeding had resided, the patient had an elevated act. it was recommended that the heparin concentration be reduced by 50% due to the elevated activated clotting time (act). | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleeding issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
29360 | 1220648-2024-08338 | 18940369 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18940369&pc=OZD | 03/04/2024 | 04/23/2024 |  | 03/20/2024 |  |  | 03/04/2024 | 03/20/2024 |  |  | 08/31/2025 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025443770 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/04/2024 |  |  |  |  | The user facility reported a 57-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed ischemia in their left leg during impella cp support. a distal perfusion cannula was added to manage the condition. support was maintained with the implanted pump following the intervention. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
29361 | 1220648-2024-08339 | 18940391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18940391&pc=OZD | 10/22/2022 | 03/10/2025 |  | 03/20/2024 |  |  | 10/22/2022 | 03/20/2024 |  |  | 07/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023257370 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/22/2022 |  |  |  |  | The user facility reported a 62-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. bleeding arose at the insertion site requiring the pump to be removed, a 14f to be placed, and the device to be reinserted while the sheath was left in place. due to the patient acutely decompensated overnight, the decision was made to remove the impella. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleeding issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
29362 | 1220648-2024-08340 | 18940427 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18940427&pc=OZD | 10/28/2022 | 05/30/2025 |  | 03/20/2024 |  |  | 10/28/2022 | 03/20/2024 |  |  | 04/30/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023218103 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/28/2022 |  |  |  |  | The user facility reported a 84-year-old female noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the patient was unable to be weaned from the device, due to unstable vitals resulting in increased vasopressor support. due to the patient being baseline anemia, 1unit of packed red blood cells (prbcs) was transfused. the left leg was noted to be cold without a palpable or doppler pulse. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events: acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported limb ischemia and access site bleeding issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
29363 | 1220648-2024-08341 | 18940459 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18940459&pc=OZD | 11/05/2022 | 03/20/2024 |  | 03/20/2024 |  |  | 11/05/2022 | 03/20/2024 |  |  | 05/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023228807 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/05/2022 |  |  |  |  | The user facility reported a 73-year-old male on impella cp support after infracting during an angiogram. the patient developed lots of oozing after the peel away sheath was removed, after the impella was sutured in place. an extra suture was added and pressure dressing was applied. post sternotomy redo the patient no longer required impella cp support and the doctor removed at bedside. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29364 | 1220648-2024-08342 | 18940474 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18940474&pc=OZD | 12/06/2022 | 03/12/2025 |  | 03/20/2024 |  |  | 12/06/2022 | 03/20/2024 |  |  | 10/31/2023 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022155882 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/06/2022 |  |  |  |  | The user facility reported a 65-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient¿s groin dressing was initially dry and intact in the cath lab; however, upon arrival to the critical care unit (ccu), the dressing was saturated. the attending physician called and instructed the perclose's be tightened, and the site re-dressed to achieve hemostasis. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29365 | 1220648-2024-08343 | 18940478 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18940478&pc=OZD | 11/02/2022 | 03/11/2025 |  | 03/20/2024 |  |  | 11/02/2022 | 03/20/2024 |  |  | 06/30/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023236786 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/02/2022 |  |  |  |  | The user facility reported a 65-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was on impella cp and extra-corporeal membrane oxygenation (ecmo) had a drop in hemoglobin as a result was given 4 units of packed red blood cells. an internal bleed is suspected. the decision to explant the impella was made after the team was unable to perform vav (triple cannulation) strategy. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | The investigation into the reported issue has been completed. no device was received for evaluation. pump failed post sterile inspection due to not having medcert approval. this issue is unrelated to the failure mode due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29366 | 1220648-2024-08344 | 18940586 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18940586&pc=OZD | 03/04/2024 | 04/18/2024 |  | 03/20/2024 |  |  | 03/04/2024 | 03/20/2024 |  |  | 06/30/2025 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024431013 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Deliver | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/04/2024 |  |  |  |  | The user facility reported a 82-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. during implantation, the impella cp became displaced and was unable to advance. the impella was removed. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause was determined to be use related as impella became displaced due to cpr. there was no product returned and no clinical information supporting patient anatomy complications.
29367 | 1220648-2024-08345 | 18940660 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18940660&pc=OZD | 11/07/2022 | 10/20/2025 |  | 03/20/2024 |  |  | 11/07/2022 | 03/20/2024 |  |  | 06/30/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023249101 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/07/2022 |  |  |  |  | The user facility reported a 69-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding from the impella cp left groin access site. manual pressure was unsuccessful. the pump was removed and re-inserted through a new sheath. in light of intermittent loss of pulsitility in left lower extremity, the decision was made to wean and remove the impella cp device. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
29368 | 1220648-2024-08346 | 18940711 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18940711&pc=OZD | 09/01/2022 | 03/20/2024 |  | 03/20/2024 |  |  | 09/01/2022 | 03/20/2024 |  |  | 05/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023229692 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 09/01/2022 |  |  |  |  | The user facility reported a 76-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient developed bleeding at the impella insertion site when the reposition sheath was placed. as a result of the bleeding, the impella was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
29369 | 1220648-2024-08347 | 18940779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18940779&pc=OZD | 07/21/2022 | 05/23/2025 |  | 03/20/2024 |  |  | 07/21/2022 | 03/20/2024 |  |  | 12/31/2023 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023181425 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 07/21/2022 |  |  |  |  | The user facility reported a 71-year-old female was implanted with an impella cp device for mechanical circulatory support due to acute myocardial infarction and cardiogenic shock. while on support, the surgeon attempted multiple times to reopen the left anterior descending artery; however, due to vascular damage the right femoral artery was opened. the patient kept losing flow. after placing the stent, some flow was restored to circumflex. | Family and heart care team withdrew all care. patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08347 was provided in sections b2, b5,d6, h1, and h6. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29370 | 1220648-2024-08348 | 18940868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18940868&pc=OZD | 10/08/2022 | 04/04/2025 |  | 03/20/2024 |  |  | 10/08/2022 | 03/20/2024 |  |  | 03/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023211277 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/08/2022 |  |  |  |  | The user facility reported a 69-year-old female was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the volume was turned to p1 due to chugging/suction. the physician tried to reposition impella cp away from the mitral valve, but was unsuccessful and the pump was exchanged for a new one. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29371 | 1220648-2024-08349 | 18940876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18940876&pc=OZD | 11/08/2022 | 03/12/2025 |  | 03/20/2024 |  |  | 11/08/2022 | 03/20/2024 |  |  | 08/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023264041 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/08/2022 |  |  |  |  | The user facility reported a 77-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella cp flows were intermittently poor along with waveforms. under fluoroscopy, it was determined impella cp was under papillary muscle. despite multiple attempts, the impella cp was not successfully repositioned. the impella cp was removed and ran underwater and placed in surgical mode. the impella cp was successfully reimplanted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29372 | 1220648-2024-08350 | 18941028 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18941028&pc=OZD | 03/04/2024 | 04/23/2024 |  | 03/20/2024 |  |  | 03/04/2024 | 03/20/2024 |  |  | 10/31/2025 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024422635 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 24 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/04/2024 |  |  |  |  | The user facility reported a 24-year-old female with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed bleeding at the groin area. femstop was applied, impella angle matched and re-sutured performed, and two units of red blood cells were given as treatment. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical details provided was sufficient to determine the root cause. the root cause of the access site bleeding was most likely due to use issue as the impella angle matched and resutured as per the clinical description.
29373 | 1220648-2024-08351 | 18941061 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18941061&pc=OZD | 11/08/2022 | 03/20/2024 |  | 03/20/2024 |  |  | 11/08/2022 | 03/20/2024 |  |  | 04/30/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023215989 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/08/2022 |  |  |  |  | The user facility reported a 41-year-old female on impella cp support due to left anterior descending artery disease. while on support, the patient experienced access site bleeding. the impella cp was weaned and removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29374 | 1220648-2024-08352 | 18941131 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18941131&pc=OZD | 03/04/2024 | 04/11/2024 |  | 03/20/2024 |  |  | 03/04/2024 | 03/20/2024 |  |  |  |  | 03/20/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1411424 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/04/2024 |  |  |  |  | The user facility reported they had to replace the automated impella controller (aic) during a support due to a yellow alarm. the instructions for use (ifu) were followed and a backup controller was used. no patient harm was reported. | [5 empty] | The investigation for the reported console (aic) yellow alarm issue has been completed. the aic was returned for evaluation. the console was tested and the reported failure was unable to reproduced by updating the system time while running the optical pump. found the front panel optical connector contaminated, unable to remove the debris upon cleaning (unrelated to the reported failure mode). data logs were reviewed which revealed the console received the controller error (loss of comm (pbm2)) ¿ alarm event set #1025, which confirms the reported failure mode. event set #1025 was never resolved, and the console was shut down five minutes after the alarm. the root cause for the intermittent loss of communication between pbm 2 and the 4-in-1 was not determined due to the inability to reproduce. added-d9 device return date. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29375 | 1220648-2024-08353 | 18941157 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18941157&pc=OZD | 07/22/2022 | 03/12/2025 |  | 03/20/2024 |  |  | 07/22/2022 | 03/20/2024 |  |  | 02/29/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023196170 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Inaccurate Flow Rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/22/2022 |  |  |  |  | The user facility reported a 54-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. there was continuous suction during impella support. it was recommended to perform angle matching with the pump, however, even when the pump was pulled back, suction alarms persisted. recommendations on how to properly shut down the aic were going to be discussed, however, it had been decided at that point to withdraw care. impella was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29376 | 1220648-2024-08354 | 18941167 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18941167&pc=OZD | 12/06/2022 | 03/12/2025 |  | 03/20/2024 |  |  | 12/06/2022 | 03/20/2024 |  |  | 08/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023267973 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/06/2022 |  |  |  |  | The user facility reported a 79-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was on impella cp support due to left bundle branch block (lbbb). while on support, suction alarms started. after several attempts to reposition, a clot was noted. patient was given anticoagulant, and impella was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29377 | 1220648-2024-08355 | 18941210 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18941210&pc=OZD | 07/22/2022 | 03/20/2024 |  | 03/20/2024 |  |  | 07/22/2022 | 03/20/2024 |  |  | 03/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023210796 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/22/2022 |  |  |  |  | The user facility reported a 67-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported while on support, the patient experienced a severe right groin hematoma. pressure was held to the site and the patient received packed red blood cells (prbc). the impella cp device was explanted. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29378 | 1220648-2024-08357 | 18941251 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18941251&pc=OZD | 07/22/2022 | 03/20/2024 |  | 03/20/2024 |  |  | 07/22/2022 | 03/20/2024 |  |  | 03/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023209551 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/22/2022 |  |  |  |  | The user facility reported a 73-year-old female was implanted with an impella cp device for mechanical circulatory support. it was reported while on support, the patient experienced large hematoma and was bleeding from the access site. pressure was held to the site and the patient received 2 units of packed red blood cells (prbc). the impella cp device was explanted | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29379 | 1220648-2024-08356 | 18941263 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18941263&pc=OZD | 10/08/2022 | 03/12/2025 |  | 03/20/2024 |  |  | 10/08/2022 | 03/20/2024 |  |  | 06/30/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023249497 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Impairment | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/08/2022 |  |  |  |  | The user facility reported a 64-year-old male was implanted with an impella device for mechanical circulatory support. the patient developed hemolysis that was not resolved with repositioning the impella cp. the patient had hemolyzed labs and no urine output. the patient had tolerated sustained low-efficiency dialysis (sled) dialysis treatment. the patient was escalated to an impella 5.5. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29380 | 1220648-2024-08358 | 18941322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18941322&pc=OZD | 11/09/2022 | 05/21/2025 |  | 03/20/2024 |  |  | 11/09/2022 | 03/20/2024 |  |  | 08/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023262285 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/09/2022 |  |  |  |  | The user facility reported a 58-year-old male on impella cp for a bypass surgery support. coming off bypass surgery, suction would occur any time the p-level was increased higher than p-1. the position was optimal, volume was adequate, and the right ventricle function was great. the impella cp was explanted and perclose suture x1 tightened and manual pressure held for 15 minutes with good hemostasis noted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29381 | 1220648-2024-08359 | 18941334 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18941334&pc=OZD | 07/23/2022 | 03/10/2025 |  | 03/20/2024 |  |  | 07/23/2022 | 03/20/2024 |  |  | 04/30/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023214032 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/23/2022 |  |  |  |  | The complainant reported a patient, race and ethnicity unknown, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported while on support, the patient experienced cardiac arrhythmia that required defibrillation to resolve. additionally, the patient experienced hemolysis while on support. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29382 | 1220648-2024-08360 | 18941431 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18941431&pc=OZD | 07/24/2022 | 05/30/2025 |  | 03/20/2024 |  |  | 07/24/2022 | 03/20/2024 |  |  | 01/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023187662 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/24/2022 |  |  |  |  | The user facility reported a 74-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed a hematoma at the impella access site. manual pressure was held at bedside and the patient was given blood products. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29383 | 1220648-2024-08361 | 18941449 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18941449&pc=OZD | 03/05/2024 | 04/05/2024 |  | 03/20/2024 |  |  | 03/05/2024 | 03/20/2024 |  |  |  |  | 03/20/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1756407 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/05/2024 |  |  |  |  | The complainant reported a patient undergoing coronary artery bypass was implanted with an impella 5.5 device for mechanical circulatory support. the automated impella controller (aic) shut down on two separate incidences. it was completely shutdown and rebooted. the impella function was back to normal after both shutdowns but temporary support was lost when shutdown occurred. a white alarm occurred on the screen after auto reboot. the controller was replaced with a back up controller. the patient survived and a new aic is running with the 5.5 pump successfully to date. | [5 empty] | The investigation into the aic - shutdown or restart issue has been completed since the initial report was submitted. the console was returned, visually inspected, and tested. the reported failure mode was not confirmed. no unexpected shutdowns or irregularities in performance were observed. the cause of the aic issue was most likely a defective 4in1 board or pbm board. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
29384 | 1220648-2024-08363 | 18941553 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18941553&pc=OZD | 03/05/2024 | 04/24/2024 |  | 03/20/2024 |  |  | 03/05/2024 | 03/20/2024 |  |  | 06/30/2025 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024372535 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/05/2024 |  |  |  |  | The user facility reported a patient of unknown age, gender, race, and ethnicity noted as high-risk percutaneous coronary intervention was to be implanted with an impella cp device for mechanical circulatory support. the impella cp pump was unable to cross the patient's new valve during attempted delivery. due to concerns about the interaction between the pump and the new valve, the decision was made to abort the implant. an intra-aortic balloon pump (iabp) was subsequently placed for patient support. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue was most likely patient condition related, as the impella was unable to pass through the new valve as per the clinical description provided.
29385 | 1220648-2024-08362 | 18941555 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18941555&pc=OZD | 10/09/2022 | 03/12/2025 |  | 03/20/2024 |  |  | 10/09/2022 | 03/20/2024 |  |  | 02/29/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200305 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/09/2022 |  |  |  |  | The user facility reported a 51-year-old female was implanted with an impella cp device for mechanical circulatory support. during support, the flows dropped and there was a suction alarm. the physician decided to remove the impella cp device, and the clot was seen on the impella cp device, a second impella cp device was open and inserted in normal fashion with normal flows. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29386 | 1220648-2024-08364 | 18941639 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18941639&pc=OZD | 03/05/2024 | 04/24/2024 |  | 03/20/2024 |  |  | 03/05/2024 | 03/20/2024 |  |  | 11/30/2025 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024434317 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/05/2024 |  |  |  |  | The user facility reported a 74-year-old female patient post cardiotomy cardiogenic shock/low cardiac output syndrome was to be implanted with an impella 5.5 for mechanical circulatory support. the impella 5.5 placement was aborted due to patient¿s severe vessel calcification. an impella cp was placed instead. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, the impella was unable to pass through due to severe calcification of the vessel as per the clinical description provided.
29387 | 1220648-2024-08365 | 18941714 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18941714&pc=OZD | 11/09/2022 | 03/12/2025 |  | 03/20/2024 |  |  | 11/09/2022 | 03/20/2024 |  |  | 06/30/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023236756 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/09/2022 |  |  |  |  | The user facility reported a 66-year-old male was being implanted with an impella cp device for mechanical circulatory support. attempts to place large bore access was not successful due to heavy calcium burden bilaterally and previous endografts. attempts were abandoned and the patient was rushed for emergent coronary artery bypass graft. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29388 | 1220648-2024-08366 | 18941882 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18941882&pc=OZD | 03/07/2024 | 04/24/2024 |  | 03/20/2024 |  |  | 03/07/2024 | 03/20/2024 |  |  | 11/30/2025 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024431422 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/07/2024 |  |  |  |  | The user facility reported a 58-year-old with cardiomyopathy was to be implanted with an impella 5.5 for mechanical circulatory support. it was reported that the pump was unable to advance past the anastomosis due to the patient's anatomy. as a result of the resistance, the decision was made to abort the impella 5.5 implant. an impella cp pump subsequently implanted for patient support. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue was most likely patient condition related, the impella was unable to pass through as the vessel was too small as per the clinical description provided.
29389 | 1220648-2024-08367 | 18941978 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18941978&pc=OZD | 10/09/2022 | 03/20/2024 |  | 03/20/2024 |  |  | 10/09/2022 | 03/20/2024 |  |  | 05/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231211 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/09/2022 |  |  |  |  | The user facility reported a 87-year-old male was implanted with an impella cp device for mechanical circulatory support. the patient developed bleeding at the impella insertion site. the dressing was removed and a hematoma was also noted. pressure was held and the right side of the sheath was re-sutured. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29390 | 1220648-2024-08368 | 18942016 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18942016&pc=OZD | 07/26/2022 | 05/21/2025 |  | 03/20/2024 |  |  | 07/26/2022 | 03/20/2024 |  |  | 04/30/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023214985 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/26/2022 |  |  |  |  | The user facility reported a 47-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the physician attempted to engage the coronaries but was unsuccessful. aortography revealed aortic dissection (type a). in the ccl the patient coded multiple times, and they were unable to intubate patient and emergent cric performed. transthoracic echocardiogram was also performed to evaluate the extent of dissection. the decision was made to explant the impella cp due to aortic dissection and severe aortic insufficiency. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29391 | 1220648-2024-08369 | 18942041 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18942041&pc=OZD | 03/08/2024 | 04/29/2024 |  | 03/20/2024 |  |  | 03/08/2024 | 03/20/2024 |  |  | 11/30/2025 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024423198 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Malfunction | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/08/2024 |  |  |  |  | The user facility reported a 77-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the complainant reported that following a patient sheet change and turn in bed, the patient became very unstable requiring increases in pressers. the patient¿s hemoglobin was found to have dropped to 3.8 and required a transfusion. the patient was rushed to the catheterization lab and extravasation was found in the common femoral on the impella cp side. the patient was extremely impella dependent so the decision was made to place a new impella on the right side and explant the impella on the left side using a manta device. the physician was able to use a single access up and over technique for a balloon tamponade. the manta device failed closure. the team was able to place a covered stent. during the procedure, the patient arrested requiring cardiopulmonary resuscitation as well as a total of ten units of red blood cells. | [5 empty] | The investigation into the vascular damage - peripheral vessel issue has been completed since the original report was filed. the pump was returned, tested, and evaluated. no physical abnormalities were found. the re-position sheath was inserted in adjunction to the peel away sheath which is against the impella ifu. the peel away was left in during patient transfer which is against the impella ifu. per the clinical information provided, the root cause of the vascular damage - peripheral vessel issue was most likely use related to improper technique. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
29392 | 1220648-2024-08370 | 18942179 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18942179&pc=OZD | 10/09/2022 | 03/20/2024 |  | 03/20/2024 |  |  | 10/09/2022 | 03/20/2024 |  |  | 07/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023253283 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/09/2022 |  |  |  |  | The user facility reported a 35-year-old female was implanted with an impella cp device for mechanical circulatory support after presenting with pulmonary edema and becoming hypotensive. the patient had blood in the urine (hematuria). due to the hemolysis, the heart team decided to wean and remove the impella cp. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
29393 | 1220648-2024-08371 | 18942250 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18942250&pc=OZD | 03/09/2024 | 04/12/2024 |  | 03/20/2024 |  |  | 03/09/2024 | 03/20/2024 |  |  |  |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025443301 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/09/2024 |  |  |  |  | The user facility reported a 81-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient developed bleeding overnight after a big turn. sutures were replaced and three units of red blood cells were given. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical details provided was sufficient to determine the root cause. the root cause of the access site bleeding was most likely use issue since the sutures at site were not intact on impella site.
29394 | 1220648-2024-08372 | 18942310 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18942310&pc=OZD | 10/10/2022 | 03/12/2025 |  | 03/20/2024 |  |  | 10/10/2022 | 03/20/2024 |  |  | 06/30/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023249103 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/10/2022 |  |  |  |  | The user facility reported a 62-year-old male was implanted with an impella cp device for mechanical circulatory support. the patient had some bleeding and a hematoma at the insertion site. forward tension sutures were applied, angle matching was established, and then manual compression was applied with resolution of the bleeding and hematoma. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29395 | 1220648-2024-08373 | 18942364 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18942364&pc=OZD | 07/26/2022 | 03/12/2025 |  | 03/20/2024 |  |  | 07/26/2022 | 03/20/2024 |  |  | 03/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023212788 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/26/2022 |  |  |  |  | The user facility reported a 45-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient had an elevated temperature. there was no source of infection noted. a new impella cp was inserted due to fevers. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿infusion filter the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29396 | 1220648-2024-08374 | 18942384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18942384&pc=OZD | 03/07/2024 | 04/19/2024 |  | 03/20/2024 |  |  | 03/09/2024 | 03/20/2024 |  |  |  |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025443036 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/09/2024 |  |  |  |  | The user facility reported a 65-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that a hematoma at the impella access site was evacuated. the patient has not required blood products since the hematoma was evacuation. the patient is stable. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
29397 | 1220648-2024-08375 | 18942387 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18942387&pc=OZD | 10/10/2022 | 03/20/2024 |  | 03/20/2024 |  |  | 10/10/2022 | 03/20/2024 |  |  | 07/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023252247 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/10/2022 |  |  |  |  | The user facility reported a 62-year-old male was implanted with an impella device for mechanical circulatory support. during support the patient received blood products from loss of blood during procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records.
29398 | 1220648-2024-08376 | 18942443 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18942443&pc=OZD | 09/01/2022 | 03/20/2024 |  | 03/20/2024 |  |  | 09/01/2022 | 03/20/2024 |  |  | 09/30/2023 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022144150 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/01/2022 |  |  |  |  | The user facility reported a 65-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient experienced significant bleeding from the insertion site. the activated clotting time did not fall within range since adding heparin to the purge. the physician noted a decrease in hemoglobin and decided to remove the impella. hemostasis could not be achieved after manual pressure and the patient decompensated and required cpr and blood products. return of spontaneous circulation was obtained and femstop was applied to groin. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
29399 | 1220648-2024-08377 | 18942541 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18942541&pc=OZD | 03/10/2024 | 04/22/2024 |  | 03/20/2024 |  |  | 03/10/2024 | 03/20/2024 |  |  | 10/31/2025 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024414979 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/10/2024 |  |  |  |  | The user facility reported a 79-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that during support, there was an observation made of saturated flows after three days. the pump was explanted but not replaced as the patient admitted cardiogenic shock had recovered hemodynamically. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported saturated flow has been completed. no product was returned for a visual inspection. data log analysis confirmed a gradual decrease in purge flow and gradual increase in pump flow followed by purge pressure spikes. according to the clinical, on the third day of impella 5.5 support saturated flow was recorded. no attempts were made to resolve the issue, and the decision was made to explant the pump to avoid a potential pump stop. upon explant biomaterial was observed. the most likely cause of the saturated flow was biomaterial. the instructions for use states ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿
29400 | 1220648-2024-08378 | 18942568 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18942568&pc=OZD | 09/01/2022 | 03/20/2024 |  | 03/20/2024 |  |  | 09/01/2022 | 03/20/2024 |  |  | 03/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023207350 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/01/2022 |  |  |  |  | The user facility reported a 72-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient had previously placed stents in left anterior descending artery were occluded. the patient was administered dobutamine which lead to ventricular tachycardia. the patient was withdrawn from care. | [5 empty] | The investigation for the reported issue has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
29401 | 1220648-2024-08379 | 18942593 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18942593&pc=OZD | 07/26/2022 | 03/12/2025 |  | 03/20/2024 |  |  | 07/26/2022 | 03/20/2024 |  |  | 04/30/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023220168 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/26/2022 |  |  |  |  | The user facility reported a 57-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support due to stemi. while on support, the patient¿s lower leg was cool to touch, and pulses were not heard on doppler. the doctor decided to explant impella due to no flow to lower leg and patient constant movement. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29402 | 1220648-2024-08380 | 18942602 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18942602&pc=OZD | 09/01/2022 | 03/20/2024 |  | 03/20/2024 |  |  | 09/01/2022 | 03/20/2024 |  |  | 04/30/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023217311 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/01/2022 |  |  |  |  | The user facility reported a 76-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. after a successful implant procedure, heparin was stopped due to oozing at the impella cp insertion site and elevated partial thromboplastin time. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
29403 | 1220648-2024-08381 | 18942652 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18942652&pc=OZD | 10/10/2022 | 03/20/2024 |  | 03/20/2024 |  |  | 10/10/2022 | 03/20/2024 |  |  | 06/30/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023247297 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/10/2022 |  |  |  |  | The user facility reported a 57-year-old male was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced some site oozing. the patient received two units of packed red blood cells (prbcs). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29404 | 1220648-2024-08382 | 18942705 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18942705&pc=OZD | 12/07/2022 | 03/20/2024 |  | 03/20/2024 |  |  | 12/07/2022 | 03/20/2024 |  |  | 08/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023262946 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/07/2022 |  |  |  |  | The user facility reported a 80-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient had poor distal perfusion and an external femorofemoral bypass was performed. patient care had persistent gastrointestinal(gi) bleed and the patient was administered14 units of packed red blood cells units (prbc). flow was restored, and the patient regained pulsatility. the patient was successfully weaned off the impella cp. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
29405 | 1220648-2024-08383 | 18942731 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18942731&pc=OZD | 09/01/2022 | 09/30/2025 |  | 03/20/2024 |  |  | 09/01/2022 | 03/20/2024 |  |  | 05/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231100 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/01/2022 |  |  |  |  | The user facility reported a 64-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient experienced hemolysis. four units were administered. additionally, the patient experienced arrythmia and the impella cp was repositioned. | [5 empty] | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined. d6a and d6b revised from the initial manufacturer device report 1220648-2024-08383 in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2024-08383 in accordance with updated procedures. g1. reporting contact email revised on manufacturer device report 1220648-2024-08383 in accordance with updated procedures. g1. reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2024-08383. h6. component code revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
29406 | 1220648-2024-08384 | 18942759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18942759&pc=OZD | 09/02/2022 | 03/12/2025 |  | 03/20/2024 |  |  | 09/02/2022 | 03/20/2024 |  |  | 06/30/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235168 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/02/2022 |  |  |  |  | The user facility reported a 79-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed significant bleeding from the impella cp insertion site. femostop, angle matching and quick clot were all used. the patient was also provided one unit of blood. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29407 | 1220648-2024-08385 | 18942781 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18942781&pc=OZD | 07/28/2022 | 05/21/2025 |  | 03/20/2024 |  |  | 07/28/2022 | 03/20/2024 |  |  | 03/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023225898 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/28/2022 |  |  |  |  | The user facility reported a 79-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after the removal of the impella sheath the patient became hypotensive. contralateral access gained via the left femoral artery. femoral perforation visualized. two perclose sutures were placed followed by short cp sheath. the pump was placed, and flows optimized. left anterior descending (lad) wired and several passes with orbital atherectomy. long lad stent placed, the pressures dropped several times with inflations but returned to baseline. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29408 | 1220648-2024-08386 | 18942797 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18942797&pc=OZD | 09/02/2022 | 03/12/2025 |  | 03/20/2024 |  |  | 09/02/2022 | 03/20/2024 |  |  | 01/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023189791 |  | 0048-0073 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/02/2022 |  |  |  |  | The user facility reported a(b)(6)-year-old male noted as high risk percutaneous coronary intervention was implanted with impella cp for mechanical circulatory support. after the impella was successfully implanted, there was difficultly with achieving optimal flows. fluid bolus was administered twice. flows improved with slight repositioning of impella. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29409 | 1220648-2024-08387 | 18942799 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18942799&pc=OZD | 03/10/2024 | 03/26/2024 |  | 03/20/2024 |  |  | 03/10/2024 | 03/20/2024 |  |  |  |  | 03/20/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1426098 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/10/2024 |  |  |  |  | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that the automated impella controller (aic) had a self-diagnosis failed screen. the aic was replaced successfully. | [5 empty] | The investigation into the aic - system self-check error has been completed since the original report was filed. the console was returned and tested after arrival and the failure mode was reproduced; when the console was turned on, it immediately output a system self-check fail. the cause of the aic issue was contamination of a communication line. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed.
29410 | 1220648-2024-08388 | 18943038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18943038&pc=OZD | 12/07/2022 | 03/12/2025 |  | 03/20/2024 |  |  | 12/07/2022 | 03/20/2024 |  |  | 08/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023261462 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/07/2022 |  |  |  |  | The user facility reported a 69-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. during insertion, the patient went into complete heart block but subsequently regained heart rhythm spontaneously. the impella cp started with 3.5 l/ min of flow and a transvenous pacer was placed for protection. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29411 | 1220648-2024-08389 | 18943061 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18943061&pc=OZD | 03/10/2024 | 05/01/2024 |  | 03/20/2024 |  |  | 03/10/2024 | 03/20/2024 |  |  | 11/30/2025 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024430964 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 25 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/10/2024 |  |  |  |  | The user facility reported a 25-year-old male with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that purge pressure high alarms occurred intermittently with purge system blocked alarms. the cassette was changed recently and the purge bag change to heparin purge temporarily resolved the alarm, but the purge flow did not go higher than 2.8 ml/hr. alarms returned. the decision was made by the team to switch pump to a different console. after the switch, the pump was restarted, and purge flows and pressure resolved. it was further reported that stopping and restarting the pump is what ended up clearing the issue that was affecting the purge system. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the high purge pressure has been completed since the original report was submitted. the device was not returned for investigation. data log analysis confirmed a gradual increase in purge pressure and gradual decrease in purge flow. heparin was not used during this time. while attempting to resolve these issues the purge cassette change took 1:56 and two hours later the new bag spike took 1:36. the most likely cause of the high purge pressure was biomaterial. impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.6 code 3211 was reported in the investigation findings on the previous reported report. this code was incorrect. h.6 has been revised to show code 3221 which should of been on the previous reported report.
29412 | 1220648-2024-08390 | 18943095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18943095&pc=OZD | 12/07/2022 | 03/12/2025 |  | 03/20/2024 |  |  | 12/07/2022 | 03/20/2024 |  |  | 08/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023269970 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Dislodged or Dislocated | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/07/2022 |  |  |  |  | The user facility reported a 75-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. the patient woke up at the end of the procedure and dislodged impella with their leg. impella cp was subsequently removed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29413 | 1220648-2024-08391 | 18943192 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18943192&pc=OZD | 12/07/2022 | 03/12/2025 |  | 03/20/2024 |  |  | 12/07/2022 | 03/20/2024 |  |  | 07/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023254802 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/07/2022 |  |  |  |  | The user facility reported a 62-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that there was access site bleeding during explant as both perclose closures and angio-seal closures had failed. as a result of pump explant complications, surgical closure was subsequently required post escalation to impella 5.5. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29414 | 1220648-2024-08392 | 18943230 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18943230&pc=PYX | 03/10/2024 | 04/05/2024 |  | 03/20/2024 |  |  | 03/10/2024 | 03/20/2024 |  |  | 10/31/2025 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024418614 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/10/2024 |  |  |  |  | The user facility reported a 52-year-old male undergoing cardiac surgery was implanted with an impella rp flex device for mechanical circulatory support. the complainant had placed an impella rp flex for support of an acute myocardial infarction patient. the patient had extracorporeal membrane oxygenation decannulated and the patient had impella support running. the pump was explanted after just under 51 hours due to concerns of clot ingestion with low flows. the issue was seen after swan manipulation. | [5 empty] | The investigation into the low pump flow has been completed since the original report was filed. the pump was returned, tested, and evaluated. upon investigation of the pump, a large clot was observed around the impeller. the root cause of the low pump flow was determined to be biomaterial. | The impella device was received from the customer and therefore.an evaluation of the device is underway. upon investigation closure, a supplemental mdr will be filed.
29415 | 1220648-2024-08393 | 18943277 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18943277&pc=OZD | 10/10/2022 | 09/11/2024 |  | 03/20/2024 |  |  | 10/10/2022 | 03/20/2024 |  |  | 04/30/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023221054 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Vascular Dissection | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/10/2022 |  |  |  |  | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 61-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support the patient experienced a large hematoma at the impella site. the patient was transfused with eight units of packed red blood cells (prbc). the medical team decided to remove impella and consult vascular surgery for right femoral arteriotomy repair, prior to removal the patient declined and began decompensating. the decision was made to withdraw care and the patient expired. there are no allegations of impella cp device causing or contributing to death. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08393 was provided in sections b2, b5, h1, and h6. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ this report is being filed as part of a retrospective review of historical records.
29416 | 1220648-2024-08394 | 18943303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18943303&pc=OZD | 12/07/2022 | 03/20/2024 |  | 03/20/2024 |  |  | 12/07/2022 | 03/20/2024 |  |  | 03/08/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023270471 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/07/2022 |  |  |  |  | The user facility reported a 62-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported the patient had pulsatile arterial blood around the repositioning sheath. the impella cp was removed from the right central femoral artery and the right central femoral artery was successfully closed via perclose x2. a new impella cp was then successfully placed via the left central femoral artery. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29417 | 1220648-2024-08395 | 18943319 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18943319&pc=OZD | 07/29/2022 | 05/19/2025 |  | 03/20/2024 |  |  | 07/29/2022 | 03/20/2024 |  |  | 04/30/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023220171 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/29/2022 |  |  |  |  | The user facility reported a 70-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed retroperitoneal bleeding. the patient was given blood products and exploratory lap was going to be performed in the operating room. right heart failure was discussed as being a sign/symptom of hemolysis. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29418 | 1220648-2024-08396 | 18943354 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18943354&pc=OZD | 12/08/2022 | 03/20/2024 |  | 03/20/2024 |  |  | 12/08/2022 | 03/20/2024 |  |  | 08/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023262874 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/08/2022 |  |  |  |  | The user facility reported a 67-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported the patient developed a hematoma at the access site. systemic heparin was withdrawn and the patient received 3 units of packed red blood cells (prbc). | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29419 | 1220648-2024-08397 | 18943522 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18943522&pc=OZD | 03/10/2024 | 06/05/2024 |  | 03/20/2024 |  |  | 03/11/2024 | 03/20/2024 |  |  | 11/30/2025 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024426961 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Activation Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/11/2024 |  |  |  |  | Upon abiomed's medical safety officer review, there is not enough information to associate the use of the device with the patient's outcome. | The user facility reported a 73-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported the patient was on impella cp support and the pump connection was defective and an error message "please replace" appeared during priming. the impella was replaced successfully. |  |  |  |  | The investigation into the priming issue has been completed since the original report was filed. the pump was returned for evaluation. upon visual inspection, no abnormalities were found. electrical testing was performed and no open lines were found and electrical lines were all within normal range. the pump was plugged into a lab automated impella controller (aic) to attempt to reproduce "impella defective" alarm. no alarms were triggered when plugged into console and the aic prompted to begin case start. the pump was then connected to multiple aic's to attempt to reproduce the alarm. six of the six aic's did not alarm when the pump was plugged in and aic's prompted case start. the data logs from field were not returned. the clinical details noted that the impella defective alarm occurred when the priming pump and new pump were able to run successfully on the same aic. the root cause of the impella defective alarms could not be determined due to field logs not being returned and limited clinical details. d.6a and d.6b revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29420 | 1220648-2024-08398 | 18943545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18943545&pc=OZD | 12/08/2022 | 03/20/2024 |  | 03/20/2024 |  |  | 12/08/2022 | 03/20/2024 |  |  | 04/30/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023216731 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/08/2022 |  |  |  |  | The user facility reported a 42-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported that while on the impella cp support, the patient experienced multiple episodes of ventricular tachycardia that required defibrillation. the impella cp pump was explanted and replaced with an impella 5.5. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29421 | 1220648-2024-08399 | 18943564 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18943564&pc=OZD | 10/05/2022 | 05/21/2025 |  | 03/20/2024 |  |  | 10/05/2022 | 03/20/2024 |  |  | 04/30/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023224767 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  |  |  | Injury | Yes | Yes | Improper Flow or Infusion | [6 empty] | Ischemia | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/05/2022 |  |  |  |  | The user facility reported a 60-year-old female was implanted with an impella cp device for mechanical circulatory support. it was reported while being placed on impella cp support, the pump flows were unable to reach above 1.3 upon insertion. bedside echocardiography was performed and notable pericardial effusion was visualized. pericardiocentesis was performed and after 500ml of bloody drainage was removed, both ecmo and impella flows improved. additionally, the patient was experiencing a lack of distal perfusion to right leg and the vascular team was contacted for a reperfusion catheter. the patient was transfused with platelets prior to the removal of the impella device due to thrombocytopenia. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29422 | 1220648-2024-08400 | 18943612 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18943612&pc=OZD | 07/29/2022 | 05/14/2025 |  | 03/20/2024 |  |  | 07/29/2022 | 03/20/2024 |  |  | 03/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023211190 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Female |  |  |  | Death | Yes | Yes | Difficult to Insert | Patient-Device Incompatibility | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Low Blood Pressure/ Hypotension | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 07/29/2022 |  |  |  |  | The user facility reported a female of unknown age was being implanted with an impella cp device for mechanical circulatory support. there was difficulty encountered when the pump was inserted as it was unable to advance past the proximal portion of the patient's aneurysm. the pump was subsequently removed, and an intra-aortic balloon pump was attempted instead with no success. | Patient become hypotensive and continued to decompensate and expired in cath lab. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08400 was provided. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29423 | 1220648-2024-08401 | 18943691 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18943691&pc=OZD | 09/28/2022 | 05/20/2025 |  | 03/20/2024 |  |  | 09/28/2022 | 03/20/2024 |  |  | 05/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023229222 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/28/2022 |  |  |  |  | The user facility reported a 68-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. after removing the peel away sheath, the patient decompensated further requiring central extracorporeal membrane oxygenation. the impella position was lost when manipulating the heart and was not able to recross the aortic valve. the left ventricle remained decompressed so the physician made the decision to explant the device. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29424 | 1220648-2024-08402 | 18943732 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18943732&pc=OZD | 08/17/2023 | 03/20/2024 |  | 03/20/2024 |  |  | 08/17/2023 | 03/20/2024 |  |  | 04/30/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023209937 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 08/17/2023 |  |  |  |  | The user facility reported an 86-year-old patient needing mechanical circulatory support. it was reported that the patient experienced cold sensation of the lower limbs during impella cp support. there was difficulty inserting the pump through the introducer due to the patient anatomy. the original introducer was swapped for a 14 fr stiff dilator and the pump was successfully implanted. however, following a successful percutaneous coronary intervention (pci), it was noted that the patient's lower limbs were cold to the touch. as a result, the decision was made to remove the pump. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
29425 | 1220648-2024-08403 | 18944236 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18944236&pc=OZD | 07/25/2022 | 03/12/2025 |  | 03/20/2024 |  |  | 07/25/2022 | 03/20/2024 |  |  | 02/29/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195279 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/25/2022 |  |  |  |  | The user facility reported a 64-year-old male on impella cp for mechanical circulatory support. it was reported that there was access site bleeding noted by the hematoma and some oozing at the insertion site. the pump was explanted as a result. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29426 | 1220648-2024-08404 | 18944253 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18944253&pc=OZD | 07/30/2022 | 03/20/2024 |  | 03/20/2024 |  |  | 07/30/2022 | 03/20/2024 |  |  | 03/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023207356 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/30/2022 |  |  |  |  | The user facility reported a 63-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed a stable hematoma at the impella site. a unit of blood was given overnight due to a slight drop in hemoglobin and hematocrit. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29427 | 1220648-2024-08405 | 18944267 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18944267&pc=OZD | 03/11/2024 | 04/30/2024 |  | 03/20/2024 |  |  | 03/11/2024 | 03/20/2024 |  |  |  |  | 03/20/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1122977 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Component Incompatible | Component or Accessory Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/11/2024 |  |  |  |  | It was reported that the automated impella controller (aic) was not accepting cassettes. the cassettes were replaced twice to no avail. the aic was then replaced. | [5 empty] | The investigation for the reported console (aic) purge issue has been completed. the aic was returned for evaluation. this console was tested and was unable to reproduce the event (b)(6) by running the test pump and the cassette for more than 12 hours. purge pressure sensor accuracy was good. a thin dextrose layer was found in the stepper motor gasket. the data logs were reviewed and confirmed a purge flow decreased ¿ alarm, event(b)(6). after 1 hour, the console received the purge pressure high ¿ alarm, event (b)(6), and purge system blocked ¿ alarm, event (b)(6), which never resolved until the purge cassette was removed. the root cause of the high purge pressure and the blocked purge system was not determined due to the inability to reproduce. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed.
29428 | 1220648-2024-08406 | 18944323 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18944323&pc=OZD | 07/25/2022 | 03/20/2024 |  | 03/20/2024 |  |  | 07/25/2022 | 03/20/2024 |  |  | 02/29/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023199921 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/25/2022 |  |  |  |  | The user facility reported a 66-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported while on support, the patient experienced bleeding and hemolysis. the impella cp device was explanted. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29429 | 1220648-2024-08407 | 18944475 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18944475&pc=OZD | 08/03/2022 | 04/08/2025 |  | 03/20/2024 |  |  | 08/03/2022 | 03/20/2024 |  |  | 08/31/2023 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2022130345 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/03/2022 |  |  |  |  | The user facility reported a 66-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced a cerebral vascular accident (cva). the medical team and family made the decision to withdraw care and the patient expired. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08407 was provided. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29430 | 1220648-2024-08408 | 18944511 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18944511&pc=OZD | 03/11/2024 | 03/20/2024 |  | 03/20/2024 |  |  | 03/11/2024 | 03/20/2024 |  |  |  |  | 03/20/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1079313 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 03/11/2024 |  |  |  |  | The user facility reported a 54-year-old male with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that automated impella controller (aic) was exchanged for a backup due to a yellow controller alarm. no patient harm was reported and the impella 5.5 support continued. | [5 empty] | The investigation has been completed. console logs from the reported day of event were consistent with the complaint and confirmed the reported failure mode given there was a controller error alarms triggered during the clinical procedure. real time logs confirmed that all signals received from optical bench were for approximately four minutes interpreted as -16384 values. the console was investigated. the failure was unable to be reproduced throughout testing and was characterized by a temporary loss of optical data and triggering of a controller error alarm, that has a low probability of reoccurrence. if needed, patient hemodynamics and impella position can be monitored with imaging. the cause of the issue was not determined since the issue could not be reproduced.
29431 | 1220648-2024-08409 | 18944647 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18944647&pc=OZD | 08/02/2022 | 03/20/2024 |  | 03/20/2024 |  |  | 08/02/2022 | 03/20/2024 |  |  | 10/31/2023 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023158852 |  | 0048-0032 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/02/2022 |  |  |  |  | The user facility reported a 80-year-old male in acute myocardial infarction/cardiogenic shock, and was implanted with an impella cp device for mechanical circulatory support, the patient developed hemolysis during impella cp support. the pump was repositioned the following day, but due to the overall condition of the patient the decision was made to withdraw care. the patient was weaned from the impella cp and the device was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
29432 | 1220648-2024-08410 | 18944678 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18944678&pc=OZD | 03/10/2024 | 04/04/2024 |  | 03/20/2024 |  |  | 03/11/2024 | 03/20/2024 |  |  |  |  | 03/20/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1133145 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/11/2024 |  |  |  |  | The user facility reported a 34-year-old male of unknown race and ethnicity in st- segment elevation myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the purge disc not detected alarm occurred when changing the cassette. the automated impella controller (aic) was exchanged, and the alarm resolved. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the initial report was submitted. the console was returned, visually inspected, and tested. the issue with detecting the purge disc was reproduced, and the purge flag was confirmed to be broken. the root cause of this issue was a broken/missing purge flag. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
29433 | 1220648-2024-08411 | 18944825 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18944825&pc=OZD | 08/02/2022 | 03/12/2025 |  | 03/20/2024 |  |  | 08/02/2022 | 03/20/2024 |  |  | 12/31/2023 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023174729 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/02/2022 |  |  |  |  | The user facility reported a 64-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced bleeding at the access site near the groin area. also, constant suction was observed; therefore, the impella was repositioned. however, after experiencing major bleeding and arrhythmia, the impella was explanted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29434 | 1220648-2024-08412 | 18945407 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18945407&pc=OZD | 08/02/2022 | 03/20/2024 |  | 03/20/2024 |  |  |  | 03/20/2024 |  |  | 03/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204558 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/02/2022 |  |  |  |  | The user facility reported a 81-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient had a hematoma at the site. the patient was unable to maintain adequate pressure despite best efforts, ultimately suffering multisystem organ failure and expiring on support. although the patient died while on impella support, no allegations were made towards the impella for cause of death. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29435 | 1220648-2024-08413 | 18946203 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18946203&pc=OZD | 10/04/2022 | 05/23/2025 |  | 03/20/2024 |  |  |  | 03/20/2024 |  |  | 06/30/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235251 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | Thrombosis/Thrombus | Heart Failure/Congestive Heart Failure | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/04/2022 |  |  |  |  | The user facility reported a 32-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. as thestaff was preparing the patient for intensive care unit, the patient rapidly decompensated, impella flows decreased quickly. the second repeat echocardiogram showed pericardial effusion. the staff completed pericardiocentesis but were unable to stabilize the patient. they were unable to identify the perforation or rupture location. cardio-pulmonary resuscitation (cpr) plus additional pericardiocentesis efforts were unable to stabilize the patient. the patient passed on the table. no allegations were made towards the impella for cause of death. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08413 was provided. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29436 | 1220648-2024-08414 | 18946290 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18946290&pc=OZD | 10/15/2022 | 03/12/2025 |  | 03/20/2024 |  |  |  | 03/20/2024 |  |  | 01/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023187764 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/15/2022 |  |  |  |  | The user facility reported a 52-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient had bleeding at the groin site and multiple suction events. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29437 | 1220648-2024-08415 | 18946492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18946492&pc=OZD | 09/28/2022 | 03/20/2024 |  | 03/20/2024 |  |  |  | 03/20/2024 |  |  | 12/31/2023 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023170766 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/28/2022 |  |  |  |  | The user facility reported a 49-year-old male was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding at the impella site and developed a hematoma. the medical sfaff held pressure and performed angle matching, but the issue did not resolve. femostop was utilized and this helped to improve the bleeding to a light-to-moderate continuous oozing. the physician elected to stop integrillin. 4 units of packed cells were administered. the patient also exhibited signs of hemolysis with dark red urine. the impella cp was explanted via cutdown and artery repair per vascular surgery. the hematoma was evacuated, and a wound vac was placed at the right groin. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis. ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records.
29438 | 1220648-2024-08416 | 18946605 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18946605&pc=OZD | 09/29/2022 | 03/12/2025 |  | 03/20/2024 |  |  |  | 03/20/2024 |  |  | 06/30/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023248980 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/29/2022 |  |  |  |  | The user facility reported a 77-year-old female was being implanted with an impella cp device for mechanical circulatory support. it was reported that it was very difficult to position the pump in mid ventricular space, possibly due to septal wall hypokinesia. many attempts were made to reposition the pump. the decision was made to use re-access port to replace impella with new device. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29439 | 1220648-2024-08417 | 18946675 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18946675&pc=OZD | 08/02/2022 | 04/17/2025 |  | 03/20/2024 |  |  | 08/02/2022 | 03/20/2024 |  |  | 03/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2023215974 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Decrease in Pressure | Pressure Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/02/2022 |  |  |  |  | The user facility reported a 87-year-old female noted as high risk percutaneous coronary intervention (hrpci) was implanted with an impella cp device for mechanical circulatory support. prior to inserting the device there was a low purge pressure alarm noted. purge cassette and line checked for leaks or disconnects. the purge was re-primed and connected. purge pressure quickly fell to low 100¿s and low alarm noted. the purge cassette was changed and primed with the same results. after multiple attempts to correct the purge pressure low alarm a new cp device was utilized. the impella was advanced and support initiated. single access technique used and hrpci performed to left main with drug eluding stent x 1. the impella support was weened down and device removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29440 | 1220648-2024-08418 | 18946694 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18946694&pc=OZD | 08/02/2022 | 03/12/2025 |  | 03/20/2024 |  |  |  | 03/20/2024 |  |  | 03/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023212716 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/02/2022 |  |  |  |  | The user facility reported a 67-year-old male in acute non-stelevated myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was brought to the cath lab for trans-esophogeal echo and impella explant. the patient was positive for bacteremia. the physician was worried that the impella site could contribute to further infection. the impella was later explanted and the patient was monitored until cleared for bypass surgery. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29441 | 1220648-2024-08419 | 18946786 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18946786&pc=OZD | 08/03/2022 | 03/12/2025 |  | 03/20/2024 |  |  |  | 03/20/2024 |  |  | 02/29/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023196161 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/03/2022 |  |  |  |  | The user facility reported a 82-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, ischemia was detected in the impella leg during patient support. the decision was made to use swan and the healthcare professional was able to wean the cp pump flow to 1 l/min. the impella cp pump was subsequently explanted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29442 | 1220648-2024-08420 | 18946833 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18946833&pc=OZD | 08/03/2022 | 03/13/2025 |  | 03/20/2024 |  |  |  | 03/20/2024 |  |  | 04/30/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023215029 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/03/2022 |  |  |  |  | The user facility reported a 61-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support, due to cardiac arrest. while on support the patient had significant oozing. the patient was given 1 unit of packed red blood cells and systemic heparin was withheld. impella was explanted without complication. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29443 | 1220648-2024-08421 | 18946840 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18946840&pc=OZD | 09/30/2022 | 03/13/2025 |  | 03/20/2024 |  |  |  | 03/20/2024 |  |  | 08/31/2023 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED |  |  | US |  | 2022146955 |  | 005062-RO |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/30/2022 |  |  |  |  | The user facility reported a 72-year-old female in post cardiotomy cardiogenic shock /low cardiac output syndrome. the impella 5.0 was inserted into the vessel; however, the device was unable to traverse the vessel, due to tortuous anatomy. therefore, the impella 5.0 was removed and replaced with the impella cp. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29444 | 1220648-2024-08422 | 18946853 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18946853&pc=OZD | 07/10/2022 | 03/20/2024 |  | 03/20/2024 |  |  |  | 03/20/2024 |  |  | 12/31/2023 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023178836 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/10/2022 |  |  |  |  | The user facility reported a 58-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient developed access site bleeding prior to transfer which had been resolved with manual pressure and suture placement. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
29445 | 1220648-2024-08423 | 18946858 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18946858&pc=OZD | 07/11/2022 | 03/13/2025 |  | 03/20/2024 |  |  |  | 03/20/2024 |  |  | 03/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204360 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/11/2022 |  |  |  |  | The user facility reported a 42-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported while on support, the patient experienced distal limb ischemia. the impella cp device was removed to resolve the distal limb ischemia. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia" this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29446 | 1220648-2024-08424 | 18946866 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18946866&pc=OZD | 08/04/2022 | 03/20/2024 |  | 03/20/2024 |  |  |  | 03/20/2024 |  |  | 04/30/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023215045 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 08/04/2022 |  |  |  |  | The user facility reported a 73-year-old patient needing mechanical circulatory support. it was reported that the patient was placed on impella cp support after the patient went into cardiac arrest with immediate resuscitation. while inserting the impella the patient experienced ventricular tachycardia (vt). the patient was defibrillated with success and the proper position was verified. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿correct impella catheter position ¿catheter angled toward the left ventricular apex away from the heart wall and not curled up or blocking the mitral valve¿ this report is being filed as part of a retrospective review of historical records.
29447 | 1220648-2024-08425 | 18946893 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18946893&pc=OZD | 07/12/2022 | 10/24/2025 |  | 03/20/2024 |  |  |  | 03/20/2024 |  |  | 04/30/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023216704 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Difficult to Insert | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/12/2022 |  |  |  |  | The user facility reported a 79-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that the repositioning sheath could not be advanced into the artery; therefore, a .035 wire was inserted through the repositioning port and access was regained to the artery. the decision was then made to remove the impella and insert another 14fr sheath and a second impella cp. | A patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device, experienced complications and subsequently expired the same day of implant. the patient presented with shortness of breath and underwent a left heart catheterization that revealed calcified stenosis of the left anterior descending artery (lad). surgery was consulted and the decision was made for percutaneous coronary intervention (pci). during the pci, the patient formed thrombus in the circumflex artery and left anterior descending artery post shockwave. the decision made at this point to implant an impella cp device as bailout through the right femoral artery in order proceed with aspiration of the clot. proper position was verified by fluoroscopy and the impella cp was flowing 2.7l/min at p6. after impella insertion, the physician continued with the procedure, aspirating the clot and placing drug-eluting stent to the distal lad noted with "great angiographic results." postprocedure it was decided to remove the 14f peel-away sheath and advance the repositioning sheath. due to the patient's body habitus and a possible hematoma, the repositioning sheath could not be advanced into the artery. therefore, a .035 wire was inserted through the repositioning port and access was regained to the artery. the decision was then made to remove the initially placed impella cp and insert another 14fr sheath and a new impella cp. the new impella cp was opened and prepped. prior to inserting new impella cp, an angiogram of the right femoral artery was taken, and it was shown that the 14fr sheath was occlusive. it was recommended to do an external fem-fem bypass; however, the physician decided against it. a dry closure technique was used to downsize the 14fr to a 9fr sheath using one proglide. the patient became bradycardic and lost pulsatility. a heroic effort was made to stabilize the patient but to no avail. the patient expired. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined. medical safety review of the event was completed. there is not enough information to exclude the impella cp pump 1 as an associated factor to the patient¿s outcome. regarding the introducer, there was ischemia and the physician decided against a fem-fem bypass. however, this is a serious injury attributed to the introducer." note: the related manufacturing report number for the 14fr introducer sheath will be submitted in a following supplemental report. b2 death and date was inadvertently omitted on the initial manufacturer device report number 1220648-2024-08425. b5 additional information and patient outcome was inadvertently omitted on the initial manufacturer device report number 1220648-2024-08425. b7 was inadvertently omitted on the initial manufacturer device report number 1220648-2024-08425. d6a and d6b revised from the initial manufacturer device report 1220648-2024-08425 in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2024-08425 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-08425 in accordance with updated procedures. g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2024-08425. h1 was erroneously selected on the initial manufacturer device report number 1220648-2024-08425. h6 health effect - impact code 1802 was inadvertently omitted on the initial manufacturer device report number 1220648-2024-08425.
29448 | 1220648-2024-08426 | 18946921 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18946921&pc=OZD | 08/04/2022 | 05/28/2025 |  | 03/20/2024 |  |  |  | 03/20/2024 |  |  | 03/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023205054 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup |  |  | 08/04/2022 |  |  |  |  | The user facility reported a(b)(6)patient needing mechanical circulatory support. it was reported that the patient was placed on impella cp support for hemodynamics support and to stabilize the patient post procedure. the patient developed oozing at the impella cp insertion site and pressure dressing was applied to achieve hemostasis. the patient had reoccurring bouts of ventricular tachycardia (vt) and ventricular fibrillation (vf). patient has been shocked over 15 times with a reduction in ejection fraction to 10%. the physician requested a transfer for a higher level of care and possible upgrade to a 5.5 pump. | Additional information noted the patient expired on support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08426 was provided. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29449 | 1220648-2024-08427 | 18946924 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18946924&pc=OZD | 12/28/2023 | 05/10/2024 |  | 03/20/2024 |  |  |  | 03/20/2024 |  |  | 10/31/2025 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024418707 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Atrial Fibrillation | Sepsis | Mitral Valve Insufficiency/ Regurgitation | Unspecified Gastrointestinal Problem | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | Followup | 02/19/2024 |  |  |  |  | Additional information received via loqi (abiomed¿s long-term outcome and quality indicator impella registry), on (b)(6) 2024 indicates this patient also endured mesenteric ischemia with a "possible" relationship to the impella device. intervention was required via 1u packed red blood cells, cpr due to subsequent hypotension leading to lost pulsatility, and a wound vac placed. the colon was dilated, fluid-filled throughout, laparotomy and right colectomy was performed. additionally, "on (b)(6) 2024, ex lap: removal of quikclot, end ileostomy and fascial closure with staples. patient remained with ileostomy post procedure. (drainage montiored). h/h baseline was unknown. however at admission, 12.2/36.3. nadir levels 6.9/22. (b)(6) 2024 had ileostomy placed." | We received additional information from our medical safety team via loqi (abiomed¿s long-term outcome and quality indicator impella registry), pertaining to the above case. it was noted that supraventricular ventricular tachycardia, ventricular tachycardia, concern for sepsis and severe mitral valve regurgitation with a possible relationship to the impella 5.5 device used on this patient. | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry regarding a 57 year old male patient presenting with post cardiotomy cardiogenic shock for impella support. during support, it was alleged that the patient had endured atrial fibrillation with a "possible" relationship to the impella device. a cardioversion was performed, as treatment. |  |  |  | Additional information received via loqi study. see b5 for additional narrative. upon completion of investigation, a supplemental mdr will be filed. | Additional informations received ; section b5 and h6 with health effect - clinical code was updated accordingly. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for this record has been updated to include evaluation of sepsis, heart valve damage, and mesenteric ischemia and has been completed. after review of the clinical details, the root cause of the infection and the vascular damage could not be determined without the actual device returned for analysis. however, clinical details provided was sufficient to determine the root cause of the heart valve damage. as a result, the root cause of the heart valve damage was determined to be patient condition, because it was found that the patient had pre-existing condition of mvr(mitral valve regurgitation) during tee studies. | The investigation for the reported atrial arrhythmia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the atrial arrhythmia could not be determined. the instructions for use states ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
29450 | 1220648-2024-08428 | 18946926 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18946926&pc=OZD | 07/12/2022 | 10/01/2025 |  | 03/20/2024 |  |  |  | 03/20/2024 |  |  | 02/29/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023196922 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/12/2022 |  |  |  |  | The user facility reported a 63-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that the impella stopped. the care team determined that swapping the new impella would be the best course of action. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. b1 adverse event was inadvertently omitted on the initial manufacturer device report number. b2 other serious injury was inadvertently omitted on the initial manufacturer device report number. d6a and d6b revised from the initial manufacturer device report number in accordance with updated procedures f6/f8-should have been left blank on the initial manufacturer device report number in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number. h1 type of report was erroneously selected on the initial manufacturer device report number. h6 health effect - clinical code 4582 was erroneously entered on the initial manufacturer device report number. h6 health effect - impact code 4614 was inadvertently omitted on the initial manufacturer device report number.
29451 | 1220648-2024-08429 | 18946932 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18946932&pc=OZD | 08/04/2022 | 03/20/2024 |  | 03/20/2024 |  |  |  | 03/20/2024 |  |  | 04/30/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023223971 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/04/2022 |  |  |  |  | The user facility reported a 80-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed bleeding/hematoma at the impella cp insertion site. milrinone was restarted. there was no purge or systemic heparin. the impella cp was explanted, and the patient was escalated to the impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury" this report is being filed as part of a retrospective review of historical records.
29452 | 1220648-2024-08430 | 18946954 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18946954&pc=OZD | 07/13/2022 | 03/13/2025 |  | 03/20/2024 |  |  |  | 03/20/2024 |  |  | 03/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023209890 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Insert | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/13/2022 |  |  |  |  | The user facility reported a 72-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. while attempting to install support, the pigtail was difficult to insert; however, it was successfully crossed aortic valve. 0.018" wire was placed without difficulty. the impella cp advanced successfully through the sheath but would not advance through the iliac artery where multiple attempts were performed. due to patient instability, the impella cp was removed then percutaneous coronary intervention was initiated. at the end of the case, the patient had significant left femoral artery bleeding that required medical intervention for vascular repair of the left femoral artery. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29453 | 1220648-2024-08431 | 18946971 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18946971&pc=OZD | 07/14/2022 | 03/13/2025 |  | 03/20/2024 |  |  |  | 03/20/2024 |  |  | 02/29/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023199417 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/14/2022 |  |  |  |  | The user facility reported a 74-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. pulses were unable to be obtained. the decision was to move the patient back to table to review leg perfusion. antegrade sheath was placed in right femoral artery and an arterial sheath was placed in the left femoral artery (lfa) for reperfusion. the patient developed a hematoma and bleeding at the impella cp insertion site and pressure was held and fem stop was applied. the patient continued to bleed from the groin site despite many therapies/attempts to stop bleeding. the impella cp was removed in the or and fem stop was placed. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29454 | 1220648-2024-08432 | 18946982 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18946982&pc=OZD | 07/14/2022 | 03/13/2025 |  | 03/20/2024 |  |  | 07/14/2022 | 03/20/2024 |  |  | 03/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023211494 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/14/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding that required removal of the device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
29455 | 1220648-2024-08433 | 18947003 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947003&pc=OZD | 07/14/2022 | 03/13/2025 |  | 03/20/2024 |  |  | 07/14/2022 | 03/20/2024 |  |  | 03/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023207812 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/14/2022 |  |  |  |  | The user facility reported a 63-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the impella cp was inserted and crossed the valve and ventricular placement alarms occurred. the physician attempted to pull back and the impella cp ¿shot¿ back across the valve, and the device could not advance. the physician decided to remove the impella cp and rewire. the physician wanted to use a 45cm sheath to get passed the portion of the iliacs as it was believed to be causing too much friction on the catheter. the physician was able to gain access in the right femoral artery. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29456 | 1220648-2024-08435 | 18947008 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947008&pc=OZD | 08/04/2022 | 03/20/2024 |  | 03/20/2024 |  |  | 08/04/2022 | 03/20/2024 |  |  | 04/30/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213453 |  | 0048-0045 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/04/2022 |  |  |  |  | The user facility reported a 64-year-old female in actue myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory. while on support, the patient experienced an access site bleeding near the groin area, therefore the groin area was redressed and re-sutured. the cause of the bleeding was unknown. the patient remained on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29457 | 1220648-2024-08434 | 18947010 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947010&pc=OZD | 07/15/2022 | 04/07/2025 |  | 03/20/2024 |  |  | 07/15/2022 | 03/20/2024 |  |  | 03/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023209323 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/15/2022 |  |  |  |  | The user facility reported a 69-year-old female was implanted with an impella cp device for mechanical circulatory support. it was reported the patient¿s leg pulses are very faint on dopplerl the leg was noted to be warm and lightly mottled. an external femoral bypass was performed and subsequently all pulses were identified. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29458 | 1220648-2024-08437 | 18947026 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947026&pc=OZD | 08/08/2022 | 03/20/2024 |  | 03/20/2024 |  |  | 08/08/2022 | 03/20/2024 |  |  | 04/30/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023226604 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/08/2022 |  |  |  |  | The user facility reported a 69-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support due to a thrombectomy. while on support, the patient experienced access site bleeding and had to be taken to the coronary care unit for unspecified intervention. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records.
29459 | 1220648-2024-07968 | 18947028 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947028&pc=OZD | 08/30/2022 | 03/13/2025 |  | 03/20/2024 |  |  | 08/30/2022 | 03/20/2024 |  |  | 05/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231017 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Arrhythmia | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/30/2022 |  |  |  |  | The user facility reported a 51-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while patient was on impella cp support, the pump alarmed for "impella in ventricle" and had alarms for suction and placement/low signal. transthoracic echocardiogram (tte) showed impella deep and towards mitral valve (mv). the impella was repositioned successfully. the patient experienced red urine and laboratory samples were reported as hemolyzed. it was also reported after the impella was positioned with tee the patient experienced persistent ectopy. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29460 | 1220648-2024-07970 | 18947029 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947029&pc=OZD | 08/18/2022 | 05/23/2025 |  | 03/20/2024 |  |  | 08/18/2022 | 03/20/2024 |  |  | 05/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023228536 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  |  | Death | Yes | Yes | Pressure Problem | [6 empty] | Stroke/CVA | Ischemia | Thrombocytopenia | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/18/2022 |  |  |  |  | The user facility reported a 51-year-old male in cardiomyopathy was implanted with an impella cp as a bridge to heart transplant. unfortunately, when the thrombus was removed following the initial stroke, the patient suffered a second stroke and it was noted the purge pressure was low; however, the problem was resolved before administering. both lower legs of the patient were mottled and cool to the touch. roughly two days later, the staff called with concerns regarding heparin induced thrombocytopenia. the staff was advised on a heparin free purge alternative in response to the condition. the following day, the staff called and reported that the pump was now experiencing high purge pressure. the staff was advised to run a tpa protocol to manage the high purge pressure. | Additional information noted care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07970 was provided. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29461 | 1220648-2024-07993 | 18947030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947030&pc=OZD | 11/01/2022 | 03/11/2025 |  | 03/20/2024 |  |  | 11/01/2022 | 03/20/2024 |  |  | 06/30/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2023235062 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/01/2022 |  |  |  |  | The user facility reported a 36-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient was on impella cp support and the patient experienced a loss of pulses in bilateral lower extremities. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29462 | 1220648-2024-07930 | 18947031 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947031&pc=OZD | 09/11/2022 | 05/30/2025 |  | 03/20/2024 |  |  | 09/11/2022 | 03/20/2024 |  |  | 03/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023225710 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/11/2022 |  |  |  |  | The user facility reported a patient of unknown gender was being implanted with an impella cp device for mechanical circulatory support, and that the impella cp was unable to cross the valve. ultimately the medical team aborted the procedure. there was no patient harm reported. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29463 | 1220648-2024-08141 | 18947032 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947032&pc=OZD | 08/03/2022 | 03/14/2025 |  | 03/20/2024 |  |  | 08/03/2022 | 03/20/2024 |  |  | 05/31/2023 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022089962 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/03/2022 |  |  |  |  | The user facility reported a 62-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support, due to rapid ventricular response, atrial fibrillation, hypotension, and cardiogenic shock. it was reported that while on support, bedside echocardiography was performed and found the impella was too deep into the left ventricle. the impella cp device was pulled back in the left ventricle and repositioned. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29464 | 1220648-2024-08210 | 18947033 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947033&pc=OZD | 08/06/2022 | 05/21/2025 |  | 03/20/2024 |  |  | 08/06/2022 | 03/20/2024 |  |  | 04/30/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023215544 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/06/2022 |  |  |  |  | The user facility reported a 59-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported after insertion of the introducer it was found that the femoral artery was blocked and had a tear. the femoral artery was surgically repaired and closed. the impella cp pump was explanted and replaced with an impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29465 | 1220648-2024-08130 | 18947034 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947034&pc=OZD | 03/07/2024 | 04/24/2024 |  | 03/20/2024 |  |  | 03/07/2024 | 03/20/2024 |  |  | 10/31/2025 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024424931 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Break | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/07/2024 |  |  |  |  | The user facility reported a 89-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the complainant reported that complications were encountered during attempted delivery of an impella cp pump. it was initially noted that the pump was unable to advance over the arch during implant. when attempting to remove the pump, the guidewire also came back and got caught in the peel away sheath. upon pulling the guidewire, the wire broken. the sheath, wire and pump were all replaced as a result of the noted issues. there was no patient harm. | [5 empty] | The investigation for the reported delivery issues and introducer damage has been completed. the pump, guidewire, and introducer were not returned for investigation. the root cause of the delivery issues was most likely patient condition related due to the patient tortuosity. the root cause of the product damage guidewire damager peripheral vessels issue was not determined since product was not returned. the root cause of the introducer damage - mechanical issue was not determined since product was not returned. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29466 | 1220648-2024-07992 | 18947035 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947035&pc=OZD | 03/02/2024 | 04/19/2024 |  | 03/20/2024 |  |  | 03/02/2024 | 03/20/2024 |  |  |  |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025438026 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/02/2024 |  |  |  |  | The user facility reported a 59-year-old female patient with cardiomyopathy was to be implanted with an impella 5.5 device for mechanical circulatory support. it was reported that impella 5.5 was unable to advance through the innominate artery due to the patient¿s tortuous anatomy. an impella cp was placed instead. there was no patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition due to patient's tortuous anatomy as per the clinical description provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29467 | 1220648-2024-07969 | 18947036 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947036&pc=OZD | 02/28/2024 | 04/29/2024 |  | 03/20/2024 |  |  | 02/28/2024 | 03/20/2024 |  |  |  |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025443892 |  | 0048-0014 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  |  | Malfunction | No | Yes | Material Integrity Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/28/2024 |  |  |  |  | The user facility reported a 73-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the complainant reported the sheath was noted to have kinked. the peeled 14 french sheath was used to push the impella forward. there was no reported patient harm. | [5 empty] | The investigation for the reported product damage has been completed. the impella device was not returned for investigation. the cause of introducer damage was most likely sheath kink, but the failure type is unknown since product was not returned and due to insufficient clinical data. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29468 | 1220648-2024-07991 | 18947037 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947037&pc=OZD | 12/02/2022 | 03/14/2025 |  | 03/20/2024 |  |  | 12/02/2022 | 03/20/2024 |  |  | 07/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023251492 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/02/2022 |  |  |  |  | The user facility reported a 71-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient experienced no distal flow. an external fem-fem bypass was performed successfully. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. the cause of the limb ischemia was not determined since product was not returned.
29469 | 1220648-2024-08230 | 18947038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947038&pc=OZD | 03/02/2024 | 09/13/2024 |  | 03/20/2024 |  |  | 03/02/2024 | 03/20/2024 |  |  |  |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025439532 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  |  | Death | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that an echocardiogram showed a mural thrombus in the apex. the decision was made leave the impella cp inserted as the patient was in cardiogenic shock and the impella flows and waveforms were good. the clot was being monitored for clot ingestion. several days later, the impella was deep with an active placement signal low alarm. the pump was unable to be pulled back due to left ventricular thrombus around the pigtail. a motor current spike was followed by pump stopping. the impella cp was explanted. | Patient did not survive after thrombectomy surgery. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrected information for the initial mfr 1220648-2024-08230 was provided. | The investigation for the reported thrombus and pump stop has been completed. the pump was not returned for investigation. data logs were reviewed, and motor current spike of 1442 ma was seen with high motor current-pump stop alarm at the end. saturated flows prior to pump stop at p-5 was found. no other issues were observed. the cause of the pump stop issue was most likely biomaterial per clinical and log review. as the necessary information was not provided, the root cause of the thrombus was not determined."
29470 | 1220648-2024-08212 | 18947039 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947039&pc=OZD | 08/06/2022 | 03/18/2024 |  | 03/20/2024 |  |  | 08/06/2022 | 03/20/2024 |  |  | 02/29/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023197964 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/06/2022 |  |  |  |  | The user facility reported a 63-year-old male was implanted with an impella cp device for mechanical circulatory support. while on support the patient developed excessive oozing at the impella site. the impella cp was explanted, and manual pressure was applied. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
29471 | 1220648-2024-08295 | 18947040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947040&pc=OZD | 10/19/2022 | 03/11/2025 |  | 03/20/2024 |  |  | 10/19/2022 | 03/20/2024 |  |  | 06/30/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023236010 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/19/2022 |  |  |  |  | The user facility reported a 76-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support it was also noted brief suction alarms, the patient was having ectopy issues at that time and intermittent ¿impella position unknown¿ and ¿placement signal low¿ alarms. patient was escalated to impella 5.5. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29472 | 1220648-2024-08211 | 18947041 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947041&pc=OZD | 08/06/2022 | 05/27/2025 |  | 03/20/2024 |  |  | 08/06/2022 | 03/20/2024 |  |  | 02/29/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023201803 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/06/2022 |  |  |  |  | The user facility reported a 59-year-old female patient was implanted with an impella cp device for mechanical circulatory support. the patient¿s right foot was not moving. a right calf fasciotomy was performed, and the leg appeared more molted. additionally, the patient was started on continuous renal replacement therapy. the patient experienced dark urine. the impella cp device was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29473 | 1220648-2024-08292 | 18947042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947042&pc=OZD | 07/16/2022 | 03/19/2024 |  | 03/20/2024 |  |  | 07/16/2022 | 03/20/2024 |  |  | 03/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023211350 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/16/2022 |  |  |  |  | The user facility reported a 72-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient was on impella cp support due to intervention performed upon observation of elevated troponins. while on support, the patient experienced ventricular tachycardia. additionally, the patient experienced access site bleeding that resolved without any intervention noted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29474 | 1220648-2024-08291 | 18947043 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947043&pc=OZD | 08/28/2022 | 03/14/2025 |  | 03/20/2024 |  |  | 08/28/2022 | 03/20/2024 |  |  | 12/31/2023 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023180735 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 92 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/28/2022 |  |  |  |  | The user facility reported a (b)(6) male on impella cp for mechanical circulatory support. the patient experienced limb ischemia and bleeding near groin access site area. the pump was repositioned, and hemodynamics stabilized as flow restored. j-wire prolapsed on valve at final position revealed good pump position. echo performed in procedural area prior to transport to intensive care unit. the patient was successfully weaned and explanted in cath lab. hemostasis was achieved with perclose x2. distal pulses were dopplerable post case. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29475 | 1220648-2024-08131 | 18947044 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947044&pc=OZD | 07/27/2022 | 03/14/2025 |  | 03/20/2024 |  |  | 07/27/2022 | 03/20/2024 |  |  | 04/30/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023214941 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 07/27/2022 |  |  |  |  | The user facility reported a 38-year-old patient was implanted with an impella cp device for mechanical circulatory support. it was reported while on support, the patient experienced oozing at the access site. the patient was transfused with one unit of packed red blood cells and discontinued bivalirudin infusion. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29476 | 1220648-2024-08231 | 18947045 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947045&pc=OZD | 08/27/2022 | 03/14/2025 |  | 03/20/2024 |  |  | 08/27/2022 | 03/20/2024 |  |  | 03/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023207298 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/27/2022 |  |  |  |  | The user facility reported an 80-year-old male on impella cp for mechanical circulatory support. it was reported the patient was brought to the cardiac catheterization laboratory (ccl). the impella cp was implanted post intra-aortic balloon pump (iabp) removal. impella flows on auto did not go above 2.5l/min, initially there was significant ectopy causing runs of ventricular tachycardia (vt). with suction at p5, the patient again had runs of vt. the health care professional (hcp) was pleased with the 2.5l/min flow at p4 as mean arterial pressure (map) increased. echocardiogram (echo) performed in ccl. echo shows flail leaflet not obstructing the impella inflow. suction alarms and ectopy were resolved. impella weaned overnight and the device was explanted. | [5 empty] | The impella device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. the root cause of vf/vt/af/at could not be determined as insufficient clinical details were provided.
29477 | 1220648-2024-08296 | 18947046 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947046&pc=OZD | 10/05/2022 | 03/19/2024 |  | 03/20/2024 |  |  | 10/05/2022 | 03/20/2024 |  |  | 07/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023249718 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Fracture | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/05/2022 |  |  |  |  | The user facility reported a 79-year-old male was implanted with an impella device for mechanical circulatory support. the physician reported experiencing very difficult impella cp insertion and bleeding from the introducer sheath. the physician believed that the introducer cracked during insertion. the decision made to remove the peel-away sheath and advance the repositioning sheath. a hematoma was observed and appeared to be enlarging. one unit of prbc was administered and the impella cp device was removed. despite all intervention, the patient expired. no allegations were made towards the impella for cause of or contribution to death. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29478 | 1220648-2024-08293 | 18947047 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947047&pc=PYX | 03/09/2024 | 04/26/2024 |  | 03/20/2024 |  |  | 03/09/2024 | 03/20/2024 |  |  |  |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC |  |  | US |  | 2025443738 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/09/2024 |  |  |  |  | The user facility reported a 62-year-old male was implanted with an impella rp for mechanical circulatory support. it was reported that the impella rp motor current spiked and the impella stopped and restarted three times. the physician removed the impella. the patient was reported as stable. | [5 empty] | The investigation for the reported pump stop has been completed. the impella device was returned for evaluation. upon inspection of the device, biomaterial was seen in the outflow cage and throughout the cannula. the device was able to be run after removal of biomaterial. clinical summary reported biomaterial observed in the outflow after explant and that pump stop occurred during patient move. the data logs confirm a sudden motor current spike and a rise in purge pressure and multiple #13 alarms and #18 impella stopped alarms. the root cause of the temporary pump stop was determined to be biomaterial. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
29479 | 1220648-2024-08214 | 18947048 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947048&pc=OZD | 08/13/2022 | 03/18/2024 |  | 03/20/2024 |  |  | 08/13/2022 | 03/20/2024 |  |  | 03/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023205999 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/13/2022 |  |  |  |  | The user facility reported a 75-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported that while on the impella cp support, the patient experienced hematuria. the hemolyzed labs indicated the patient had developed hemolysis. the impella cp pump was repositioned. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
29480 | 1220648-2024-08290 | 18947049 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947049&pc=OZD | 03/08/2024 | 04/25/2024 |  | 03/20/2024 |  |  | 03/08/2024 | 03/20/2024 |  |  | 11/30/2025 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024427030 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient Device Interaction Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/08/2024 |  |  |  |  | The user facility reported a 68-year-old male post cardiotomy cardiogenic shock/low cardiac output syndrome was to be implanted with an impella 5.5 for mechanical circulatory support. it was reported that after multiple attempts to place the impella 5.5, the implant was unsuccessful and was aborted. | [5 empty] | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue was most likely due to patient condition as difficulty was met at the innominate aortic junction and the axillary vessel was found to be too small for insertion. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.". | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29481 | 1220648-2024-08213 | 18947050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947050&pc=OZD | 08/06/2022 | 03/18/2024 |  | 03/20/2024 |  |  | 08/06/2022 | 03/20/2024 |  |  | 12/31/2023 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023174845 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/06/2022 |  |  |  |  | The user facility reported a 80-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported that while being placed on impella cp support, the patient developed an ischemic leg on the impella side. the impella cp pump was explanted and replaced with an impella 5.5. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records.
29482 | 1220648-2024-08294 | 18947051 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947051&pc=OZD | 09/01/2022 | 03/14/2025 |  | 03/20/2024 |  |  | 09/01/2022 | 03/20/2024 |  |  | 04/30/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023225615 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/01/2022 |  |  |  |  | The user facility reported a 76-year-old female was implanted with an impella cp device for mechanical circulatory support. the impella cp was placed in the left femoral artery and could not pass despite the multiple attempts made to balloon. during the third attempt to pass, the iliac was dissected. the physician was able to get access to contralateral side and a covered stent was placed and intra-aortic balloon pump was then placed. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29483 | 1220648-2024-08299 | 18947052 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947052&pc=OZD | 08/29/2022 | 03/14/2025 |  | 03/20/2024 |  |  | 08/29/2022 | 03/20/2024 |  |  | 05/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023227911 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/29/2022 |  |  |  |  | The user facility reported a 54-year-old male on impella cp for mechanical circulatory support. exploratory surgery was performed to assess distal limb perfusion. the patient's right leg had increased mottle with only dopplerable popliteal. the patient continued support on the impella cp. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. the root cause of limb ischemia could not be determined as insufficient clinical details were provided.
29484 | 1220648-2024-08327 | 18947053 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947053&pc=OZD | 10/05/2022 | 03/20/2024 |  | 03/20/2024 |  |  | 10/05/2022 | 03/20/2024 |  |  | 03/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023212395 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/05/2022 |  |  |  |  | The user facility reported a 69-year-old female was implanted with an impella cp device for mechanical circulatory support. the patient developed bleeding at the impella cp insertion site. the patient's activated clotting time (act) result was out of range and as a result, the introducer sheath was removed and perclose sutures were tied. following deployment of both sutures, the patient still had significant oozing and as a result, protamine was given to reverse the affects of heparin and 45 minutes of manual pressure was held to resolve the bleeding. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29485 | 1220648-2024-07900 | 18947054 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947054&pc=OZD | 09/06/2022 | 03/13/2024 |  | 03/20/2024 |  |  | 09/06/2022 | 03/20/2024 |  |  | 06/30/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231341 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/06/2022 |  |  |  |  | The user facility reported a 72-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. when placing the impella cp, the patient was in torsades for 45 minutes but able to maintain a mean arterial pressure (map) of 60. the impella position was checked via echo and adjusted accordingly. once in position and after dilation, the patient converted to sinus rhythm. all pressors were removed and patient's map was reported at 68. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29486 | 1220648-2024-07885 | 18947056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947056&pc=OZD | 08/26/2022 | 03/14/2025 |  | 03/20/2024 |  |  | 08/26/2022 | 03/20/2024 |  |  | 05/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023230116 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/26/2022 |  |  |  |  | The user facility reported a 68-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced loss of pulsatility in their leg adjacent to impella cp device. an angiogram showed occlusion of the common femoral artery and slowly filling profounda. the impella cp device was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. the root cause of limb ischemia could not be determined as insufficient clinical details were provided.
29487 | 1220648-2024-07901 | 18947057 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947057&pc=OZD | 02/21/2023 | 03/13/2024 |  | 03/20/2024 |  |  | 02/21/2023 | 03/20/2024 |  |  |  |  | 03/20/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1205049 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Difficult to Program or Calibrate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 02/21/2023 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) experienced an unspecified purge cassette related failure. the aic was removed from service. there was no patient harm. | [5 empty] | The investigation into the automated impella controller (aic) purge issue-other has been completed. console ic5367 was inspected in danvers, but no irregularities of the purge system were discovered, and the device operated within specification. the cause of the issue was most likely use related. this report is being filed as part of a retrospective review of historical records.
29488 | 1220648-2024-07887 | 18947058 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947058&pc=OZD | 07/26/2022 | 03/13/2025 |  | 03/20/2024 |  |  | 07/26/2022 | 03/20/2024 |  |  | 02/29/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202117 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hematoma | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/26/2022 |  |  |  |  | The user facility reported a 59-year-old male was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient had hematoma and bleeding at the access site. this was treated by applying manual pressure and sutures. a decline in hemoglobin levels was also noted so two units of blood were given to the patient. additionally, through a computer tomography (ct) test it was found that the patient had an intracranial hemorrhage. in response to the hemorrhage, the decision was made to withdraw care. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29489 | 1220648-2024-07967 | 18947059 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947059&pc=OZD | 02/28/2024 | 05/01/2024 |  | 03/20/2024 |  |  | 02/28/2024 | 03/20/2024 |  |  | 08/31/2025 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024388377 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/28/2024 |  |  |  |  | The user facility reported a 65-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient¿s leg was ¿marbled¿ and not perfused. this lead to the impella cp being escalated to an impella 5.5. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ to conform with updated procedures, sections d6 & h10 have been revised with updated information.
29490 | 1220648-2024-08298 | 18947060 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947060&pc=OZD | 07/16/2022 | 03/19/2024 |  | 03/20/2024 |  |  | 07/16/2022 | 03/20/2024 |  |  | 02/29/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2023195337 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  |  | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/16/2022 |  |  |  |  | The user facility reported a 68-year-old male, race and ethnicity unknown, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient was on impella cp support due to an unspecified cardiac incident. while on support, the patient experienced access site bleeding that appears to have resolved without any additional care. additionally, the patient experienced low pump flow that required replacement with a new pump. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29491 | 1220648-2024-07886 | 18947061 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947061&pc=OZD | 03/03/2024 | 04/01/2024 |  | 03/20/2024 |  |  | 03/04/2024 | 03/20/2024 |  |  |  |  | 03/20/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1464036 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 21 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/04/2024 |  |  |  |  | It was reported that an automated impella controller (aic) displayed a controller failure alarm during patient support. the aic was swapped in response to the issue. there was no patient involved. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - controller failure (red alarm) has been completed since the initial report was submitted. the console was returned, visually inspected, and tested. the failure was unable to be reproduced throughout testing and is characterized by a temporary loss of optical data and triggering of a controller error alarm, has a low probability of reoccurrence. the cause of the issue was not determined since issue could not be reproduced.
29492 | 1220648-2024-08297 | 18947063 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947063&pc=OZD | 09/01/2022 | 03/13/2025 |  | 03/20/2024 |  |  | 09/01/2022 | 03/20/2024 |  |  | 04/30/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023224574 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/01/2022 |  |  |  |  | The user facility reported a 61-year-old male was implanted with an impella cp device for mechanical circulatory support. while on impella cp support, the patient developed bleeding at the access site and concerns of left lower leg ischemia. the patient received multiple units of packed red blood cells (prbc) and systemic heparin was removed. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29493 | 1220648-2024-08301 | 18947064 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947064&pc=OZD | 09/01/2022 | 03/19/2024 |  | 03/20/2024 |  |  | 09/01/2022 | 03/20/2024 |  |  | 02/29/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200899 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/01/2022 |  |  |  |  | The user facility reported a 40-year-old male was implanted with an impella cp device for mechanical circulatory support and that stitches were placed to control bleeding. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29494 | 1220648-2024-08438 | 18947193 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947193&pc=OZD | 08/14/2022 | 03/20/2024 |  | 03/20/2024 |  |  | 08/14/2022 | 03/20/2024 |  |  | 04/30/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023223899 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Female |  | Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/14/2022 |  |  |  |  | The user facility reported a 45-year-old female patient on an impella cp for mechanical circulatory support. the patient experienced groin complications prior to and following impella cp implant. the pump was implanted via right femoral access after complications in the groin were thought to have stabilized. after implant, a significant hematoma was found at the site of the original complication and hemoglobin levels began trending downward. as a result of the patient¿s worsening condition, the decision was made to withdraw care. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29495 | 1220648-2024-08436 | 18947195 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947195&pc=OZD | 12/26/2023 | 04/04/2024 |  | 03/20/2024 |  |  | 02/19/2024 | 03/20/2024 |  |  | 10/31/2025 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024411186 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 43 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/19/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry regarding a 43 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, hemolysis was alleged with a "possible" relationship to the impella device. unspecified intervention was required. | [5 empty] | The investigation into the hemolysis issue has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
29496 | 1220648-2024-08439 | 18947208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947208&pc=OZD | 08/14/2022 | 05/27/2025 |  | 03/20/2024 |  |  | 08/14/2022 | 03/20/2024 |  |  | 02/29/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195608 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/14/2022 |  |  |  |  | The user facility reported a 69-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was placed on the impella cp for hemodynamic support. during post op, the patient was turned and appeared to have some pink tinged urine. it was reported they had position issues and the impella needed to be repositioned several times. the impella device was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29497 | 1220648-2024-08440 | 18947245 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947245&pc=OZD | 08/14/2022 | 03/07/2025 |  | 03/20/2024 |  |  | 08/14/2022 | 03/20/2024 |  |  | 03/31/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023204906 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/14/2022 |  |  |  |  | The user facility reported a 73-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient had a weak doppler pulse in the left foot at night and the day shift had a hard time finding a pulse. bilateral feet were dusky. two units of packed red blood cells (prbcs) were ordered to help increase the p-level. the family withdrew care and the patient expired. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
29498 | 1220648-2024-08441 | 18947262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18947262&pc=OZD | 08/21/2022 | 03/13/2025 |  | 03/20/2024 |  |  | 08/21/2022 | 03/20/2024 |  |  | 02/29/2024 |  | 03/20/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023203101 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/21/2022 |  |  |  |  | The user facility reported a 65-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the pump had to be repositioned multiple times due to continuous suction/malposition issues. the pump was repositioned and found to be in the ventricle. the impella was explanted. | [5 empty] | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29499 | 1220648-2024-08215 | 18928655 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18928655&pc=OZD | 08/04/2022 | 03/11/2025 |  | 03/19/2024 |  |  | 08/04/2022 | 03/19/2024 |  |  | 04/30/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023217737 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/04/2022 |  |  |  |  | The user facility reported a 55-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced bleeding at the impella access site. the patient was transfused with 1 unit of packed red blood cells and heparin was discontinued. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | The root cause of the access site bleeding could not be determined as insufficient clinical details were provided.
29500 | 1220648-2024-08216 | 18928662 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18928662&pc=OZD | 08/23/2022 | 03/19/2024 |  | 03/19/2024 |  |  | 08/23/2022 | 03/19/2024 |  |  | 07/31/2023 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022121428 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/23/2022 |  |  |  |  | The user facility reported a 73-year-old female patient on an impella 5.5 for mechanical circulatory support. it was reported that during support the patient experienced a hematoma at the impella site. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29501 | 1220648-2024-08217 | 18928689 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18928689&pc=OZD | 08/25/2022 | 03/19/2024 |  | 03/19/2024 |  |  | 08/25/2022 | 03/19/2024 |  |  | 04/30/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023219264 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/25/2022 |  |  |  |  | The user facility reported a 57-year-old female patient was implanted with an impella cp for mechanical circulatory support after experiencing a non-st elevated myocardial infarction. it was reported that the patient experienced bleeding and loss of distal pulses. the patient was weaned down to p2 and did well. the pump was removed at bedside. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
29502 | 1220648-2024-08218 | 18928767 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18928767&pc=OZD | 08/04/2022 | 03/12/2025 |  | 03/19/2024 |  |  | 08/04/2022 | 03/19/2024 |  |  | 04/30/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023215557 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/04/2022 |  |  |  |  | The user facility reported a 54-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. patient was placed on impella cp support after the patient experienced ventricular fibrillation (vfib) and had three rounds of epinephrine and multiple defibrillations before return of spontaneous circulation was achieved. it was reported that while on impella cp support, the patient experienced vfib arrest. the medical team performed compressions and shocked the patient and return of spontaneous circulation was achieved. the impella cp was repositioned with echocardiography to ensure it was kept shallow in the ventricle and to avoid touching ventricle any deeper. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The root cause of vt/vf could not be determined as insufficient clinical details were provided. no logs or device was returned for analysis.
29503 | 1220648-2024-08219 | 18928778 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18928778&pc=OZD | 08/25/2022 | 03/11/2025 |  | 03/19/2024 |  |  | 08/25/2022 | 03/19/2024 |  |  | 05/31/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023228723 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Deliver | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/25/2022 |  |  |  |  | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient was being placed on impella cp support, the impella sheath was placed and .018 wire placed in left ventricle. the physician had difficulty advancing past the iliac bifurcation requiring multiple attempts. the patient started to stabilize, and the physician decided to abort the impella placement due to patient vascular anatomy. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29504 | 1220648-2024-08220 | 18928799 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18928799&pc=OZD | 08/25/2022 | 03/10/2025 |  | 03/19/2024 |  |  | 08/25/2022 | 03/19/2024 |  |  | 05/31/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231157 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/25/2022 |  |  |  |  | The user facility reported a 64-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported during placement the pigtail was removed and impella cp was backloaded onto the .018 wire. dr. drew an abg from oscor sheath and clot was noted in aspirated blood. it was reported that no pulses were noted in feet. potassium remains over 7 so decision made to perform dialysis to wash and clean patient blood but not pull off volume at this time. patient was given 2 units of prbcs and 1 pack of platelets. left leg mottling was much worse and now right leg was starting to mottle as well. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29505 | 1220648-2024-08221 | 18928952 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18928952&pc=OZD | 07/15/2022 | 03/19/2024 |  | 03/19/2024 |  |  | 07/15/2022 | 03/19/2024 |  |  | 01/31/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023190556 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/15/2022 |  |  |  |  | The user facility reported a patient, demographics unknown, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient was on impella cp support due to st-elevated myocardial infarction. while on support, the patient experienced access site bleeding that required vascular surgery to resolve. the patient needed a cutdown for complete hemostasis. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29506 | 1220648-2024-08222 | 18928991 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18928991&pc=OZD | 02/23/2024 | 04/11/2024 |  | 03/19/2024 |  |  | 02/23/2024 | 03/19/2024 |  |  | 11/30/2025 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024431085 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | PHYSICIAN ASSISTANT | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/23/2024 |  |  |  |  | The user facility reported a 28-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the patient had sternal bleeding and anticoagulation was discontinued. a sternal washout was performed and there was bloody gastrointestinal output. the patient received three units of packed red blood cells. furthermore, the patient had increasing pericardial effusion. clots were also noted in the left anterior artery. additionally, the patient had rapid atrial fibrillation (af) while on support. | [5 empty] | The investigation into the reported issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the ventricle perforation cannot be determined as the time and cause of the ventricle perforation is not known. as the necessary information was not provided, the root cause of the atrial fibrillation (vf/vt/af/at) was not determined. the cause of the optical signal issue was most likely a damaged optical fiber line due to torque. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿
29507 | 1220648-2024-08223 | 18929050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18929050&pc=OZD | 07/15/2022 | 03/12/2025 |  | 03/19/2024 |  |  | 07/15/2022 | 03/19/2024 |  |  | 03/31/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023205974 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/15/2022 |  |  |  |  | The user facility reported a 76-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient was on impella cp support in association with experiencing gallstones, chest pains, and liver dysfunction. while on support, the pump pulled back into the aorta and was unable to be repositioned. the pump was replaced to resolve the issue. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The root cause of the positioning issues could not be determined as insufficient clinical details were provided.
29508 | 1220648-2024-08224 | 18929071 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18929071&pc=OZD | 02/26/2024 | 04/15/2024 |  | 03/19/2024 |  |  | 02/26/2024 | 03/19/2024 |  |  | 11/30/2025 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024426479 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/26/2024 |  |  |  |  | The user facility reported a 69-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was on impella cp support due to left anterior descending artery stenosis. while on support, the patient had access site bleeding and red blood cells were provided. additionally, there were multiple attempts to reposition the pump that were unsuccessful. the staff decided to remove the impella cp and replace the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation for the reported access site bleed and the positioning issues has been completed. the product was not returned for review. based on clinical description, the root cause of positioning issue was most likely use related. the cause of the access site bleeding issue was most likely patient condition related due to bleeding at several sites per clinical.
29509 | 1220648-2024-08225 | 18929330 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18929330&pc=OZD | 07/15/2022 | 05/23/2025 |  | 03/19/2024 |  |  | 07/15/2022 | 03/19/2024 |  |  | 03/31/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023214679 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 07/15/2022 |  |  |  |  | The user facility reported a 88-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient's right coronary artery was noted to be occluded, percutaneous coronary intervention of right coronary artery with no reflow. upon decompensation, impella cp was placed, and the decision was made to leave left anterior descending (lad) lesion alone at this time due to low viability and chronic in nature. it was noted that the patient's urine was pink tinged, and as hemolysis was still present, discussion was to potentially reposition impella cp. echocardiogram (echo) was performed, and the left impella was repositioned. slack removed, sheath secured, stick angle maintained. | Additional information noted care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08225 was provided. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29510 | 1220648-2024-08226 | 18929332 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18929332&pc=OZD | 10/04/2022 | 05/27/2025 |  | 03/19/2024 |  |  | 10/04/2022 | 03/19/2024 |  |  | 07/31/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023251715 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/04/2022 |  |  |  |  | The user facility reported a 86-year-old male was being implanted with an impella cp device for mechanical circulatory support. it was reported during insertion, the impella cp was inserted through the sheath but was unable to be delivered to the patient's left ventricle, and the insertion was aborted due to tortuous and small anatomy. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29511 | 1220648-2024-08227 | 18929954 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18929954&pc=OZD | 07/19/2022 | 03/12/2025 |  | 03/19/2024 |  |  | 07/19/2022 | 03/19/2024 |  |  | 12/31/2023 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023172203 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/19/2022 |  |  |  |  | The user facility reported a 39-year-old female was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced limb ischemia that was resolved with a distal perfusion sheath. also, the impella cp was removed, and the patient was escalated to impella 5.5. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The root cause of limb ischemia could not be determined as insufficient clinical details were provided.
29512 | 1220648-2024-08228 | 18930088 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18930088&pc=OZD | 02/28/2024 | 05/02/2024 |  | 03/19/2024 |  |  | 02/28/2024 | 03/19/2024 |  |  | 11/30/2025 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024430237 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 36 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during impella cp support for a 36-year-old female patient, a left ventricle perforation requiring a covered stent and/or balloon tamponade occurred. the perforation was detected 30 minutes after the impella cp was implanted. the impella cp was explanted as a result. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ | The investigation into the ventricle perforation issue has been completed since the original report was submitted. the device was not returned by the user facility for investigation. clinical details noted that perforation of ventricle occurred when cpr was being performed. the root cause cannot be determined as the pump and data logs were not returned for evaluation and limited clinical details were provided. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
29513 | 1220648-2024-08229 | 18930510 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18930510&pc=OZD | 08/26/2022 | 03/07/2025 |  | 03/19/2024 |  |  | 08/26/2022 | 03/19/2024 |  |  | 03/31/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023206493 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/26/2022 |  |  |  |  | The user facility reported a 34-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the device came out of position within the aorta, post cpr. the physician was able to cross the valve; however, it was under papillary muscle which caused significant mitral regurgitation (mr) on the echocardiogram (echo). the decision was made to explant the pump. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29514 | 1220648-2024-08232 | 18931087 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18931087&pc=OZD | 07/15/2022 | 03/19/2024 |  | 03/19/2024 |  |  | 07/15/2022 | 03/19/2024 |  |  | 01/31/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2023187119 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/15/2022 |  |  |  |  | The user facility reported a 40-year-old male patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the complainant reported a large left ventricle thrombus was present on echocardiogram (echo). the pump was weaned and explanted from the patient. intra-aortic balloon pump was placed to bridge the patient to possible left ventricular assist device or transplant as the patient continued to recover from shock. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ this report is being filed as part of a retrospective review of historical records.
29515 | 1220648-2024-08233 | 18931172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18931172&pc=OZD | 07/20/2022 | 02/18/2025 |  | 03/19/2024 |  |  | 07/20/2022 | 03/19/2024 |  |  | 04/30/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023216669 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic |  | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/20/2022 |  |  |  |  | The user facility reported a 78-year-old female, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the purge tubing was damaged upon insertion and required replacement due to leakage and low purge pressure. additionally, the patient experienced access site bleeding and a large hematoma that required vascular surgery to remove the pump and repair the groin. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29516 | 1220648-2024-08234 | 18931328 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18931328&pc=OZD | 03/02/2024 | 04/18/2024 |  | 03/19/2024 |  |  | 03/02/2024 | 03/19/2024 |  |  | 06/30/2025 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024375526 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/02/2024 |  |  |  |  | The user facility reported a 77-year-old male in st-segment elevation myocardial infarction was to be implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp pump was unable to advance through the vasculature of the patient during attempted implant. it was noted that the vessel was too tortuous and calcified, which created resistance during the delivery. the implant was subsequently aborted, and an intra-aortic balloon pump was placed for patient support. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition since the vessel too tortuous and highly calcified.
29517 | 1220648-2024-08235 | 18931485 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18931485&pc=OZD | 08/02/2022 | 03/12/2025 |  | 03/19/2024 |  |  | 08/02/2022 | 03/19/2024 |  |  | 04/30/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213920 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/02/2022 |  |  |  |  | The user facility reported a 64-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient had bedside echocardiography performed that revealed shallow pump position. the impella cp device was repositioned and advanced forward. the patient also experienced distal limb ischemia during impella support and the medical team decided to explant the device to resolve. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29518 | 1220648-2024-08236 | 18931512 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18931512&pc=OZD | 10/18/2022 | 05/13/2025 |  | 03/19/2024 |  |  | 10/18/2022 | 03/19/2024 |  |  | 06/30/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023237688 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/18/2022 |  |  |  |  | The user facility reported a 75-year-old female was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, an aorta alarm occurred. it was determined under fluoroscopy that the impella cp was prolapsed in the ascending aorta. the physician tried to reposition the impella cp back across the aortic valve but was unsuccessful, so the pump was exchanged for a new one. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29519 | 1220648-2024-08237 | 18932126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18932126&pc=OZD | 10/18/2022 | 03/19/2024 |  | 03/19/2024 |  |  | 10/18/2022 | 03/19/2024 |  |  | 06/30/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023236703 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/18/2022 |  |  |  |  | The user facility reported a 62-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported that during impella support, the patient developed bleeding at the impella access site. as a result of the noted issue, heparin was discontinued, sutures and manual pressure were applied. the impella cp device was explanted, and the patient was escalated to an impella 5.5 device. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29520 | 1220648-2024-08240 | 18932160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18932160&pc=OZD | 10/18/2022 | 03/19/2024 |  | 03/19/2024 |  |  | 10/18/2022 | 03/19/2024 |  |  | 03/31/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023206492 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 10/18/2022 |  |  |  |  | The user facility reported a 63-year-old needing mechanical circulatory support. it was while on support, patient had significant amount of oozing from the groin site and a large hematoma. therefore, the impella was removed, and the patient was hemodynamically stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records.
29521 | 1220648-2024-08238 | 18932170 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18932170&pc=OZD | 03/04/2024 | 05/06/2024 |  | 03/19/2024 |  |  | 03/04/2024 | 03/19/2024 |  |  | 10/31/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023272567 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported impella cp insertion was aborted due to the patient's anatomy. the device was unable to be advanced through the calcified vasculature. the resistance was felt trying to access the femoral artery. there was no reported patient harm. | [5 empty] | The investigation for the reported delivery and removal issues has been completed. no product was returned for evaluation. clinical details noted that pump was unable to pass through a constriction in the femoral artery due to calcification. the root cause of the delivery issues was due to patient's calcified arteries. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29522 | 1220648-2024-08239 | 18932188 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18932188&pc=OZD | 07/20/2022 | 03/19/2024 |  | 03/19/2024 |  |  | 07/20/2022 | 03/19/2024 |  |  | 04/30/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213639 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/20/2022 |  |  |  |  | The user facility reported a 48-year-old male, was implanted with an impella cp device for mechanical circulatory support. the patient had some oozing at the entry site. increasing the angle of entry, tension sutures and applied dressings were able to stop the bleeding. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29523 | 1220648-2024-08241 | 18932247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18932247&pc=OZD | 10/18/2022 | 03/19/2024 |  | 03/19/2024 |  |  | 10/18/2022 | 03/19/2024 |  |  | 09/30/2023 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022146015 |  | 0048-0045 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Suction Problem | [6 empty] | Hematoma | Diminished Pulse Pressure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/18/2022 |  |  |  |  | The user facility reported a 58-year-old male patient noted as high-risk percutaneous coronary intervention, was implanted with an impella cp for mechanical circulatory support. while on support, it was reported the patient experienced dampening in pressure while stinting and received fluids. in addition, the patient experienced suction alarms and developed a hematoma at the access site that was resolved with pressure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29524 | 1220648-2024-08242 | 18932267 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18932267&pc=OZD | 08/05/2022 | 03/07/2025 |  | 03/19/2024 |  |  | 08/05/2022 | 03/19/2024 |  |  | 04/30/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023218833 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/05/2022 |  |  |  |  | The user facility reported a 70-year-old female in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported during impella cp support, the patient experienced a lack of distal pulses. the medical team decided to remove the impella cp device because of the limb ischemia. during removal of the impella device, the femoral artery was found to be occluded and the surgeon performed cutdown for the impella and thrombectomy removal. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events: acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Analysis summary: the failure mode of "thrombus¿ is to be used when a thrombus is noted but cannot be connected to any other adverse events or decrease in pump performance. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions description of the location and characteristics of the thrombus (e.g., size, shape, consistency). relevant laboratory test results, such as coagulation profiles and platelet counts imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). medications or treatments that the patient was receiving at the time of thrombus formation. surgical or procedural details if applicable (e.g, cardiac catheterization) any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition. any concomitant medication. vascular size or variations thereof. detailed description of the symptoms experienced by the patient. physical examination findings, including pulse assessment and visual examination of the affected limb. diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. laboratory test results, including blood tests for markers of inflammation or clotting disorders. any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
29525 | 1220648-2024-08244 | 18932393 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18932393&pc=OZD | 10/04/2022 | 03/12/2025 |  | 03/19/2024 |  |  | 10/04/2022 | 03/19/2024 |  |  | 05/31/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023230084 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Malfunction | Yes | Yes | Failure to Deliver | Failure to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/04/2022 |  |  |  |  | The user facility reported a 65-year-old female patient was attempting to be implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp pump was unable to advance due to high calcification within the patient's aortic arch. therefore, the impella cp implant was aborted and an impella 2.5 was subsequently placed for patient support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29526 | 1220648-2024-08246 | 18932407 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18932407&pc=OZD | 07/20/2022 | 03/12/2025 |  | 03/19/2024 |  |  | 07/20/2022 | 03/19/2024 |  |  | 04/30/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023215037 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/20/2022 |  |  |  |  | The user facility reported a 65-year-old male was implanted with an impella cp device for mechanical circulatory support. while on impella cp support, the ¿impella stopped - motor current high¿ alarm was triggered, the impella was restarted three times but was unsuccessful. the physician was made aware that impella would not be able to assist with cardioplegia delivery unless clamp was repositioned. once cross clamp was removed, flows were able to be increased without the ¿impella alarm,¿ however, after trying to increase flow above p2 the impella alarmed continuous suction alarms. transesophageal echocardiography (tee) was performed and the impella cp noted to be very shallow and stuck in papillary. the pump was repositioned, and flows were then able to be increased. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29527 | 1220648-2024-08243 | 18932418 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18932418&pc=OZD | 03/04/2024 | 04/18/2024 |  | 03/19/2024 |  |  | 03/04/2024 | 03/19/2024 |  |  |  |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025440986 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the patient was on impella cp support due to having cardiogenic shock. the doctor found that the the impella cp was deep at 5.5 centimeters. the impella cp was withdrawn and came across the atrioventricular with the pigtail still in the left ventricle. the staff attempted to advance the pump but, it came entangled in the papillary. the impella cp was therefore replaced. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the positioning issue was use related.
29528 | 1220648-2024-08245 | 18932454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18932454&pc=OZD | 07/16/2022 | 03/19/2024 |  | 03/19/2024 |  |  | 07/16/2022 | 03/19/2024 |  |  | 03/31/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023207138 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/16/2022 |  |  |  |  | The user facility reported a female, age unknown, in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that a patient was placed on impella cp support, and the patient developed bleeding at the impella cp insertion site. surgical cut down was required to fix the bleeding issue at impella site. as the patient was not pump dependent, the device was removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29529 | 1220648-2024-08247 | 18932636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18932636&pc=OZD | 07/20/2022 | 03/19/2024 |  | 03/19/2024 |  |  | 07/20/2022 | 03/19/2024 |  |  | 01/31/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023186753 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/20/2022 |  |  |  |  | The user facility reported a 70-year-old male was implanted with an impella cp device for mechanical circulatory support. the patient developed bleeding from the impella cp insertion site and was switched to sodium bicarbonate. the site was redressed, and angle matched. the oozing resolved with tightening of pro-glide. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29530 | 1220648-2024-08248 | 18933041 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18933041&pc=OZD | 07/20/2022 | 05/13/2025 |  | 03/19/2024 |  |  | 07/20/2022 | 03/19/2024 |  |  | 04/30/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023217591 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/20/2022 |  |  |  |  | The user facility reported a 69-year-old male was implanted with an impella cp device for mechanical circulatory support. during support, the patient was experiencing a hematoma at the access site with manual pressure performed. the decision was made to wean the impella and explant with manual pressure. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29531 | 1220648-2024-08249 | 18933100 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18933100&pc=OZD | 03/04/2024 | 04/19/2024 |  | 03/19/2024 |  |  | 03/04/2024 | 03/19/2024 |  |  | 02/28/2025 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024337609 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/04/2024 |  |  |  |  | The user facility reported a 75-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that after placement of impella cp patient developed large hematoma, contralateral access gained, and groin shot visualized small dissection of common femoral artery. manual compression was held for an hour and 30 minutes, contralateral site utilized again to view femoral artery and bleeding had stopped. the physician decided to leave impella at this point. it was further reported that when explanting the impella, vascular surgeon closed the impella access site. the patient is stable. | [5 empty] | The investigation for the reported vessel damage has been completed. the device was not returned for evaluation. however, clinical details provided was sufficient to determine the root cause. the root cause of the vessel damage was most likely due to use issue. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
29532 | 1220648-2024-08250 | 18933195 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18933195&pc=OZD | 03/05/2024 | 04/23/2024 |  | 03/19/2024 |  |  | 03/05/2024 | 03/19/2024 |  |  | 11/30/2025 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024430096 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient pulled the impella and repositioning sheath completely out of the leg. manual pressure was held over the access site until the physician arrived for emergent groin closure. existing perclose sutures secured with hemostasis achieved and light manual pressure was held over the access site. distal pulse was present on doppler. two units of packed red blood cells were given. the patient was hemodynamically stable. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical details provided was sufficient to determine the root cause. the root cause of the access site bleeding was most likely due to use issue be patient movement as it was reported that the patient pulled the repositioning sheath out of the leg. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
29533 | 1220648-2024-08251 | 18933202 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18933202&pc=OZD | 10/04/2022 | 05/29/2025 |  | 03/19/2024 |  |  | 10/04/2022 | 03/19/2024 |  |  | 07/31/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023250090 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/04/2022 |  |  |  |  | The user facility reported a 61-year-old female was being implanted with an impella cp device for mechanical circulatory support. it was reported the physician was unable to deliver the pump to the left ventricle due to patient's anatomy. the impella cp implant procedure was aborted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29534 | 1220648-2024-08252 | 18933353 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18933353&pc=OZD | 03/05/2024 | 05/02/2024 |  | 03/19/2024 |  |  | 03/05/2024 | 03/19/2024 |  |  |  |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025451004 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 87 YR |  | Male |  |  |  | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following impella cp implant procedure for 87-year-old male patient, no flow was established. the left ventricle was noted to have good filling. the impella was therefore explanted and a new impella was placed to continue the high risk percutaneous coronary intervention procedure without issue. the explanted impella was tested in water bath and no flow could be established there either. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported low pump flow has been completed. the cp pump was returned and visually inspected. no biomaterial nor broken blades were found. a foreign object was observed around impeller and a cannula kink was observed near the outlet cage. the cause of the low pump flow was interaction with other device due to foreign object ingested before or during insertion that obstructed the flow. the cannula kink may have occurred during implant or explant. to conform to updated procedures, section d6 was revised with updated information.
29535 | 1220648-2024-08254 | 18933391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18933391&pc=OZD | 03/04/2024 | 04/19/2024 |  | 03/19/2024 |  |  | 03/04/2024 | 03/19/2024 |  |  | 07/31/2025 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024380227 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/04/2024 |  |  |  |  | The user facility reported a 55-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. while on support, hemolysis was experienced. the impella cp was positioned low and the patient's urine was tea colored followed by anuria post impella implant. the patient received one unit of packed red blood cells and was started on continuous veno-venous hemofiltration. the impella was pulled back. | [5 empty] | The investigation for the reported hemolysis issue has been completed. the cp pump was not returned. data logs were reviewed and several positioning alarms in aorta and ventricle were observed within first 20 hours of impella placement. no other abnormalities were noted. as per clinical hemolysis likely resolved after pump repositioning and blood being given. the cause of the hemolysis issue was most likely improper positioning of the device per clinical and log review. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
29536 | 1220648-2024-08253 | 18933396 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18933396&pc=OZD | 08/31/2022 | 05/20/2025 |  | 03/19/2024 |  |  | 08/31/2022 | 03/19/2024 |  |  | 04/30/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023224906 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Fluid/Blood Leak | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/31/2022 |  |  |  |  | The user facility reported a 61-year-old male was implanted with an impella cp device for mechanical circulatory support. the patient experienced oozing at all catheter access sites. the dressing was changed, and it was reported that the access site was clean, dry, and intact. the patient was on sodium bicarbonate in the purge with extension tubing and systemic heparin. the staff changed the purge cassette three (3) times within twenty-four (24) hours due to leaking fluid from the cassette. the patient was escalated to an impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29537 | 1220648-2024-08255 | 18933561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18933561&pc=OZD | 07/22/2022 | 03/12/2025 |  | 03/19/2024 |  |  | 07/22/2022 | 03/19/2024 |  |  | 04/30/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023214965 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/22/2022 |  |  |  |  | The user facility reported a 79-year-old female in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump could not cross the aortic valve, so another pump was placed. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29538 | 1220648-2024-08256 | 18933577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18933577&pc=OZD | 03/05/2024 | 04/08/2024 |  | 03/19/2024 |  |  | 03/05/2024 | 03/19/2024 |  |  | 11/30/2025 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024427522 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/05/2024 |  |  |  |  | The user facility reported a 54-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient was on impella 5.5 support due to needing a bridge to recovery. while on support, there was a hematoma at the access site and a washout was performed. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided.
29539 | 1220648-2024-08257 | 18933585 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18933585&pc=OZD | 07/20/2022 | 04/17/2025 |  | 03/19/2024 |  |  | 07/20/2022 | 03/19/2024 |  |  | 02/29/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200837 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/20/2022 |  |  |  |  | The user facility reported a 65-year-old female was implanted with an impella cp device for mechanical circulatory support, and that the guard wire was lost during the dilator exchange, manual pressure was held, and hemostasis was achieved without complications. the impella cp device was aborted, and the patient was escalated to an impella 5.5 device. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29540 | 1220648-2024-08258 | 18933596 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18933596&pc=OZD | 07/22/2022 | 03/19/2024 |  | 03/19/2024 |  |  | 07/22/2022 | 03/19/2024 |  |  | 12/31/2023 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023172895 |  | 0048-0045 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/22/2022 |  |  |  |  | The user facility reported a male patient on impella cp for mechanical circulatory support. while on support, the patient experienced limb ischemia and an access site bleeding near the groin area. there was unspecified intervention for the bleeding. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29541 | 1220648-2024-08259 | 18933614 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18933614&pc=OZD | 10/19/2022 | 04/07/2025 |  | 03/19/2024 |  |  | 10/19/2022 | 03/19/2024 |  |  | 07/31/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023253600 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup |  |  | 10/19/2022 |  |  |  |  | The user facility reported a 46-year-old patient needing mechanical circulatory support. it was reported that the patient was placed on impella cp support. left heart catheterization was performed showing both stents had gone down due to thrombus. | Additional information noted that patient was send for emergency cabg but efforts were unsuccessful and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrected information for the initial mfr 1220648-2024-08259 was provided in sections b2, b5, d6, h1, and h6. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29542 | 1220648-2024-08260 | 18933756 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18933756&pc=OZD | 10/19/2022 | 03/10/2025 |  | 03/19/2024 |  |  | 10/19/2022 | 03/19/2024 |  |  | 03/31/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023209365 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 45 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/19/2022 |  |  |  |  | The user facility reported a 45-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the physician reported that the patient was weaned off the impella due to the oozing from the groin site. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleeding issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
29543 | 1220648-2024-08261 | 18933761 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18933761&pc=OZD | 07/24/2022 | 03/12/2025 |  | 03/19/2024 |  |  | 07/24/2022 | 03/19/2024 |  |  | 01/31/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2023190539 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  |  |  | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/24/2022 |  |  |  |  | The user facility reported a 61-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient experienced access site bleeding that required manual compression and preclose with perclose proglide sutures. the patient also experienced a positioning issue where the device measured at 3.9cm and shallow left ventricle depth and the pigtail at apex. the surgeon chose to not reposition the impella cp at that time. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29544 | 1220648-2024-08262 | 18933793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18933793&pc=OZD | 03/05/2024 | 04/17/2024 |  | 03/19/2024 |  |  | 03/05/2024 | 03/19/2024 |  |  | 11/30/2025 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024427543 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/05/2024 |  |  |  |  | The user facility reported a 69-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient was on impella 5.5 support due to needing a bridge. while on support, the patient experienced a gastrointestinal bleed (gib) on dual antiplatelet therapy. the patient was not a candidate for revascularization procedure. the decision was made to wean the impella and explant. two units of packed red blood cells were provided and the pump was removed. | [5 empty] | The investigation into the vascular damage - peripheral vessel issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the vascular damage gi bleed and its relation to impella were not determined since product was not returned and there is a lack of clinical details. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
29545 | 1220648-2024-08263 | 18933820 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18933820&pc=OZD | 10/20/2022 | 03/10/2025 |  | 03/19/2024 |  |  | 10/20/2022 | 03/19/2024 |  |  | 07/31/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023250551 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Pumping Stopped | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/20/2022 |  |  |  |  | The user facility reported a 59-year-old female was implanted with an impella cp device for mechanical circulatory support. the pump's motor stopped 20 minutes into procedure. the physician tried restarting the pump at the same p-level and failed. two more attempts of restarting at p-2 and failed. after numerous tries of restarting the pump, the device was exchanged for another and the exchange was a success. perforation requiring covered stent and/or balloon tamponade occurred. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events : acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ "if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported pump stop and vascular damage - peripheral vessel issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
29546 | 1220648-2024-08264 | 18933844 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18933844&pc=OZD | 10/20/2022 | 03/19/2024 |  | 03/19/2024 |  |  | 10/20/2022 | 03/19/2024 |  |  | 06/30/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  | 2023237434 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 10/20/2022 |  |  |  |  | The user facility reported an 80-year-old patient needing mechanical circulatory support. during insertion, it was reported that the distal flow was compromised after leg angio was performed. so, an antegrade 5 for sheath was placed on impella side then connected to impella sheath side-port for distal limb perfusion. the decision to remove impella due to distal limb ischemia. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
29547 | 1220648-2024-08265 | 18933887 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18933887&pc=OZD | 03/05/2024 | 04/22/2024 |  | 03/19/2024 |  |  | 03/05/2024 | 03/19/2024 |  |  | 09/30/2025 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025446038 |  | 1000008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Defective Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/05/2024 |  |  |  |  | The user facility reported a 61-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp was prepped and ready to be placed but the purge cassette failed. the impella insertion was aborted, and a new cp set was opened. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the purge cassette failure has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the priming issue was not determined since product was not returned.
29548 | 1220648-2024-08266 | 18933914 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18933914&pc=OZD | 12/17/2022 | 10/08/2025 |  | 03/19/2024 |  |  | 12/17/2022 | 03/19/2024 |  |  | 09/30/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023274413 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/17/2022 |  |  |  |  | The user facility reported an impella cp implanted in a (b)(6) female patient for mechanical circulatory support. it was reported the patient was on impella cp support and the patient coded on support. the patient went into runs of ventricular tachycardia and then converted into atrial fibrillation with rapid ventricular response. lidocaine and amino drips were started to treat rhythm issues. additionally, the patient experienced excessive bleeding at the area of repo sheath insertion ad was transfused with 1 unit of packed red blood cells. after discussion with physician and reviewing the stability of the patient the decision was made by the family to withdraw care and remove the impella. no allegations were made towards the impella for cause of death. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | Us complainant reported the patient was on impella cp support and the patient coded on support. the patient went into runs of ventricular tachycardia and then converted into atrial fibrillation with rapid ventricular response. lidocaine and amio drips were started to treat rhythm issues. additionally, the patient experienced excessive bleeding at the area of repo sheath insertion. after discussion with physician and reviewing the stability of the patient the decision was made by the family to withdraw care and remove the impella. patient was on impella support for 149.68 before patient expired. no allegations were made towards the impella for cause of death. reportable due to patient's excess bleeding at the insertion site. patients' family withdrew care, and it was noted care team was unable to wean patient off impella support. |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08266 was provided in sections b2, b5, d6, h1, and h6. | The root cause of the access site bleeding cannot be determined since the product was not returned and enough clinical info was not provided. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29549 | 1220648-2024-08267 | 18933992 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18933992&pc=OZD | 07/23/2022 | 05/13/2025 |  | 03/19/2024 |  |  | 07/23/2022 | 03/19/2024 |  |  | 02/29/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195017 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/23/2022 |  |  |  |  | The user facility reported a 66-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported during impella support, the patient experienced red urine output. an echocardiogram was performed, and the pump was repositioned. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29550 | 1220648-2024-08268 | 18934088 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18934088&pc=OZD | 08/31/2022 | 05/20/2025 |  | 03/19/2024 |  |  | 08/31/2022 | 03/19/2024 |  |  | 03/31/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023206879 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/31/2022 |  |  |  |  | The user facility reported a 76-year-old male was implanted with an impella cp device for mechanical circulatory support after admission to emergency for being found unresponsive at home. during implantation of the impella cp inotropes/pressors were escalated throughout procedure; and multiple shocks for ventricular fibrillation were performed. the patient required many treatments; however, deterioration continues. the family withdrew care and the patient expired shortly after. no allegations were made towards the impella for cause of death. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation for the reported ventricular fibrillation has been completed. the impella device was discarded by the customer. as the necessary clinical information was not provided and the device was not returned, the root cause of the ventricular fibrillation was not determined. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08268 was provided in sections b2, b5, d6, h1, and h6. | Analysis summary: "cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. ¿timing of the arrythmia event in relation to impella support (during or post-implant). ¿electrocardiogram (ekg) recordings of the arrhythmia episodes. ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. ¿response to any antiarrhythmic treatments or interventions during the event. ¿any documented device-related issues or complications during impella support. ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. ¿review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined."
29551 | 1220648-2024-08269 | 18934089 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18934089&pc=OZD | 08/26/2022 | 03/12/2025 |  | 03/19/2024 |  |  | 08/26/2022 | 03/19/2024 |  |  | 03/31/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2023213182 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/26/2022 |  |  |  |  | The user facility reported a 65-year-old female with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported the patient was transfused with blood products due to bleeding while on impella cp support. it was also reported the patient experienced limb ischemia in both legs. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
29552 | 1220648-2024-08270 | 18934090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18934090&pc=OZD | 03/06/2024 | 05/09/2024 |  | 03/19/2024 |  |  | 03/06/2024 | 03/19/2024 |  |  |  |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025450741 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  |  |  | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/06/2024 |  |  |  |  | The user facility reported a 62year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the impella cp was laid. activated clotting time was above 250 before removing the dilator. when the pump was started, it did not start, only 0.6 liters per minute. the patient's position and volume status were checked immediately. everything was within the normal range. the recommendation was therefore made to replace the pump with a new one. the new pump ran without any problems. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow issue has been completed since the original report was filed. the pump was returned for analysis. visual inspection found no issues on the pump. the failure mode could not be reproduced as pump ran without issue under bench test. based on clinical description, the root cause of low flow was most likely biomaterial ingestion as biomaterial observed during explant. in accordance with updated procedures, section d6 has been revised from the initial submission.
29553 | 1220648-2024-08271 | 18934257 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18934257&pc=OZD | 08/31/2022 | 03/07/2025 |  | 03/19/2024 |  |  | 08/31/2022 | 03/19/2024 |  |  | 05/31/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023231255 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/31/2022 |  |  |  |  | The user facility reported a 52-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the patient had very little pulsatility and was unable to get pulse distally in both limbs. the vascular was consulted for mottling of right foot. per vascular team, no further intervention and discoloration will improve with lowering vasopressin and levophed drips. additionally, there is a low suspicion for compartment syndrome given non doppler pulses in the right lower extremity status post impella cp. the right groin was oozing (bleeding) and a small hematoma and was compressed with gauze and multiple tegaderm/ under leg immobilizer before redressing was used. the patient was treated with 2 units of packed red blood cells (prbcs) and another 4 units of packed red blood cells (prbcs) the next day for additional treatment. heparin remained in purge solution. eventually the patient was escalated to impella 5.5 for native heart recovery. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
29554 | 1220648-2024-08272 | 18934265 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18934265&pc=OZD | 03/06/2024 | 04/24/2024 |  | 03/19/2024 |  |  | 03/06/2024 | 03/19/2024 |  |  | 11/30/2025 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024422916 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/06/2024 |  |  |  |  | The user facility reported a 78-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was on impella cp support until the patient could have transcatheter aortic valve replacement done. while on support, there was access site bleeding. forward tension sutures were applied, and angle matching was performed. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding issue was use related, as angle matching, and forward tension sutures were done as per the clinical description. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
29555 | 1220648-2024-08273 | 18934266 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18934266&pc=OZD | 07/24/2022 | 04/22/2025 |  | 03/19/2024 |  |  | 07/24/2022 | 03/19/2024 |  |  | 03/31/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204977 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 07/24/2022 |  |  |  |  | The user facility reported a 78-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had bleeding at the access site and also experienced ventricular tachycardia and sinus tachycardia elevated myocardial infraction. correct position of impella cp was verified via trans-thoracic echo. the impella cp was explanted. | Discontinued therapy after discussion with family, patient got visibly worse and just passed away. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. corrected information for the initial mfr 1220648-2024-08273 was provided in sections b2, b5, d6, h1, and h6. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29556 | 1220648-2024-08274 | 18934340 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18934340&pc=OZD | 07/25/2022 | 05/29/2025 |  | 03/19/2024 |  |  | 07/25/2022 | 03/19/2024 |  |  | 04/30/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022078802 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/25/2022 |  |  |  |  | The user facility reported a 74-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the patient's right leg was mottled post insertion and no flow was seen distal to common femoral artery (cfa). patient¿s leg was assessed and it was determined profunda was providing limited distal perfusion. decision to escalate to 5.5 was made. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the root cause of limb ischemia could not be determined as insufficient clinical details were provided. this report is being filed as part of a retrospective review of historical records.
29557 | 1220648-2024-08275 | 18934482 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18934482&pc=OZD | 10/11/2022 | 04/07/2025 |  | 03/19/2024 |  |  | 10/11/2022 | 03/19/2024 |  |  | 06/30/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023240866 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup |  |  | 10/11/2022 |  |  |  |  | The user facility reported a 68-year-old patient needing circulatory support. it was reported that the patient was placed on impella cp after going into cardiac arrest. the patient developed a major brain bleed and neuro was getting worse. the family decided to withdraw care due to no brain function. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08275 was provided.
29558 | 1220648-2024-08276 | 18934503 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18934503&pc=OZD | 08/31/2022 | 03/19/2024 |  | 03/19/2024 |  |  | 08/31/2022 | 03/19/2024 |  |  | 03/31/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023214178 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/31/2022 |  |  |  |  | The user facility reported a 70-year-old female was implanted with an impella cp device for mechanical circulatory support and the patient developed a hematoma at the access site. the patient was taken to the cardiac catheterization laboratory for a covered stent placement in the left femoral artery due to the patient requiring multiple transfusions and growing groin hematoma. additionally, upon repositioning of the impella cp, the device slipped out of the left ventricle and was unable to be pushed back across the atrioventricular node wirelessly. the impella cp was replaced with an impella rp. heparin was withdrawn, manual pressure was applied, and the patient received 7 units of packed red blood cells (prbc). | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29559 | 1220648-2024-08277 | 18934520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18934520&pc=OZD | 07/25/2022 | 03/12/2025 |  | 03/19/2024 |  |  | 07/25/2022 | 03/19/2024 |  |  | 12/31/2023 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023180154 |  | 0048-0071 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/25/2022 |  |  |  |  | The user facility reported a 58-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the medical doctor (md) accidentally pulled impella cp into the aorta. md used guide wire re-access port, removed impella cp, and inserted new sheath. they flushed and manually re-wired impella cp and re-inserted the pump. the patient developed hematoma in the right groin (impella side). impella was discontinued post percutaneous coronary intervention (pci) due to hematoma development. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29560 | 1220648-2024-08278 | 18934592 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18934592&pc=OZD | 07/25/2022 | 03/10/2025 |  | 03/19/2024 |  |  | 07/25/2022 | 03/19/2024 |  |  | 10/31/2023 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2022152396 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/25/2022 |  |  |  |  | The user facility reported a 51-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient encountered a traumatic foley event which caused the urine output to have red blood. the next day the urine was still red and the potassium results came back hemolyzed. it was recommended that the pump be repositioned. eventually the device was successfully weaned. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29561 | 1220648-2024-08279 | 18934688 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18934688&pc=OZD | 03/07/2024 | 04/08/2024 |  | 03/19/2024 |  |  |  | 03/19/2024 |  |  | 09/30/2025 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025444510 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/07/2024 |  |  |  |  | The user facility reported a 48-year-old female in st-segment elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced oozing and bleeding from the access site. the site had to be re-dressed several times, tension via perclose and angle matching reinforcement was done, and received three units of packed red blood cells as treatment. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation into the access site bleeding ¿ major issue has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was most likely due to use issue since hemostasis was achieved by angle matching reinforcement.
29562 | 1220648-2024-08280 | 18934796 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18934796&pc=OZD | 10/21/2022 | 03/19/2024 |  | 03/19/2024 |  |  |  | 03/19/2024 |  |  | 06/30/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023240387 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 10/21/2022 |  |  |  |  | The user facility reported a 58-year-old patient needing mechanical circulatory support. it was reported that the patient developed a clot in the guide catheter. more anticoagulation was given, post angio showed slow flow so the case was aborted. guide and wires were removed. patient remained stable and impella was weaned and removed without incident. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records.
29563 | 1220648-2024-08281 | 18934801 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18934801&pc=OZD | 09/01/2022 | 03/13/2025 |  | 03/19/2024 |  |  |  | 03/19/2024 |  |  | 04/30/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218300 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/01/2022 |  |  |  |  | The user facility reported a 70-year-old male was implanted with an impella cp device for mechanical circulatory support. the patient experienced limb ischemia in the right leg. the physician added nitro paste and warm blankets to right foot and continued to monitor. the impella cp was weaned and explanted and pulses could be heard with the doppler. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury" this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the root cause of limb ischemia could not be determined as insufficient clinical details were provided. this report is being filed as part of a retrospective review of historical records.
29564 | 1220648-2024-08282 | 18934821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18934821&pc=OZD | 03/08/2024 | 04/19/2024 |  | 03/19/2024 |  |  | 03/08/2024 | 03/19/2024 |  |  | 11/30/2025 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024431208 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/08/2024 |  |  |  |  | The user facility reported a 35-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the patient developed a hematoma at the impella insertion site. heparin was held and a washout was done. the patient was stable. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
29565 | 1220648-2024-08283 | 18934872 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18934872&pc=OZD | 07/26/2022 | 05/14/2025 |  | 03/19/2024 |  |  | 07/26/2022 | 03/19/2024 |  |  | 02/29/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023194832 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 07/26/2022 |  |  |  |  | The user facility reported a 67-year-old female was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding at the impella site. pressure was held and hemostasis was achieved after two hours. | After 2 attempts of cpr patient would not stabilize and eventually coded again. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08283 was provided in sections b2, b5,d6, h1, and h6. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29566 | 1220648-2024-08284 | 18934900 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18934900&pc=OZD | 07/27/2022 | 04/09/2025 |  | 03/19/2024 |  |  |  | 03/19/2024 |  |  | 03/31/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023211257 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/27/2022 |  |  |  |  | The user facility reported a 67-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient on impella cp had access site bleeding. the impella device was weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29567 | 1220648-2024-08285 | 18934908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18934908&pc=OZD | 10/15/2022 | 04/07/2025 |  | 03/19/2024 |  |  |  | 03/19/2024 |  |  | 04/30/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023217252 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup |  |  | 10/15/2022 |  |  |  |  | The user facility reported a 67-year-old patient needing circulatory support. it was reported that the patient was placed on impella cp after intra-aortic balloon pump removal. the patient developed a brain bleed and care was withdrawn. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08285 was provided in sections b2, b5,d6, h1, and h6.
29568 | 1220648-2024-08286 | 18934932 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18934932&pc=OZD | 03/08/2024 | 04/26/2024 |  | 03/19/2024 |  |  | 03/08/2024 | 03/19/2024 |  |  | 10/31/2025 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024415318 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Device Handling Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/08/2024 |  |  |  |  | The user facility reported a 73-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient on support was taken to the catheterization lab and impella cp was repositioned under fluoroscopy due to ¿impella in aorta¿ alarms. the patient had the 14 french peel-away sheath in. this was removed. the impella cp was positioned appropriately and a reposition sheath was inserted into the arteriotomy, and forward tension was sutured. the patient was in stable condition. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the cause of the positioning issue was use related issue due to improper technique as the pump pulled out. potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29569 | 1220648-2024-08287 | 18934939 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18934939&pc=OZD | 10/21/2022 | 03/19/2024 |  | 03/19/2024 |  |  | 10/21/2022 | 03/19/2024 |  |  | 06/30/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023237442 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 10/21/2022 |  |  |  |  | The user facility reported a 56-year-old patient needing mechanical circulatory support. it was reported that during impella insertion the right axillary artery was dissected while attempting to get vascular access. vascular surgery was performed to graft and repair the right axillary. the impella cp was then placed without issue into the left ventricle. decision to explant was made due to loss of pulse in the right leg. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
29570 | 1220648-2024-08288 | 18934964 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18934964&pc=OZD | 09/01/2022 | 05/16/2025 |  | 03/19/2024 |  |  | 09/01/2022 | 03/19/2024 |  |  | 06/30/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231306 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/01/2022 |  |  |  |  | The user facility reported a 79-year-old male was implanted with an impella cp device for mechanical circulatory support and that the transition from peel away to repo sheath, the pump pulled back into aorta and was unable to readvance. the dilator and peel away sheath was reinserted and the impella cp pump was removed and replaced without issue. additionally, the patient experienced a loss of pulses in the right lower extremity, but pulses were restored bilaterally. the patient was successfully weaned off the impella cp. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury." this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29571 | 1220648-2024-08289 | 18935041 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18935041&pc=OZD | 07/27/2022 | 03/19/2024 |  | 03/19/2024 |  |  | 07/27/2022 | 03/19/2024 |  |  | 04/30/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  |  |  | 2023219495 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/27/2022 |  |  |  |  | The user facility reported a 57-year-old patient needing mechanical circulatory support. the device was mispositioned so the physician withdrew the impella and reassessed the left ventricle. upon removal of the impella, a small clot was found on the inlet. the clot was flushed and cleared. the surgeon felt it was safe to reinsert the pump. angiomax bolus was administered/drip started and act was 238. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29572 | 1220648-2024-08300 | 18936238 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18936238&pc=OZD | 08/29/2022 | 03/19/2024 |  | 03/19/2024 |  |  |  | 03/19/2024 |  |  | 05/31/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2023228577 |  | 0048-0058 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/29/2022 |  |  |  |  | The user facility reported a 68-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient clenched their abdomen, resulting in the bleeding at the pump insertion site. manual pressure was held and an attempt at deploying a perclose in a post close fashion was unsuccessful. the device was removed while maintaining access to artery and eventually a 20 fr. a gore dry sheath was inserted and once hemostasis was obtained; the pump was re-inserted and started with appropriate waveforms. peelaway sheath was removed and gwrs advanced into dry seal. there was a slight increase in oozing (bleeding) at the patient's groin site the next day, due to repositioning of the device. eventually the pump was removed by a 2 perclose parallel wire technique which did not result in hemostasis. ultimately, all access points closed with good hemostasis. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. instructions for use for the related event are as follows: ¿potential adverse events: acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
29573 | 1220648-2024-08302 | 18936926 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18936926&pc=OZD | 08/29/2022 | 03/13/2025 |  | 03/19/2024 |  |  |  | 03/19/2024 |  |  | 01/31/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023190655 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/29/2022 |  |  |  |  | The user facility reported a 78-year-old female noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the patient experienced oozing at the groin. the groin area was treated by tightening the perclose, adding a suture, re-angle matching, and re-dressing with manual pressure. however, oozing at the groin persisted and hematoma was found. the hematoma was treated with manual pressure and the oozing was treated by giving the patient three units of blood. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29574 | 1220648-2024-08303 | 18936955 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18936955&pc=OZD | 07/17/2022 | 03/19/2024 |  | 03/19/2024 |  |  |  | 03/19/2024 |  |  | 03/31/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023210502 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/17/2022 |  |  |  |  | The user facility reported a 49-year-old female patient on an impella cp for mechanical circulatory support. while on support systemic heparin was withheld due to oozing at the access site that was controlled. the patient was weaned off of norepinephrine and vasopression and the impella was successfully implanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29575 | 1220648-2024-08304 | 18937195 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18937195&pc=OZD | 11/02/2022 | 05/21/2025 |  | 03/19/2024 |  |  |  | 03/19/2024 |  |  | 07/31/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | CP IMPELLA |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023249525 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/02/2022 |  |  |  |  | The user facility reported a 54-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was transferred for possible aortic dissection and pneumopericardium. the impella cp was weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29576 | 1220648-2024-08305 | 18937214 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18937214&pc=OZD | 08/30/2022 | 03/14/2025 |  | 03/19/2024 |  |  |  | 03/19/2024 |  |  | 03/31/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023209690 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/30/2022 |  |  |  |  | The user facility reported a 54-year-old male on impella cp for mechanical circulatory support. it was reported the patient experienced a hematoma at the impella cp access site. pressure was held to stop the site bleed and integrilin was turned off. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29577 | 1220648-2024-08306 | 18937273 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18937273&pc=OZD | 08/30/2022 | 03/14/2025 |  | 03/19/2024 |  |  |  | 03/19/2024 |  |  | 10/31/2023 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022152673 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/30/2022 |  |  |  |  | The user facility reported a 59-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient experienced a large amount of oozing at the access site following removal of the peel away sheath. the pump was temporarily removed and a 16 fr gore dry seal sheath was placed to manage the condition. the pump was then flushed and re-inserted without issue. however, following placement in the left ventricle, the pump was pulled back across the aortic valve during removal of the guidewire. due to the positioning issue, the decision was made to explant the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not deter.
29578 | 1220648-2024-08307 | 18937282 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18937282&pc=OZD | 11/02/2022 | 03/14/2025 |  | 03/19/2024 |  |  |  | 03/19/2024 |  |  | 11/30/2023 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023160176 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/02/2022 |  |  |  |  | The user facility reported a 62-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, it was noted a left ventricle thrombus showed on echocardiogram. the impella cp was weaned and explant. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. the cause of the thrombus was not determined since product was not returned.
29579 | 1220648-2024-08308 | 18937370 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18937370&pc=OZD | 08/30/2022 | 10/13/2025 |  | 03/19/2024 |  |  |  | 03/19/2024 |  |  | 04/03/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023226205 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/30/2022 |  |  |  |  | The user facility reported a 59-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. patient had no pulse in distal limb and the leg was mottled, due to increased sedation causing a pink foot. decision was made to leave the device within the patient and find a sedation level that controls patients withdrawals and maintains perfusion of distal limb. eventually the device was weaned and explanted. . | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. d6a and d6b revised from the initial manufacturer device report 1220648-2024-08308 in accordance with updated procedures f6/f8 should have been left blank on the initial manufacturer device report number 1220648-2024-08308 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-08308 in accordance with updated procedures. g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2024-08308. h6 component code revised from the initial submission in accordance with updated procedures.
29580 | 1220648-2024-08309 | 18937383 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18937383&pc=OZD | 12/04/2022 | 03/14/2025 |  | 03/19/2024 |  |  |  | 03/19/2024 |  |  | 06/30/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023232803 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/04/2022 |  |  |  |  | The user facility reported a 65-year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that while the patient was on impella 5.5 support, the patient experienced some bleeding from the groin insertion site, requiring packed red blood cells (prbcs). the staff switched the patient to bicarbonate and withheld the systemic heparin drip. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29581 | 1220648-2024-08310 | 18937424 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18937424&pc=OZD | 08/31/2022 | 03/14/2025 |  | 03/19/2024 |  |  |  | 03/19/2024 |  |  | 03/31/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023207363 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Positioning Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/31/2022 |  |  |  |  | The user facility reported a 60-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the patient was doing much better hemodynamically, but continued to have issues with position and some signs of hemolysis. the patient had stopped producing urine. the physician had opted to explant the impella (to manual pressure) instead of adjusting the position of the device. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29582 | 1220648-2024-08311 | 18937485 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18937485&pc=OZD | 08/31/2022 | 03/10/2025 |  | 03/19/2024 |  |  |  | 03/19/2024 |  |  | 07/31/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023231516 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/31/2022 |  |  |  |  | The user facility reported a 37-year-old male in cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. the patient developed severe hemolysis and was given 1 unit of blood. the physician made the decision to escalate the patient to impella 5.5. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29583 | 1220648-2024-08312 | 18937507 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18937507&pc=OZD | 07/18/2022 | 03/19/2024 |  | 03/19/2024 |  |  |  | 03/19/2024 |  |  | 04/30/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218254 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/18/2022 |  |  |  |  | The complainant reported the patient, a 58 year old male, was on impella cp support due to mitral valve replacement and aortic valve regurgitation. while on support the patient was not on any anticoagulation due to bleeding from the bilateral groin, which was resolved. it was also reported that the patient continued to experience atrial fibrillation (afib). when afib occurred, the patient had a drop in pressure. medication in an iv drip resolved the episodes of afib. the patient also experienced partial thromboplastin time. the patient was started on heparin. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records.
29584 | 1220648-2024-08313 | 18937533 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18937533&pc=OZD | 07/18/2022 | 03/12/2025 |  | 03/19/2024 |  |  |  | 03/19/2024 |  |  | 03/31/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2023213194 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Failure to Deliver | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/18/2022 |  |  |  |  | The user facility reported a 82-year-old female patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient was on impella cp support due to left main and circumflex coronary artery disease with ejection fraction 20%. during attempted delivery, staff was unable to pass the impella on initial access attempt on left side due to tortuosity and required going into the right groin. additionally, while on support, the patient experienced access site bleeding that required transfusion of one unit of red blood cells at the end of the case. patient transferred to intensive care unit for close monitoring. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29585 | 1220648-2024-08314 | 18937599 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18937599&pc=OZD | 11/13/2022 | 03/19/2024 |  | 03/19/2024 |  |  |  | 03/19/2024 |  |  | 03/31/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023210853 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/13/2022 |  |  |  |  | The user facility reported a 59-year-old male patient was implanted with an impella cp for mechanical circulatory support. placement of the impella was checked via echocardiogram indicating the pigtail was close to the mitral valve. the impella was repositioned. while on support, the patient experienced septic shock that was resolved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29586 | 1220648-2024-08315 | 18937621 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18937621&pc=OZD | 11/15/2022 | 05/21/2025 |  | 03/19/2024 |  |  |  | 03/19/2024 |  |  | 07/31/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023254831 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility |  |  |  | 65 YR |  | Male |  | Non Hispanic |  | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/15/2022 |  |  |  |  | The user facility reported a 65-year-old male patient on an impella cp for mechanical circulatory support. while on support, the patient was unable to get hemostasis with repositioning the sheath. a femostop was placed, which resulted in loss of pulses to the leg. the impella pump was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29587 | 1220648-2024-08316 | 18937648 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18937648&pc=OZD | 11/22/2022 | 04/07/2025 |  | 03/19/2024 |  |  |  | 03/19/2024 |  |  | 06/30/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023247287 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/22/2022 |  |  |  |  | The user facility reported a 82-year-old male patient was implanted with an impella cp for mechanical circulatory support due to a st-elevation myocardial infarction. when the impella was explanted a fibrin shell was noted on the pigtail. the groin was held for 45 minutes and femstop was applied. the impella was successfully explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29588 | 1220648-2024-08317 | 18937693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18937693&pc=OZD | 11/28/2022 | 03/12/2025 |  | 03/19/2024 |  |  |  | 03/19/2024 |  |  | 06/30/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023236516 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/28/2022 |  |  |  |  | The user facility reported a 78-year-old male on impella cp for mechanical circulatory support. it was reported that prior to the patient being placed was on impella cp, the patient presented with an inferior stemi. while on impella cp support, there was some oozing at the groin site. the dressing was changed and the oozing stopped. heparin was discontinued systemically due to platelets dropping. the patient was successfully weaned and explanted of the impella cp. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29589 | 1220648-2024-08318 | 18937717 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18937717&pc=OZD | 11/29/2022 | 10/20/2025 |  | 03/19/2024 |  |  |  | 03/19/2024 |  |  | 07/31/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023252974 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/29/2022 |  |  |  |  | The user facility reported a 66-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported that while on impella cp support the patient developed oozing and hematoma at the access site. the purge was switched to bicarb to help with partial thromboplastin time (ptt) levels and heparin was discontinued. patient continued to show improvement and the impella was explanted. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined. g1: updated manf. email address.
29590 | 1220648-2024-08319 | 18937742 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18937742&pc=OZD | 11/29/2022 | 03/19/2024 |  | 03/19/2024 |  |  |  | 03/19/2024 |  |  | 06/30/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023268416 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 11/29/2022 |  |  |  |  | The user facility reported a 68-year-old patient needing mechanical circulatory support. it was reported that after being placed on impella cp support, the patient experienced bleeding at the left groin site. the site was redressed with best practices. a figure 8 stitch was also applied and the oozing stopped. the purge fluid was changed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records.
29591 | 1220648-2024-08320 | 18937802 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18937802&pc=OZD | 12/05/2022 | 05/13/2025 |  | 03/19/2024 |  |  |  | 03/19/2024 |  |  | 08/31/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023269840 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/05/2022 |  |  |  |  | The user facility reported a 87-year-old female on impella cp for mechanical circulatory support. it was reported that during support, there was oozing at groin site, so site was redressed. the patient required fluids due to intermittent suctioning. hemoglobin was low so a unit of blood was given. the patient was stabilized and the impella cp was weaned and explanted. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29592 | 1220648-2024-08321 | 18937819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18937819&pc=OZD | 12/05/2022 | 03/12/2025 |  | 03/19/2024 |  |  |  | 03/19/2024 |  |  | 06/30/2024 |  | 03/19/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023248090 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/05/2022 |  |  |  |  | The user facility reported 56 year old female was placed on impella cp after coding. the impella was placed via the right femoral artery. external fem-fem bypass was in place with a contralateral stick. mottling began to take over the patient¿s lower extremities. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29593 | 1220648-2024-08123 | 18920053 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18920053&pc=OZD | 03/06/2024 | 04/24/2024 |  | 03/18/2024 |  |  | 03/06/2024 | 03/18/2024 |  |  |  |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025438725 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/06/2024 |  |  |  |  | The user facility reported a 72-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that a kink in the impella cp and the impella cp was removed and replaced. no patient harm was reported. | [5 empty] | The investigation into the product damage (cannula kink) issue has been completed since the original report was filed. the pump was returned, tested, and evaluated. visual inspection found no issues on the pump. blood was found inside the cannula & pump housing indicated that pump has been placed inside the patient body. the root cause of kinked cannula was most likely use related. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed.
29594 | 1220648-2024-08124 | 18920122 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18920122&pc=OZD | 03/06/2024 | 04/24/2024 |  | 03/18/2024 |  |  | 03/06/2024 | 03/18/2024 |  |  | 12/31/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024297512 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/06/2024 |  |  |  |  | The user facility reported a 60-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was on support due to having a left anterior descending artery occlusion. while on support, the pump advanced and the pump angled towards the mitral apparatus. the doctor attempted to torque out of the mitral/papillary multiple times which was unsuccessful. the pump was left shallow at three centimeters. the doctor was able to get the inlet away from the anterior leaflet but was still angled towards the mitral apparatus. upon impella cp removal, there was significant bleeding from the femoral artery. there was a femoral artery perforation that required vascular surgery to be called for assistance in repairing artery. one unit of packed red blood cells was provided. | [5 empty] | The investigation into the vascular damage - peripheral vessel and the positioning issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the vascular damage peripheral vessel issue was most likely use related as it occurred during pump cut down and removal. the root cause of the positioning issues was not determined since product was not returned and there is a lack of clinical detail. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
29595 | 1220648-2024-08125 | 18920154 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18920154&pc=OZD | 03/07/2024 | 04/18/2024 |  | 03/18/2024 |  |  | 03/07/2024 | 03/18/2024 |  |  | 08/31/2025 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024395694 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/07/2024 |  |  |  |  | The user facility reported an 81-year-old female patient of unknown ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed a hematoma and experienced oozing from the impella access site. manual pressure was applied to manage the condition and the two units packed red blood cells were transfused. angle matching dressing access was verified and support with the implanted pump was maintained. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding issue was not determined since product was not returned and not much information is provided in the clinical description.
29596 | 1220648-2024-08126 | 18920173 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18920173&pc=OZD | 03/06/2024 | 04/24/2024 |  | 03/18/2024 |  |  | 03/07/2024 | 03/18/2024 |  |  | 11/30/2025 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024424912 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Insufficient Flow or Under Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/07/2024 |  |  |  |  | The user facility reported a 50-year-old female in st-segment elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support, there were continuous suction alarms. when the impella was explanted, a thrombus was noted on the tip of the catheter. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow has been completed since the original report was submitted. the device was not returned for investigation. per the data log review, motor current spikes and suction alarms occurred. the cause of the low pump flow issue was most likely biomaterial ingestion as the clinical data confirmed thrombus seen after impella explanted.
29597 | 1220648-2024-08127 | 18920223 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18920223&pc=OZD | 03/07/2024 | 04/01/2024 |  | 03/18/2024 |  |  | 03/07/2024 | 03/18/2024 |  |  |  |  | 03/18/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1613085 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/07/2024 |  |  |  |  | The user facility biomed department reported that the automated impella controller (aic) had a broken purge clip and was removed from service. no patient involvement or harm was reported. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the initial report was submitted. the console was returned, visually inspected, and tested. upon examining aic for any visible defects, it was evident that the blue purge disc retainer was broken. the root cause of this issue was a broken purge disc retainer. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
29598 | 1220648-2024-08128 | 18920252 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18920252&pc=OZD | 03/07/2024 | 04/24/2024 |  | 03/18/2024 |  |  | 03/07/2024 | 03/18/2024 |  |  |  |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025443052 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/07/2024 |  |  |  |  | The user facility reported a 70-year-old female patient in acute myocardial infarction/cardiogenic shock was to be implanted with an impella 5.5 device for mechanical circulatory support. it was reported that resistance was encountered during attempted delivery of an impella 5.5 pump. it was noted that due to the patient smaller anatomy, the inserted 5.5 pump was unable to advance to the intended location. due to the noted complication, the impella 5.5 pump was removed and an impella cp pump was placed for patient support. there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was returned for evaluation and no abnormalities were found upon inspection. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely due to patient condition since the pump could not be delivered due to small tortuous anatomy.
29599 | 1220648-2024-08129 | 18920299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18920299&pc=OZD | 03/07/2024 | 04/24/2024 |  | 03/18/2024 |  |  | 03/07/2024 | 03/18/2024 |  |  | 11/30/2025 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024436333 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/07/2024 |  |  |  |  | The user facility reported an 83-year-old male patient noted as high-risk percutaneous coronary intervention (hrpci) was implanted with an impella cp device for mechanical circulatory support when the surgery team turned down the elderly patient with underlying sepsis and acute renal insufficiency with multivessel coronary artery disease. the pump was placed but the patients condition deteriorated, and a code was called. the team explanted the pump with concerns of pump thrombosis. the team did not replace the pump nor proceed with the percutaneous coronary intervention. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the thrombus issue has been completed since the original report was filed. the pump was returned, tested, and evaluated. no issue found with the pump. it was reported that patient developed bradycardia, lost pulsatility, and ended up in asystole. to determine the true root cause of the vt/vf, related clinical information would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined.
29600 | 1220648-2024-08132 | 18920452 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18920452&pc=OZD | 03/08/2024 | 06/05/2024 |  | 03/18/2024 |  |  | 03/08/2024 | 03/18/2024 |  |  |  |  | 03/18/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1725570 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/08/2024 |  |  |  |  | It was reported that a controller failure alarm occurred. the automated impella controller (aic) was exchanged. there was no patient involved. | [5 empty] | The investigation for the reported console (aic) controller failure alarm has been completed. the aic was returned for evaluation. the console was visually inspected and found to have no observable defect. the power supply from pbm to the pud was found to be nominal at about 5v. the failure was reproduced on boot up of the console in the lab. it was found that the alarm had no observable effect on the performance of the purge system when a known good pump and purge line were plugged into the console. the pud firmware was extracted and compared with production code. it was found to be corrupted. data logs were reviewed which confirm that the controller failure (pud sw corruption) alarm was issued on date of occurrence the root cause for the controller failure was a purge unit driver with corrupted firmware. in accordance with updated procedures, section d2 has been revised from the initial submission. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed.
29601 | 1220648-2024-08133 | 18920472 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18920472&pc=OZD | 07/27/2022 | 03/18/2025 |  | 03/18/2024 |  |  | 07/27/2022 | 03/18/2024 |  |  | 01/31/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023174135 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/27/2022 |  |  |  |  | The user facility reported a 82-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient developed thrombocytopenia and was later resolved when the impella cp device was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
29602 | 1220648-2024-08134 | 18920531 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18920531&pc=OZD | 02/03/2024 | 04/25/2024 |  | 03/18/2024 |  |  | 03/08/2024 | 03/18/2024 |  |  | 06/30/2025 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024431032 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/08/2024 |  |  |  |  | The user facility reported a 61-year-old male undergoing cardiac surgery was to be implanted with an impella cp device for mechanical circulatory support. it was reported that a 14 french sheath was unable to be delivered past a third of the right femoral. the team evaluated the contralateral leg and it too was calcified/tortuous. the team instead placed an intra-aortic balloon pump (iabp) for support. no harm or injury was reported from the failed delivery. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the difficulty inserting/removing introducer has been completed since the original report was submitted. the device was not returned for investigation. according to the clinical, before beginning impella cp support the 14fr short sheath could only be advanced 1/3 of the way due to heavy calcium and tortuosity. the most likely cause of the difficult insertion was the patient's condition.
29603 | 1220648-2024-08135 | 18920620 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18920620&pc=OZD | 07/28/2022 | 03/18/2025 |  | 03/18/2024 |  |  | 07/28/2022 | 03/18/2024 |  |  | 03/31/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204975 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/28/2022 |  |  |  |  | The user facility reported a 70-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support, patient experienced oozing at groin site. the patient's systemic heparin was discontinued to resolve bleeding. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29604 | 1220648-2024-08136 | 18920627 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18920627&pc=OZD | 03/08/2024 | 05/01/2024 |  | 03/18/2024 |  |  | 03/08/2024 | 03/18/2024 |  |  | 04/30/2025 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024339427 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Life Threatening |  |  |  |  | EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/08/2024 |  |  |  |  | The user facility reported an 80-year-old male undergoing coronary artery bypass grafting was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that while on extracorporeal membrane oxygenation (ecmo) and impella support, the patient had a stroke. heparin was paused after the stroke. it was noted that a massive clot was found in the ecmo circuit. the family withdrew care. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ d.6a and d.6b revised from the initial submission in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures.
29605 | 1220648-2024-08137 | 18920737 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18920737&pc=OZD | 07/30/2022 | 03/18/2025 |  | 03/18/2024 |  |  | 07/30/2022 | 03/18/2024 |  |  | 04/30/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023221088 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/30/2022 |  |  |  |  | The user facility reported that the 52-year-old male patient was on impella cp for hemodynamic support. while on support, the patient experienced severe hemolysis and the impella was explanted to resolve. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29606 | 1220648-2024-08138 | 18920807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18920807&pc=OZD | 03/08/2024 | 04/24/2024 |  | 03/18/2024 |  |  | 03/08/2024 | 03/18/2024 |  |  | 10/31/2025 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024421278 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/08/2024 |  |  |  |  | The user facility reported a 65-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that following impella cp implant, the patient lost distal pulse in their right leg. the loss in pulse coincided with symptoms of pain in the lower extremity. vascular surgery was performed to initially address the condition, but the decision was ultimately made to explant the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
29607 | 1220648-2024-08139 | 18921443 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18921443&pc=OZD | 08/02/2022 | 03/18/2025 |  | 03/18/2024 |  |  | 08/02/2022 | 03/18/2024 |  |  | 03/31/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023211468 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Difficult to Remove | Material Split, Cut or Torn | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/02/2022 |  |  |  |  | The user facility reported a 79-year-old female noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that during implantation of impella cp, the sheath was split and was exchanged out for cook 14f sheath it was reported that the impella cp device was difficult to remove. it was reported during removal the sheath was pinched due to a calcium blockage but was successfully removed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29608 | 1220648-2024-08140 | 18921733 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18921733&pc=OZD | 08/02/2022 | 03/18/2025 |  | 03/18/2024 |  |  | 08/02/2022 | 03/18/2024 |  |  | 03/31/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023210415 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/02/2022 |  |  |  |  | The user facility reported a 58-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. patient experienced oozing from the groin access site. the patient's heparin was decreased in the purge fluid bag, and dressings were reinforced. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29609 | 1220648-2024-08143 | 18922693 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18922693&pc=OZD | 08/06/2022 | 03/10/2025 |  | 03/18/2024 |  |  | 08/06/2022 | 03/18/2024 |  |  | 11/30/2023 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023165706 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/06/2022 |  |  |  |  | The user facility reported a 53-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported there was minimal to no blood flow to the right leg. impella was secured appropriately and redressed. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events: acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported limb ischemia ¿ reversible issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
29610 | 1220648-2024-08144 | 18922712 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18922712&pc=OZD | 10/28/2022 | 03/18/2025 |  | 03/18/2024 |  |  | 10/28/2022 | 03/18/2024 |  |  | 06/30/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023248346 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/28/2022 |  |  |  |  | The user facility reported a 78-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. vessel occlusion was noted and there was no flow distal to the impella sheath. after multiple attempts to place the antegrade sheath, vascular surgery was consulted and an antegrade sheath was obtained. external fem-fem bypass was put in place with distal perfusion noted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
29611 | 1220648-2024-08145 | 18922723 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18922723&pc=OZD | 08/03/2022 | 03/11/2025 |  | 03/18/2024 |  |  | 08/03/2022 | 03/18/2024 |  |  | 02/29/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023198455 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/03/2022 |  |  |  |  | The user facility reported a 87-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that when the impella cp was removed, the access site bled and a mynx closure device and manual pressure was used to resolve. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29612 | 1220648-2024-08147 | 18922865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18922865&pc=OZD | 10/29/2022 | 03/11/2025 |  | 03/18/2024 |  |  | 10/29/2022 | 03/18/2024 |  |  | 08/31/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023260069 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/29/2022 |  |  |  |  | The user facility reported a 39-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient was placed on impella cp to unload the left ventricle after labs indicated coagulopathy when the patient was found unresponsive. while on the impella, a head ct was performed, and the patient tested positive for multiple hemorrhagic processes. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29613 | 1220648-2024-08148 | 18922944 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18922944&pc=OZD | 10/30/2022 | 03/11/2025 |  | 03/18/2024 |  |  | 10/30/2022 | 03/18/2024 |  |  | 03/31/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023209879 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/30/2022 |  |  |  |  | The user facility reported a 69-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient had merlot to cabernet colored urine. it was then reported that the patient was still having gross hemolysis even after the device was repositioned. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29614 | 1220648-2024-08149 | 18923014 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18923014&pc=OZD | 11/01/2022 | 03/11/2025 |  | 03/18/2024 |  |  | 11/01/2022 | 03/18/2024 |  |  | 03/31/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023210522 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/01/2022 |  |  |  |  | The user facility reported a 62-year-old male in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. impella was measuring shallow, so was advanced under echocardiogram. during support, computer tomography angiography (cta) found clots in multiple branches of the retropharyngeal abscess (rpa). the clots in the lungs were removed in the catheterization laboratory. the patient also experienced impella site bleeding steadily. the perclose were tightened and bleeding slowed down with pressure dressing. however, every time the patient was turned it resulted in significant bleeding from the site. during ecmo it was discovered the patient had large femoral deep vein thrombosis (dvt) and pulmonary embolisms (pe). two units of blood were auto transfused back to the patient. reportedly on (b)(6) 2022 the impella cp also appeared to have pigtail wrapped around papillary muscle. multiple attempts to release pigtail and reposition were made. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29615 | 1220648-2024-08150 | 18923081 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18923081&pc=OZD | 08/07/2022 | 03/10/2025 |  | 03/18/2024 |  |  | 08/07/2022 | 03/18/2024 |  |  | 04/30/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023225128 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/07/2022 |  |  |  |  | The user facility reported a 52-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that a foley was placed causing the urine to become cherry red and there was a discussion about the hemolysis. heparin was stopped and there was a decreased in the rpms as long as the mean arterial pressure (map) was maintained. the patient gained pulsitility and made urine which was currently greater than 30 cc/hr & yellow. however, the impella cp was explanted and the impella 5.5 was implanted into the patient. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events : acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported hemolysis and access site bleeding issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
29616 | 1220648-2024-08151 | 18923104 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18923104&pc=OZD | 08/04/2022 | 03/28/2025 |  | 03/18/2024 |  |  | 08/03/2022 | 03/18/2024 |  |  | 12/31/2023 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023173608 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/03/2022 |  |  |  |  | The user facility reported a 67-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that while the patient was on impella cp support, the patient experienced a hematoma at the impella site. systemic heparin was discontinued as a result of the hematoma. | The user facility reported a 67-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that while the patient was on impella cp support, the patient was noted to have been febrile and was treated with antibiotics. the physician noted that the likely source of infection was pneumonia, however no known source of infection was provided. the patient remained on impella support until a conversion from femoral artery to axillary artery was performed during a planned extracorporeal membrane oxygenation decannulation. the patient was successfully supported on the new impella cp. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. section h6 component code 925 was added. the impella cp device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records. the following has been corrected from the initial mfr 1220648-2024-08151: section b3: date of event has been corrected. section b5: description of event has been corrected. section g1 reporting contact email has been updated. section h6: health effect - clinical code corrected to 2067; health effect - impact code corrected to 4644; investigation findings code corrected to 3221 only; investigation conclusions code corrected to 4315.
29617 | 1220648-2024-08152 | 18923219 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18923219&pc=OZD | 08/08/2022 | 03/18/2024 |  | 03/18/2024 |  |  | 08/08/2022 | 03/18/2024 |  |  | 01/31/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023189058 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/08/2022 |  |  |  |  | The user facility reported a 44-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported that the pump had leakage from the impella purge connection and then the pump began to experience pump saturation. the pump was replaced. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29618 | 1220648-2024-08153 | 18923333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18923333&pc=OZD | 02/17/2024 | 04/16/2024 |  | 03/18/2024 |  |  | 02/17/2024 | 03/18/2024 |  |  | 10/31/2025 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024420452 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/17/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 as a bridge to transplant. while on support, there was a catheter fracture at the red impella plug and catheter junction. the doctor stated that the pump appeared to be damaged by the patient. the pump was replaced. | Additional information was received. the engineer reported that there was no issue with the pump serial number (b)(6). the issue was with pump serial number (b)(6). the representative confirmed that there was no issue with pump serial number (b)(6). |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | B.5 report number 1220648-2024-08153 is not reportable as there was no issue with this pump. pump serial number (b)(6) was the issue and all information for this pump has been reported in report number 1220648-2024-07120.
29619 | 1220648-2024-08154 | 18923372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18923372&pc=OZD | 08/03/2022 | 03/11/2025 |  | 03/18/2024 |  |  | 08/03/2022 | 03/18/2024 |  |  | 01/31/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023188089 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/03/2022 |  |  |  |  | The user facility reported 81-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient on impella cp had hematoma. the peel away sheath was removed and repositioning sheath was inserted and secured with forward tension sutures, and perclose was deployed successfully. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | Correction to the initial mfr 1220648-2024-08154 from a reportable serious injury to a non-reportable complaint. upon review of the complaint, there was no treatment provided to the patient beyond standard practice of pressure held at the site and a change of dressing. therefore. this event should not have been reported.
29620 | 1220648-2024-08155 | 18923456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18923456&pc=OZD | 08/08/2022 | 03/18/2024 |  | 03/18/2024 |  |  | 08/08/2022 | 03/18/2024 |  |  | 04/30/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023225806 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/08/2022 |  |  |  |  | The user facility reported a 78-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported before implanting impella cp, the pump had a kink, therefore another pump was used for patient support. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29621 | 1220648-2024-08156 | 18923524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18923524&pc=OZD | 08/09/2022 | 03/07/2025 |  | 03/18/2024 |  |  | 08/09/2022 | 03/18/2024 |  |  | 03/31/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023209572 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 23 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/09/2022 |  |  |  |  | The user facility reported a 23-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the ct of the patient's head showed there was a large watershed infarct. subsequently, the pump was explanted. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | Strokes/cvas can occur during or after impella support, and can be either ischemic or hemorrhagic. to make a determination as to the cause of the stroke, the following clinical information must be considered: -patient's medical history, including any previous strokes, cardiovascular disease, or risk factors such as hypertension, diabetes, or atrial fibrillation -timing of the stroke in relation to impella support (during or post-implant) -detailed description of the symptoms experienced by the patient -neurological examination findings and any focal deficits observed -diagnostic imaging results, such as ct scan or mri of the brain, to identify the type and location of the stroke (ischemic or hemorrhagic) -laboratory test results, including coagulation profiles and cardiac markers -any relevant procedural or device-related factors, such as duration and settings of impella support, presence of embolic protection devices, or any documented procedural complications -response to any previous anticoagulation or antiplatelet therapies -evaluation of potential alternative causes of stroke, such as arrhythmias, embolic sources, or underlying vascular pathology as this clinical information was not provided, the true root cause of the stroke/cva could not be determined.
29622 | 1220648-2024-08158 | 18923618 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18923618&pc=OZD | 11/11/2022 | 06/09/2025 |  | 03/18/2024 |  |  | 11/11/2022 | 03/18/2024 |  |  | 07/31/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023251484 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Break | Failure to Advance | [5 empty] | Tachycardia | Hemodynamic instability | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/11/2022 |  |  |  |  | The user facility reported a 82-year-old male patient needing an impella cp for mechanical circulatory support. while attempting to place the impella, the left ventricle (lv) wire access was lost, and the wire was removed. while crossing the lesion, the stent came off the delivery system. while attempting to retrieve the stent, another interventional balloon broke from the shaft. advanced cardiac life support protocol was initiated. the impella was unable to be crossed into the lv and efforts were stopped. | The complainant reported the insertion of impella cp device as a bailout was unsuccessful. the patient presented with non-st-segment elevation myocardial infarction. angiogram showed a "tight" left anterior descending artery lesion. percutaneous coronary intervention (pci) was initiated. while crossing the lesion, the stent came off the delivery system. during an attempt to retrieve the stent with a balloon catheter, its catheter shaft "broke," and the patient began coding. the decision was made at this point to urgently place an impella cp device for mechanical circulatory support (mcs). while attempting to place the impella cp device during cardiopulmonary resuscitation (cpr), the left ventricle guidewire was lost. the guidewire was removed, and physician attempted to cross wirelessly as a last effort. the patient was under cpr between 35 to 40 minutes; the impella cp device was unable to cross into the left ventricle, efforts were discontinued, and the patient would subsequently expire from the cardiac arrest event that started during pci stent dislodgement retrieval attempt. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | Medical safety officer had reviewed the event and determined a serious injury type of reportable event; the use of the device cannot conclusively be associated with the outcome (patient's demise). this report was filed as part of a retrospective review of historical records. during review of the complaint file it was noted a follow-up report was required to address missed and erroneous (unwittingly provided) information. therefore, updates are captured the respective sections of this supplemental medwatch report accordingly: section b1: adverse event and/or product problem section b2: outcomes section b5: event description section h1: type of reportable event section h6: health effect/medical device/component codes section h10: additional manufacturer narrative note: investigation outcome statement(s) and its h6 coding (type, findings and conclusions) remains unchanged as previously reported under follow-up 1 report.
29623 | 1220648-2024-08157 | 18923625 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18923625&pc=OZD | 02/24/2024 | 04/11/2024 |  | 03/18/2024 |  |  | 02/24/2024 | 03/18/2024 |  |  | 11/30/2025 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024435348 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/24/2024 |  |  |  |  | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during impella cp support, the patient was experiencing hemolysis. pro-teck setting was increased overnight and impella flow as well, but no signs of improvement. pressors were unable to be weaned. albumin and blood products were given. the impella was removed. | [5 empty] | The investigation for the reported hemolysis has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hemolysis could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
29624 | 1220648-2024-08159 | 18923666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18923666&pc=OZD | 08/09/2022 | 03/10/2025 |  | 03/18/2024 |  |  | 08/09/2022 | 03/18/2024 |  |  | 04/30/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023220566 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/09/2022 |  |  |  |  | The user facility reported a 70-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient was taken to vascular surgery for pulseless right upper extremity to evacuate brachial hematoma. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events: acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported limb ischemia ¿ reversible issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
29625 | 1220648-2024-08160 | 18923682 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18923682&pc=OZD | 11/11/2022 | 03/11/2025 |  | 03/18/2024 |  |  | 11/11/2022 | 03/18/2024 |  |  | 08/31/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023265399 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/11/2022 |  |  |  |  | The user facility reported a 66-year-old female was implanted with an impella cp and the patient developed a large expanding hematoma on right groin of the impella cp access site without any bleeding. the dressing was dry and intact. manual pressure was applied. fem stop was applied above the insertion site. the patient was given two units packed red blood cells (prbc). | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29626 | 1220648-2024-08161 | 18923697 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18923697&pc=OZD | 08/09/2022 | 03/11/2025 |  | 03/18/2024 |  |  | 08/09/2022 | 03/18/2024 |  |  | 04/30/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218596 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/09/2022 |  |  |  |  | The user facility reported a 73-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. a large thrombus was noted in the left ventricle (lv). subsequently, impella was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29627 | 1220648-2024-08162 | 18923719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18923719&pc=OZD | 11/11/2022 | 03/18/2024 |  | 03/18/2024 |  |  | 11/11/2022 | 03/18/2024 |  |  | 07/31/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023257311 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/11/2022 |  |  |  |  | The user facility reported a 74-year-old male patient was implanted with an impella cp for mechanical circulatory support. the patient developed slight bleeding at the impella cp insertion site. the patient had bicarbonate in the purge and the bleeding was being actively managed by the bedside nurse. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29628 | 1220648-2024-08163 | 18923754 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18923754&pc=OZD | 02/24/2024 | 04/12/2024 |  | 03/18/2024 |  |  | 02/24/2024 | 03/18/2024 |  |  |  |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024436410 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Separation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/24/2024 |  |  |  |  | The user facility reported a 78-year-old female undergoing coronary artery bypass grafting was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the sterile sleeve came apart from the tuoghy. the doctor was aware and re-secured with tape after cleaning. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the product damage (sterile sleeve damage) issue has been completed since the original report was submitted. the device was not returned for investigation. the cause of the sterile sleeve product damage could not be determined because no product was returned, and limited clinical information was provided.
29629 | 1220648-2024-08164 | 18923773 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18923773&pc=OZD | 11/29/2022 | 03/18/2024 |  | 03/18/2024 |  |  | 11/29/2022 | 03/18/2024 |  |  | 06/30/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023242212 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Difficult to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/29/2022 |  |  |  |  | The user facility reported an 67-year-old female was implanted with an impella cp. the complainant reported that the impella cp implantation was aborted due to the patient¿s anatomy. the device was unable to pass through femoral artery. | [5 empty] | The investigation for the reported delivery issues has been completed. however, clinical information was sufficient to determine the root cause. the impella device was not returned for evaluation. the root cause of the delivery issue was due to patient condition. this report is being filed as part of a retrospective review of historical records.
29630 | 1220648-2024-08166 | 18923777 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18923777&pc=OZD | 08/30/2022 | 03/11/2025 |  | 03/18/2024 |  |  | 08/30/2022 | 03/18/2024 |  |  | 05/31/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023227376 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Activation, Positioning or Separation Problem | Improper Flow or Infusion |  |  |  |  | Arrhythmia | Hemolysis | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/30/2022 |  |  |  |  | The user facility reported a 62-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced dark red urine, episodes of ventricular tachycardia (vt), and loss of temperature and color in the impella leg. the medical team reported suction alarms and reduced left ventricle diastolic waveform. echocardiography showed the impella cp device was not in the proper position, and physician was advised to reposition the device to resolve vt, hemolysis, and suction issues. before physician could perform a repositioning of the device, the patient moved and the impella repositioned to a more optimal location in the left ventricle. the physician declined to reposition the device further because of the vt issues the patient had been experiencing. the patient was escalated to venoarterial extracorporeal membrane oxygenation (va ecmo). | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29631 | 1220648-2024-08165 | 18923778 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18923778&pc=OZD | 08/11/2022 | 03/11/2025 |  | 03/18/2024 |  |  | 08/11/2022 | 03/18/2024 |  |  | 04/30/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023226050 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/11/2022 |  |  |  |  | The user facility reported a 67 year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced positioning alarms, the impella cp was placed too deep. hematoma was noted at groin site and manual pressure was held. bleeding at insertion site had also increased. perclose sutures were tightened and bleeding stopped. impella was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29632 | 1220648-2024-08167 | 18923838 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18923838&pc=OZD | 02/28/2024 | 04/09/2024 |  | 03/18/2024 |  |  | 02/28/2024 | 03/18/2024 |  |  | 10/31/2025 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024415385 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/28/2024 |  |  |  |  | The user facility reported a 63-year-old male noted as high-risk percutaneous coronary intervention was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that while on support, the impella 5.5 was pulled into the aorta when the cardiac surgeon was attempting to reposition. the pump was urgently removed. the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the positioning issue could not be determined.
29633 | 1220648-2024-08168 | 18923895 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18923895&pc=OZD | 08/31/2022 | 03/11/2025 |  | 03/18/2024 |  |  | 08/31/2022 | 03/18/2024 |  |  | 04/30/2022 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023216099 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/31/2022 |  |  |  |  | The user facility reported a 75-year-old female was implanted with an impella cp device for mechanical circulatory support. the patient developed hemolysis and thrombocytopenia during impella cp pump support. treatment information is unknown. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29634 | 1220648-2024-08169 | 18923906 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18923906&pc=OZD | 07/23/2022 | 03/11/2025 |  | 03/18/2024 |  |  | 07/23/2022 | 03/18/2024 |  |  | 02/29/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023198866 |  | 0048-0069 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/23/2022 |  |  |  |  | The user facility reported a 72-year-old female with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support patient experienced access site bleeding and hemolysis. the patient's dressing was immediately removed to assess site with active bleed noted. angle of entry elevated with immediate cessation of bleeding. site was redressed with angle match dressing. no further oozing noted with soft groin. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: hemolysis: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ access site bleeding: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29635 | 1220648-2024-08170 | 18923947 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18923947&pc=OZD | 03/08/2024 | 04/24/2024 |  | 03/18/2024 |  |  | 03/08/2024 | 03/18/2024 |  |  |  |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025443332 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/08/2024 |  |  |  |  | The user facility reported a 47-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the patient was on impella cp support due to the patient becoming tachypneic and hypotensive requiring intubation. while on support, the staff was planning on escalating the patient to an impella 5.5. however, a left ventricle thrombus was present. the impella cp was weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombus could not be determined.
29636 | 1220648-2024-08171 | 18923967 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18923967&pc=OZD | 07/12/2022 | 03/11/2025 |  | 03/18/2024 |  |  | 07/12/2022 | 03/18/2024 |  |  | 04/30/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023216597 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/12/2022 |  |  |  |  | The user facility reported a 79-year-old male patient noted as high risk percutaneous coronary intervention, needed an impella cp for mechanical circulatory support. during implantation of the impella, the physician was unable to cross the aortic valve and the procedure was aborted. the left formal artery was manually closed. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29637 | 1220648-2024-08172 | 18923975 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18923975&pc=OZD | 07/21/2022 | 03/11/2025 |  | 03/18/2024 |  |  | 07/21/2022 | 03/18/2024 |  |  | 03/31/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023211689 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/21/2022 |  |  |  |  | The user facility reported a 77-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient experienced persistent groin bleed and reported that a hematoma developed following impella implant. in attempt to manage their condition, heparin was held, and one unit of blood was given to treat the bleed and manage condition. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29638 | 1220648-2024-08173 | 18924089 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18924089&pc=OZD | 07/18/2022 | 03/11/2025 |  | 03/18/2024 |  |  | 07/18/2022 | 03/18/2024 |  |  | 03/31/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023203569 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/18/2022 |  |  |  |  | The user facility reported a 50-year-old male implanted with an impella cp device for mechanical circulatory support due to having chest pain. while on support the patient experienced stenosis in the left anterior descending coronary artery branch, the patient required a thrombectomy, drug eluding stent x2 to the left anterior descending coronary artery branch. the patient also experienced ventricular tachycardia while on support, and was shocked twice. amiodarone bolus was ordered via iv drip. the impella cp was successfully weaned and explanted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29639 | 1220648-2024-08174 | 18924125 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18924125&pc=OZD | 03/09/2024 | 04/30/2024 |  | 03/18/2024 |  |  | 03/09/2024 | 03/18/2024 |  |  | 10/31/2025 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024426521 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/09/2024 |  |  |  |  | The user facility reported a 74-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced persistent subcutaneous bleeding during support. as a result of the condition, the pump was explanted and a 16 french medikit sheath was placed from the ipsilateral side to stop the bleeding. an intra-aortic balloon pump was then placed for ongoing support. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. clinical details states that the bleeding was resolved by holding heparin. therefore, the root cause of the access site bleeding was most likely due to anticoagulation management since the act at the time of bleed was 200 which is above the range limit. in accordance with updated procedures, sections d6, g3, and h10 have been revised from the initial submission.
29640 | 1220648-2024-08175 | 18924280 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18924280&pc=OZD | 07/19/2022 | 03/18/2024 |  | 03/18/2024 |  |  | 07/19/2022 | 03/18/2024 |  |  | 03/31/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023209293 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/19/2022 |  |  |  |  | The user facility reported a male (age unknown) was implanted with an impella device for mechanical circulatory support. the patient had a hematoma noted at the insertion site of the right femoral. as right groin hematoma increased in size, there was difficulty with return to right femoral extracorporeal membrane oxygenation (ecmo). a foley catheter was placed to decompress bladder and decrease pressure on right groin hematoma. the patient had 2 units of blood and mattress suture in place. patient also had thrombocytopenia. the patient remained on impella cp support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29641 | 1220648-2024-08176 | 18924286 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18924286&pc=OZD | 03/04/2024 | 04/22/2024 |  | 03/18/2024 |  |  | 03/04/2024 | 03/18/2024 |  |  | 09/30/2025 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025444233 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Inaccurate Delivery | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that resistance was encountered during attempted delivery of an impella cp pump. both groins were accessed while trying to implant the pump, but the catheter failed to cross via either access point. it was noted that the patient's tortuosity coinciding with calcification and stenosis contributed to the complication. as a result, the decision was made to abort the implant. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issues is determined to be patient condition because of the tortuous anatomy of the patient and was unable to pass through even when attempting from both sides and resistance was met at bifurcation even from both the sides.
29642 | 1220648-2024-08177 | 18924289 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18924289&pc=OZD | 06/09/2022 | 03/11/2025 |  | 03/18/2024 |  |  | 06/09/2022 | 03/18/2024 |  |  | 02/29/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023194829 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 71 YR |  | Male |  |  |  | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/09/2022 |  |  |  |  | The user facility reported a 71-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during insertion, the physician lost a wire in the left ventricle. the pump was removed and a new pump was placed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29643 | 1220648-2024-08178 | 18924301 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18924301&pc=OZD | 08/04/2022 | 03/18/2024 |  | 03/18/2024 |  |  | 08/04/2022 | 03/18/2024 |  |  | 03/31/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023207133 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/04/2022 |  |  |  |  | The user facility reported a 77-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced oozing at the groin access site. the bleeding was treated by placing forward facing sutures, an angle matching dressing, and pressure was held at the access site. after there was no improvement to the bleeding, the dressing was removed and sutures were replaced with larger bites of forward facing sutures. the bleeding than subsided. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records.
29644 | 1220648-2024-08179 | 18924387 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18924387&pc=OZD | 07/21/2022 | 03/11/2025 |  | 03/18/2024 |  |  | 07/21/2022 | 03/18/2024 |  |  | 03/31/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023209616 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/21/2022 |  |  |  |  | The user facility reported a 69-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, perclose failed, attempted 8fr sheath, and angioseal, but there was too much bleeding around the sheath. the physician decided to keep 14fr in and pull when activated clotting factor went below 160. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29645 | 1220648-2024-08180 | 18924433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18924433&pc=OZD | 07/21/2022 | 03/11/2025 |  | 03/18/2024 |  |  | 07/21/2022 | 03/18/2024 |  |  | 03/31/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023206460 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/21/2022 |  |  |  |  | The user facility reported a 54 year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding that was resolved with transfusion of blood products. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29646 | 1220648-2024-08181 | 18924467 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18924467&pc=OZD | 07/19/2022 | 03/11/2025 |  | 03/18/2024 |  |  | 07/19/2022 | 03/18/2024 |  |  | 03/31/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023207339 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/19/2022 |  |  |  |  | The user facility reported a 48-year-old male was implanted with an impella cp device for mechanical circulatory support. the patient developed a cold leg and with no dopplerable pluses and complained of pain and numbness in the impella leg. the pump was explanted, and the leg was reportedly warmer after the removal. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29647 | 1220648-2024-08182 | 18924482 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18924482&pc=OZD | 03/05/2024 | 05/10/2024 |  | 03/18/2024 |  |  | 03/05/2024 | 03/18/2024 |  |  | 09/30/2025 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024404568 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a male patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. a long priming process was observed but after a few seconds, liquid was exiting the outflow area. after the pump was implanted and an attempt to turn the device on, impella black alarm appeared. despite attempts to reset were not possible to restart the support. the pump stopped directly after starting. a new pump was implanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump did not start issue has been completed since the original report was filed. the pump, purge cassette and data logs were returned and investigated. there were no physical abnormalities seen on the pump and cassette. no biomaterial residue was seen in the outflow cage or around the impellers. the field data logs were review revealed that the pump was activated prior to being inserted into the patient. the root cause of the pump did not start issue was use related to activating the pump prior to inserting into the patient.
29648 | 1220648-2024-08183 | 18924507 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18924507&pc=OZD | 08/13/2022 | 03/18/2024 |  | 03/18/2024 |  |  | 08/13/2022 | 03/18/2024 |  |  | 03/31/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023207855 |  | 0048-0045 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/13/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 59-year-old male patient for mechanical circulatory support. the impella was placed via the patient¿s right femoral artery, it was noted that the impella cp was under the papillary muscle and measuring deep within. the physician declined repositioning, the physician elected to wean and explant the impella device. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29649 | 1220648-2024-08184 | 18924519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18924519&pc=OZD | 07/21/2022 | 03/11/2025 |  | 03/18/2024 |  |  | 07/21/2022 | 03/18/2024 |  |  | 04/30/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023221029 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/21/2022 |  |  |  |  | The user facility reported an 80-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient developed a small hematoma at the impella cp site. the hematoma was expressed in the catheterization laboratory and femostop was placed. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29650 | 1220648-2024-08185 | 18924587 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18924587&pc=OZD | 08/04/2022 | 03/11/2025 |  | 03/18/2024 |  |  | 08/04/2022 | 03/18/2024 |  |  | 10/31/2022 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2021004430 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/04/2022 |  |  |  |  | The user facility reported a 53-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced bleeding at the impella access site and lower limb ischemia. the patient was transfused with red blood cells, fresh frozen plasma, and platelets. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29651 | 1220648-2024-08187 | 18924689 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18924689&pc=OZD | 08/19/2022 | 03/10/2025 |  | 03/18/2024 |  |  | 08/19/2022 | 03/18/2024 |  |  | 05/31/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023228721 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/19/2022 |  |  |  |  | The user facility reported a 73-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported patient had some bleeding from peel away that was left in place and leaking with sheath was more noticeable when flushing sheath. staff was able to ween patient successfully off of impella support. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. the instructions for use states ¿potential adverse events: acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleeding issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
29652 | 1220648-2024-08189 | 18924755 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18924755&pc=OZD | 07/19/2022 | 03/11/2025 |  | 03/18/2024 |  |  | 07/19/2022 | 03/18/2024 |  |  | 03/31/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023209259 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/19/2022 |  |  |  |  | The user facility reported a 70-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. a left heart catheterization was performed, due to lesion morphology, lesion was located in ostial left anterior descending, and large size of left main. the patient was transported to the intensive care unit in critical condition. a computer tomography scan of the head demonstrated a massive cerebral bleed. after discussion with the family, the decision to withdraw care was made. the patient expired, no allegations were made towards the impella for cause of death. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29653 | 1220648-2024-08188 | 18924756 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18924756&pc=OZD | 08/04/2022 | 03/11/2025 |  | 03/18/2024 |  |  | 08/04/2022 | 03/18/2024 |  |  | 12/31/2023 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023170786 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/04/2022 |  |  |  |  | The user facility reported a 76-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced bleeding at the impella access site. the positioning sheath was exchanged for a peel away and required additional advancement of the blue hub inserted into puncture site and forward sutures for the arteriotomy to stop bleeding. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29654 | 1220648-2024-08190 | 18924758 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18924758&pc=OZD | 07/12/2022 | 03/11/2025 |  | 03/18/2024 |  |  | 07/12/2022 | 03/18/2024 |  |  | 02/29/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202860 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/12/2022 |  |  |  |  | The user facility reported a 46-year-old female patient on an impella cp for mechanical circulatory support. while on support, the patient developed a small hematoma at the insertion site, however, it was noted as being stable. additionally, the patient¿s right leg (impella leg) was unable to palpate pulses. the patient¿s bilateral lower extremities demonstrated mottled coloring and were cool. the patient was escalated to the impella 5.5. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29655 | 1220648-2024-08191 | 18924770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18924770&pc=OZD | 08/20/2022 | 03/18/2024 |  | 03/18/2024 |  |  | 08/20/2022 | 03/18/2024 |  |  | 04/30/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023226199 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/20/2022 |  |  |  |  | The user facility reported a 52-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported while being placed on impella cp support, the patient developed ischemic leg on impella side. additionally, the impella cp had a placement signal low alarm. the impella cp pump was explanted and replaced with an impella 5.5. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
29656 | 1220648-2024-08192 | 18924832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18924832&pc=OZD | 07/12/2022 | 03/11/2025 |  | 03/18/2024 |  |  | 07/12/2022 | 03/18/2024 |  |  | 12/31/2023 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023178921 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Failure to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/12/2022 |  |  |  |  | The user facility reported a 43-year-old female on impella cp for mechanical circulatory support. it was reported that the patient was placed on impella cp after the patient presented with chest pains when admitted, and went into vt post diagnostic cath. the impella cp was advanced deep into the left ventricle which resulted in reduced flows. repositioning helped, however, the pigtail appeared to not return previous flows. the physician decided to explant the impella cp. no patient harm reported. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29657 | 1220648-2024-08193 | 18924905 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18924905&pc=OZD | 07/12/2022 | 03/12/2025 |  | 03/18/2024 |  |  | 07/12/2022 | 03/18/2024 |  |  | 03/31/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023214222 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Difficult to Insert | Failure to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/12/2022 |  |  |  |  | The user facility reported a 71-year-old female on impella cp for mechanical circulatory support. it was reported that the patient was placed on impella cp after presenting with nausea, vomiting, and palpitations and a st elevated myocardial infarction diagnosis. the impella cp was explanted early due to the inability to reposition off the inferior posterior infarcted wall. no patient harm was reported. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29658 | 1220648-2024-08194 | 18924998 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18924998&pc=OZD | 07/19/2022 | 03/18/2024 |  | 03/18/2024 |  |  | 07/19/2022 | 03/18/2024 |  |  | 04/30/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023214856 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/19/2022 |  |  |  |  | The user facility reported a 49-year-old male was implanted with an impella cp device for mechanical circulatory support for high risk percutaneous coronary intervention for non- sinus tachycardia elevated myocardial infarction. the patient developed a small hematoma at the impella site and femostop was placed. the patient¿s hemoglobin dropped 7.7 from 9.4 and the patient received one (1) unit of packed red blood cells. hemodynamics recovered, and the impella was successfully explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
29659 | 1220648-2024-08195 | 18925576 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18925576&pc=OZD | 07/19/2022 | 05/21/2025 |  | 03/18/2024 |  |  | 07/19/2022 | 03/18/2024 |  |  | 02/29/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023201373 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/19/2022 |  |  |  |  | The user facility reported an 84-year-old male was being implanted with an impella cp device for mechanical circulatory support. the pump was inserted into the left ventricle and was started, but there was no flow and no motor current. attempts to restart were un successful and the decision was made to replace the pump. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29660 | 1220648-2024-08196 | 18926117 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18926117&pc=OZD | 07/19/2022 | 03/18/2024 |  | 03/18/2024 |  |  | 07/19/2022 | 03/18/2024 |  |  | 10/31/2023 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022154867 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/19/2022 |  |  |  |  | The user facility reported a 75-year-old male was implanted with an impella cp device for mechanical circulatory support. the patient developed bleeding and a large hematoma at the insertion site. the impella cp was removed, and the patient underwent vascular surgery to repair the vessel. an impella 5.5 was implanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29661 | 1220648-2024-08197 | 18926511 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18926511&pc=OZD | 07/19/2022 | 03/18/2024 |  | 03/18/2024 |  |  | 07/19/2022 | 03/18/2024 |  |  | 04/30/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023217713 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/19/2022 |  |  |  |  | The user facility reported a 36-year-old female was implanted with an impella cp device for mechanical circulatory support. the patient developed bleeding at the impella cp insertion site. impella site was still a little oozy with peel away in place and partial thromboplastin time 127.9 so systemic heparin was put on hold and the patient was given bicarbonate. heparin was later started and the impella cp device was removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29662 | 1220648-2024-08198 | 18926567 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18926567&pc=OZD | 07/21/2022 | 03/18/2024 |  | 03/18/2024 |  |  | 07/21/2022 | 03/18/2024 |  |  | 04/30/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218609 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/21/2022 |  |  |  |  | The user facility reported an impella cp implanted in a 56-year-old male patient for mechanical circulatory support. it was reported while on impella cp support, the patient was experienced a hematoma at the groin site. a standard bleeding protocol was followed to treat the bleeding. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29663 | 1220648-2024-08199 | 18926598 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18926598&pc=OZD | 07/13/2022 | 03/11/2025 |  | 03/18/2024 |  |  | 07/13/2022 | 03/18/2024 |  |  | 02/29/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023201416 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Difficult to Insert | Failure to Advance | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/13/2022 |  |  |  |  | The user facility reported a 46-year-old male implanted with an impella cp for mechanical circulatory support. it was reported that the patient was on the impella cp support due to placed st elevated myocardial infarction, acute myocardial infarction, cardiogenic shock and multiple rounds of cardio-pulmonary resuscitation. while on support, the patient experienced hemolysis due to positioning issues with the impella cp. no patient harm was reported. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29664 | 1220648-2024-08200 | 18926618 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18926618&pc=OZD | 07/21/2022 | 03/18/2024 |  | 03/18/2024 |  |  | 07/21/2022 | 03/18/2024 |  |  | 04/30/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023215154 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/21/2022 |  |  |  |  | The user facility reported a 76-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported while attempting to place the patient on support, the patient was observed having iliac and aorta torturous vessels making inserting the device difficult. the patient was on support for 33 minutes prior to explanting the impella with perclose and manual pressure was applied. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29665 | 1220648-2024-08201 | 18926712 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18926712&pc=OZD | 07/15/2022 | 03/11/2025 |  | 03/18/2024 |  |  | 07/15/2022 | 03/18/2024 |  |  | 03/31/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023214251 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/15/2022 |  |  |  |  | The user facility reported an 85-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient was on impella cp support due to intervention for non-st elevated myocardial infarction. while on support, the patient experienced access site bleeding that required transfusion of blood products and surgical wound closure. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29666 | 1220648-2024-08202 | 18926738 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18926738&pc=OZD | 07/21/2022 | 03/18/2024 |  | 03/18/2024 |  |  | 07/21/2022 | 03/18/2024 |  |  | 04/30/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023217231 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/21/2022 |  |  |  |  | The user facility reported a 73-year-old male patient was on impella cp support due to coronary artery disease and coronary artery bypass graft. while on support, the patient experienced bleeding at the groin. femostop was placed at low pressure above the impella site. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29667 | 1220648-2024-08203 | 18926810 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18926810&pc=OZD | 07/15/2022 | 03/11/2025 |  | 03/18/2024 |  |  | 07/15/2022 | 03/18/2024 |  |  | 04/30/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023216608 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 26 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/15/2022 |  |  |  |  | The user facility reported a 26-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient was placed on impella cp support after being found unresponsive and coding. the patient developed loss of perfusion to distal extremity, which was treated using an external fem-fem bypass. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29668 | 1220648-2024-08204 | 18926833 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18926833&pc=OZD | 07/21/2022 | 03/18/2024 |  | 03/18/2024 |  |  | 07/21/2022 | 03/18/2024 |  |  | 12/31/2023 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023178137 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/21/2022 |  |  |  |  | The user facility reported a 54-year-old female patient on an impella cp for mechanical circulatory support. while on support the patient experienced oozing at the impella site. redress and angle matching did not help, however, a skin suture helped with the bleeding. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29669 | 1220648-2024-08205 | 18926884 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18926884&pc=OZD | 07/21/2022 | 03/18/2024 |  | 03/18/2024 |  |  | 07/21/2022 | 03/18/2024 |  |  | 04/30/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023216622 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/21/2022 |  |  |  |  | The user facility reported a 65-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced mottled feet and pulses were absent. the impella cp was weaned and explanted and the patient regained pulsatility. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
29670 | 1220648-2024-08206 | 18926889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18926889&pc=OZD | 07/15/2022 | 03/11/2025 |  | 03/18/2024 |  |  | 07/15/2022 | 03/18/2024 |  |  | 02/29/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023199047 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/15/2022 |  |  |  |  | The user facility reported a male patient was placed on an impella cp for mechanical circulatory support as a percutaneous coronary intervention after presenting with st elevated myocardial infarction. the decision was made to explant the impella cp early and replace it due to the patient experiencing hemolysis and the device was sub-valvular. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29671 | 1220648-2024-08207 | 18926907 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18926907&pc=OZD | 07/21/2022 | 03/12/2025 |  | 03/18/2024 |  |  |  | 03/18/2024 |  |  | 10/31/2023 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022157522 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/21/2022 |  |  |  |  | The user facility reported a 63-year-old female was implanted with an impella cp device for mechanical circulatory support. the patient had multiple episodes of ventricular tachycardia (vtach) while on impella cp support which required defibrillations. an echocardiogram was performed, and the pump was repositioned. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29672 | 1220648-2024-08208 | 18926938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18926938&pc=OZD | 07/15/2022 | 03/12/2025 |  | 03/18/2024 |  |  |  | 03/18/2024 |  |  | 04/30/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023217687 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/15/2022 |  |  |  |  | The user facility reported a male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. this report is a continuation of case (b)(4). the complainant reported that a patient was implanted with a new impella cp due to poor positioning and continued hemolysis (complaint (b)(4)). during the explanting of the replacement impella cp, a dissection occurred, which required surgical stenting to close. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29673 | 1220648-2024-08209 | 18926993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18926993&pc=OZD | 07/15/2022 | 03/18/2024 |  | 03/18/2024 |  |  |  | 03/18/2024 |  |  | 03/31/2024 |  | 03/18/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023205291 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/15/2022 |  |  |  |  | The user facility reported a 84-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient developed bleeding/oozing at the impella cp insertion site. an additional 8 french angio seal was needed for hemostasis. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
29674 | 1220648-2024-08116 | 18919102 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18919102&pc=OZD | 03/05/2024 | 04/23/2024 |  | 03/17/2024 |  |  | 03/05/2024 | 03/17/2024 |  |  | 11/30/2025 |  | 03/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024431087 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/05/2024 |  |  |  |  | The user facility reported a 77-year-old male undergoing coronary artery bypass grafting was to be implanted with an impella 5.5 device for mechanical circulatory support. it was reported that while advancing the pump, the catheter kinked due to the patient's challenging anatomy and the pump was subsequently removed. a second pump was prepped and implanted for patient support. there was no patient harm. | [5 empty] | The investigation into the delivery issues has been completed since the original report was filed. the pump was returned, tested, and evaluated. a catheter and cannula kink were observed. the root cause of delivery issue is determined to be use issue because the pump was removed and a visual kinking/buckling of the catheter shaft at the level of the motor was seen. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29675 | 1220648-2024-08117 | 18919103 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18919103&pc=OZD | 03/05/2024 | 06/12/2024 |  | 03/17/2024 |  |  | 03/05/2024 | 03/17/2024 |  |  |  |  | 03/17/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1721787 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Erratic or Intermittent Display | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/05/2024 |  |  |  |  | The user facility biomed department reported that the automated impella controller (aic) screen was flickering. the aic was removed from service. no patient involvement or harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - display issue has been completed since the initial report was submitted. the console was returned, tested, and visually inspected. no flicker had been observed. screen components were inspected but no anomalies were found. reported issue could not be confirmed or reproduced. the cause of the issue was not determined since issue could not be reproduced.
29676 | 1220648-2024-08118 | 18919104 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18919104&pc=OZD | 03/05/2024 | 04/09/2024 |  | 03/17/2024 |  |  | 03/05/2024 | 03/17/2024 |  |  |  |  | 03/17/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1250417 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/05/2024 |  |  |  |  | The complainant reported that the automated impella controller (aic) had a blank screen and triggered a pump stop during patient support. the aic was swapped in response to the noted failure. there was no known patient harm. | [5 empty] | The investigation for the reported console (aic) optical signal issues has been completed. data logs were reviewed which revealed multiple errors due to an invalid crc to the optical bench. it also showed fluctuating pressures on the from the raw optical sensor followed by a pump stop alarm, which occurred when no pump is inserted but pump disconnection detection was disturbed by the out of bounds snr preceding the pump removal. the aic was returned for evaluation and functional testing was performed. the failure was unable to be reproduced throughout testing and is characterized by a temporary loss of optical data (placement, snr, contrast, lamp, gain) and triggering of a controller error alarm, has a low probability of reoccurrence. if needed, patient hemodynamics and impella position can be monitored with imaging. the cause of the issue was not determined since issue could not be reproduced. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed.
29677 | 1220648-2024-08119 | 18919105 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18919105&pc=OZD | 03/05/2024 | 04/24/2024 |  | 03/17/2024 |  |  | 03/05/2024 | 03/17/2024 |  |  | 07/31/2025 |  | 03/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024392123 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/05/2024 |  |  |  |  | The user facility reported a 48-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported the patient was on support due to the patient coding. while on support, there was access site bleeding. dressing was addressed, a figure 8 stitch was applied and powder was used around the site. a sand bag was also applied. additionally, there was low pump flow and two units of blood products were provided. | [5 empty] | The impella device was discarded and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation into the access site bleeding ¿ major and the low pump flow has been completed since the original report was submitted. the device was not returned for investigation. the cause of the access site bleeding was anticoagulation management issue due to improper anticoagulation: high acts (act was 287 at time of bleeding). as the necessary information was not provided, and the device was not returned, the root cause of the low pump flow could not be determined.
29678 | 1220648-2024-08120 | 18919107 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18919107&pc=OZD | 03/05/2024 | 04/24/2024 |  | 03/17/2024 |  |  | 03/05/2024 | 03/17/2024 |  |  | 09/30/2025 |  | 03/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024402073 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/05/2024 |  |  |  |  | The user facility reported a 53-year-old male with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported that a clot ingestion was causing constant suction requiring the impella cp to be exchanged. on explant, a visual clot was seen in the catheter. the patient was reported to be stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the thrombus has been completed. the impella device has not been completed. the data logs were reviewed and confirm multiple suction alarms and motor current spikes, which is consistent with thrombus ingestions. the root cause of the low pump flow was most likely biomaterial. the instructions for use states ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
29679 | 1220648-2024-08121 | 18919109 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18919109&pc=OZD | 03/05/2024 | 04/17/2024 |  | 03/17/2024 |  |  | 03/05/2024 | 03/17/2024 |  |  | 02/28/2025 |  | 03/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024338629 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  |  |  | Malfunction | No | Yes | Material Twisted/Bent | Device Handling Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/05/2024 |  |  |  |  | The user facility reported a 61-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that the patient was on impella cp support and the patient coded. cardio pulmonary resuscitation was performed and caused the cannula to kink in the aorta. the physician was able to straighten the cannula, but elected to replace the pump instead. the clinical representative confirmed the patient coding was not due to the impella. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the product damage (cannula kink) issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the cannula kink was determined to be use related as this occurred while staff actively performed cpr.
29680 | 1220648-2024-08122 | 18919110 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18919110&pc=OZD | 03/05/2024 | 04/24/2024 |  | 03/17/2024 |  |  | 03/05/2024 | 03/17/2024 |  |  | 11/30/2025 |  | 03/17/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024434190 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/05/2024 |  |  |  |  | The user facility reported a 74-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during placement, the impella 5.5 was not able to advance into the ascending aorta despite stiff buddy wires. the decision was then made to proceed with placing an impella cp through the right axillary graft. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because the doctor was unable to advance impella 5.5 into the ascending aorta despite stiff buddy wires but they were able insert impella cp later without any issues. so the root cause is determined to be patient condition.
29681 | 1220648-2024-08110 | 18918944 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18918944&pc=OZD | 12/28/2023 | 09/11/2024 |  | 03/16/2024 |  |  | 02/15/2024 | 03/16/2024 |  |  | 11/30/2025 |  | 03/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024423437 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/15/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry regarding a 60 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient endured an impella site hematoma that prompted a blood transfusion. a "possible" relationship to the impella device was made: | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08104 was provided in sections b2, b5, h1, and h6. | The investigation into the access site bleeding ¿ major issue has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was unable to be determined as no product was returned and insufficient clinical details were provided.
29682 | 1220648-2024-08112 | 18919001 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18919001&pc=OZD | 10/19/2023 | 08/30/2024 |  | 03/16/2024 |  |  | 02/16/2024 | 03/16/2024 |  |  | 07/31/2025 |  | 03/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024377867 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 02/16/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry regarding a 58 year old male patient presenting with acute myocardial infarction and cardiogenic shock for hemodynamic support via the impella device. during support, this patient endured hemolysis with a possible relationship to the impella device. one unit of blood products were delivered, in response to a drop in hemoglobin. | Additional information received via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured lle acute limb ischemia with a "possible" relationship to the impella device used: ¿after impella removal, immediately l leg was pulseless, cold, with pain from thigh down to left toes. vascular and cardiac team evaluated. team suspected thrombus formation on impella device with embolization down the left leg during removal. vascular performed embolectomy/thrombectomy and fasciotomy on l iliac artery. the patient recovered with no sequelae. |  |  |  |  | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. added h6 component code 925 d4-corrected model number added d6a/d6b. f6/f8 should have been left blank in the initial report. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions: including catheter position, pre-existing medical conditions, and small left ventricular volumes / may also play a role in patient susceptibility to hemolysis.¿ | Section b5 and h6 were updated as per additional information received. | The investigation into the hemolysis issue has been completed since the initial report was submitted. the product was not returned for investigation. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review.
29683 | 1220648-2024-08113 | 18919020 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18919020&pc=OZD | 12/16/2023 | 05/20/2024 |  | 03/16/2024 |  |  | 02/16/2024 | 03/16/2024 |  |  | 09/30/2025 |  | 03/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024398228 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 36 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 02/16/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry regarding a 36 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient endured hemolysis with a "possible" relationship to the impella device. unspecified intervention was required, in response. | Additional informations received from our medical safety team via loqi (abiomed¿s long-term outcome and quality indicator impella registry), pertaining to the above case. it was noted that worsening ventricular tachycardia with a possible relationship to the impella 5.5 device used on this patient. impella was moved back to address the issue. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the hemolysis and the vf/vt/af/at (worsening ventricular tachycardia) issues has been completed since the original report was submitted. the pump was not returned for investigation. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. as the necessary information was not provided, the root cause of the vt/vf was not determined. d6a, d6b. | Section b5 and h6 were updated as per additional information received. | The investigation for the reported access site bleed and hemolysis has been completed. the device nor sufficient clinical information was returned for evaluation. data logs revealed some impella in aorta alarms and impella stop-mc high alarm close to end of the case but they were resolved quickly. however, without additional clinical details, the root cause of the access site bleed and hemolysis could not be determined.
29684 | 1220648-2024-08114 | 18919055 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18919055&pc=OZD | 08/02/2022 | 03/16/2024 |  | 03/16/2024 |  |  | 08/02/2022 | 03/16/2024 |  |  | 03/31/2024 |  | 03/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023212178 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/02/2022 |  |  |  |  | The user facility reported a 69-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient presented with severe left main stenosis and severe right coronary artery stenosis. patient decompensated overnight and was intubated, after which the impella cp was placed. while on support the patient had significant oozing around the impella. femstop was used to resolve the hematoma, and the device was eventually explanted without further reported complications. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
29685 | 1220648-2024-08115 | 18919056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18919056&pc=OZD | 08/12/2022 | 03/16/2024 |  | 03/16/2024 |  |  | 08/12/2022 | 03/16/2024 |  |  | 04/30/2024 |  | 03/16/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023221023 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Diminished Pulse Pressure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 08/12/2022 |  |  |  |  | The user facility reported a 51-year-old patient needing mechanical circulatory support. it was reported that the patient decompensated after being treated for afib and shortness of breath (sob). the decision was made to place the patient on impella cp. during support, low pulsatility, low urine output and hemolysis was observed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
29686 | 1220648-2024-08056 | 18911184 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18911184&pc=OZD | 07/31/2022 | 03/15/2024 |  | 03/15/2024 |  |  | 07/31/2022 | 03/15/2024 |  |  | 01/31/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023187142 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/31/2022 |  |  |  |  | The user facility reported a 75-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. due to aortic valve stenosis, the impella could not pass post the proximal part. case aborted after several attempts using platinum plus wire and body wire. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29687 | 1220648-2024-08057 | 18911225 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18911225&pc=OZD | 08/27/2022 | 03/18/2025 |  | 03/15/2024 |  |  | 08/27/2022 | 03/15/2024 |  |  | 03/31/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023206461 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/27/2022 |  |  |  |  | The user facility reported a 61-year-old male on a impella cp for mechanical circulatory support. it was reported that immediately after peeling away and inserting the repositioning sheath, the patient¿s groin area began bleeding profusely. after twelve (12) minutes of manual pressure, the physician elected to explant the impella cp. units of blood were provided to the patient due to significant blood loss. a femstop was placed as well as a foley catheter. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29688 | 1220648-2024-08058 | 18911253 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18911253&pc=OZD | 09/08/2022 | 03/15/2024 |  | 03/15/2024 |  |  | 09/08/2022 | 03/15/2024 |  |  | 08/31/2023 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022134530 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 09/08/2022 |  |  |  |  | The user facility reported a 73-year-old patient needing mechanical circulatory support. it was reported that there was oozing from the impella site as an echo verified that the pump was slightly deep. the decision was made to stop heparin. noted plans to pull pump back 1 cm and to also place a central line. perfusion notified to do console swap. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
29689 | 1220648-2024-08059 | 18911287 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18911287&pc=OZD | 10/25/2022 | 03/18/2025 |  | 03/15/2024 |  |  | 10/25/2022 | 03/15/2024 |  |  | 06/30/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023237614 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/25/2022 |  |  |  |  | The user facility reported a 60-year-old female with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the patient experienced immediate suction and low flows. the impella appeared to be in correct position per fluoro. a trans-esophageal echocardiogram probe revealed a large thrombus in the left ventricle. the impella was explanted and clots were verified in the right atrium, right ventricle, left atrium, and left ventricle per echo. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29690 | 1220648-2024-08060 | 18911313 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18911313&pc=OZD | 10/25/2022 | 03/18/2025 |  | 03/15/2024 |  |  | 10/25/2022 | 03/15/2024 |  |  | 07/31/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023256322 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/25/2022 |  |  |  |  | The user facility reported a 84-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support the peel away sheath was removed and the perclose sutures were tightened. minimal arterial bleeding from site was noted. the doctor chose to place a 5 fr sheath to assist with hemostasis. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
29691 | 1220648-2024-08061 | 18911330 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18911330&pc=OZD | 10/26/2022 | 03/18/2025 |  | 03/15/2024 |  |  | 10/26/2022 | 03/15/2024 |  |  | 02/29/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023194065 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/26/2022 |  |  |  |  | The user facility reported a 86-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced a hematoma. the hematoma was expressed, and the physician decided to pull pump and apply manual pressure and a fem-stop. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29692 | 1220648-2024-08062 | 18911360 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18911360&pc=OZD | 07/28/2022 | 03/18/2025 |  | 03/15/2024 |  |  | 07/28/2022 | 03/15/2024 |  |  | 11/30/2023 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023165997 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/28/2022 |  |  |  |  | The user facility reported a 83-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. contralateral left femoral artery was closed with angioseal but did not hold. manual pressure held but a large left sided hematoma was observed. protamine was administered to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29693 | 1220648-2024-08063 | 18911373 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18911373&pc=OZD | 07/29/2022 | 03/18/2025 |  | 03/15/2024 |  |  | 07/29/2022 | 03/15/2024 |  |  | 01/31/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023192891 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/29/2022 |  |  |  |  | The user facility reported a 81-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. impella cp was repositioned because the catheter was caught in echo. however, due to suction alarm and being hemodynamically unstable, the impella was explanted. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29694 | 1220648-2024-08064 | 18911407 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18911407&pc=OZD | 07/31/2022 | 03/18/2025 |  | 03/15/2024 |  |  | 07/31/2022 | 03/15/2024 |  |  | 03/31/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023212475 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/31/2022 |  |  |  |  | Us complainant reported that a 76 year old male patient was placed on impella cp support after presenting with a st elevated myocardial infarction. it was noted that the patient had frank blood red colored urine. it was observed to have occurred sometime after the balloon pump was removed. the care team was attempting to determine if the hematuria was caused by hemolysis or bleeding in the foley. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
29695 | 1220648-2024-08065 | 18911421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18911421&pc=OZD | 07/31/2022 | 03/18/2025 |  | 03/15/2024 |  |  | 07/31/2022 | 03/15/2024 |  |  | 04/30/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023220155 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/31/2022 |  |  |  |  | The user facility reported a 61-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. the complainant reported that there was access site bleeding noted by oozing at the site. two units of blood were administered. impella cp was removed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29696 | 1220648-2024-08066 | 18911457 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18911457&pc=OZD | 08/01/2022 | 03/18/2025 |  | 03/15/2024 |  |  | 08/01/2022 | 03/15/2024 |  |  | 02/29/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023203036 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 19 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/01/2022 |  |  |  |  | The user facility reported a 19-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient started bleeding at the impella cp insertion site. as a result, systematic heparin was off. the bleeding at the insertion site stopped. additionally, the medical team was unable to detect a dopplerable pulse below the popliteal artery. the console was exchanged. all of the patient¿s labs hemolyzed and there was no urine output and the patient was placed on continuous renal replacement therapy (ccrt). | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29697 | 1220648-2024-08067 | 18911473 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18911473&pc=OZD | 08/01/2022 | 03/18/2025 |  | 03/15/2024 |  |  | 08/01/2022 | 03/15/2024 |  |  | 04/30/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023219119 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 27 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/01/2022 |  |  |  |  | The user facility reported a 27-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the pump was unable to go above p-2 due to impella in aorta alarms. the impella was later explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29698 | 1220648-2024-08068 | 18911491 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18911491&pc=OZD | 08/02/2022 | 03/18/2025 |  | 03/15/2024 |  |  | 08/02/2022 | 03/15/2024 |  |  | 02/29/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023197971 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/02/2022 |  |  |  |  | The user facility reported a 59-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. a clot was present after the removal of impella cp. impella weaned and successfully explanted as normal with manual pressure held for 45 minutes. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency) ¿relevant laboratory test results, such as coagulation profiles and platelet counts ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states) ¿medications or treatments that the patient was receiving at the time of thrombus formation ¿surgical or procedural details if applicable (e.g, cardiac catheterization) ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events) ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. this report is being filed as part of a retrospective review of historical records.
29699 | 1220648-2024-08069 | 18911929 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18911929&pc=OZD | 09/14/2022 | 03/18/2025 |  | 03/15/2024 |  |  | 09/14/2022 | 03/15/2024 |  |  | 05/31/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023228963 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/14/2022 |  |  |  |  | The user facility reported a 77-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported that during the impella cp pump insertion, the pump was unable to pass through the abdominal aortic aneurysm (aaa) stent struts. the medical team removed pump and rewired with a non-tapered wire, and using a buddy wire was able to successfully cross the left ventricle and place the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29700 | 1220648-2024-08072 | 18911937 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18911937&pc=OZD | 09/14/2022 | 03/18/2025 |  | 03/15/2024 |  |  | 09/14/2022 | 03/15/2024 |  |  | 01/31/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023188335 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/14/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 64-year-old female patient for mechanical circulatory support. it was reported while on impella cp support, the patient developed bleeding at the impella access site. bicarbonate was added to the purge, heparin was withdrawn, and the patient received blood products. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29701 | 1220648-2024-08071 | 18912069 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18912069&pc=OZD | 03/03/2024 | 09/13/2024 |  | 03/15/2024 |  |  | 03/03/2024 | 03/15/2024 |  |  | 06/30/2025 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024371811 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  |  |  | Death | Yes | Yes | Failure to Advance | Activation, Positioning or Separation Problem | [5 empty] | Cardiac Arrest | [12 empty] | Death | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/03/2024 |  |  |  |  | The user facility reported a 76-year-old male patient of unknown race and ethnicity in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp was attempted to be placed, however the patient arrested during the impella placement. cardiopulmonary resuscitation was performed and the impella was pulled out of the sheath. after multiple attempts to reinsert the impella over a wire and second device was opened but the patient expired before the pump could be inserted. no allegations were made towards the impella for cause of death. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the reported issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue cannot be determined since the product was not returned and insufficient clinical details are provided. as the necessary clinical information was not provided and the device was not returned, the root cause of the vf/vt/af/at (ventricular tachycardia) was not determined. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrected information for the initial mfr 1220648-2024-08071 was provided.
29702 | 1220648-2024-08076 | 18912073 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18912073&pc=OZD | 08/20/2022 | 03/10/2025 |  | 03/15/2024 |  |  | 08/20/2022 | 03/15/2024 |  |  | 03/31/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023208068 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/20/2022 |  |  |  |  | The user facility reported a 41-year-old female in cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. while on support, the patient¿s urine turned red and the labs were hemolyzed. currently at p9, there was a discussion to decrease the p-level to p8 to reduce the hemolysis. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. the instructions for use state ¿potential adverse events:acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported hemolysis issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
29703 | 1220648-2024-08074 | 18912116 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18912116&pc=OZD | 03/04/2024 | 04/23/2024 |  | 03/15/2024 |  |  | 03/04/2024 | 03/15/2024 |  |  | 07/31/2025 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024388707 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  |  | White | Malfunction | Yes | Yes | Material Integrity Problem | Device Handling Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/04/2024 |  |  |  |  | The user facility reported a 69-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the patient was on impella cp support due to going into cardiac arrest. while on support, the patient went into ventricular fibrillation (vf) and the patient was shocked out of it. additionally, there was significant access site bleeding. the site was resutured and angle matching was performed. additionally, the drive shaft kinked during repositioning. there was a motor spike and a drop in flows. the pump was explanted. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation into the reported issues has been completed since the original report was submitted. the device was not returned for investigation. based on clinical description and the data log review, the cause of the positioning issue most likely was use related issue as the impella pulled back across the valve, unable to re-cross and catheter kinked due to excessive force during repositioning. the cause of the low pump flow issue was not determined since product was not returned for investigation.
29704 | 1220648-2024-08070 | 18912132 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18912132&pc=OZD | 03/03/2024 | 04/22/2024 |  | 03/15/2024 |  |  | 03/03/2024 | 03/15/2024 |  |  |  |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025441814 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Insufficient Flow or Under Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/03/2024 |  |  |  |  | The user facility reported a 30-year-old male patient undergoing coronary artery bypass grafting was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there was clot ingestion concerns due to the low flows. the low flows had been treated by blood volume delivery. the patient had a known small left ventricle (lv). the low flows prompted the team to explant and replace the pump. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the pump stop issue has been completed since the original report was filed. the pump was returned, tested, and evaluated. it was reported that motor current spikes and suction alarms occurred on the date of the event. per the data log review, the cause of the low pump flow issue was biomaterial ingestion.
29705 | 1220648-2024-08077 | 18912188 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18912188&pc=OZD | 08/20/2022 | 03/10/2025 |  | 03/15/2024 |  |  | 08/20/2022 | 03/15/2024 |  |  | 04/30/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023211646 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/20/2022 |  |  |  |  | The user facility reported a female in cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding and hematoma at the access site. a blood transfusion was required. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. the instructions for use states¿ potential adverse events (united states):acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleeding issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
29706 | 1220648-2024-08073 | 18912284 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18912284&pc=OZD | 08/19/2022 | 03/07/2025 |  | 03/15/2024 |  |  | 08/19/2022 | 03/15/2024 |  |  | 03/31/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023207121 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/19/2022 |  |  |  |  | The user facility reported a 79-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient coded while in the cath lab and the impella was pulled into the aorta during cpr. the physician opted to remove the pump and cannulate extracorporeal membrane oxygenation through the impella site due to the emergent situation. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29707 | 1220648-2024-08075 | 18912333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18912333&pc=OZD | 08/20/2022 | 03/07/2025 |  | 03/15/2024 |  |  | 08/20/2022 | 03/15/2024 |  |  | 05/31/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023230392 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/20/2022 |  |  |  |  | The user facility reported a 78-year-old female in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the device was placed in the right femoral artery (rfa) and when removing the long sheath the pump came across atrioventricular (av). the impella was removed and a new sheath was placed, while the atrioventricular was re-crossed with the same pump that was reinserted. the patient experienced multiple non pulsatility runs; however, a mean arterial pressure(map) was maintained as well suction events occurring. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29708 | 1220648-2024-08078 | 18912390 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18912390&pc=OZD | 08/21/2022 | 03/10/2025 |  | 03/15/2024 |  |  | 08/21/2022 | 03/15/2024 |  |  | 05/31/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023229632 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/21/2022 |  |  |  |  | The user facility reported a 83-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed a large hematoma at the insertion site of the pump, requiring manual pressure to be held as well as 2units of packed red blood cells(prbcs) for treatment. the groin was resutured and dressed. however, another 1 unit of packed red blood cells were given to the patient in the afternoon. the next day the patient's groin was found to be still oozing at the insertion sit. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. the instructions for use states¿ potential adverse events: acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleeding ¿ major issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
29709 | 1220648-2024-08079 | 18912439 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18912439&pc=OZD | 08/21/2022 | 03/11/2025 |  | 03/15/2024 |  |  | 08/21/2022 | 03/15/2024 |  |  | 04/30/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023217250 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/21/2022 |  |  |  |  | The user facility reported a 76-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced some oozing at the pump insertion site and the site was redressed and angle matched. the current partial thromboplastin time was over 200 and systemic purge heparin was withheld until in goal range. the pump was successfully weaned and explanted from the patient's bedside with 40 minutes of manual pressure. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29710 | 1220648-2024-08080 | 18912441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18912441&pc=OZD | 08/21/2022 | 03/10/2025 |  | 03/15/2024 |  |  | 08/21/2022 | 03/15/2024 |  |  | 02/29/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023198217 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/21/2022 |  |  |  |  | The user facility reported a 49-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient was given over 20 units of blood, due to bleeding from all sites and no heparin was given to the patient. femostop was removed from left femoral artery and no active bleeding from impella site was observed at this time. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. the instructions for use states¿ potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleeding ¿ major issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
29711 | 1220648-2024-08082 | 18912654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18912654&pc=OZD | 10/07/2022 | 03/18/2025 |  | 03/15/2024 |  |  | 10/07/2022 | 03/15/2024 |  |  | 07/31/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023252377 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 10/07/2022 |  |  |  |  | The user facility reported a 77-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, a dissection of the left anterior descending artery (lad) occurred while treating the lad portion of a trifurcation lesion. a pericardiocentesis was performed and the patient required a blood transfusion as a result of the dissection of the lad. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29712 | 1220648-2024-08083 | 18913357 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18913357&pc=OZD | 07/22/2022 | 03/10/2025 |  | 03/15/2024 |  |  | 07/22/2022 | 03/15/2024 |  |  | 03/31/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023210405 |  | 0048-0068 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/22/2022 |  |  |  |  | The user facility reported a 49-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient lost pulses and coded. catheter revealed chronic total occlusion of right coronary artery, 100% circ and left anterior descending lesion. the patient had a significant drop in platelets and hemoglobin. heparin was removed. surgeon opted to remove the impella cp. patient was monitored and attempted to titrate pressors. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29713 | 1220648-2024-08084 | 18913610 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18913610&pc=OZD | 08/31/2022 | 03/18/2025 |  | 03/15/2024 |  |  | 08/31/2022 | 03/15/2024 |  |  | 12/31/2023 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023178108 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | Improper Flow or Infusion | [5 empty] | Ischemia | Gastrointestinal Hemorrhage | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/17/2022 |  |  |  |  | The user facility reported a 43-year-old female was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced mottled feet, gi bleeding, intermittent suction alarms from the pump, and pump stoppage. the suction alarms were reported to have been treated with transfusion of blood products and albumen. the patient's gi bleeding was treated with transfusion of blood products. there was no additional information provided for treatment of mottled feet. the impella cp device stopped operating after 19 days of implantation, and was removed. the patient remained on veno-arterio-venous extracorporeal membrane oxygenation (vav ecmo). | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29714 | 1220648-2024-08085 | 18914090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18914090&pc=OZD | 10/14/2022 | 03/18/2025 |  | 03/15/2024 |  |  | 10/14/2022 | 03/15/2024 |  |  | 07/31/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023252251 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/14/2022 |  |  |  |  | The user facility reported a 67-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported patient was brought to cardiac cath lab for emergent left heart catheterization where the patient coded, and it was determined that the patient had a dissection that required a pericardial drain and blood products. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29715 | 1220648-2024-08087 | 18914230 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18914230&pc=OZD | 03/01/2024 | 04/18/2024 |  | 03/15/2024 |  |  | 03/01/2024 | 03/15/2024 |  |  | 09/30/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023275356 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella purge flow dropped, and purge pressure climbed. sterile water purge solution was administered, and the issue was resolved within a few hours, with only half of the purge solution being administered. the purge solution was then swapped back to the standard. the impella cp was performing well 24 hours later. discussed the potential of future pump stoppage if this was to occur again. the physician decided to let the existing pump remain and to monitor the patient¿s progress. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation completion, a supplemental mdr will be filed. | The investigation into the low pump flow has been completed since the original report was filed. the pump was returned, tested, and evaluated. visually inspected the pump and no abnormalities were observed, and the issue was not reproduceable during testing. the cause of the high purge pressure issue was most likely biomaterial as, per clinical, tpa resolved the issue during support.
29716 | 1220648-2024-08088 | 18914312 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18914312&pc=OZD | 11/12/2022 | 05/16/2025 |  | 03/15/2024 |  |  | 11/12/2022 | 03/15/2024 |  |  | 08/31/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023259567 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/12/2022 |  |  |  |  | The user facility reported a 78-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while on impella cp support there were no pulses in the patients left leg and the patient experienced red urine. the p-level was lowered on the impella cp device, patient was started on continuous renal replacement therapy and the patient was transfused with 2 units of blood. the limb ischemia was treated by a distal perfusion. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29717 | 1220648-2024-08089 | 18914313 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18914313&pc=OZD | 03/05/2024 | 04/17/2024 |  | 03/15/2024 |  |  | 03/05/2024 | 03/15/2024 |  |  | 08/31/2025 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025443755 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/05/2024 |  |  |  |  | The user facility reported a 69-year-old male patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that the p level was reduced to p9 and after that the impella spontaneously stopped for a few seconds and then restarted on its own. after a few minutes more, the left ventricle systolic pressure started rising to a pedal pressure of around 260 with a left ventricle diastolic pressure being -114. as the pressures increased above normal range, the physician attempted to reposition the impella but the impella abruptly stopped. the pump was replaced with a new impella cp and used without incidence or complications. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop issue has been completed since the original report was submitted. the device was not returned for investigation. according to the data lag and the clinical information reviewed, the root cause of the pump stop was due to the clipped purge line on purge cassette tubing causing high purge pressure leading to the pump stop.
29718 | 1220648-2024-08090 | 18914490 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18914490&pc=OZD | 03/05/2024 | 03/15/2024 |  | 03/15/2024 |  |  | 03/05/2024 | 03/15/2024 |  |  |  |  | 03/15/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1663431 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 03/05/2024 |  |  |  |  | The user facility reported a 65-year-old male undergoing a possible mitral valve replacement was implanted with an impella 5.5 device for mechanical circulatory support pre-operatively. it was reported that while on support, the automated impella controller (aic) had an ¿unable to detect purge fluid¿ alarm. the cassette was replaced twice to no avail. the aic was replaced, and the error resolved. no patient harm was reported. | [5 empty] | The impella device was received from the customer and an investigation regarding the returned device has been completed. logs were reviewed and the issues with the aic not being able to detect the purge fluid was confirmed. there were numerous instances of the purge fluid not detected prompt found in the logs on the reported occurrence date. the console was tested. the purge fluid not detected issue was able to be reproduced, and the purge disc retainer was found to be broken. a portion of the retainer that was supposed to be fastened to the purge assembly was no longer intact. because the purge disc retainer was not fully flush with the purge assembly, the purge transducer could not detect any pressure going to the purge disc which would result in this purge fluid not detected prompt. the root cause of this issue was a broken blue purge disc retainer.
29719 | 1220648-2024-08091 | 18914678 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18914678&pc=OZD | 03/01/2024 | 05/29/2024 |  | 03/15/2024 |  |  | 03/02/2024 | 03/15/2024 |  |  | 11/30/2025 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024424533 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a female patient on impella cp support was bleeding from the access site and a hematoma formed. the patient had a drop in hemoglobin and was given three units of packed red blood cells and pressure dressing was placed above the site. the bleeding was noted to have subsided. | The patient's outcome was reviewed by abiomed's medical safety officer (mso). it was determined that there was not enough information to associate use of device with outcome. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ added- mso review in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
29720 | 1220648-2024-08092 | 18914690 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18914690&pc=OZD | 06/09/2022 | 09/12/2024 |  | 03/15/2024 |  |  | 06/09/2022 | 03/15/2024 |  |  | 03/31/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023199355 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  |  | Death | Yes | Yes | Positioning Failure | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/09/2022 |  |  |  |  | The user facility reported a 61-year-old male, was implanted with an impella cp device for mechanical circulatory support. while on support, the impella dislodged from the patient under transport and patient restlessness. replacing the impella was not able to be achieved, the patient expired before a second impella could be implanted. no allegations were made towards the impella for cause of death. | It was reported by medical staff that the patient expired within 3 hours of admission to the unit. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08092 was provided in sections b1, b2, b5, h1, and h6.
29721 | 1220648-2024-08093 | 18914768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18914768&pc=OZD | 03/02/2024 | 10/03/2025 |  | 03/15/2024 |  |  | 03/02/2024 | 03/15/2024 |  |  |  |  | 03/15/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1163150 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Death | Yes | Yes | Connection Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that during impella cp support, the automated impella controller (aic) did not recognize the pump. the aic was therefore exchanged and the issue resolved. there was no reported patient harm. | The patient's outcome was reviewed by abiomed's medical safety officer (mso). it was determined that there was no association between impella use and outcome. | The patient expired while on support. |  |  |  | The automated impella controller was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console (aic) optical issue has been completed. data logs were confirmed that when the pump was initially connected and detected by the console on the date of the complaint, an optical sensor system error message was displayed on the case start wizard screen. right data logs from the console on the day the snr was significantly low immediately after the pump was plugged in. the console was returned and tested. an oscilloscope was used to get the fringe pattern with and without a pump connected showing no observable anomaly or distortion. the issue was confirmed. the root cause of the optical sensor error was most probably due to an air gap. the root cause for the optical signal issue was due to low snr which could be caused by an airgap at the front optical connector. added- b5 mso review, d9 device return date d2-updated common device name. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 adverse event updated. b2 updated to death. b5 updated to include death. h1 updated to death. h6 updated to include optical signal coding. d10 concomitant medical products and therapy dates updated. g1 reporting contact email updated.
29722 | 1220648-2024-08094 | 18914885 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18914885&pc=OZD | 07/27/2022 | 03/18/2025 |  | 03/15/2024 |  |  | 07/27/2022 | 03/15/2024 |  |  | 04/30/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023219147 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/27/2022 |  |  |  |  | The user facility reported a 74-year-old male was implanted with an impella cp device for mechanical circulatory support and developed a slight hematoma at the groin. the hematoma was compressible and light belt pressure was applied. heparin was held until heparin xa (heparin assay) was less than one (1). the impella device was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records.
29723 | 1220648-2024-08095 | 18915056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18915056&pc=PYX | 03/02/2024 | 04/19/2024 |  | 03/15/2024 |  |  | 03/02/2024 | 03/15/2024 |  |  | 08/31/2025 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024397374 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 41-year-old male noted as high-risk percutaneous cardiac insertion was implanted with an impella rp flex device for mechanical circulatory support. the patient was on impella rp support and the patient had a thrombus. the impella rp flex was entrained with a clot. the patient was provided with anticoagulation. the impella rp flex was removed and replaced with protek duo. | [5 empty] | The investigation for the reported hemolysis and thrombocytopenia has been completed. the impella device was returned for evaluation. upon evaluation, a clot was observed around the impeller. the sterile sleeve was detached on one side. there were no cannula kinks. data logs were reviewed which showed positioning alarms throughout the case, but the pump's position was confirmed to be correct via echo. there are placement signal low alarms throughout the case. the root cause of the hemolysis was most likely biomaterial. the root cause of the thrombocytopenia was not determined due to insufficient clinical details. the instructions for use states ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected low s or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29724 | 1220648-2024-08096 | 18915168 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18915168&pc=OZD | 10/17/2022 | 09/12/2024 |  | 03/15/2024 |  |  | 10/17/2022 | 03/15/2024 |  |  | 06/30/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023237437 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Thrombocytopenia | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/17/2022 |  |  |  |  | The user facility reported a 64-year-old male was implanted with an impella device for mechanical circulatory support after experiencing afib and worsening shortness of breath/leg swelling over the past few months. the patient developed bleeding at the impella cp insertion site. the patient received two (2) units of packed red blood cells due to low platelets and was switched to bicarb in the purge and had not been provided heparin. however, the patient had a cangrelor drip. additionally, the patient had bloody urine output, bleeding in the gi tract and was on continuous renal replacement therapy (crrt). the patient was escalated to the impella 5.5. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08096 was provided in sections b2, b5, h1, and h6.
29725 | 1220648-2024-08097 | 18915225 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18915225&pc=OZD | 03/02/2024 | 04/08/2024 |  | 03/15/2024 |  |  | 03/02/2024 | 03/15/2024 |  |  | 07/31/2025 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025442582 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an 82-year-old female patient was on an impella cp due to needing hemodynamic support and high risk percutaneous coronary intervention. while on support, the cutdown site was swollen with pressure dressing. one unit of packed red blood cells was provided. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation into the access site bleeding ¿ major issue has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided.
29726 | 1220648-2024-08098 | 18915376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18915376&pc=OZD | 03/03/2024 | 05/29/2024 |  | 03/15/2024 |  |  | 03/03/2024 | 03/15/2024 |  |  | 10/31/2025 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024412652 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Female |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support for 52-year-old female patient, there was a leak noted on the purge side arm. there was no reported patient harm. | The patient expired while on support. the patient's outcome was reviewed by abiomed's medical safety officer (mso). it was determined that there was not enough information to associate use of device with outcome. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation into the purge sidearm leak has been completed. the device was not returned for evaluation. however, the clinical report has been reviewed. the root cause of the purge leak was unable to be determined because the product was not returned so the location of the leak could not be identified. the instructions for use reported in the initial report is from ifu for impella cp with smart assist system, section: cleaning. added- b5- patient outcome and mso review added- d6a/d6b.
29727 | 1220648-2024-08099 | 18915708 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18915708&pc=OZD | 10/23/2022 | 03/15/2024 |  | 03/15/2024 |  |  | 10/23/2022 | 03/15/2024 |  |  | 05/31/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023230356 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/23/2022 |  |  |  |  | The user facility reported a 73-year-old male was implanted with an impella cp device for mechanical circulatory support. the patient experienced access site bleeding noted by the hematoma at the groin site. this was treated by re-suturing, placing gauze under the repositioning sheath to maintain the angle and manually applying pressure to the site. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
29728 | 1220648-2024-08100 | 18916267 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18916267&pc=OZD | 10/25/2022 | 09/12/2024 |  | 03/15/2024 |  |  | 10/25/2022 | 03/15/2024 |  |  | 04/30/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218257 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ischemia | Hematuria | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/25/2022 |  |  |  |  | The user facility reported a 89-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. an occlusion was noted to the left leg and pedal pulse was not present. a fem-fem bypass was performed with good results. hematuria was noted along with bleeding around insertion site, so the site dressing was changed. it was also noted that when ic crossed the valve, the patient went into complete heart block. therefore, a temporary pacer was placed. the patient's family decided to withdraw care and the patient later expired. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08100 was provided in sections b2, b5, h1, and h6.
29729 | 1220648-2024-08101 | 18918707 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18918707&pc=OZD | 07/24/2022 | 03/15/2024 |  | 03/15/2024 |  |  | 07/24/2022 | 03/15/2024 |  |  | 01/31/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023184683 |  | 0048-0070 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/24/2022 |  |  |  |  | The user facility reported a 74-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while attempting to implant the impella, the peel away sheath kinked in a tortuous vessel rendering the impella unable to pass. the impella sheath was inserted with serial dilation with additional heparin administered. the peel-away was exchanged for a 16f cook. the impella was later weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29730 | 1220648-2024-08102 | 18918736 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18918736&pc=OZD | 07/25/2022 | 03/15/2024 |  | 03/15/2024 |  |  | 07/25/2022 | 03/15/2024 |  |  | 05/31/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023216868 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 38 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/25/2022 |  |  |  |  | The user facility reported a 38-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported while being placed on impella cp support, the patient developed ischemic leg on impella side. the patient had small vessels and had to do the antegrade bypass for the leg ischemia. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury" this report is being filed as part of a retrospective review of historical records.
29731 | 1220648-2024-08103 | 18918766 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18918766&pc=OZD | 07/27/2022 | 03/18/2025 |  | 03/15/2024 |  |  | 07/27/2022 | 03/15/2024 |  |  | 03/31/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023206570 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 24 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/27/2022 |  |  |  |  | The user facility reported a 24-year-old female patient on an impella cp for mechanical circulatory support. while on support the patient had some bleeding at the access site that was stopped by manual compression. the patient also developed thrombus. the impella cp was removed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29732 | 1220648-2024-08104 | 18918768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18918768&pc=OZD | 07/27/2022 | 09/12/2024 |  | 03/15/2024 |  |  | 07/27/2022 | 03/15/2024 |  |  | 04/30/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023216056 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Suction Problem | [6 empty] | Vascular Dissection | Thrombosis/Thrombus | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/27/2022 |  |  |  |  | The user facility reported a 62-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. us complainant reported that while the patient was on impella cp support, there was clot ingestion at the pump site. also, the patient experienced melena and hemoptysis and heparin was withheld. additionally, the impella had suction events and blood products were provided. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08104 was provided in sections b1, b2, b5, h1, and h6.
29733 | 1220648-2024-08105 | 18918781 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18918781&pc=OZD | 07/28/2022 | 03/15/2024 |  | 03/15/2024 |  |  | 07/28/2022 | 03/15/2024 |  |  | 01/31/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023181708 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 07/28/2022 |  |  |  |  | The user facility reported a 67-year-old patient needing mechanical circulatory support. it was reported that while the patient was on the impella cp experienced limb ischemia which improved by to inserting an antegrade sheath. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
29734 | 1220648-2024-08106 | 18918783 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18918783&pc=OZD | 08/02/2022 | 09/26/2024 |  | 03/15/2024 |  |  | 08/02/2022 | 03/15/2024 |  |  | 03/31/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023198194 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 45 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/02/2022 |  |  |  |  | The user facility reported a 45-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was in hemodynamic stable condition after the pump's placement. cerebral bleeding was found in cct scan and anticoagulation and impella therapy was stopped. the device was explanted. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08106 was provided in sections b2, b5, h1, and h6.
29735 | 1220648-2024-08107 | 18918786 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18918786&pc=OZD | 08/08/2022 | 09/20/2024 |  | 03/15/2024 |  |  | 08/08/2022 | 03/15/2024 |  |  | 04/30/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2023220016 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/08/2022 |  |  |  |  | The user facility reported a 42-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient was placed on impella cp support after the patient became tachycardiac and decompensated while a central line was inserted into the patient. the patient developed a brain bleed after having a stroke earlier in the week. family withdrew care. | It was reported the patient expired after care was withdrawn. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08107 was provided in sections b2, b5, d6a, d6b, h1, and h6.
29736 | 1220648-2024-08108 | 18918789 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18918789&pc=OZD | 08/18/2022 | 02/18/2025 |  | 03/15/2024 |  |  | 08/18/2022 | 03/15/2024 |  |  | 04/30/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023220054 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Positioning Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/18/2022 |  |  |  |  | The user facility reported a 85-year-old female noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. during installation of the pump the impella positioning was suboptimal and repositioning was attempted. suction alarms continued to persist above p2 while coming off the pump and attempting to fill the left ventricle; however, continual bleeding around the posterior portion of the left ventricle as well as a large left ventricle rupture were noted. due to the bleeding not ceasing, the device was explanted. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08108 was provided in sections b2, b5, h1, and h6. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29737 | 1220648-2024-08109 | 18918800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18918800&pc=OZD | 08/10/2022 | 03/18/2025 |  | 03/15/2024 |  |  | 08/10/2022 | 03/15/2024 |  |  | 01/31/2024 |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2023189831 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/10/2022 |  |  |  |  | The user facility reported a 63-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported patient was placed on impella cp support after removal of first pump. impella had to be repositioned and the patient experienced hemolysis with very dark urine. impella was exchanged due to low flows and the hemolysis resolved. patient experienced multiple codes prior to exchange. the pump was explanted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29738 | MW5152900 | 18920563 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18920563&pc=OZD | 03/12/2024 | 03/13/2024 |  | 03/15/2024 |  |  |  | 03/15/2024 |  |  |  |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | ABIOMED IMPELLA |  |  | OZD | [5 empty] | ABIOMED, INC. | [8 empty] | Yes | [9 empty] | Voluntary report | Not Available | Other | [6 empty] | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | [7 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | [9 empty] | Pt in ccu2 was on (b)(4) when it had a controller error at 11:44:37 and resolved at 11:44:54. decision was made to switch to a backup console. pt was switched from (b)(4). (b)(4)encountered an immediate purge failure where purge motor was not running. a new cassette was inserted (b)(6) without resolution of purge failure. switched back to (b)(4) until a new backup console was retrieved. pt then switched to (b)(4) without issue. abiomed ¿ please pull the logs from (b)(4) to determine why the controller error occurred. and let us know when we can expect the loaners to arrive. reference report: mw5152901.
29739 | MW5152901 | 18920589 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18920589&pc=OZD | 03/12/2024 | 03/13/2024 |  | 03/15/2024 |  |  |  | 03/15/2024 |  |  |  |  | 03/15/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | ABIOMED IMPELLA |  |  | OZD | [5 empty] | ABIOMED, INC. | [8 empty] | Yes | [9 empty] | Voluntary report | Not Available | Other | [6 empty] | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | [7 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | [9 empty] | Pt in ccu2 was on (b)(4) when it had a controller error at 11:44:37 and resolved at 11:44:54. decision was made to switch to a backup console. pt was switched from (b)(4). (b)(4)encountered an immediate purge failure where purge motor was not running. a new cassette was inserted (b)(6) without resolution of purge failure. switched back to (b)(4) until a new backup console was retrieved. pt then switched to (b)(4) without issue. abiomed ¿please pull the logs from (b)(4) to determine why the controller error occurred. and let us know when we can expect the loaners to arrive. reference report: mw5152900.
29740 | 1220648-2024-07962 | 18900086 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18900086&pc=OZD | 10/25/2022 | 03/18/2025 |  | 03/14/2024 |  |  | 10/25/2022 | 03/14/2024 |  |  | 06/30/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023247959 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/25/2022 |  |  |  |  | The user facility reported a 64-year-old male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. while on support, the patient experienced recurrence of ventricular tachycardia and hematuria. the impella was successfully weaned and removed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Analysis summary: cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. timing of the arrythmia event in relation to impella support (during or post-implant). electrocardiogram (ekg) recordings of the arrhythmia episodes. evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. response to any antiarrhythmic treatments or interventions during the event. any documented device-related issues or complications during impella support. evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. review of any concomitant medications that may influence cardiac electrophysiology. with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. this report is being filed as part of a retrospective review of historical records.
29741 | 1220648-2024-07963 | 18901661 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18901661&pc=OZD | 02/21/2024 | 05/02/2024 |  | 03/14/2024 |  |  | 02/27/2024 | 03/14/2024 |  |  | 10/31/2025 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024415147 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Difficult to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/27/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was being implanted with an impella cp device for mechanical circulatory support. the complainant reported the 0.018 guidewire was unable to be inserted into the easy guide lumen during the priming phase of the impella cp. there was no reported patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. impella device was returned and investigated. the red lumen was still in the pump when returned. the red lumen seemed stretched from 2 areas narrowing the passage for the 0.018" guidewire. when tried to pass guidewire during investigation, there was a lot of friction and the wire had to pushed and manipulated to pass through. the root cause of the delivery issue was a red lumen issue resulting in the red lumen interference issue resulting in the red lumen to stretch upon gw insertion. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
29742 | 1220648-2024-07964 | 18901721 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18901721&pc=OZD | 02/21/2024 | 04/01/2024 |  | 03/14/2024 |  |  | 02/27/2024 | 03/14/2024 |  |  | 11/30/2025 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024431433 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Leak/Splash | [6 empty] | Stroke/CVA | Embolism/Embolus | [11 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/27/2024 |  |  |  |  | The user facility reported a 52-year-old female was implanted with an impella 5.5 as a bridge to a transplant. it was reported that roughly seven days into impella 5.5 support, suspected bubbles were seen via echocardiogram in the left ventricle. the same issue was noted six days later, and it was further reported at that time that the abnormality was only seen post impella implant. a slow purge leak was discovered the following day at the luer lock, but both purge pressure and purge flows remained unaffected. support was maintained with the pump until planned explant following completion of the scheduled heart transplant. one day following explant, the patient was unable to move their right side. a computed tomography scan showed that the patient had suffered a basilar stroke. a thrombectomy was scheduled to treat the condition. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation into the stroke/cva, the embolism, and the purge leak ¿ pump has been completed since the initial report was submitted. the device was returned, visually inspected, and tested. the purge leak did not reproduce and there were no cracks noted. the root cause of the embolism was not determined since there were no abnormalities on the returned product and no purge system leaks were observed. the root cause of the strokes/cva could not be determined as clinical information related to the event was not provided.
29743 | 1220648-2024-07965 | 18901781 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18901781&pc=OZD | 02/28/2024 | 04/29/2024 |  | 03/14/2024 |  |  | 02/28/2024 | 03/14/2024 |  |  |  |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025442421 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/28/2024 |  |  |  |  | The user facility reported a 60-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the complainant reported during support, the impella cp stopped. the left ventricle curve was noted to be upwards prior to stop. the pump was attempted to be restarted without success. the pump was explanted. there was no reported patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation for the reported pump stop has been completed. no data logs were returned. the impella pump was returned for evaluation. upon inspection of the product, there was a large impeller purge gap observed. the root cause of the pump stop was determined to be bearing wear. the pump stop occurred on day 10, beyond the 8 day impella cp indication for use. added- d6a/d6b in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
29744 | 1220648-2024-07966 | 18901812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18901812&pc=OZD | 08/18/2022 | 03/10/2025 |  | 03/14/2024 |  |  | 08/18/2022 | 03/14/2024 |  |  | 05/31/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023226171 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/18/2022 |  |  |  |  | The user facility reported a 79-year-old male noted as high risk percutaneous cardiac insertion had difficulties with the impella cp passing through the torturous femoral artery due to a kink in the introducer sheath. the pump was placed in normal fashion and started on auto with normal flows. the long 14fr introducer sheath was removed and replaced with the repositioning (shorter) sheath and the pump was implanted successfully. angle matching with forward sutures were placed. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported difficulty inserting/removing introducer issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
29745 | 1220648-2024-07971 | 18901957 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18901957&pc=OZD | 02/28/2024 | 03/22/2024 |  | 03/14/2024 |  |  | 02/29/2024 | 03/14/2024 |  |  |  |  | 03/14/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1540253 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Defective Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/29/2024 |  |  |  |  | The complainant reported the automated impella controller (aic) displayed a purge-disc failure. the aic was exchanged. there was no reported patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the console was returned and tested after arriving in fs. the issue with detecting the purge disc was reproduced, and the purge flag was confirmed to be broken. the root cause of low flow was a kinked cannula. the root cause of this issue was a broken/missing purge flag.
29746 | 1220648-2024-07972 | 18902008 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18902008&pc=OZD | 09/02/2022 | 03/18/2025 |  | 03/14/2024 |  |  | 09/02/2022 | 03/14/2024 |  |  | 05/31/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023219563 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  |  |  | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/02/2022 |  |  |  |  | The user facility reported a 80-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that there were several ps low alarms. pump was repositioning by transesophageal echocardiography (tee). additionally, it was reported that the patient had signs of hemolysis as reported by laboratory results. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29747 | 1220648-2024-07973 | 18902055 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18902055&pc=OZD | 02/29/2024 | 05/02/2024 |  | 03/14/2024 |  |  | 02/29/2024 | 03/14/2024 |  |  | 10/31/2025 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024420271 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/29/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the complainant reported that during impella cp support, "mc too high", "retrograde flow", and "pump stop" alarms occurred. the patient's activated clotting time was 250. the physician was noted to have deactivated the retrograde flow and the impella was restarted again. after a day of support, the pump was successfully explanted. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation for the reported pump stop has been completed. data logs were not returned for evaluation. however, without the device nor sufficient clinical information for evaluation, the exact root cause of the pump stop could not be determined. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ to conform to updated procedures, section d6 & h10 were revised with updated information.
29748 | 1220648-2024-07974 | 18902112 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18902112&pc=OZD | 02/29/2024 | 05/29/2024 |  | 03/14/2024 |  |  | 02/29/2024 | 03/14/2024 |  |  | 11/30/2025 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024425742 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/29/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the complainant reported that during impella cp support, a large hematoma formed and spread to the thigh. the patient also did not have palpable foot pulse, and multiple suction alarms were noted. volume was administered. the patient ultimately ended up expiring due to multi-organ failure. | There is not enough information to associate use of device with outcome. |  |  |  |  | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation into the access site bleeding ¿ major and the limb ischemia - irreversible issues and the unrelated death has been completed since the original report was submitted. the cause of the vascular damage issue was not determined since product was not returned and sufficient clinical information was not provided. as all the necessary information and returned product were not provided, the root cause of the limb ischemia was not determined. the death was confirmed by the physician to be unrelated to the impella device. b5, d6a, d6b.
29749 | 1220648-2024-07975 | 18902224 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18902224&pc=OZD | 02/29/2024 | 05/02/2024 |  | 03/14/2024 |  |  | 02/29/2024 | 03/14/2024 |  |  |  |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025438412 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | Material Puncture/Hole | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/29/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the complainant reported that during impella cp support, a few drops of blood were noted to be leaking from the red impella plug. there were no error messages displayed. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the product damage (hole in catheter) has been completed since the original report was filed. the pump was returned and evaluated. upon inspection of the product, a hole was found at about the 80cm mark on the catheter and blood was in the red handle. the root cause of the product damage (hole in the catheter) was determined to be use related as a puncture was found in the catheter. to conform to updated procedures, section d6 was revised with updated information.
29750 | 1220648-2024-07976 | 18902297 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18902297&pc=OZD | 02/29/2024 | 04/04/2024 |  | 03/14/2024 |  |  | 02/29/2024 | 03/14/2024 |  |  |  |  | 03/14/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1424423 |  | 0042-0000-EU |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  | White | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/29/2024 |  |  |  |  | The user facility reported during impella 5.5 support for 60-year-old male patient, the automated impella controller (aic) showed a red alarm. the issue was attempted to be solved but the aic was ultimately exchanged without issue. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - controller failure (red alarm) has been completed since the initial report was submitted. the console was returned, visually inspected, and tested. the reported issue with this console¿s inability to detect the purge disc was able to be reproduced. after visual inspection of the console, it was discovered that there was a significant crack on the left side of the retainer. the root cause of the issue was a cracked purge disc retainer.
29751 | 1220648-2024-07977 | 18902329 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18902329&pc=OZD | 08/19/2022 | 03/10/2025 |  | 03/14/2024 |  |  | 08/19/2022 | 03/14/2024 |  |  | 04/30/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023216594 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/19/2022 |  |  |  |  | The user facility reported a 40-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that there were reduced pedal pulses in both lower limbs while the partial thromboplastin time was being adjusted as needed. inotropes/pressors were continued to be reduced which caused the pedal pulses to reduce in the bilateral lower extremities. the next morning absent pedal pulses in the bilateral lower extremities were warmer to the touch and the left and right ventricular assist devices were placed. the device was successfully removed. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported limb ischemia issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
29752 | 1220648-2024-07978 | 18902411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18902411&pc=OZD | 08/19/2022 | 03/10/2025 |  | 03/14/2024 |  |  | 08/19/2022 | 03/14/2024 |  |  | 03/31/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023205656 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/19/2022 |  |  |  |  | The user facility reported a 78-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that a patient began to show signs of hemolysis during impella cp support. due to the noted condition, the decision was made to explant the pump. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: potential adverse events: acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported hemolysis issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
29753 | 1220648-2024-07979 | 18902484 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18902484&pc=OZD | 10/28/2022 | 03/14/2024 |  | 03/14/2024 |  |  | 10/28/2022 | 03/14/2024 |  |  | 07/31/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023256570 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/28/2022 |  |  |  |  | The user facility reported a 53-year-old female was implanted with an impella cp device for mechanical circulatory support. the patient was having to have blood replaced due to blood loss and a echocardiogram (echo) was performed to reposition the device. the source of blood loss was not specified. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records.
29754 | 1220648-2024-07980 | 18902507 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18902507&pc=OZD | 03/01/2024 | 04/30/2024 |  | 03/14/2024 |  |  | 03/01/2024 | 03/14/2024 |  |  |  |  | 03/14/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1228186 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/01/2024 |  |  |  |  | The user facility reported the physician wanted to exchange the automated impella controller (aic) during impella cp support for patient of unknown age and gender, as the selector knob was defective. the aic was exchanged without issue and there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console (aic) knob issues has been completed. the aic was returned for evaluation. functional testing was performed with the rotatory push knob and found the menu options were skipped, which made the user report as a defective selector know (menu_options_skipped_fm_reproduced.zip). however, no issue was found in the rotary encoder (push button). the issue was resolved upon replacing the impellatronic board with the known good board. data logs were reviewed and found no explicit evidence to confirm the reported failure mode. the cause of the aic issue was a defective impellatronics board. added- d9 device return date in accordance with updated procedures, sections d6, d8, d9, g3, and h10 have been revised from the initial submission.
29755 | 1220648-2024-07982 | 18902544 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18902544&pc=OZD | 10/27/2022 | 03/18/2025 |  | 03/14/2024 |  |  | 10/27/2022 | 03/14/2024 |  |  | 07/31/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023260007 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/27/2022 |  |  |  |  | The user facility reported a 74-year-old male with cardiomyopathy and myocarditis was implanted with an impella cp for mechanical circulatory support. the medical providers were unable to assess a distal pulse in both the patient's the right and left lower extremities. the patient also experienced a gi bleed and the flexiseal output remained dark. the patient received 3 units of packed red blood cells. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29756 | 1220648-2024-07983 | 18902582 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18902582&pc=OZD | 10/26/2022 | 03/11/2025 |  | 03/14/2024 |  |  | 10/26/2022 | 03/14/2024 |  |  | 08/31/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023257628 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/26/2022 |  |  |  |  | The user facility reported a 81-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced a large arterial bleed while on impella cp support. a mattress suture and a fem-stop was placed to resolve the bleeding. it was also reported the right foot was purple and mottled but had dopplerable pulse. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29757 | 1220648-2024-07984 | 18902609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18902609&pc=OZD | 11/01/2022 | 03/18/2025 |  | 03/14/2024 |  |  | 11/01/2022 | 03/14/2024 |  |  | 04/30/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023224130 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/01/2022 |  |  |  |  | The user facility reported a 70-year-old female was implanted with an impella cp device for mechanical circulatory support. the patient developed a hemorrhagic stroke requiring rapid impella cp wean and removal. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Strokes/cvas can occur during or after impella support, and can be either ischemic or hemorrhagic. to make a determination as to the cause of the stroke, the following clinical information must be considered: patient's medical history, including any previous strokes, cardiovascular disease, or risk factors such as hypertension, diabetes, or atrial fibrillation. timing of the stroke in relation to impella support (during or post-implant). -detailed description of the symptoms experienced by the patient. neurological examination findings and any focal deficits observed. diagnostic imaging results, such as ct scan or mri of the brain, to identify the type and location of the stroke (ischemic or hemorrhagic). laboratory test results, including coagulation profiles and cardiac markers. any relevant procedural or device-related factors, such as duration and settings of impella support, presence of embolic protection devices, or any documented procedural complications. response to any previous anticoagulation or antiplatelet therapies. evaluation of potential alternative causes of stroke, such as arrhythmias, embolic sources, or underlying vascular pathology. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. this report is being filed as part of a retrospective review of historical records.
29758 | 1220648-2024-07985 | 18902690 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18902690&pc=OZD | 03/02/2024 | 04/08/2024 |  | 03/14/2024 |  |  | 03/02/2024 | 03/14/2024 |  |  | 11/30/2025 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024435584 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/02/2024 |  |  |  |  | The user facility reported a 60-year-old male patient of unknown race and ethnicity noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that the implanted impella cp came out of the left ventricle. the impella cp was turned off, removed, re-inserted and started. once again, the impella cp came out of the left ventricle. the impella cp was ultimately explanted due to being unable to advance back into position. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed since the initial report was submitted. the device was not returned for investigation based on clinical description & data analysis; the root cause of positioning issue was most likely use related.
29759 | 1220648-2024-07986 | 18902709 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18902709&pc=OZD | 10/30/2022 | 03/18/2025 |  | 03/14/2024 |  |  | 10/30/2022 | 03/14/2024 |  |  | 06/30/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235781 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/30/2022 |  |  |  |  | The user facility reported a 60-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding at the impella cp access site and perfusion concerns. the impella cp was removed after increasing dobutamine and levophed. two perclose was used for hemostasis and the patient¿s groin was soft and without oozing by end of procedure. after impella was removed, perfusion was restored, under fluoroscopy. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
29760 | 1220648-2024-07987 | 18902763 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18902763&pc=OZD | 09/03/2022 | 03/31/2025 |  | 03/14/2024 |  |  | 09/03/2022 | 03/14/2024 |  |  | 03/31/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023206445 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/03/2022 |  |  |  |  | The user facility reported a 67-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported during insertion procedure, the physician went through the artery three times before successfully placing the impella device. this resulted in a large hematoma that required surgical repair. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | The investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29761 | 1220648-2024-07988 | 18902770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18902770&pc=OZD | 10/28/2022 | 03/10/2025 |  | 03/14/2024 |  |  | 10/28/2022 | 03/14/2024 |  |  | 07/31/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023259244 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/28/2022 |  |  |  |  | The user facility reported a 57-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported placement signal low advisory on the impella cp while a patient was on support and were unable to obtain non-invasive blood pressure (bp) on patient. bilateral lower extremity (b/l le) mottled and unable to obtain doppler signals. patient developed bloody stools and partial thromboplastin time (ptt) found to be supratherapeutic. the site was redressed. clean, dry, and intact (c/d/l) with no bleeding noted after dressing change and first therapeutic ptt with redraw was completed. bloody stools were resolved. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29762 | 1220648-2024-07990 | 18902838 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18902838&pc=OZD | 03/02/2024 | 04/18/2024 |  | 03/14/2024 |  |  | 03/02/2024 | 03/14/2024 |  |  | 02/28/2025 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024339410 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/02/2024 |  |  |  |  | The user facility reported a 71-year-old female patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed an access site hematoma, that required a washout to be performed and a drain to be placed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis.
29763 | 1220648-2024-07989 | 18902840 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18902840&pc=OZD | 11/01/2022 | 03/17/2025 |  | 03/14/2024 |  |  | 11/01/2022 | 03/14/2024 |  |  | 06/30/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023236540 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/01/2022 |  |  |  |  | The user facility reported a 75-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support the patient went into ventricular tachycardia requiring defibrillation. impella in aorta alarm presented after the patient arrived in the intensive care unit, so the pump was repositioned. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29764 | 1220648-2024-07994 | 18903001 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18903001&pc=OZD | 03/02/2024 | 04/19/2024 |  | 03/14/2024 |  |  | 03/02/2024 | 03/14/2024 |  |  | 09/30/2025 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025445429 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/02/2024 |  |  |  |  | The user facility reported a 75-year-old female in st-segment elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that after 42 hours, the pump was explanted due to limb ischemia. lost pulses at the impella accessed limb were observed. | Additional information states that the patient had hemolysis and the impella was repositioned. it is unknown if the repositioning resolved the hemolysis. the hemolysis occurred prior to the reported limb ischemia. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury". | The investigation for the reported hemolysis and limb ischemia has been completed. the impella device was not returned for evaluation. data logs were reviewed the data logs were retuned and there were suction alarms around the time of occurrence of the issue. there were no placement signal alarms and no motor current spike that were noticed. the root cause of the hemolysis was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. the instructions for use states ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected low s or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿
29765 | 1220648-2024-07995 | 18903008 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18903008&pc=OZD | 10/25/2022 | 03/18/2025 |  | 03/14/2024 |  |  | 10/25/2022 | 03/14/2024 |  |  | 07/31/2023 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022110791 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/25/2022 |  |  |  |  | The user facility reported a 78-year-old female in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella support, the patient developed an ischemic leg. an external bypass from the right radial to the right femoral was performed. the patient did not wish any further treatment and requested the impella be removed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | Limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition . ¿ any concomitant medication . ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient ¿ physical examination findings, including pulse assessment and visual examination of the affected limb . ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography . ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders . ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension . ¿ patient's response to previous treatments or interventions for vascular conditions . with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
29766 | 1220648-2024-07996 | 18903109 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18903109&pc=OZD | 10/26/2022 | 03/18/2025 |  | 03/14/2024 |  |  | 10/26/2022 | 03/14/2024 |  |  | 05/31/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231025 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/26/2022 |  |  |  |  | The user facility reported a 61-year-old female had severely calcified triple vessel disease was implanted with an impella cp device for mechanical circulatory support. the decision was made to remove the impella cp due to no flow detected down the patient's leg. the impella was successfully weaned and removed. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
29767 | 1220648-2024-07997 | 18903155 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18903155&pc=OZD | 12/02/2022 | 03/18/2025 |  | 03/14/2024 |  |  | 12/02/2022 | 03/14/2024 |  |  | 08/31/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023260473 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/02/2022 |  |  |  |  | The user facility reported a (b)(6) male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed bleeding/oozing at the impella cp insertion site. bicarb was in the purge, no anticoagulants were provided and manual pressure was held. additionally, the perclose was tightened and the oozing stopped. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
29768 | 1220648-2024-07998 | 18903227 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18903227&pc=OZD | 12/04/2022 | 03/18/2025 |  | 03/14/2024 |  |  | 12/04/2022 | 03/14/2024 |  |  | 07/31/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023254978 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 24 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/04/2022 |  |  |  |  | The user facility reported a 24-year-old female was implanted with an impella cp. the patient bled through three impella site dressings. the patient subsequently received 1 unit of platelets. the physician placed another suture to the blue suture wings which resolved the bleeding. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
29769 | 1220648-2024-07999 | 18903259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18903259&pc=OZD | 10/26/2022 | 04/15/2025 |  | 03/14/2024 |  |  | 10/26/2022 | 03/14/2024 |  |  | 07/31/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023255937 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/26/2022 |  |  |  |  | The user facility reported a 88-year-old male was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding issues and the impella cp was explanted. proglide and an angio-seal vascular closure device was used and hemostasis was achieved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29770 | 1220648-2024-08000 | 18903454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18903454&pc=OZD | 10/26/2022 | 03/18/2025 |  | 03/14/2024 |  |  | 10/26/2022 | 03/14/2024 |  |  | 07/31/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023251508 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/26/2022 |  |  |  |  | The user facility reported a 65-year-old female was implanted with an impella cp device for mechanical circulatory support after a motor vehicle accident. during support, a suction alarm occurred and the pump was repositioned to resolve the issue. also, the patient lost pulse in distal extremity and the decision was made to explant the device. the patient was successfully weaned from the impella cp device. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29771 | 1220648-2024-08001 | 18903646 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18903646&pc=OZD | 12/07/2022 | 03/18/2025 |  | 03/14/2024 |  |  | 12/07/2022 | 03/14/2024 |  |  | 12/31/2023 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023171620 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/07/2022 |  |  |  |  | The user facility reported a 67-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during insertion of the impella cp, the patient went into ventricular fibrillation arrest and required defibrillation. return of spontaneous circulation was obtained. the patient was noted to have a thrombus as well. the patient was able to be weaned off levophed after opening vessel. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29772 | 1220648-2024-08002 | 18903658 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18903658&pc=OZD | 10/28/2022 | 03/14/2024 |  | 03/14/2024 |  |  | 10/28/2022 | 03/14/2024 |  |  | 06/30/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023237565 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/28/2022 |  |  |  |  | The user facility reported a 83-year-old male was implanted with an impella device for mechanical circulatory support for a high risk percutaneous coronary intervention after experiencing a substernal chest pain. the patient developed bleeding at the impella access site. heparin was withdrawn and the patient received two (2) units of packed red blood cells (prbc). after the impella site was redressed, no bleeding or hematoma was noted from the impella site. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records.
29773 | 1220648-2024-08003 | 18903670 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18903670&pc=OZD | 06/16/2022 | 03/18/2025 |  | 03/14/2024 |  |  | 06/16/2022 | 03/14/2024 |  |  | 01/31/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023187822 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/16/2022 |  |  |  |  | The user facility reported a 56-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed oozing/bleeding at the impella cp insertion site (groin). bicarbonate was added to the purge to help control bleeding concerns and a systemic heparin drip was added. the bleeding subsided and the site re-sutured. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29774 | 1220648-2024-08004 | 18903732 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18903732&pc=OZD | 10/27/2022 | 03/18/2025 |  | 03/14/2024 |  |  | 10/27/2022 | 03/14/2024 |  |  | 07/31/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023255574 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/27/2022 |  |  |  |  | The user facility reported a 73-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while patient was on impella cp support, the patient's leg lost color and become mottled in appearance. the impella was reposition and distal perfusion cannulation was performed for the left leg. although the color improved, the skin remained mottled and capillary refill was less than 10 seconds. the decision was made to explant the impella cp to resolve. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29775 | 1220648-2024-08005 | 18903784 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18903784&pc=OZD | 08/13/2022 | 03/14/2024 |  | 03/14/2024 |  |  | 08/13/2022 | 03/14/2024 |  |  | 05/31/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023227918 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/13/2022 |  |  |  |  | The user facility reported a 75-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient lost pulses and the limbs were cooler than previous assessments. a decision was made to wean the impella. immediately after removal, pulses came back and the foot was warmer to touch. | [5 empty] | The impella device was discarded by the customer, therefore, evaluation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. this report is being filed as part of a retrospective review of historical records.
29776 | 1220648-2024-08006 | 18903809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18903809&pc=OZD | 08/19/2022 | 03/18/2025 |  | 03/14/2024 |  |  | 08/19/2022 | 03/14/2024 |  |  | 04/30/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023225099 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/19/2022 |  |  |  |  | The user facility reported a 65-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. upon receiving the pump, the next day the physician accidentally explanted the device during repositioning. there was continuous repositioning with the pump; however, there was no success in re-crossing the valve and the impella was explanted. it was advised by the abiomed representative not to have the device cleaned for reuse; however, the physician proceeded to have the device cleaned and prep for re-implantation. after cleaning the pump, the physician noted there was a clot within the inflow and cannula and was advised to utilize a new impella. no issues explanting or implanting new device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29777 | 1220648-2024-08007 | 18903856 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18903856&pc=OZD | 08/25/2022 | 03/18/2025 |  | 03/14/2024 |  |  | 08/25/2022 | 03/14/2024 |  |  | 01/31/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023184724 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/25/2022 |  |  |  |  | The user facility reported a 51-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella stopped. a decision was made to explant the pump. no further details were provided. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29778 | 1220648-2024-08008 | 18903934 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18903934&pc=OZD | 08/30/2022 | 05/23/2025 |  | 03/14/2024 |  |  | 08/30/2022 | 03/14/2024 |  |  | 05/31/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231365 |  | 0048-445 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemolysis | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/30/2022 |  |  |  |  | The user facility reported a 71-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was having signs of hemolysis, such as hemolysed labs and red urine. it was decided to reduce pressure level and mean arterial pressure was stable at pressure level 4. additionally, the patient was having myoclonic jerking. neurological consult and computer tomography were ordered. it was determined that the patient had an anoxic injury leading to brain death. family withdrew care. | Additional information noted care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08008 was provided. | No product was returned. the root cause of the limb ischemic reversible issue was most likely patient condition related. this report is being filed as part of a retrospective review of historical records.
29779 | 1220648-2024-08009 | 18904002 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18904002&pc=OZD | 02/09/2024 | 04/01/2024 |  | 03/14/2024 |  |  | 02/16/2024 | 03/14/2024 |  |  |  |  | 03/14/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1404666 |  | 0048-0024 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 35 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/16/2024 |  |  |  |  | The user facility reported a 35-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during repositioning, the displayed impella flow dropped to 0l/min. the issue persisted for roughly one to two minutes until the system was successfully restarted. patient support was maintained at the time of the noted issue with the assistance of ecmo (extracorporeal membrane oxygenation). upon return to functionality, no further issues were noted for the remainder of support. there was no patient harm. | [5 empty] | The impella device was received from the customer on 14-mar-2024 and therefore, an evaluation of the device is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console (aic) pump stop has been completed. the aic was returned and found that there were two unexpected reboots due to calculator process crashes during functional testing. data logs were reviewed finding that the console had a reboot triggered by the calculator watchdog. after two heartbeat checks that confirmed that the calc process had crashed, a partial restart is initiated. the console reboots and initially has no calculator issues. the cause of the unexpected reboot was the calculator crashing but the failure was not reproduced so the cause of the calculator crash is undetermined. the instructions for use states ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
29780 | 1220648-2024-08010 | 18904030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18904030&pc=OZD | 10/29/2022 | 05/19/2025 |  | 03/14/2024 |  |  | 10/29/2022 | 03/14/2024 |  |  | 07/31/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023258929 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/29/2022 |  |  |  |  | The user facility reported a 59-year-old male was implanted with an impella cp device for mechanical circulatory support emergently prior to stent procedure. due to oozing around the sheath, a femostop was placed. the patient's pulses were doppler with cool extremity and dusky toes, and the patient showed signs of hemolysis. the physician decided to explant the impella cp device and heparin was stopped. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29781 | 1220648-2024-08011 | 18904076 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18904076&pc=OZD | 08/31/2022 | 03/18/2025 |  | 03/14/2024 |  |  | 08/31/2022 | 03/14/2024 |  |  | 05/31/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231130 |  | 0048-445 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/31/2022 |  |  |  |  | The user facility reported a 41-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient exhibited signs of hemolysis, but the patient¿s urine was clearing up. the physician and nurse were advised to optimize impella flows and reduce hemolysis with a mean arterial pressure between 60-80. the patient stablized and subsequently the impella device was weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29782 | 1220648-2024-08012 | 18904313 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18904313&pc=OZD | 11/01/2022 | 03/18/2025 |  | 03/14/2024 |  |  | 11/01/2022 | 03/14/2024 |  |  | 07/31/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023254593 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/01/2022 |  |  |  |  | The user facility reported a 56-year-old male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the complainant reported the patient was on impella cp support and the patient experienced a loss of pulse in bilateral lower extremities. the patient later expired while on the impella cp support. no allegations were made towards the impella cp for cause of death | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
29783 | 1220648-2024-08013 | 18904421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18904421&pc=OZD | 10/29/2022 | 03/18/2025 |  | 03/14/2024 |  |  | 10/29/2022 | 03/14/2024 |  |  | 07/31/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023252272 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/29/2022 |  |  |  |  | The user facility reported a 35-year-old male was implanted with an impella cp device for mechanical circulatory support to vent the left ventricle (lv). the patient was found down in a park and emergency medical services (ems) performed advanced cardiac life support (acls) en route to the emergency department. multiple hemorrhagic processes found on the patient's head computed tomography (ct). | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Strokes/cvas can occur during or after impella support, and can be either ischemic or hemorrhagic. to make a determination as to the cause of the stroke, the following clinical information must be considered: patient's medical history, including any previous strokes, cardiovascular disease, or risk factors such as hypertension, diabetes, or atrial fibrillation. timing of the stroke in relation to impella support (during or post-implant). detailed description of the symptoms experienced by the patient. neurological examination findings and any focal deficits observed. diagnostic imaging results, such as ct scan or mri of the brain, to identify the type and location of the stroke (ischemic or hemorrhagic). laboratory test results, including coagulation profiles and cardiac markers any relevant procedural or device-related factors, such as duration and settings of impella support, presence of embolic protection devices, or any documented procedural complications. response to any previous anticoagulation or antiplatelet therapies. evaluation of potential alternative causes of stroke, such as arrhythmias, embolic sources, or underlying vascular pathology. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. this complaint has been identified as part of a retrospective review.
29784 | 1220648-2024-08014 | 18904507 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18904507&pc=OZD | 08/31/2022 | 03/27/2024 |  | 03/14/2024 |  |  | 08/31/2022 | 03/14/2024 |  |  | 04/30/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023224007 |  | 0048-445 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/31/2022 |  |  |  |  | The user facility reported a 72-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. after the impella was weaned and explanted, manta vascular closure device was deployed but bleeding was observed on angiography. a balloon was utilized to tamponade which resolved bleeding. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleeding has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the access site bleeding was most likely due to anticoagulation management since the activated clotting time noted was 309 which is above the range limit. this report is being filed as part of a retrospective review of historical records.
29785 | 1220648-2024-08015 | 18904647 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18904647&pc=OZD | 10/30/2022 | 03/14/2024 |  | 03/14/2024 |  |  | 10/30/2022 | 03/14/2024 |  |  | 07/31/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023256630 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/30/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 83-year-old female patient for mechanical circulatory support and the patient developed oozing and hematoma at the impella access site. as a result of the noted issue, sutures were applied, the impella was weaned, and a right heart catheterization (rhc) was performed. additionally, the impella cp was explanted. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29786 | 1220648-2024-08016 | 18904765 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18904765&pc=OZD | 09/07/2022 | 03/18/2025 |  | 03/14/2024 |  |  | 09/07/2022 | 03/14/2024 |  |  | 06/30/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023229762 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/07/2022 |  |  |  |  | The user facility reported a 53-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that during impella cp support, there was inadequate leg perfusion on impella access site leading to pump explant. patient was noted to be stable after the explant procedure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition. any concomitant medication. vascular size or variations thereof. detailed description of the symptoms experienced by the patient. physical examination findings, including pulse assessment and visual examination of the affected limb. diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. laboratory test results, including blood tests for markers of inflammation or clotting disorders. any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
29787 | 1220648-2024-08017 | 18904889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18904889&pc=OZD | 09/07/2022 | 09/24/2024 |  | 03/14/2024 |  |  | 09/07/2022 | 03/14/2024 |  |  | 05/31/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023228667 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/07/2022 |  |  |  |  | The user facility reported a 55 year-old male patient on impella cp support had cerebral edema which was found due to hypoxic brain injury. the catecholamines were reduced. the patient expired. the injury was noted to not be attributed to the impella cp. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. the initial mfr 1220648-2024-08017 reported a patient experienced a hypoxic brain injury while on impella cp support. after review by medical safety, it is noted the hypoxic brain injury occurred prior to impella insertion, and therefore the device did not cause or contribute to the patient's injury.
29788 | 1220648-2024-08018 | 18904964 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18904964&pc=OZD | 01/19/2021 | 03/14/2024 |  | 03/14/2024 |  |  | 01/26/2024 | 03/14/2024 |  |  | 08/31/2022 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1496670 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/26/2024 |  |  |  |  | The user facility reported a 71-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the complainant reported in (b)(6) 2021, the patient developed hemolysis. treatment details were not provided. (b)(6) 2024, the team in japan learned that the patient was given haptoglobin to treat the hemolysis. | [5 empty] | The impella device was not received from the customer. the root cause of the hemolysis was determined to be unrelated to the use of impella as no evidence was provided to suggest a causal relationship between the suspected hemolysis and impella. no product or data was returned for this investigation. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
29789 | 1220648-2024-08019 | 18905066 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18905066&pc=OZD | 11/01/2022 | 05/21/2025 |  | 03/14/2024 |  |  | 11/01/2022 | 03/14/2024 |  |  | 08/31/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023259590 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/01/2022 |  |  |  |  | The user facility reported a 75-year-old male was implanted with an impella cp device for mechanical circulatory support. the impella was explanted, and a covered stent was placed to the right femoral artery and a new impella cp was successfully inserted in left femoral artery, however an angiogram was done and found an open extravasation. a femstop was applied, angle matched, and was redressed. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29790 | 1220648-2024-08020 | 18905311 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18905311&pc=OZD | 11/10/2022 | 04/15/2025 |  | 03/14/2024 |  |  | 11/10/2022 | 03/14/2024 |  |  | 05/31/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023223365 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/10/2022 |  |  |  |  | The user facility reported a 70-year-old male patient implanted with an impella cp for mechanical circulatory support. while on support the patient experienced bleeding problems at the right groin requiring two units of packed red blood cells. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29791 | 1220648-2024-08021 | 18905340 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18905340&pc=OZD | 07/27/2022 | 05/13/2025 |  | 03/14/2024 |  |  | 07/27/2022 | 03/14/2024 |  |  | 09/30/2023 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022143539 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/27/2022 |  |  |  |  | The user facility reported a 66-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced hemolysis. the patient received blood. the patient remained on support for one additional day. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29792 | 1220648-2024-08022 | 18905404 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18905404&pc=OZD | 11/01/2022 | 03/07/2025 |  | 03/14/2024 |  |  | 11/01/2022 | 03/14/2024 |  |  | 06/30/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023247176 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/01/2022 |  |  |  |  | The user facility reported a 73-year-old male was implanted with an impella cp device for mechanical circulatory support. the patient experienced red urine output. the impella was repositioned, and the p level was lowered, but laboratory testing reported hemolyzed samples. another echo was done, and again the impella was repositioned slightly. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
29793 | 1220648-2024-08023 | 18905466 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18905466&pc=OZD | 02/15/2022 | 03/14/2024 |  | 03/14/2024 |  |  | 01/30/2024 | 03/14/2024 |  |  | 09/30/2022 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1503409 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that 58-year-old male patient on impella cp support had limb ischemia. there were no lasting effects. additional information was received indicating that there was an intervention applied to treat. the patient was placed on bypass and was noted to have recovered. | [5 empty] | The investigation has been completed. no product was returned for investigation. the root cause of the limb ischemia was most likely patient condition. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
29794 | 1220648-2024-08024 | 18905725 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18905725&pc=OZD | 07/27/2022 | 03/14/2024 |  | 03/14/2024 |  |  | 07/27/2022 | 03/14/2024 |  |  | 03/31/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023212175 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/27/2022 |  |  |  |  | The user facility reported a 56-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support, due to non-disclosed cardiovascular medical history. while on support, the patient experienced bleeding from multiple sites and had a retroperitoneal bleed. patient received three units of blood and albumin. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records.
29795 | 1220648-2024-08025 | 18905736 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18905736&pc=OZD | 11/02/2022 | 04/08/2025 |  | 03/14/2024 |  |  | 11/02/2022 | 03/14/2024 |  |  | 06/30/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023234973 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Thrombosis/Thrombus | [11 empty] | Other | Required Intervention | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/02/2022 |  |  |  |  | The user facility reported a 56-year-old male was implanted with an impella cp device for mechanical circulatory support. while on impella cp support, the patient¿s kidney function declined, impella sheath likely clotted, and patient went into ventricular tachycardia (vt) arrest. family decided to discontinue care. | Additional information noted care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-08025 was provided. | Medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29796 | 1220648-2024-08027 | 18905739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18905739&pc=OZD | 10/29/2022 | 03/18/2025 |  | 03/14/2024 |  |  | 10/29/2022 | 03/14/2024 |  |  | 02/29/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200130 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/29/2022 |  |  |  |  | The user facility reported a 56-year-old female, race and was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient on impella cp had bleeding and was given bicarbonate instead of the heparin in the purge fluid and received blood products. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29797 | 1220648-2024-08026 | 18905741 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18905741&pc=OZD | 10/28/2021 | 03/14/2024 |  | 03/14/2024 |  |  | 02/05/2024 | 03/14/2024 |  |  | 07/31/2023 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2022111254 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/05/2024 |  |  |  |  | The user facility reported a 63-year-old female patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced hemolysis the entirety of the support, from insertion to removal. no treatments were reported. | [5 empty] | The pump was not returned for investigation. the cause of the hemolysis was not determined since the product was not returned. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes may also play a role in patient susceptibility to hemolysis.¿
29798 | 1220648-2024-08028 | 18905907 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18905907&pc=OZD | 08/03/2022 | 03/18/2025 |  | 03/14/2024 |  |  | 08/03/2022 | 03/14/2024 |  |  | 01/31/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023185046 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/03/2022 |  |  |  |  | The user facility reported a 77-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the patient developed ischemia to left leg. the patient was escalated to where impella 5.5 and impella cp cutdown and removed successfully. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | Limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition . ¿ any concomitant medication . ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient . ¿ physical examination findings, including pulse assessment and visual examination of the affected limb . ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography . ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders . ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension . ¿ patient's response to previous treatments or interventions for vascular conditions . with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
29799 | 1220648-2024-08029 | 18905961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18905961&pc=OZD | 02/27/2024 | 11/22/2024 |  | 03/14/2024 |  |  | 02/27/2024 | 03/14/2024 |  |  |  |  | 03/14/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1461597 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Female |  |  | Asian | Malfunction | No | Yes | Material Integrity Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/27/2024 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while the patient was on support, the automated impella controller (aic) had the pressure transmitter mounting component malfunction. the device was replaced. there was no patient harm. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the console (aic) purge issue has been completed. logs were not relevant to this failure. the console was returned for evaluation. upon visual inspection, the device was missing the purge flag. therefore the cause of the issue was a missing purge flag. corrections to the initial mfr report: d2 common device name; f6/f8 should have been left blank.
29800 | 1220648-2024-08030 | 18905977 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18905977&pc=OZD | 08/13/2022 | 03/14/2024 |  | 03/14/2024 |  |  | 08/13/2022 | 03/14/2024 |  |  | 01/31/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023186326 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/13/2022 |  |  |  |  | The user facility reported a 80-year-old male was implanted with an impella cp for percutaneous coronary intervention (pci) it was reported that the patient with low platelets developed continuous oozing at the access site of the left femoral artery. a pressure dressing was placed with angle insertion maintained. two rbc (red blood cell) units were required. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
29801 | 1220648-2024-08031 | 18905986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18905986&pc=OZD | 11/02/2022 | 03/14/2024 |  | 03/14/2024 |  |  | 11/02/2022 | 03/14/2024 |  |  | 01/31/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191797 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/02/2022 |  |  |  |  | The user facility reported a 73-year-old female in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient developed access site bleeding. the femstop was removed, and the site was angle matched to control the bleeding. additionally, the patient developed severe peripheral arterial disease of the lower extremities and required mechanical circulatory support for the device access site. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
29802 | 1220648-2024-08032 | 18906021 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18906021&pc=OZD | 03/03/2024 | 04/19/2024 |  | 03/14/2024 |  |  | 03/03/2024 | 03/14/2024 |  |  | 08/31/2025 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025441090 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/03/2024 |  |  |  |  | The user facility reported a 60-year-old male post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the patient's bilirubin tripled in value over the last three days during impella 5.5 support. due to the noted condition, the decision was made to explant the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ . ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿. | The investigation into the hemolysis issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of hemolysis was not able to be determined since product was not returned and there was limited information from data log & clinical review.
29803 | 1220648-2024-08033 | 18906027 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18906027&pc=OZD | 08/29/2022 | 03/18/2025 |  | 03/14/2024 |  |  | 08/29/2022 | 03/14/2024 |  |  | 03/31/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023210948 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/29/2022 |  |  |  |  | The user facility reported a 82-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the medical team reported positioning issues. the impella cp was repositioned with no further issues. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29804 | 1220648-2024-08034 | 18906041 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18906041&pc=OZD | 11/29/2022 | 03/14/2024 |  | 03/14/2024 |  |  | 11/29/2022 | 03/14/2024 |  |  | 09/24/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023259137 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/29/2022 |  |  |  |  | The user facility reported a 74-year-old male was implanted with an impella 5.5. the complainant reported that resistance was met while attempting to implant an impella 5.5 pump for patient support. it was documented that the pump was unable to pass the turn into the aorta during the delivery. as a result, the pump was removed and direct access was utilized to place the device. there was no patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. however, clinical information was sufficient to determine the root cause. the impella device was not returned for evaluation. the root cause of the delivery issue was due to patient condition. this report is being filed as part of a retrospective review of historical records.
29805 | 1220648-2024-08035 | 18906069 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18906069&pc=OZD | 08/14/2022 | 03/14/2024 |  | 03/14/2024 |  |  | 08/14/2022 | 03/14/2024 |  |  | 03/31/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023210858 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Remove | Insufficient Flow or Under Infusion | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/14/2022 |  |  |  |  | The user facility reported a 67-year-old male was implanted with an impella cp. during support, patient had groin bleeding issues after removal of 14fr peel away sheath. suction alarms and muted lv waveform were presented. a new cp catheter resolved suction and waveform issues. the t-b was tightened down with repositioning sheath and sutured in place. during removal of pumps the physician tried to perclose using the two perclose but was unsuccessful. then tried and 8fr angio seal which did not take. the third perclose was used but was unable to achieve hemostasis. the decision was made to place a 14fr sheath back in and call cv surgery for surgical removal and closure of arteriotomy. rp impella was successfully weaned and removed. site was successfully closed with two box sutures. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ the ifu also caution users to handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time. this report is being filed as part of a retrospective review of historical records.
29806 | 1220648-2024-08036 | 18906198 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18906198&pc=OZD | 08/16/2022 | 05/12/2025 |  | 03/14/2024 |  |  | 08/16/2022 | 03/14/2024 |  |  | 02/29/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023193824 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/16/2022 |  |  |  |  | The user facility reported a 69-year-old female patient was implanted with an impella cp for mechanical circulatory support. the impella was placed across the valve and the patient developed ventricular fibrillation (vfib). multiple shocks and cardiopulmonary resuscitation (cpr) were administered with a return of spontaneous circulation (rosc). during support, oozing was noted from all lines. sutures were placed for hemostasis. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined
29807 | 1220648-2024-08037 | 18906208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18906208&pc=OZD | 07/27/2022 | 04/16/2025 |  | 03/14/2024 |  |  | 07/27/2022 | 03/14/2024 |  |  | 04/30/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP - EU |  |  | OZD |  |  | IMPELLA CP - EU |  |  | ABIOMED, INC. |  |  | US |  | 2023218914 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 49 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/27/2022 |  |  |  |  | The user facility reported a 49-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support, due to non-disclosed cardiovascular medical history. while on support, the patient experienced groin bleeding due to very high partial thromboplastin time (ppt). heparin perfusion was paused. the patient remained on support for two additional days. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29808 | 1220648-2024-08038 | 18906225 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18906225&pc=OZD | 03/01/2024 | 04/18/2024 |  | 03/14/2024 |  |  | 03/01/2024 | 03/14/2024 |  |  |  |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025443206 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  |  | Malfunction | No | Yes | Improper Flow or Infusion | Pressure Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/01/2024 |  |  |  |  | The user facility reported a 68-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the impella 5.5 had a high purge pressure alarm and it evolved into a purge system blocked alarm. tissue plasminogen activator was started. the waveforms and motor current remained consistent, but the purge system blocked alarm persisted. the impella 5.5 was exchanged and a new 5.5 was successfully placed. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation into the high purge pressure issue has been completed since the original report was submitted. the device was not returned for investigation. data logs were investigated, and it was found that there were numerous purge pressure high and purge system blocked alarms after 10 hours till the end of the case which suggests biomaterial buildup. as per the clinical information and data log review, the root cause of the high purge pressure issue was most likely biomaterial.
29809 | 1220648-2024-08039 | 18906277 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18906277&pc=OZD | 08/16/2022 | 03/14/2024 |  | 03/14/2024 |  |  | 08/16/2022 | 03/14/2024 |  |  | 01/31/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183967 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/16/2022 |  |  |  |  | The user facility reported a 59-year-old male patient was implanted with an impella cp for mechanical circulatory support. during support it was noted that the right lower extremity (rle) mottled and pulses were not dopplerable. the patient received 1l bicarb, 500cc albumin and 1 unit of packed red blood cell (prbc) transfusion. plans were made for the vascular surgery team to re-evaluate the right leg for potential revascularization. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29810 | 1220648-2024-08040 | 18906290 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18906290&pc=OZD | 11/10/2022 | 03/14/2024 |  | 03/14/2024 |  |  | 11/10/2022 | 03/14/2024 |  |  | 07/31/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023254954 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/10/2022 |  |  |  |  | The user facility reported while the male patient was being implanted with the impella cp, the impella was unable to get past the iliac. as a result, the procedure was aborted. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issue was most likely patient condition related, the impella was unable to pass past iliac as per the clinical description provided.
29811 | 1220648-2024-08041 | 18906581 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18906581&pc=OZD | 03/01/2024 | 04/18/2024 |  | 03/14/2024 |  |  | 03/01/2024 | 03/14/2024 |  |  | 09/30/2025 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024396440 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/01/2024 |  |  |  |  | The user facility reported a 71-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the patient developed a clot on the end of the impella 5.5 catheter coinciding with a period of sub-therapeutic anticoagulation following a planned procedure. in response to the noted biomaterial, plans were made to resume anticoagulation as tolerated. support with the pump was maintained. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the access site bleeding and the thrombus issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of access site bleeding is not determined because the location of the bleed is unknown and was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined.
29812 | 1220648-2024-08042 | 18906782 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18906782&pc=OZD | 08/19/2022 | 03/10/2025 |  | 03/14/2024 |  |  | 08/19/2022 | 03/14/2024 |  |  | 04/30/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023218270 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/19/2022 |  |  |  |  | The user facility reported a 70-year-old female in post cardiotomy cardiogenic shock /low cardiac output syndrome was implanted with an impella cp device for mechanical circulatory support. it was reported that the device was repositioned and advancement of hematuria occurred overnight. the patient received 1 unit of packed red blood cells (prbcs) over night, 2 units of packed red blood cells (prbcs) the next morning and 1 unit of platelets. it was confirmed hemolysis was present and the patience's urine was pink-red. the pump was repositioned under echo guidance and the p-level decreased. the device was explanted. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events: acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported hemolysis issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
29813 | 1220648-2024-08043 | 18906881 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18906881&pc=OZD | 09/07/2022 | 03/18/2025 |  | 03/14/2024 |  |  | 09/07/2022 | 03/14/2024 |  |  | 08/31/2023 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022124680 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/07/2022 |  |  |  |  | The user facility reported a 51-year-old male patient was implanted with an impella cp for mechanical circulatory support for percutaneous coronary intervention. the patient developed heavy bleeding at the impella cp insertion site. pressure was held. the physician chose to explant the impella to exchange for a 16 french sheath to control the bleeding. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29814 | 1220648-2024-08044 | 18907119 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18907119&pc=OZD | 08/19/2022 | 03/10/2025 |  | 03/14/2024 |  |  | 08/19/2022 | 03/14/2024 |  |  | 03/31/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023204881 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/19/2022 |  |  |  |  | The user facility reported a 60-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. upon receiving the pump, the physicians were concern about the flow down the left leg and a redraw of numbers revealed pa sat of 51% and ci of 1.5. the decision was made to leave in the device and place external fem-fem bypass to perfuse down the left leg. the impella was increased up to p9 and the antegrade sheath placed in the left femoral artery and connected to a retrograde sheath in the right femoral artery for distal limb perfusion. eventually the impella cp was explanted and the impella 5.5 was implanted successfully. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events: acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported limb ischemia ¿ reversible issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
29815 | 1220648-2024-08045 | 18907210 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18907210&pc=OZD | 07/27/2022 | 03/14/2024 |  | 03/14/2024 |  |  | 07/27/2022 | 03/14/2024 |  |  | 02/29/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195206 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Hemorrhagic Stroke | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/27/2022 |  |  |  |  | The user facility reported a 71-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support, due to a non-disclosed cardiovascular medical history. while on support, the patient experienced access site bleeding at the groin after placement of the repositioning sheath. bleeding at the groin could not be controlled and ct confirmed a hemorrhage stroke. once the patients pressure was stable, decision was made to remove the impella. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29816 | 1220648-2024-08046 | 18907453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18907453&pc=OZD | 09/07/2022 | 03/18/2025 |  | 03/14/2024 |  |  | 09/07/2022 | 03/14/2024 |  |  | 02/29/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023199250 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/07/2022 |  |  |  |  | The user facility reported a 80-year-old female patient was implanted with an impella cp for mechanical circulatory support. it was reported while on impella cp support, patient experienced oozing at the groin access site. the groin was redressed and angle match adjusted. the patient was transfused with 3 units of packed red blood cells and 1 unit of platelets. the patient also experienced a loss of pedal pulses, the limb cool to the touch, and pain in the leg. the patient's leg regained color and temperature without intervention, but pulses were not palpable. the medical team made decision to wean the patient and explant the impella cp device, but patient become unstable and expired prior to explant. no allegation of device contributing to death of patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
29817 | 1220648-2024-08047 | 18907822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18907822&pc=OZD | 09/12/2022 | 03/18/2025 |  | 03/14/2024 |  |  | 09/12/2022 | 03/14/2024 |  |  | 12/31/2023 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023177729 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 09/12/2022 |  |  |  |  | The user facility reported a 70-year-old male needing mechanical circulatory support. it was reported that the patient on impella cp had slow ooze at the access site. the site was redressed and heparin was discontinued overnight until bleeding was resolved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29818 | 1220648-2024-08048 | 18908960 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18908960&pc=OZD | 07/22/2022 | 03/18/2025 |  | 03/14/2024 |  |  | 07/22/2022 | 03/14/2024 |  |  | 04/30/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023217196 |  | 0048-0065 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/22/2022 |  |  |  |  | The user facility reported a 68-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding that was resolved with administering of blood products. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29819 | 1220648-2024-08049 | 18909132 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18909132&pc=OZD | 07/22/2022 | 03/18/2025 |  | 03/14/2024 |  |  |  | 03/14/2024 |  |  | 04/30/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023217469 |  | 0048-0067 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/22/2022 |  |  |  |  | The user facility reported a 60-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was slight oozing at the impella site. the site was redressed and two units of blood were administered. in addition, cannot get pulses in left femoral artery site. the patient went to the cath lab for placement of a distal perfusion catheter. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
29820 | 1220648-2024-08050 | 18909241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18909241&pc=OZD | 07/24/2022 | 09/22/2024 |  | 03/14/2024 |  |  |  | 03/14/2024 |  |  | 04/30/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023216614 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  |  | Death | No | Yes | Material Twisted/Bent | [6 empty] | Cardiac Arrest | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/24/2022 |  |  |  |  | The user facility reported a 63-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. a kink was noted on impella catheter by outlet. attempts to reposition impella to remove the kink were unsuccessful. re-access port was wired and exchanged for 14fr introducer placed. | Additional information noted that the procedure outcome was that patient expired due to cardiac arrest. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. as noted in the impella ifu: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrected information for the initial mfr 1220648-2024-08050 was provided in sections b2, b5, d6, h1, and h6.
29821 | 1220648-2024-08051 | 18909543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18909543&pc=OZD | 08/02/2022 | 03/14/2024 |  | 03/14/2024 |  |  | 08/02/2022 | 03/14/2024 |  |  | 02/29/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202885 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/02/2022 |  |  |  |  | The user facility reported a 58-year-old male patient in acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support post stent surgery. prior to implant, the patient became hypotensive and had low cardiac output. the physician then performed percutaneous coronary intervention and implanted the impella cp. the patient later developed oozing at the impella cp insertion site (groin). the physician wanted heparin held until the activated clotting time (act) was around 160 due to the oozing. dressing was changed and there were no signs of oozing/bleeding. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
29822 | 1220648-2024-08052 | 18909558 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18909558&pc=OZD | 08/02/2022 | 05/13/2025 |  | 03/14/2024 |  |  | 08/02/2022 | 03/14/2024 |  |  | 07/31/2023 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022122161 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/02/2022 |  |  |  |  | The user facility reported a 75-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on the impella cp, the patient developed hemolysis that was not resolved despite attempts at repositioning the pump. the impella cp was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29823 | 1220648-2024-08053 | 18909571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18909571&pc=OZD | 08/03/2022 | 03/14/2024 |  | 03/14/2024 |  |  | 08/03/2022 | 03/14/2024 |  |  | 01/31/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023187095 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/03/2022 |  |  |  |  | The user facility reported a 63-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding near the groin area. the cause of bleeding was unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29824 | 1220648-2024-08054 | 18909611 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18909611&pc=OZD | 08/23/2022 | 03/14/2024 |  | 03/14/2024 |  |  | 08/23/2022 | 03/14/2024 |  |  | 05/31/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023227775 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 87 YR |  | Male |  |  | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/23/2022 |  |  |  |  | The user facility reported a 87-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported that it was difficult to advance the sheath due to kinking. an attempt was made with buddy wire but impella insertion was not possible. therefore, the physician switched sides and a second impella cp was successfully used. no harm to the patient was reported. the patient was successfully weaned from the impella and survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29825 | 1220648-2024-08055 | 18909615 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18909615&pc=OZD | 08/24/2022 | 03/14/2024 |  | 03/14/2024 |  |  | 08/24/2022 | 03/14/2024 |  |  | 05/31/2024 |  | 03/14/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023229246 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 42 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/24/2022 |  |  |  |  | The user facility reported a 42-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was suspected that the patient developed hemolysis. the patient was successfully weaned from the impella after 5 days of support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
29826 | 1220648-2024-07882 | 18892839 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18892839&pc=OZD | 02/26/2024 | 04/29/2024 |  | 03/13/2024 |  |  | 02/26/2024 | 03/13/2024 |  |  |  |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025439815 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/26/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. during impella 5.5 support, suction and aortic placement signal and left ventricle placement signal were decoupled. the impella position was four and a half centimeters to the aortic valve. the impella was pushed another one centimeter into the left ventricle but suction did not change. the impella was pulled back into the aorta. as proper flow was unable to be achieved, the impella was explanted and a new one was placed. the explanted pump was noted to have biomaterial in the cannula. | [5 empty] | The investigation into the low pump flow issues has been completed since the original report was filed. the pump was returned, tested, and evaluated. a big chunk of biomaterial observed in the inlet cage. there were no broken blades found and no cannula kink observed. the cause of the low pump flow was biomaterial ingestion. to conform to updated procedures, section d6 has revised with updated information. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
29827 | 1220648-2024-07883 | 18892921 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18892921&pc=OZD | 02/29/2024 | 09/05/2024 |  | 03/13/2024 |  |  | 02/29/2024 | 03/13/2024 |  |  |  |  | 03/13/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1103515 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/29/2024 |  |  |  |  | The user facility reported a 70-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device to support a patient transferred from the community to tertiary hospital for three vessel coronary artery bypass grafting (cabg) due to multivessel coronary artery disease (cad). the pump was placed via direct insertion but after approximately a half hour, the pump stopped. there was an alarm followed closely by the pump stopping. there was also a brief controller failure alert. the staff performed the appropriate automated impella controller (aic) swap but the pump still would not start. the pump was explanted and the patient was placed on an intra-aortic balloon pump (iabp) and then escalated to extracorporeal membrane oxygenation (ecmo). | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed. | According to the engineer the reported issue was related to pump only. no fault with the automated impella controller (aic). a regulatory report is no longer necessary. please see manufacturer device report 1220648-2024-07716 for investigation for the pump.
29828 | 1220648-2024-07884 | 18893004 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18893004&pc=OZD | 03/01/2024 | 04/18/2024 |  | 03/13/2024 |  |  | 03/01/2024 | 03/13/2024 |  |  | 07/31/2025 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025443613 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 03/01/2024 |  |  |  |  | The user facility reported a 80-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the impella cp was placed but flows were not optimal. the impella was seen under fluoroscopy to have a partial kink in place. the medical director decided to continue with the percutaneous coronary intervention procedure under impella support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow has been completed since the original report was filed. the pump was returned, tested, and evaluated. visual inspection found a kink on the cannula. the cause of the low pump flow issue was patient condition related since cannula kinked due to patient anatomical issues.
29829 | 1220648-2024-07888 | 18893052 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18893052&pc=OZD | 08/26/2022 | 03/13/2024 |  | 03/13/2024 |  |  | 08/26/2022 | 03/13/2024 |  |  | 03/31/2023 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2022068962 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 27 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/26/2022 |  |  |  |  | The user facility reported a 27-year-old male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported while on impella cp support, the patient experienced ectopy and the impella cp was explanted to resolve. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29830 | 1220648-2024-07889 | 18893084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18893084&pc=OZD | 03/04/2024 | 10/02/2024 |  | 03/13/2024 |  |  | 03/04/2024 | 03/13/2024 |  |  |  |  | 03/13/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1381389 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | Computer Operating System Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/04/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. while on impella support, the automated impella controller (aic) had a controller error (yellow alarm). the controller was replaced successfully. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | Device received by manufacturer once investigation is complete a follow up report will be submitted. d9 date device returned to manufacturer added. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-07889. e2 health professional changed to no. f6/f8-should have been left blank. g1 reporting contact fax number missing. h6 code 4629, 2898, and 403 were reported incorrectly. new code was added to health effect - impact code and medical device problem code
29831 | 1220648-2024-07890 | 18893125 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18893125&pc=OZD | 03/04/2024 | 04/22/2024 |  | 03/13/2024 |  |  | 03/04/2024 | 03/13/2024 |  |  | 11/30/2025 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024424804 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/04/2024 |  |  |  |  | The user facility reported a 63-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient had access site bleeding. the physician decided to remove the impella device and hold manual pressure. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
29832 | 1220648-2024-07892 | 18893264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18893264&pc=OZD | 03/04/2024 | 04/22/2024 |  | 03/13/2024 |  |  | 03/04/2024 | 03/13/2024 |  |  | 11/30/2025 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024434342 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/04/2024 |  |  |  |  | The user facility reported a 58-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. impella 5.5 pump support was ongoing for a patient admitted as a high-risk surgical patient. the pump support was for a total of just under five days when the pump was explanted with thrombus burden seen. there was no harm or injury was noted. | [5 empty] | The investigation for the reported thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombus could not be determined. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29833 | 1220648-2024-07891 | 18893265 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18893265&pc=OZD | 08/15/2022 | 03/13/2024 |  | 03/13/2024 |  |  | 08/15/2022 | 03/13/2024 |  |  | 03/31/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2023210610 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/15/2022 |  |  |  |  | The user facility reported a 61-year-old male patient on an impella cp for mechanical circulatory support. while on support the patient experienced constrictive pericarditis and ventricular fibrillation. the patient was given multiple procedures of cardiopulmonary resuscitation with shock and was intubated. it was noted the patient was profoundly hypotensive and diagnosed as having a stroke. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29834 | 1220648-2024-07893 | 18893332 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18893332&pc=OZD | 08/15/2022 | 03/13/2024 |  | 03/13/2024 |  |  | 08/15/2022 | 03/13/2024 |  |  | 05/31/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2023229004 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/15/2022 |  |  |  |  | The user facility reported a 78-year-old female patient with chronic congestive heart failure implanted with an impella cp for mechanical circulatory support. while on support, there was a hematoma noted at the impella site and concerns of an ischemic leg. a decision was made to explant the impella. extravasation noted on post impella explant that required a balloon tamponade and transfusion of one unit of blood. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29835 | 1220648-2024-07894 | 18893349 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18893349&pc=OZD | 03/04/2024 | 03/13/2024 |  | 03/13/2024 |  |  | 03/04/2024 | 03/13/2024 |  |  |  |  | 03/13/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1133147 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 03/04/2024 |  |  |  |  | The user facility reported an unspecified undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that an automated impella controller (aic) displayed a controller error alarm during patient support. the aic was replaced in response to the noted failure. there was no patient harm. | [5 empty] | The investigation was completed. the device was returned. console logs from the reported day of event were consistent with the complaint and showed an controller error alarm, preceded by loss of placement signal as well as all data received from optical bench. the issue was unable to be reproduced throughout testing. the cause of the issue was not determined since the issue could not be reproduced.
29836 | 1220648-2024-07895 | 18893447 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18893447&pc=OZD | 12/13/2022 | 03/13/2024 |  | 03/13/2024 |  |  | 12/13/2022 | 03/13/2024 |  |  | 07/31/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023254088 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/13/2022 |  |  |  |  | The user facility reported a 62-year-old female with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the patient experienced oozing at the groin site while in the icu. the medical staff applied manual pressure and the purge solution was switched to sodium bicarbonate. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
29837 | 1220648-2024-07896 | 18893508 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18893508&pc=OZD | 03/06/2024 | 04/24/2024 |  | 03/13/2024 |  |  | 03/06/2024 | 03/13/2024 |  |  | 10/31/2025 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024424363 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/06/2024 |  |  |  |  | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported the patient was on impella cp support due to having biventricular failure. while on support, a hematoma was noted. the pump was explanted. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was filed. the pump was returned, tested, and evaluated. no abnormalities were found. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
29838 | 1220648-2024-07897 | 18893536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18893536&pc=OZD | 12/13/2022 | 03/13/2024 |  | 03/13/2024 |  |  | 12/13/2022 | 03/13/2024 |  |  | 07/31/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023255810 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/13/2022 |  |  |  |  | The user facility reported a 87-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that superficial oozing was noted on the gauze dressing at the left groin access site. the site was redressed and after multiple attempts the catheter was angle matched at 40 degrees with successful hemostasis. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
29839 | 1220648-2024-07898 | 18893602 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18893602&pc=OZD | 03/06/2024 | 04/24/2024 |  | 03/13/2024 |  |  | 03/06/2024 | 03/13/2024 |  |  |  |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024433400 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/06/2024 |  |  |  |  | The user facility reported a 84-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the patient was on impella cp support due to having worsening chest pain. while on support, the hemostatic valve was leaking more than usual when acquiring access with the wire and diagnostic catheters prior to placing the impella. | [5 empty] | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleed was unable to be determined as the introducer was not returned for analysis. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29840 | 1220648-2024-07902 | 18893806 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18893806&pc=OZD | 07/27/2022 | 03/12/2025 |  | 03/13/2024 |  |  | 07/27/2022 | 03/13/2024 |  |  | 04/30/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2023216572 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/27/2022 |  |  |  |  | The user facility reported a 58-year-old female was implanted with an impella cp device for mechanical circulatory support. the complainant reported that after being intubated earlier in the day, the patient was escalated to impella 5.5 due to requiring multiple pressors and losing pules in the left foot. impella cp was removed without difficulty and impella 5.5 was placed. the groin site was surgically closed to restore blood flow to the left foot. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The root cause of limb ischemia could not be determined as insufficient clinical details were provided.
29841 | 1220648-2024-07903 | 18893917 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18893917&pc=OZD | 09/07/2022 | 06/05/2024 |  | 03/13/2024 |  |  | 09/07/2022 | 03/13/2024 |  |  | 05/31/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023217188 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/07/2022 |  |  |  |  | The user facility reported a 83-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient had mitral valve insufficiency, and possible papillary muscle damage due to ischemia. it was noted that the patient was awake and slightly confused, but stable. heparin was issued; however, the device was eventually explanted. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. the instructions for use referenced in the initial is for the impella cp with smart assist system in sections: general patient care considerations, entering cardiac output, and potential adverse events (united states), which now includes the statement "cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
29842 | 1220648-2024-07904 | 18894145 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18894145&pc=OZD | 09/07/2022 | 05/30/2025 |  | 03/13/2024 |  |  | 09/07/2022 | 03/13/2024 |  |  | 05/31/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023223372 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/07/2022 |  |  |  |  | The user facility reported a patient of unknown gender was being implanted with an impella cp device due to percutaneous coronary intervention (pci) for mechanical circulatory support, and that the procedure was aborted as the device was unable to pass through vasculature due to massive kinking. straightening with stiff wire was unsuccessful. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined.
29843 | 1220648-2024-07905 | 18894295 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18894295&pc=OZD | 08/15/2022 | 03/13/2024 |  | 03/13/2024 |  |  | 08/15/2022 | 03/13/2024 |  |  | 03/31/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023211238 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/15/2022 |  |  |  |  | The user facility reported a 70-year-old male patient on an impella cp for mechanical circulatory support. the echocardiogram indicated the impella position to be shallow and pushed in. during support, hematuria, a bleed in the mouth and no pulses in the foot were observed. the patient was weaned from the impella and it was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records.
29844 | 1220648-2024-07906 | 18894300 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18894300&pc=OZD | 03/06/2024 | 04/24/2024 |  | 03/13/2024 |  |  | 03/06/2024 | 03/13/2024 |  |  | 07/31/2025 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025442800 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/06/2024 |  |  |  |  | The user facility reported a 73-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the pump began to experience high purge pressure alarms two hours after being implanted. the issue was troubleshooted, but the motor current continued to rise and the pump subsequently stopped. the pump was removed as a result of the failure. there was no patient harm. | [5 empty] | The investigation for the reported pump stop and high purge pressure has been completed. pump was returned for evaluation. upon visual inspection, lots of dried blood was present in and around purge gap. pump was soaked in warm water to remove dried blood. after soaking, biomaterial was found in purge gap. data logs were reviewed and lt logs showed gradual rise in purge pressure over a period of about an hour, then sustained high purge pressure until pump was removed. rt logs from the end of the case were reviewed and a high motor current pump stop occurred. pump was attempted to be restarted multiple times, pump was able to restart for approximately 2 minutes before high motor current pump stop occurred again. more attempts to restart pump without success. clinical details noted that no heparin was used in purge solution during case. the root cause of the pump stop was due to gradual buildup of biomaterial in the purge gap. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
29845 | 1220648-2024-07907 | 18894361 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18894361&pc=OZD | 08/16/2022 | 05/20/2025 |  | 03/13/2024 |  |  | 08/16/2022 | 03/13/2024 |  |  | 04/30/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023220595 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/16/2022 |  |  |  |  | The user facility reported a 89-year-old male patient male noted as a high risk percutaneous coronary insertion, was implanted with an impella cp for mechanical circulatory support. the patient was weaned from the impella and once hemodynamically stable, decision was made to remove the impella. the impella was removed, however, the physician was unable to remove the peel-away sheath. the vascular surgery team performed a cut-down to remove the peel-away sheath. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29846 | 1220648-2024-07908 | 18894393 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18894393&pc=OZD | 09/08/2022 | 04/17/2025 |  | 03/13/2024 |  |  | 09/08/2022 | 03/13/2024 |  |  | 06/30/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023236956 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/08/2022 |  |  |  |  | The user facility reported an impella was attempted to be implanted in a 78-year-old female patient for mechanical circulatory support and that the medical team experienced difficulty implanting the impella due to the patient's anatomy. ultimately the medical team aborted the procedure. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29847 | 1220648-2024-07910 | 18894456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18894456&pc=OZD | 08/18/2022 | 05/29/2025 |  | 03/13/2024 |  |  | 08/18/2022 | 03/13/2024 |  |  | 05/31/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023229284 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Deliver | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/18/2022 |  |  |  |  | The user facility reported a 66-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported a delivery issue with the impella pump. right and left femoral access was attempted and unable to be obtained. the attempt was aborted due to patient anatomy. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29848 | 1220648-2024-07909 | 18894459 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18894459&pc=OZD | 07/29/2022 | 03/12/2025 |  | 03/13/2024 |  |  | 07/29/2022 | 03/13/2024 |  |  | 03/31/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023205619 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/29/2022 |  |  |  |  | The user facility reported a female, age unknown, was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient was on impella cp support due to ventricular tachycardia. while on support, the patient experienced limb ischemia; however, the patient did not have issues with the impella cp. when the patient was being turned the extracorporeal membrane oxygenation (ecmo) pump stopped and they became hemodynamically decompensated. there was no allegation noted against the impella as the cause of death for the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The root cause of limb ischemia could not be determined as insufficient clinical details were provided.
29849 | 1220648-2024-07911 | 18894490 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18894490&pc=OZD | 03/06/2024 | 05/24/2024 |  | 03/13/2024 |  |  | 03/06/2024 | 03/13/2024 |  |  |  |  | 03/13/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1097384 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/06/2024 |  |  |  |  | The patient expired after the family withdrew support. after review by abiomed's medical safety officer, there is not enough information to associate outcome with the use of the device | The user facility reported a 59-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the patient had an impella cp and extracorporeal membrane oxygenation (ecmo) support ongoing when the console alarmed yellow for a controller failure. the automated impella controller (aic) was replaced with the backup controller and support proceeded without harm or injury from the interruption in support. |  |  |  |  | The investigation for the reported console (aic) controller error yellow alarm has been completed. the console was returned for evaluation. the issue was unable to be reproduced, however, replacement kbd assembly resolved the issue. data logs were reviewed which showed a controller error alarm upon boot-up, preceded by error message ¿imc-kbd error: 0xa4 0x01 0x00¿. issue did not reproduce upon consecutive boot-up(s). the cause of the aic issue was most likely a defective kbd assembly. b5-added medical safety officer review. in accordance with updated procedures, sections d2 and h10 have been revised from the initial submission. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed.
29850 | 1220648-2024-07912 | 18894607 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18894607&pc=OZD | 08/26/2022 | 03/12/2025 |  | 03/13/2024 |  |  | 08/26/2022 | 03/13/2024 |  |  | 03/31/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023214665 |  | 0048-0054 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/26/2022 |  |  |  |  | The user facility reported a 82-year-old female was being placed on impella cp for mechanical circulatory support. it was reported while on attempting the impella insertion a distal aorta/common iliac calcium lesion caused the pump to be difficult to insert. the physician had to apply significant force before the pump was able to pass through the calcified lesion and advanced up to left ventricle. the impella cp device was successfully placed. | [5 empty] | The root cause of the difficulty inserting/removing introducer could not be determined as insufficient clinical details were provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29851 | 1220648-2024-07913 | 18894623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18894623&pc=OZD | 09/08/2022 | 04/07/2025 |  | 03/13/2024 |  |  | 09/08/2022 | 03/13/2024 |  |  | 05/31/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023228968 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 09/08/2022 |  |  |  |  | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 76-year-old male was implanted with an impella cp device for mechanical circulatory support and that the patient had loss of pedal pulses when they started patient on vasopressin. patient was able to be wean from impella, and decision was made to withdraw care. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07913 was provided in sections b2, b5, h1, and h6. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29852 | 1220648-2024-07914 | 18894685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18894685&pc=OZD | 08/18/2022 | 03/11/2025 |  | 03/13/2024 |  |  | 08/18/2022 | 03/13/2024 |  |  | 04/30/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218022 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/18/2022 |  |  |  |  | The user facility reported a 74-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported that when crossing the aortic value, the impella pump¿s catheter kinked. the catheter was successfully placed but there was low flow. it was suggested that the catheter should be replaced but the doctor declined and was satisfied with the flow. patient remained stable throughout the procedure. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | The investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29853 | 1220648-2024-07915 | 18894722 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18894722&pc=OZD | 07/29/2022 | 03/12/2025 |  | 03/13/2024 |  |  | 07/29/2022 | 03/13/2024 |  |  | 04/30/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023217244 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/29/2022 |  |  |  |  | The user facility reported a 79-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported that there was difficulty encountered when the pump was inserted as it was unable to advance past the stent in the iliac artery (ia). the pump was subsequently removed, and a 2.5 pump was inserted instead. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29854 | 1220648-2024-07916 | 18894730 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18894730&pc=OZD | 03/07/2024 | 04/24/2024 |  | 03/13/2024 |  |  | 03/07/2024 | 03/13/2024 |  |  | 02/28/2025 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024321590 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/07/2024 |  |  |  |  | The user facility reported a 73-year-old male undergoing cardiac surgery was being placed on impella 5.5 support for multivessel coronary artery disease (cad) intervention. the team was unable to deliver the impella 5.5 pump despite all efforts. the native anatomy was unable to be traversed and so instead a impella cp was placed. no harm or injury. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29855 | 1220648-2024-07918 | 18894814 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18894814&pc=OZD | 09/08/2022 | 04/17/2025 |  | 03/13/2024 |  |  | 09/08/2022 | 03/13/2024 |  |  | 04/30/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023225084 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/08/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 63-year-old male patient for mechanical circulatory support. the impella was positional and had the tendency to migrate into the left ventricle. the impella was repositioned, secured and was left in its current position. | [5 empty] | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29856 | 1220648-2024-07919 | 18894817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18894817&pc=OZD | 08/17/2022 | 04/07/2025 |  | 03/13/2024 |  |  | 08/17/2022 | 03/13/2024 |  |  | 04/30/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPELEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023221589 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/17/2022 |  |  |  |  | It was noted that the patient outcome was that the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 75-year-old male on a impella cp for mechanical circulatory support. it was reported that the patient developed heparin induced thrombocytopenia during impella cp support. systemic heparin was stopped in response to the condition. support was maintained with the implanted pump. the patient was escalated to impella 5.5 from impella cp. no patient harm was reported. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07919 was provided in sections b2, b5, h1, and h6. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29857 | 1220648-2024-07920 | 18894883 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18894883&pc=OZD | 08/29/2022 | 03/12/2025 |  | 03/13/2024 |  |  | 08/29/2022 | 03/13/2024 |  |  | 03/31/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191422 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/29/2022 |  |  |  |  | The user facility reported a 66-year-old male in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported after the impella cp device was successfully placed and the patient was on support that the pump was dislocated and was not placed again. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29858 | 1220648-2024-07921 | 18894888 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18894888&pc=OZD | 03/07/2024 | 03/22/2024 |  | 03/13/2024 |  |  | 03/07/2024 | 03/13/2024 |  |  |  |  | 03/13/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1656412 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/07/2024 |  |  |  |  | It was reported that the automated impella controller (aic) had a controller error. no patient was involved. | [5 empty] | The investigation into the aic - controller failure (red alarm) has been completed since the original report was filed. the console was returned and tested after arriving in fs. the failure mode was not reproduced while running an optical pump simulator. the log entry indicated a false ¿communication stopped¿ condition. the cause of the aic issue was a software issue. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed.
29859 | 1220648-2024-07922 | 18895133 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18895133&pc=OZD | 08/18/2022 | 03/10/2025 |  | 03/13/2024 |  |  | 08/18/2022 | 03/13/2024 |  |  | 05/25/2021 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2022076031 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/18/2022 |  |  |  |  | The user facility reported a 55-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported there was some bleeding at the groin site and pressure was applied which resolved the oozing. a low platelet count (thrombocytopenia) developed and platelet treatment was given later. systemic anticoagulants were withheld and eventually the device was explanted. no patient harm reported. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleeding and thrombocytopenia issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
29860 | 1220648-2024-07923 | 18895222 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18895222&pc=OZD | 08/25/2022 | 05/23/2025 |  | 03/13/2024 |  |  | 08/25/2022 | 03/13/2024 |  |  | 03/31/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023205934 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/25/2022 |  |  |  |  | Additional information noted the patient expired on support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 50-year-old male on a impella cp for mechanical circulatory support. perclose was placed prior to insertion of the impella cp. the peel away sheath was noted to have some oozing present around it. the physician adjusted the perclose sutures to try to control the ooze. the team monitored ooze from the left groin, and likely gave patient platelet transfusion again to help control bleeding around lines. the impella cp was explanted after multiple treatments for bleeding. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07923 was provided in sections b2, b5, h1, and h6. | The impella device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29861 | 1220648-2024-07924 | 18895227 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18895227&pc=OZD | 08/18/2022 | 03/07/2025 |  | 03/13/2024 |  |  | 08/18/2022 | 03/13/2024 |  |  | 07/31/2023 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2022121190 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/18/2022 |  |  |  |  | The user facility reported a 58-year-old female noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had a significant amount of oozing around the 7fr sheath ,which caused the sheath to be removed and the pump's sheath was reassessed. a 6fr sheath was inserted with less oozing and the physician decided to give a unit of blood post procedure, due to the amount of blood lost that was oozing around the 6/7fr sheaths. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29862 | 1220648-2024-07925 | 18895236 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18895236&pc=OZD | 09/08/2022 | 03/13/2024 |  | 03/13/2024 |  |  | 09/08/2022 | 03/13/2024 |  |  | 02/29/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023201781 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/08/2022 |  |  |  |  | The user facility reported a 50-year-old male was implanted with an impella cp device for mechanical circulatory support. while the patient was on impella cp support, the patient experienced limb ischemia in the right lower extremities. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
29863 | 1220648-2024-07927 | 18895502 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18895502&pc=OZD | 08/18/2022 | 03/10/2025 |  | 03/13/2024 |  |  | 08/18/2022 | 03/13/2024 |  |  | 03/31/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023210860 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/18/2022 |  |  |  |  | The user facility reported a 61-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported there was oozing at the impella site and the patient received 3 units of packed red blood cells (prbcs). upon arrival of transfer, there was a slow ooze at the patient's groin. during the dress changing (optimal dressing technique -tegaderm surrounding impella device, no tenting or pulling to catheter) the bleeding subsided. however, bleeding around the insertion site reoccurred even though there were sutures added as well as angle matching gauze. platelets and additional packed red blood cells were given for treatment. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleeding issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
29864 | 1220648-2024-07928 | 18895506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18895506&pc=OZD | 09/10/2022 | 05/29/2025 |  | 03/13/2024 |  |  | 09/10/2022 | 03/13/2024 |  |  | 06/30/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023236792 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 25 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Patient-Device Incompatibility | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/10/2022 |  |  |  |  | The user facility reported a 25 year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. initially, the left femoral artery was accessed but insertion was aborted and pre-closed. the impella cp was then successfully inserted in the right femoral artery. the patient had a severe turn in the aortic root causing the impella cannula to kink at the turn, this caused the impella to be removed. after being run at p1 in sterile water, the impella cp was re-inserted, and kinking was still present. however, the kinking was noted to be reduced when the pump was left deep. the patient later developed right leg ischemia due to small femoral vessels. fem-fem bypass was not performed, however, the impella was explanted the following day due to small femoral vessels. the patient was stable without hemodynamic support. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29865 | 1220648-2024-07929 | 18895772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18895772&pc=OZD | 09/11/2022 | 04/17/2024 |  | 03/13/2024 |  |  | 09/11/2022 | 03/13/2024 |  |  | 05/31/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023230357 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Female | 40 KG |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/11/2022 |  |  |  |  | The user facility reported a female, age unknown, undergoing coronary artery bypass grafting was implanted with an impella cp device for mechanical circulatory support. this patient is super small at 40 kg. and support levels were unable to increase beyond p3 following a scheduled surgery. hemoglobin levels had decreased to 6.5 and two units of blood were transfused. albumin was also administered, but the patient experienced recurrent arrhythmia when support levels were increased. it was further noted that the patient's leg looked mottled. as a result of the conditions, the decision was made to explant the pump. | [5 empty] | The investigation for the reported limb ischemia and arrhythmia has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia and arrhythmia could not be determined as the device was not returned nor sufficient clinical details. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
29866 | 1220648-2024-07931 | 18896078 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18896078&pc=OZD | 08/26/2022 | 04/04/2025 |  | 03/13/2024 |  |  | 08/26/2022 | 03/13/2024 |  |  | 03/31/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023214409 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/26/2022 |  |  |  |  | The user facility reported a 63-year-old male on a impella cp for mechanical circulatory support. it was reported that the patient on impella cp had positioning issues where the flows were low. an echocardiogram (echo) showed that the impella cp was in too deep. the impella cp was repositioned, and the patient heart functions were improved, with no additional medical intervention. | [5 empty] | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. b1: initial mfr incorrectly reported as adverse event, and this has been corrected to product problem. h1: initial mfr incorrectly reported as serious injury, and this has been corrected to malfunction. | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ this report is being filed as part of a retrospective review of historical records.
29867 | 1220648-2024-07932 | 18896129 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18896129&pc=OZD | 08/29/2022 | 03/11/2025 |  | 03/13/2024 |  |  | 08/29/2022 | 03/13/2024 |  |  | 10/31/2022 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2021007158 |  | 0048-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/29/2022 |  |  |  |  | The user facility reported a 74-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while patient was on impella cp support, the patient experienced a cold and cyanotic leg with low flow under doppler echocardiography. the decision was made to explant the device as a result of no improvement to limb ischemia. | [5 empty] | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29868 | 1220648-2024-07933 | 18896273 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18896273&pc=OZD | 09/12/2022 | 03/13/2024 |  | 03/13/2024 |  |  | 09/12/2022 | 03/13/2024 |  |  | 06/30/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023237686 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/12/2022 |  |  |  |  | The user facility reported a 67-year-old male was implanted with an impella cp device for mechanical circulatory support. while the patient was on impella cp support, the patient experienced bleeding at the impella site and manual pressure was applied and the impella cp was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29869 | 1220648-2024-07934 | 18896321 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18896321&pc=OZD | 08/18/2022 | 05/27/2025 |  | 03/13/2024 |  |  | 08/18/2022 | 03/13/2024 |  |  | 03/03/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023214340 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/18/2022 |  |  |  |  | The user facility reported a 58-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella support was lowered to p4 due to suspicion of hemolysis. labs were going to be reviewed, along with volume status and impella min-max differential and discussed with the fellow to differentiate between mcs hemolysis and metabolic sequelae. there was a request to escalate the patient to the impella 5.5. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29870 | 1220648-2024-07935 | 18896330 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18896330&pc=OZD | 07/24/2022 | 03/11/2025 |  | 03/13/2024 |  |  | 07/24/2022 | 03/13/2024 |  |  | 03/31/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023209348 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/24/2022 |  |  |  |  | The user facility reported a 75-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced hemolysis due to hematuria. patient weaned to p-6 overnight, then weaned down to p-2 over 20 minutes. the impella was explanted and the access site was closed with perclose. heparin purge was administered. | [5 empty] | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
29871 | 1220648-2024-07937 | 18896459 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18896459&pc=OZD | 08/30/2022 | 04/07/2025 |  | 03/13/2024 |  |  | 08/30/2022 | 03/13/2024 |  |  | 04/30/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023217318 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/30/2022 |  |  |  |  | The user facility reported a 81-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced a slight tear in the common femoral artery during cinching of the preclose and the tear required contralateral hemostasis with 10x50 balloon for 15 mins. hemostasis obtained after 20 min manual pressure. | [5 empty] | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records.
29872 | 1220648-2024-07936 | 18896461 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18896461&pc=OZD | 08/18/2022 | 03/07/2025 |  | 03/13/2024 |  |  | 08/18/2022 | 03/13/2024 |  |  | 04/30/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023218620 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/18/2022 |  |  |  |  | The user facility reported a 70-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that an echocardiogram was performed while displaying the location of the impella under the pap muscle. while the physician was at the bedside attempting to withdraw the pump into the correct position, the device would not torque. no suction alarms were noted and the decision was made to leave the pump in the current position and continue to reduce so the device can be explanted. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29873 | 1220648-2024-07939 | 18896812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18896812&pc=OZD | 08/30/2022 | 03/11/2025 |  | 03/13/2024 |  |  | 08/30/2022 | 03/13/2024 |  |  | 05/31/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023228488 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/30/2022 |  |  |  |  | The user facility reported a 52-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported while attempting to insert the impella cp the physician had a hard time advancing the wire. the physician then turned the pump off, removed the wire and sheath, restarted the insertion process and successfully placed the impella cp device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29874 | 1220648-2024-07940 | 18896819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18896819&pc=OZD | 09/12/2022 | 05/22/2025 |  | 03/13/2024 |  |  | 09/12/2022 | 03/13/2024 |  |  | 05/31/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231060 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/12/2022 |  |  |  |  | The user facility reported an 84-year-old female was implanted with an impella cp device for mechanical circulatory support. while being placed on impella cp support, the patient began to show signs of ejection fraction (ef) of 45-50%, moderate mitral regurgitation with a calcified leaflet, and moderate tricuspid regurgitation. upon access of the right femoral artery, dissection flap was noted on tortuous anatomy. the right femoral artery was surgically repaired and closed. the patient was successfully weaned off the impella cp and hemodynamics was normalized. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29875 | 1220648-2024-07941 | 18896857 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18896857&pc=OZD | 07/24/2022 | 03/11/2025 |  | 03/13/2024 |  |  | 07/24/2022 | 03/13/2024 |  |  | 03/31/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023214763 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/24/2022 |  |  |  |  | The user facility reported a 68-year-old female with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the patient experienced oozing at the implantation site. the care team initiated anticoagulation, systemic and purge and target act 160-180. the patient was successfully weaned. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29876 | 1220648-2024-07942 | 18896963 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18896963&pc=OZD | 12/14/2022 | 03/13/2024 |  | 03/13/2024 |  |  | 12/14/2022 | 03/13/2024 |  |  | 04/30/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023221148 |  | 0048-0071 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/14/2022 |  |  |  |  | The user facility reported a 84-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the impella cp was unable to be placed due to an aortic lesion and the procedure was aborted. | [5 empty] | The investigation for the reported issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related. this report is being filed as part of a retrospective review of historical records.
29877 | 1220648-2024-07943 | 18897099 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18897099&pc=OZD | 08/30/2022 | 03/13/2024 |  | 03/13/2024 |  |  | 08/30/2022 | 03/13/2024 |  |  | 04/30/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023223987 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/30/2022 |  |  |  |  | The user facility reported a 68-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was on impella cp support and the patient experienced bleeding (oozing) at the impella cp insertion site. the site was attempted to be redressed and re-sutured. the bleeding subsided, but a large hematoma formed at the site a cut down was performed and the impella was explanted from the patient. | [5 empty] | The investigation of the reported issue has been completed. the root cause of the access site bleeding issue was not determined since product was not returned and no clear information is provided as per the clinical description. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records.
29878 | 1220648-2024-07944 | 18897104 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18897104&pc=OZD | 09/12/2022 | 03/11/2025 |  | 03/13/2024 |  |  | 09/12/2022 | 03/13/2024 |  |  | 05/31/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231281 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/12/2022 |  |  |  |  | The user facility reported a 63-year-old female was implanted with an impella cp device for mechanical circulatory support. while the patient was on impella cp support, the distal blood flow in the right leg was questionable. the patient was known to have peripheral vascular disease. the impella cp was weaned and explanted without issue. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29879 | 1220648-2024-07945 | 18897141 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18897141&pc=OZD | 08/18/2022 | 03/10/2025 |  | 03/13/2024 |  |  | 08/18/2022 | 03/13/2024 |  |  | 04/30/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023220029 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/18/2022 |  |  |  |  | The user facility reported a 79-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. while on support, one long drug-eluting stent was placed within the patient causing suction and flows to drop once more. it was noted that the interventional access had a hematoma forming which required manual pressure and it was decided to halt further intervention. the left side was closed with an angioseal and the groin site was stabilized. the pump was removed and the groin was closed through indwelling preclose. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleeding issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
29880 | 1220648-2024-07946 | 18897181 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18897181&pc=OZD | 08/18/2022 | 03/10/2025 |  | 03/13/2024 |  |  | 08/18/2022 | 03/13/2024 |  |  | 05/31/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023229269 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/18/2022 |  |  |  |  | The user facility reported a 66-year-old female in cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. the patient experience overnight oozing from the access site after the implantation. the insertion site was treated by re-suturing and angle matching which caused the oozing to cease. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleeding issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
29881 | 1220648-2024-07947 | 18897339 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18897339&pc=OZD | 02/26/2024 | 04/12/2024 |  | 03/13/2024 |  |  | 02/26/2024 | 03/13/2024 |  |  | 10/31/2025 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024412221 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/27/2024 |  |  |  |  | The user facility reported a 69-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that there was a hematoma on the left groin due to multiple puncture attempts. the physician attempted to place the impella from the left femoral side first without success and then switched to the right side. the impella was able to be placed from the right groin without issue. following percutaneous coronary intervention (pci), the patient was transferred to surgery to stop the bleeding from the left side. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical details provided was sufficient to determine the root cause. the root cause of the access site bleeding was most likely due to use issue since the bleeding occurred due to multiple puncture attempts.
29882 | 1220648-2024-07948 | 18897388 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18897388&pc=OZD | 09/13/2022 | 04/08/2025 |  | 03/13/2024 |  |  | 09/13/2022 | 03/13/2024 |  |  | 11/30/2023 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023163685 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | Thrombosis/Thrombus | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 09/13/2022 |  |  |  |  | The user facility reported a 56-year-old female was implanted with an impella cp device for mechanical circulatory support, and the patient required the impella cp to be repositioned via trans-esophageal echo guidance. the patient went into ventricular fibrillation and was shocked three times. the patient was given blood products, fresh frozen plasma (ffp), and albumin. patient was stabilized and impella cp was explanted. | Additional information noted care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07948 was provided. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29883 | 1220648-2024-07950 | 18897520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18897520&pc=OZD | 08/15/2022 | 03/13/2024 |  | 03/13/2024 |  |  | 08/15/2022 | 03/13/2024 |  |  | 03/31/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204991 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/15/2022 |  |  |  |  | The user facility reported a 81-year-old male patient on an impella cp for mechanical circulatory support. during implantation the pump was unable to advance due to kinking of the sheath at the bend in the artery. the impella was inserted past tortuosity over steel core wire into the left ventricle. it was reported that the patient had an excellent angiographic result and the impella was later weaned and removed without complication. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29884 | 1220648-2024-07949 | 18897525 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18897525&pc=OZD | 09/13/2022 | 03/17/2025 |  | 03/13/2024 |  |  | 09/13/2022 | 03/13/2024 |  |  | 06/30/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023236826 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/13/2022 |  |  |  |  | The user facility reported a 77-year-old male was implanted with an impella cp device for mechanical circulatory support. while being placed on the impella cp support, the patient experienced a perforation that required covered stent and/or balloon tamponade. the patient also experienced a perforation of left anterior descending following post-stent dilation. the patient then underwent a pericardiocentesis with auto-transfusion of blood. hemodynamics was normalized after that. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29885 | 1220648-2024-07951 | 18897669 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18897669&pc=OZD | 09/13/2022 | 03/13/2024 |  | 03/13/2024 |  |  | 09/13/2022 | 03/13/2024 |  |  | 02/29/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023199941 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/13/2022 |  |  |  |  | The user facility reported a 78-year-old female was implanted with an impella cp device for mechanical circulatory support and experienced bleeding at the access site. the access site was redressed, and position of the sheath was adjusted, and no new oozing was noted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29886 | 1220648-2024-07952 | 18897926 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18897926&pc=OZD | 08/26/2022 | 03/17/2025 |  | 03/13/2024 |  |  | 08/26/2022 | 03/13/2024 |  |  | 05/31/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023230159 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/26/2022 |  |  |  |  | The user facility reported a 60-year-old male on impella cp for mechanical circulatory support. the patient continued to bleed through transfer. heparin in purge was continued to be held due to bleeding. the patient¿s site was oozing and showing ooze through the hemostasis valve. the catheter was propped up for better alignment and the was site redressed. after the patient continued to maintain hemodynamics on his own, the impella cp was removed. hemostasis was achieved with distal perfusion/pulses intact. no patient harm was reported. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29887 | 1220648-2024-07953 | 18898114 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18898114&pc=OZD | 09/17/2022 | 03/17/2025 |  | 03/13/2024 |  |  | 09/17/2022 | 03/13/2024 |  |  | 05/31/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023230739 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/17/2022 |  |  |  |  | The user facility reported a 68-year-old male was implanted with an impella cp device for mechanical circulatory support. while the patient was on impella cp support, it was determined that the patient¿s leg was dead. the patient¿s dead leg was surgically removed and the impella cp was explanted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
29888 | 1220648-2024-07954 | 18898874 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18898874&pc=OZD | 10/24/2022 | 03/18/2025 |  | 03/13/2024 |  |  | 10/24/2022 | 03/13/2024 |  |  | 07/31/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023249734 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/24/2022 |  |  |  |  | The user facility reported a 73-year-old male was implanted with an impella cp for cardiac support after going into ventricular tachycardia and being shocked eight times. the physician attempted to reposition the impella cp but it was decided to take impella cp out and put in an intra-aortic balloon pump due to continuous heart arrhythmias. the patient was still arrhythmic post impella explant. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29889 | 1220648-2024-07955 | 18898978 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18898978&pc=OZD | 10/26/2022 | 03/13/2024 |  | 03/13/2024 |  |  | 10/26/2022 | 03/13/2024 |  |  | 07/31/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023257437 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/26/2022 |  |  |  |  | The user facility reported a 87-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding. a balloon tamponade failed and thrombin was unsuccessful. a stent was placed and the impella cp was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
29890 | 1220648-2024-07956 | 18899651 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18899651&pc=OZD | 02/16/2024 | 04/04/2024 |  | 03/13/2024 |  |  |  | 03/13/2024 |  |  | 10/31/2025 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024419252 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | Intracranial Pressure Increased | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/16/2024 |  |  |  |  | The user facility reported a 68-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient had an attempted impella 5.5 implant due to having heart failure. the impella was inserted and was unable to pass the innominate artery. several attempts were made to pass with a buddy wire and it was unsuccessful. the physician decided to switch to an impella cp. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the delivery issue is determined to be patient condition because impella 5.5 was unable to pass through due to small vessels and later, they were able to pass impella cp through the vasculature.
29891 | 1220648-2024-07957 | 18899791 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18899791&pc=OZD | 09/15/2023 | 05/28/2024 |  | 03/13/2024 |  |  | 02/12/2024 | 03/13/2024 |  |  | 05/31/2025 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024369467 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Renal Impairment | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/12/2024 |  |  |  |  | A notification was received from our medical safety team via loqi (abiomed¿s long-term outcome and quality indicator impella registry) indicating a 78 year old female patient presented for high risk percutaneous coronary intervention using the impella cp device for hemodynamic support. during support, the patient developed an access site hematoma that prompted evacuation of the site, upon device removal. | Additional information received via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured l inguinal hematoma with a "possible " relationship to the impella device used: ¿patient was diagnosed with anemia due to blood loss from the l inguinal hematoma.. patient received 5 rbc units total throughout this admission.¿ the patient recovered with no sequelae. additionally patient endured acute kidney injury wit, the possible / definite / probable" relationship to the impella device used. diuretics were held. the patient recovered with no sequelae. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | Section b5 and h6 were updated as per additional information received. | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use states ¿assess access site for bleeding and hematoma.¿ d4-updated model number.
29892 | 1220648-2024-07958 | 18899987 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18899987&pc=OZD | 07/25/2022 | 03/18/2025 |  | 03/13/2024 |  |  | 07/25/2022 | 03/13/2024 |  |  | 04/30/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023220653 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/25/2022 |  |  |  |  | The user facility reported a 69-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced a thrombus. the impella was successfully explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency). ¿relevant laboratory test results, such as coagulation profiles and platelet counts ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). ¿medications or treatments that the patient was receiving at the time of thrombus formation. ¿surgical or procedural details if applicable (e.g, cardiac catheterization) ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. this report is being filed as part of a retrospective review of historical records.
29893 | 1220648-2024-07959 | 18900016 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18900016&pc=OZD | 07/25/2022 | 03/18/2025 |  | 03/13/2024 |  |  | 07/25/2022 | 03/13/2024 |  |  | 07/07/2022 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023221595 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/25/2022 |  |  |  |  | The user facility reported a 59-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced access site bleeding. the oozing continued after the dressing changes. impella cp removed and 5.5 placed double barrel. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29894 | 1220648-2024-07960 | 18900024 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18900024&pc=OZD | 07/26/2022 | 03/18/2025 |  | 03/13/2024 |  |  | 07/26/2022 | 03/13/2024 |  |  | 04/30/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023219410 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/26/2022 |  |  |  |  | The user facility reported a 75-year-old male n acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced a thrombus that required surgical intervention to remove. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29895 | 1220648-2024-07961 | 18900046 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18900046&pc=OZD | 07/26/2022 | 03/18/2025 |  | 03/13/2024 |  |  | 07/26/2022 | 03/13/2024 |  |  | 04/30/2024 |  | 03/13/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023216734 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/26/2022 |  |  |  |  | The user facility reported a 75-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced oozing from the repo sheath site. femostop was placed above site with no added pressure, and fluids were given to elevate central venous pressure, the impella device was successfully weaned. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29896 | 1220648-2024-07842 | 18884550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18884550&pc=OZD | 09/09/2022 | 03/12/2024 |  | 03/12/2024 |  |  | 09/09/2022 | 03/12/2024 |  |  | 02/02/2024 |  | 03/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023194093 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/09/2022 |  |  |  |  | The user facility reported a 62-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. after the impella was inserted, the patient transitioned to pulseless electrical activity arrest (pea). advanced cardiac life support was provided and return of spontaneous circulation (rosc) was achieved. the patient arrested two more times during case with similar pre pea arrest arrhythmia and quick rosc. | [5 empty] | The investigation of the reported issue has been completed. ventricular fibrillation (vt), ventricular tachycardia (vf), atrial fibrillation, atrial tachycardia occurred during device support. any device contacting the heart wall in conjunction with a poor patient condition could result in vf/vt but since fluoroscopic imaging and/or sufficient clinical details were not provided, a relation between impella and the root cause vf/vt is unable to be determined. this report is being filed as part of a retrospective review of historical records.
29897 | 1220648-2024-07843 | 18884619 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18884619&pc=OZD | 09/09/2022 | 03/14/2025 |  | 03/12/2024 |  |  | 09/09/2022 | 03/12/2024 |  |  | 11/30/2022 |  | 03/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2021020298 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/09/2022 |  |  |  |  | The user facility reported a 73-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. after being placed on impella cp support, no distal pulses were found bilaterally via doppler or ultrasound. an ultrasound showed bilateral lower extremity reduced flows. the impella was repositioned without significant changes. additional information was not provided regarding the resolution of the patient¿s pulses. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29898 | 1220648-2024-07844 | 18884936 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18884936&pc=OZD | 09/09/2022 | 03/14/2025 |  | 03/12/2024 |  |  | 09/09/2022 | 03/12/2024 |  |  | 07/31/2024 |  | 03/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231521 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/09/2022 |  |  |  |  | The user facility reported a 43-year-old male was implanted with an impella cp for mechanical circulatory support. it was reported while on impella cp support the patient was transferred to another unit and the device began intermittently alarming for impella in aorta. a bedside echocardiography was performed and the pump was pulled back in the aorta and repositioned. additional alarms for positioning and suction were reported, and the pump position was confirmed with transesophageal echocardiography (tee). | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29899 | 1220648-2024-07845 | 18885099 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18885099&pc=OZD | 08/15/2022 | 03/12/2024 |  | 03/12/2024 |  |  | 08/15/2022 | 03/12/2024 |  |  | 10/31/2023 |  | 03/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022149300 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/15/2022 |  |  |  |  | The user facility reported a 80-year-old male patient was implanted with an impella cp device for mechanical circulatory support. the patient required percutaneous coronary intervention due to chronic total occlusion. the patient developed bleeding and a hematoma at the impella insertion site. the physician attempted to cannulate the left main with guide but after being informed of the hematoma, the physician made the decision to explant the impella. the site was closed and pressure applied to the hematoma, however, the bleeding was not resolved. the patient was given protamine, however, bleeding was still noted. the vascular team closed the site in the catheter lab. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29900 | 1220648-2024-07846 | 18885176 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18885176&pc=OZD | 08/09/2022 | 03/14/2025 |  | 03/12/2024 |  |  | 08/09/2022 | 03/12/2024 |  |  | 03/31/2024 |  | 03/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023205053 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/09/2022 |  |  |  |  | The user facility reported a 55-year-old female was being implanted with an impella cp device for mechanical circulatory support. the complainant reported a patient presented with st elevated myocardial infarction (stemi) in atrial fibrillation rapid ventricular response (rvr). therefore, the decision for impella cp was made. an echocardiogram at bedside measured that the impella was deep. when the pump was retracted oozing was noted at all arterial and venous sheath sites. three units packed red blood cell (prbc), albumin 250 cc x2 were administered along with applying femstop to right groin. it was also noted heavy oozing from right femoral artery (rfa) site which was managed with femstop. the patient started dialysis. while on dialysis, heavy oozing was noted from rfa which was also managed by femstop. decision to explant the impella was made. after impella sheath and pump removal a clot formation in common femoral artery (cfa) requiring intervention with peripheral penumbra and stent/balloon to right cfa. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
29901 | 1220648-2024-07847 | 18885298 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18885298&pc=OZD | 09/09/2022 | 03/14/2025 |  | 03/12/2024 |  |  | 09/09/2022 | 03/12/2024 |  |  | 04/30/2024 |  | 03/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213922 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/09/2022 |  |  |  |  | The user facility reported a 55-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient developed persistent oozing at the insertion site despite almost 20 minutes of manual pressure. a femstop was applied to the access site and systemic heparin was discontinued. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records.
29902 | 1220648-2024-07848 | 18885339 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18885339&pc=OZD | 08/27/2022 | 03/14/2025 |  | 03/12/2024 |  |  | 08/27/2022 | 03/12/2024 |  |  | 12/31/2023 |  | 03/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023173001 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/27/2022 |  |  |  |  | The user facility reported a 75-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support there were concerns of a stroke. the patient was scheduled for escalation to an impella 5.5, but the family changed the status of the patient to comfort care and support was withdrawn. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
29903 | 1220648-2024-07849 | 18885474 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18885474&pc=OZD | 08/15/2022 | 03/14/2025 |  | 03/12/2024 |  |  | 08/15/2022 | 03/12/2024 |  |  | 02/29/2024 |  | 03/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200826 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/15/2022 |  |  |  |  | The user facility reported a 70-year-old female patient was implanted with an impella cp device for mechanical circulatory support after coronary artery bypass graft (cabg) and aortic valve replacement (avr) procedure. during the impella cp support, the patient experienced suction issues despite repositioning the pump multiple times. the decision was made to explant the impella cp device early to resolve. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29904 | 1220648-2024-07850 | 18885480 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18885480&pc=OZD | 08/27/2022 | 03/28/2024 |  | 03/12/2024 |  |  | 08/27/2022 | 03/12/2024 |  |  | 05/31/2024 |  | 03/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231283 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/27/2022 |  |  |  |  | The user facility reported a 83-year-old male undergoing coronary artery bypass grafting was implanted with an impella cp device for mechanical circulatory support. the patient experienced hematoma/bleeding. the hematoma was evacuated, and additional sutures were placed. the impella cp device and the dressing were removed and pressure was applied to manage the condition and the bleeding was resolved. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. this report is being filed as part of a retrospective review of historical records. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. this report is being filed as part of a retrospective review of historical records.
29905 | 1220648-2024-07851 | 18885569 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18885569&pc=OZD | 08/12/2022 | 03/12/2025 |  | 03/12/2024 |  |  | 08/12/2022 | 03/12/2024 |  |  | 03/31/2024 |  | 03/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2023207173 |  | 0048-0047 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/12/2022 |  |  |  |  | The user facility reported a 78-year-old male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support, left heart catheterization (arterial catheterization) showed stent restenosis in left anterior descending (coronary artery branch). the clot was removed and the impella was reduced to p2. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29906 | 1220648-2024-07852 | 18885625 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18885625&pc=OZD | 08/12/2022 | 03/12/2025 |  | 03/12/2024 |  |  | 08/12/2022 | 03/12/2024 |  |  | 04/30/2024 |  | 03/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2023224529 |  | 0048-0049 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/12/2022 |  |  |  |  | The user facility reported an 65-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed a clot. the patient had surgery to remove the clot and the impella device. | [5 empty] | The investigation of the reported issue has been completed. no device or logs were returned. as the necessary information was not provided, the root cause of the thrombus was not determined. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29907 | 1220648-2024-07853 | 18885644 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18885644&pc=OZD | 08/28/2022 | 03/28/2024 |  | 03/12/2024 |  |  | 08/28/2022 | 03/12/2024 |  |  | 04/30/2024 |  | 03/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218048 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/28/2022 |  |  |  |  | The user facility reported a 50-year-old male with mitral valve disease was implanted with an impella cp device for mechanical circulatory support. while on support the patient experienced site bleeding. subsequently, heparin was withheld. the peel away sheath was removed and repositioned and the patient was reported as being stable. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. this report is being filed as part of a retrospective review of historical records.
29908 | 1220648-2024-07854 | 18885826 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18885826&pc=OZD | 08/03/2022 | 03/11/2025 |  | 03/12/2024 |  |  | 08/03/2022 | 03/12/2024 |  |  | 03/31/2024 |  | 03/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023206268 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/03/2022 |  |  |  |  | The user facility reported a 78-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient developed bleeding at the impella cp insertion site. systematic heparin was discontinued overnight due to bleeding at the site and hemoglobin being at 7.6. patient was provided with packed red blood cells. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
29909 | 1220648-2024-07855 | 18885914 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18885914&pc=OZD | 08/12/2022 | 05/28/2025 |  | 03/12/2024 |  |  | 08/12/2022 | 03/12/2024 |  |  | 05/31/2024 |  | 03/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2023207124 |  | 0048-0049 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Activation, Positioning or Separation Problem | Improper Flow or Infusion | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/12/2022 |  |  |  |  | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported an 77-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the doctor noted posterior positioning of pump under echo; however, the doctor was unable to reposition the pump. there were also issues of persistent suction. the doctor elected to run the impella at p2. the patient expired during support with no allegations made towards the device. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07855 was provided in sections b2, b5, h1, and h6. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29910 | 1220648-2024-07856 | 18885986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18885986&pc=OZD | 08/15/2022 | 05/20/2025 |  | 03/12/2024 |  |  | 08/15/2022 | 03/12/2024 |  |  | 04/30/2024 |  | 03/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023223630 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/15/2022 |  |  |  |  | The user facility reported a 79-year-old male patient that was attempted to be implanted with the impella cp for mechanical circulatory support. it was reported that the impella cp was unable to cross due to notable peripheral artery disease of the patients right iliac artery. percutaneous old balloon angioplasty and change out of the impella sheath were performed successfully. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29911 | 1220648-2024-07857 | 18886011 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18886011&pc=OZD | 08/07/2022 | 09/13/2024 |  | 03/12/2024 |  |  | 08/07/2022 | 03/12/2024 |  |  | 04/30/2024 |  | 03/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213475 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Improper Flow or Infusion | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/07/2022 |  |  |  |  | It was reported prior to transfer of patient to another hospital for impella rp device or extracorporeal membrane oxygenation, patient became too unstable and expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | The user facility reported a 53-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that a patient on impella cp had right ventricle failure noted with severe low flows on the impella. due to the low flows and condition of the patient, discussions were initiated on escalating support to include an impella rp pump. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07857 was provided in sections b1, b2, b5, d6a, d6b, h1, and h6. | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records.
29912 | 1220648-2024-07858 | 18886132 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18886132&pc=OZD | 08/09/2022 | 03/12/2025 |  | 03/12/2024 |  |  | 08/09/2022 | 03/12/2024 |  |  | 12/31/2023 |  | 03/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023176524 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/09/2022 |  |  |  |  | The user facility reported a 77-year-old male patient was being implanted with an impella cp device for mechanical circulatory support. the complainant reported patient was placed on impella cp support after coding during stent placement. a clot was seen at the leaflet on the trans-esophageal echo. there was also a possible clot under the left main artery seen bouncing around. fibrin removal was performed with a plan to change to an impella 5.5. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. this report is being filed as part of a retrospective review of historical records. | The device or the logs were not returned for analysis. the root cause of thrombus could not be determined as insufficient clinical details were provided.
29913 | 1220648-2024-07859 | 18886522 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18886522&pc=OZD | 08/09/2022 | 03/12/2025 |  | 03/12/2024 |  |  | 08/09/2022 | 03/12/2024 |  |  | 01/31/2024 |  | 03/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183636 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/09/2022 |  |  |  |  | The user facility reported a 53-year-old male was being implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient presented with chest pains and was diagnosed with non-st-elevation myocardial infarction with cardiogenic shock. the patient was placed on impella cp to support percutaneous coronary intervention (pci). during the procedure, the patient developed a large hematoma at the impella cp insertion site overnight. the impella was explanted early due to the hematoma. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ . ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29914 | 1220648-2024-07860 | 18887445 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18887445&pc=OZD | 08/09/2022 | 03/12/2025 |  | 03/12/2024 |  |  | 08/09/2022 | 03/12/2024 |  |  | 04/30/2024 |  | 03/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023224527 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Activation, Positioning or Separation Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/09/2022 |  |  |  |  | The user facility reported a 53-year-old male patient was being implanted with an impella cp device for mechanical circulatory support. the complainant reported that while the patient was on impella cp support, the patient had distal ischemia on the right lower extremity (rle). the pump was repositioned at bedside via echocardiogram. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
29915 | 1220648-2024-07861 | 18887472 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18887472&pc=OZD | 08/01/2022 | 09/11/2024 |  | 03/12/2024 |  |  | 08/01/2022 | 03/12/2024 |  |  | 03/31/2024 |  | 03/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP. |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023209280 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  |  | White | Death | Yes | Yes | Positioning Problem | [6 empty] | Insufficient Information | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/01/2022 |  |  |  |  | It was noted by medical staff the patient expired while on impella support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 48-year-male with cardiomyopathy was implanted with an impella cp device for mechanical circulatory supported. it was reported that the placement of the impella was checked under echocardiogram (echo) at bedside and the impella cp appeared to have dove into the left ventricle and was laying on mitral valve. the physician attempted to pull back impella 4cm but it appeared to be attached to a papillary and unable to steer away from mitral. per recommendations, the physician decided to use a re-access port on the repositioning sheath to wire and place a new pump. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07861 was provided in sections b1, b2, b5, h1, and h6. | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records.
29916 | 1220648-2024-07862 | 18887717 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18887717&pc=OZD | 07/28/2022 | 03/12/2024 |  | 03/12/2024 |  |  | 07/28/2022 | 03/12/2024 |  |  | 02/29/2024 |  | 03/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202136 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Positioning Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/01/2022 |  |  |  |  | The user facility reported a 80-year-old female patient on an impella cp for mechanical circulatory support. while on support, the patient experienced access site bleeding around the sheath. the patient also experienced the pump placement being too deep and required repositioning. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29917 | 1220648-2024-07864 | 18887835 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18887835&pc=OZD | 08/09/2022 | 03/12/2025 |  | 03/12/2024 |  |  | 08/09/2022 | 03/12/2024 |  |  | 04/30/2024 |  | 03/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023220118 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/09/2022 |  |  |  |  | The user facility reported a 70-year-old male was being implanted with an impella cp device for mechanical circulatory support. the complainant reported that there was an access site bleed with the impella pump. there was a groin bleed noted, specifically around the sheath. new sutures were placed, and pressure dressing was applied for the bleed. the pump was repositioned, and all was well. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
29918 | 1220648-2024-07865 | 18888346 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18888346&pc=OZD | 02/14/2024 | 04/22/2024 |  | 03/12/2024 |  |  | 02/14/2024 | 03/12/2024 |  |  |  |  | 03/12/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  |  |  | 1669290 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/14/2024 |  |  |  |  | The user facility reported a 55-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that the automated impella controller (aic) had a controller error when the patient was being transported over 9000 feet elevation. the alarm did not cause any harm or injury. | [5 empty] | The investigation for the reported console (aic) red alarm has been completed. the console was returned for evaluation. during evaluation, the console had failing contrast so the optical bench was precautionarily replaced. logs were reviewed and found that 2.5 hours into the pump running, the console alarmed controller error 1040 which resolved about 20 minutes later. controller error 1040 is triggered when the ambient pressure is out of range for at least two minutes, and according to the ifu, that range is -1000ft to 8000ft which is approximated as 550mmhg. the ambient pressure two minutes before the onset of the 1040 controller error occurred when the ambient pressure fell below 550 mmhg and resolved when the pressure came back above 550 mmhg. the cause of the issue was the console going above 9000ft dropping ambient pressure below 543 mmhg. | The impella device was received from the customer and¿an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed.
29919 | 1220648-2024-07866 | 18888775 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18888775&pc=OZD | 08/17/2022 | 05/28/2025 |  | 03/12/2024 |  |  | 08/17/2022 | 03/12/2024 |  |  | 02/29/2024 |  | 03/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202348 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Perforation of Vessels | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/17/2022 |  |  |  |  | Additional information noted care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 72-year-old male undergoing coronary artery bypass grafting was implanted with an impella cp device for mechanical circulatory support. the complaint reported the patient had repaired vascular injury and hematoma, where the patient received multiple products. support continued with the implanted pump following the intervention. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07866 was provided in sections b2, b5, h1, and h6. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29920 | 1220648-2024-07867 | 18889060 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18889060&pc=OZD | 08/15/2022 | 03/12/2024 |  | 03/12/2024 |  |  | 08/15/2022 | 03/12/2024 |  |  | 04/30/2024 |  | 03/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023215605 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/15/2022 |  |  |  |  | The user facility reported a 70-year-old male patient was implanted with an impella cp device for mechanical circulatory support. while on support the patient experienced bleeding and a cold leg postoperatively. the impella pump was removed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29921 | 1220648-2024-07868 | 18889100 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18889100&pc=OZD | 08/24/2022 | 03/12/2025 |  | 03/12/2024 |  |  | 08/24/2022 | 03/12/2024 |  |  | 02/29/2024 |  | 03/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183881 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/24/2022 |  |  |  |  | The user facility reported a 62-year-old male in cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that during support, the yellow luer connection was found to be leaking. the physician replaced the purge cassette but the issue persisted. there was no reported patient harm. the pump was explanted after 6 days of support and the patient continued to be on ecmo support. | [5 empty] | The root cause of the purge leak - pump could not be determined as insufficient clinical details were provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
29922 | 1220648-2024-07869 | 18889163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18889163&pc=OZD | 02/26/2024 | 04/30/2024 |  | 03/12/2024 |  |  | 02/26/2024 | 03/12/2024 |  |  | 11/30/2025 |  | 03/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024429992 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Optical Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/26/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. shortly after implantation, the placement signal and left ventricle disappeared with a message saying the placement control not reliable. the pump was noted to be running on auto mode at 3.2 liters per minute. the patient became unstable during the main stem intervention and the decision was made to exchange the pump and the automated impella controller (aic). no patient harm was reported. | [5 empty] | The investigation for the reported optical signal issue has been completed. the impella has been returned for evaluation. functional testing was performed and the pump passed the laser continuity test. no catheter kinks were noted. psnr alarm was reproduced on the lab console. 5s parameters were taken and they did not meet specification. removing the optical sensor revealed it to be damaged with a void in the silicone. data logs showed that there are position signal not reliable (psnr) alarms at the time when ps goes out of specification. the root cause of the optical signal issue was a determined to be a damaged sensor. added- d6a/d6b in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was received from the customer and therefore, an evaluation of the device is underway. upon investigation closure, a supplemental mdr will be filed.
29923 | 1220648-2024-07870 | 18889300 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18889300&pc=OZD | 08/26/2022 | 03/12/2024 |  | 03/12/2024 |  |  | 08/26/2022 | 03/12/2024 |  |  | 05/31/2024 |  | 03/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023227242 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/26/2022 |  |  |  |  | The user facility reported a 67-year-old female noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced bleeding from the access site. the patient¿s partial thromboplastin time (ptt) was elevated and oozing from the site was noted. the purge solution was switched from heparin to bicarbonate. once the patient's ptt normalized, systemic heparin was re-started. the access site was redressed with better angle and oozing was significantly reduced. | [5 empty] | The investigation of the reported issue has been completed. the root cause of the access site bleeding is determined to be use issue. from the clinical information, it was mentioned that the access site bleeding was reduced after redressed with better angle. hence, the root cause is determined to be use issue. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29924 | 1220648-2024-07871 | 18889365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18889365&pc=OZD | 11/08/2022 | 03/12/2024 |  | 03/12/2024 |  |  | 11/08/2022 | 03/12/2024 |  |  | 04/30/2024 |  | 03/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023216535 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/08/2022 |  |  |  |  | The user facility reported a 79-year-old male with atrial fibrillation and left anterior descending (lad) coronary artery branch disease was implanted with an impella cp support. the patient developed significant bleeding at the insertion site after the impella was implanted. the bleeding was treated with angle matching, site redressed, manual pressure and femstop. additionally, the patient was switched to bicarb for purge solution. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
29925 | 1220648-2024-07872 | 18889487 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18889487&pc=OZD | 09/03/2022 | 05/28/2025 |  | 03/12/2024 |  |  | 09/03/2022 | 03/12/2024 |  |  | 01/31/2024 |  | 03/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023181591 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 09/03/2022 |  |  |  |  | Additional information noted the patient expired while on support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 61-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after the physician removed the sheath and inserted the reposition sheath, there was squirting bleeding around the reposition sheath. unless pressure was held, the site would still bleed. the vascular surgeon came up and did a cutdown to the left femoral artery and fixed the artery. the surgeon was unable to locate any obvious tears. additionally, the patient experienced quick episodes of ventricular tachycardia and was given amiodarone. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr1220648-2024-07872 was provided in sections b2, b5, h1, and h6. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29926 | 1220648-2024-07873 | 18889653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18889653&pc=OZD | 11/08/2022 | 03/12/2024 |  | 03/12/2024 |  |  | 11/08/2022 | 03/12/2024 |  |  | 04/30/2024 |  | 03/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023225616 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/08/2022 |  |  |  |  | The user facility reported a 82-year-old male was implanted with an impella cp. after being placed on impella cp support, the patient developed leg ischemia. the impella was explanted due to the ischemia. | [5 empty] | The investigation into the reported limb ischemia has been completed. the device was not returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
29927 | 1220648-2024-07874 | 18889933 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18889933&pc=OZD | 11/22/2022 | 04/22/2024 |  | 03/12/2024 |  |  | 11/22/2022 | 03/12/2024 |  |  | 07/31/2024 |  | 03/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023240846 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/22/2022 |  |  |  |  | The user facility reported a 76-year-old male noted as elective placement was implanted with an impella cp device for mechanical circulatory support. while the patient was transitioned to a bed there was a sudden drop in the impella flow and a decrease in pulsatility. the physician decided to reposition the device on the cathlab, and a kink was clearly visible (pigtails pushing towards the lateral wall). the repositioning of the pump was a success and there was a quick recovery of the impella flow. there was no patient harm. | [5 empty] | The investigation for the reported low pump flow has been completed. the impella device was not returned for evaluation. data logs show a spike in motor current and active suction alarm, followed by low pump flow during the start of the case. according to clinical records, the pump was replaced after observing the motor current spike and decrease in pump flow. the cause of the low pump flow issue was most likely biomaterial, based on data log review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29928 | 1220648-2024-07875 | 18890391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18890391&pc=OZD | 09/03/2022 | 03/12/2025 |  | 03/12/2024 |  |  | 09/03/2022 | 03/12/2024 |  |  | 05/31/2024 |  | 03/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231364 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/03/2022 |  |  |  |  | The user facility reported a (b)(6) with treatment of a large left anterior descending artery perforation was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the patient experienced access site bleeding. blood products were provided (packed red blood cells and platelets), heparin was removed from purge and bicarb was later added to the purge. the patient¿s medical team wanted the impella 5.5 removed in the operating room by vascular due to low platelet count of 45,000. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29929 | 1220648-2024-07876 | 18890535 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18890535&pc=OZD | 12/01/2022 | 03/12/2024 |  | 03/12/2024 |  |  | 12/01/2022 | 03/12/2024 |  |  | 04/30/2024 |  | 03/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213466 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | PHYSICIAN ASSISTANT | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/01/2022 |  |  |  |  | The user facility reported a 65-year-old female was implanted with an impella cp. upon initiation of impella cp support, the device seemed to have fallen back in the aorta under fluoroscopy. the pump was removed, washed with sterile water, and reinserted independently. impella support was initiated without issue. a hematoma was also noted to the right groin but was unchanged from arrival to the unit. the access site had a good angle and no oozing under the dressing. bicarb was added to purge. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
29930 | 1220648-2024-07877 | 18890603 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18890603&pc=OZD | 12/01/2022 | 03/12/2025 |  | 03/12/2024 |  |  | 12/01/2022 | 03/12/2024 |  |  | 08/31/2024 |  | 03/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023269580 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/01/2022 |  |  |  |  | The user facility reported a 50-year-old male was implanted with an impella cp. after being placed on impella cp support, the groin site was oozing. the patient was switched to sodium bicarbonate in the purge. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | The patient had oozing from the groin and high acts after being transferred to the icu. this was addressed by stopping heparin and switching to bicarb. the root cause of the access site bleed is therefore anticoagulation management.
29931 | 1220648-2024-07878 | 18890662 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18890662&pc=OZD | 09/04/2022 | 03/12/2024 |  | 03/12/2024 |  |  | 09/04/2022 | 03/12/2024 |  |  | 12/31/2023 |  | 03/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023180196 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/04/2022 |  |  |  |  | The user facility reported a 51-year-old male was implanted with an impella 5.5 device for mechanical circulatory support after the patient underwent a pci of the circumflex. while under impella cp support, the patient was complaining of numbness and tingling in the right lower extremity. the physician placed a re-perfusion catheter on the patient. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29932 | 1220648-2024-07879 | 18890760 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18890760&pc=OZD | 12/04/2022 | 05/13/2025 |  | 03/12/2024 |  |  | 12/04/2022 | 03/12/2024 |  |  | 07/31/2024 |  | 03/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023256846 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | Positioning Problem | [5 empty] | Cardiac Perforation | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/04/2022 |  |  |  |  | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported an 86-year-old female requiring an impella cp device for mechanical circulatory support. it was reported that while on impella cp the patient experienced frequent suction alarms due to a positioning issue. the pump was unable to be repositioned as multiple attempts ended with the pump pulling back into the aorta. during a reposition attempt it was noted that patient had developed sudden pericardial effusion/cardiac tamponade. the patient was on impella support for 8.39 hours before they expired. no allegations were made towards the impella cp for cause of death. it was determined that there was a probable left ventricular rupture that caused the effusion/tamponade. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07879 was provided in sections b2, b5, h1, and h6. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29933 | 1220648-2024-07880 | 18890776 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18890776&pc=OZD | 12/13/2022 | 03/12/2025 |  | 03/12/2024 |  |  | 12/13/2022 | 03/12/2024 |  |  | 08/31/2024 |  | 03/12/2024 |  | TEMPORARY NON-ROLLER TYPE CARDIAC SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023267912 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/13/2022 |  |  |  |  | The user facility reported a 59-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while the patient was on impella cp support, the patient had active bleeding at the right groin insertion site. the impella was removed. perclose was used twice for closure. good hemostasis was achieved. the device was explanted. | [5 empty] | The root cause of the access site bleeding cannot be determined since the product was not returned and enough clinical info was not provided. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. assess access site for bleeding and hematoma. remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site. instructions for use for the related event are as follows.
29934 | 1220648-2024-07881 | 18890814 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18890814&pc=OZD | 12/13/2022 | 03/12/2025 |  | 03/12/2024 |  |  | 12/13/2022 | 03/12/2024 |  |  | 08/31/2024 |  | 03/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023270970 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/13/2022 |  |  |  |  | The user facility reported a 73-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that after the physician sutured the repo sheath in place with forward tension, a small tract ooze was noted at the access site. the physician placed another skin stitch to stop the oozing. a dressing was also placed to maintain the angle of insertion. upon arrival to the unit, a small hematoma was noted at the access site. the cardiac catheterization lab staff came to the bedside to assist in compressing it. after pressure was held, the groin site was clean, dry and intact. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
29935 | 1220648-2024-07763 | 18873361 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873361&pc=PYX | 02/26/2024 | 04/16/2024 |  | 03/11/2024 |  |  | 02/26/2024 | 03/11/2024 |  |  | 02/28/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023199599 |  | 004334 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/26/2024 |  |  |  |  | The user facility reported a 66-year-old male undergoing cardiac surgery was implanted with an impella rp device for mechanical circulatory support. it was reported that while preparing the impella rp, no purge would exit the pump. the purge was coming out of the purge line prior to connecting to the purge cassette. the impella was exchanged. | [5 empty] | H.3 revised as previously entered incorrectly. | The impella device was received from the customer and therefore, an evaluation of the device is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation into the purge cassette failure has been completed since the original report was filed. the pump was returned, tested, and evaluated. there was no issue found during the case. the device functioned/operated to specification.
29936 | 1220648-2024-07764 | 18873408 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873408&pc=OZD | 02/29/2024 | 03/21/2024 |  | 03/11/2024 |  |  | 02/29/2024 | 03/11/2024 |  |  |  |  | 03/11/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1339730 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/29/2024 |  |  |  |  | The user facility biomed department reported that the automated impella controller (aic) had a battery failure. the aic was exchanged. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the reported console (aic) red alarm issue has been completed. the aic was returned for evaluation. the battery failure alarm issue was able to be reproduced. the lcd charge level indicator on battery 1 was malfunctioning, which was determined to be caused by battery internal cell imbalance. logs were reviewed and also confirmed the battery failure on the console. therefore, the root cause of this issue was battery internal cell imbalance. h.3 revised as previously entered incorrectly.
29937 | 1220648-2024-07765 | 18873442 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873442&pc=OZD | 02/26/2024 | 04/04/2024 |  | 03/11/2024 |  |  | 02/29/2024 | 03/11/2024 |  |  |  |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025441927 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/29/2024 |  |  |  |  | The user facility reported a male of an unknown age and ethnicity with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that on day six of support, the patient developed access site bleeding requiring a drain placement and withdrawing of heparin as treatment. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was most likely due to anticoagulation management since the bleed was treated by withdrawing the heparin. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
29938 | 1220648-2024-07766 | 18873478 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873478&pc=OZD | 03/01/2024 | 04/04/2024 |  | 03/11/2024 |  |  | 03/01/2024 | 03/11/2024 |  |  |  |  | 03/11/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1096321 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/01/2024 |  |  |  |  | An observation was made by the biomed department of a red failure/battery failure. there was no patient involved. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the aic - battery failure (red alarm) has been completed since the initial report was submitted. the console was returned, visually inspected, and tested. the reported failure mode was confirmed. the failure mode was reproduced. the root cause of the battery failure alarm was cell imbalance with cell 2 of battery 2.
29939 | 1220648-2024-07767 | 18873524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873524&pc=OZD | 03/02/2024 | 04/19/2024 |  | 03/11/2024 |  |  | 03/02/2024 | 03/11/2024 |  |  | 11/30/2025 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024422136 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  |  | White | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/02/2024 |  |  |  |  | The user facility reported a 64-year-old female of unknown ethnicity in acute myocardial infarction cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that three hours after switching from heparin purge to bicarb purge the automated impella controller (aic) started to alarm for high purge pressure and low purge flow. this was in conjunction with an impella position unknown alarm. the impella 5.5 was exchanged. the patient is stable. | [5 empty] | The investigation into the high purge pressure issue has been completed since the original report was filed. the pump was returned, tested, and evaluated. there was light biomaterial found around the impeller blade. however, the failure couldn't be reproduced. the root cause of the high purge pressure issue was not determined since insufficient data logs were returned around the time of the failure. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29940 | 1220648-2024-07673 | 18873547 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873547&pc=OZD | 12/11/2022 | 05/13/2025 |  | 03/11/2024 |  |  | 12/11/2022 | 03/11/2024 |  |  | 04/30/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023221767 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Malposition of Device | [6 empty] | Hemolysis | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/11/2022 |  |  |  |  | The user facility reported a 59-year-old female had a cp support and the patient experienced hemolysis and dark red urine. the impella cp was repositioned under echo. the patient was weaned off pressors and pulses improved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined correction section h6: health effect - clinical code was inadvertently left out in error in the initial report which has been added to this report.
29941 | 1220648-2024-07679 | 18873548 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873548&pc=OZD | 11/13/2022 | 03/13/2025 |  | 03/11/2024 |  |  | 11/13/2022 | 03/11/2024 |  |  | 02/28/2023 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022054803 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 95 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/13/2022 |  |  |  |  | The user facility reported a 95-year-old male in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. after being placed on impella cp support, the patient experienced limb ischemia on the left leg due to the impella sheath. a femorofemoral bypass was performed to treat the limb ischemia. the patient continued on support until the device was weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29942 | 1220648-2024-07675 | 18873549 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873549&pc=OZD | 09/25/2023 | 03/07/2024 |  | 03/11/2024 |  |  | 09/25/2023 | 03/11/2024 |  |  | 10/31/2023 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022157652 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/25/2023 |  |  |  |  | The user facility reported an impella 5.5 implanted in a 62-year-old male patient for mechanical circulatory support. it was reported that the impella 5.5 pump was unable to pass through the subclavian, after a few attempts the decision was made to go direct aortic. impella insertion was successful. there was no patient harm. | [5 empty] | The investigation into the reported delivery issue has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine a root cause. the most likely cause of the delivery issue was due to patient condition. this report is being filed as part of a retrospective review of historical records.
29943 | 1220648-2024-07676 | 18873550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873550&pc=OZD | 12/11/2022 | 03/07/2024 |  | 03/11/2024 |  |  | 12/11/2022 | 03/11/2024 |  |  | 02/28/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2023198892 |  |  |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/11/2022 |  |  |  |  | The user facility reported a 60-year-old female had a cp support. while implanted, the impella cp was having decreased diastolic flows and competing with the intra-aortic balloon pump (iabp), so the decision was made to remove the iabp. prior to leaving the operating room, a right ventricle dysfunction was noted on trans-esophogeal echo (tee). there was also a slight hemolysis noted in the patient¿s urine and there were a few repositioning issues and suction events during placement and while the iabp was still in place. the impella dressing was also bloody. groin dressing was saturated but no active oozing and the dressing was changed. prior to changing the dressing, blood products were given. the patient continued to bleed (putting out 1500ml overnight despite no anticoagulation) and heparin induced thrombocytopenia (hit) was suspected. the patient continued to receive additional blood products. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29944 | 1220648-2024-07677 | 18873551 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873551&pc=OZD | 11/13/2022 | 03/13/2025 |  | 03/11/2024 |  |  | 11/13/2022 | 03/11/2024 |  |  | 06/30/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023246464 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/13/2022 |  |  |  |  | The user facility reported a 50-year-old female noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient developed bleeding at the groin access site and the interventional cardiologist (ic) was unable to control the bleeding in the groin. the impella cp device was removed and the sheath remained in place. vascular surgeon was consulted to remove the impella sheath and surgically close site. hemostasis was achieved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
29945 | 1220648-2024-07683 | 18873552 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873552&pc=OZD | 07/02/2023 | 03/07/2024 |  | 03/11/2024 |  |  | 07/02/2023 | 03/11/2024 |  |  |  |  | 03/11/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2024345522 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 07/02/2023 |  |  |  |  | The user facility reported a 70-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the softkeys on the automated impella controller (aic) were no longer working during patient support. the aic was swapped to resolve the issue. there was no patient harm. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not returned for evaluation. the console was visually inspected with little residue in the purge area. extensive testing couldn't reproduce the failure without the realm of briefly trying to access the softkeys while keyboard: high guard level is active. this was done by holding down multiple soft keys while trying to activate another simultaneously. the button is not available for as long as the other two are being pressed. this is within specification of the aic. the root cause for the failure was not reproduced since the failure couldn't be reproduced outside the specification of the aic. no apparent issues were found that could be related to pressing the rotary push button except for the one instance of the kbd high guard activating. as a precaution, field service was instructed to replace glass kbd and preventative maintenance was performed. this report is being filed as part of a retrospective review of historical records.
29946 | 1220648-2024-07684 | 18873553 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873553&pc=OZD | 12/29/2023 | 09/20/2024 |  | 03/11/2024 |  |  | 02/12/2024 | 03/11/2024 |  |  | 09/30/2025 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024407432 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Multiple Organ Dysfunction Syndrome | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | Followup | [5 empty] | Updated information received from the study site: the subject developed post-operatively liver and kidney failure, as well as respiratory insufficiency in combination with va-extracorporeal membrane oxygenation (ecmo); resulted in patient's death. the subject also experienced access site bleeding. | Information from the study reported a 70-year-old patient received support from an impella cp smartassist heart pump and extracorporeal membrane oxygenation (ecmo) due to acute myocardial infarction / cardiogenic shock. the patient was reported to develop postoperative liver and kidney failure as well as respiratory insufficiency. the patient was intubated and treated with medication but ultimately died. the study doctor's initial assessment is that the event may be related to the impella heart pump or the impella procedure. no further information is currently available. | Further information was reported that the patient had access site bleeding and received a transfusion. | Per medical safety officer review: use of the device cannot conclusively be associated with the outcome of death. patient's peri-procedural condition was critical. |  |  | The investigation for the reported renal failure has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the renal failure could not be determined. | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the instructions for use for the impella cp with smart assist system state the following: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). b5 updated with additional information. h6 added health effect - clinical code 1888. h6 added impact code 4641. h6 added component code 925. d6a and d6b added h6 clinical code 2041 changed to 3261 in manufacturer device report 1220648-2024-07684. h1-corrected type of reportable event in supplemental report 1220648-2024-07684-01 to death. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Updated b5 with additional information that was received. added h6 health effect - clinical code, medical device problem code, health effect - impact code h6 component code missing added 925. | B5 additional information added b2 outcomes: death should have been blank. h1-corrected type of reportable event to serious injury on manufacturer device report. h6 health effect - impact code 1802 reported incorrectly on manufacturer device report.
29947 | 1220648-2024-07686 | 18873554 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873554&pc=OZD | 02/18/2024 | 04/19/2024 |  | 03/11/2024 |  |  | 02/18/2024 | 03/11/2024 |  |  | 11/30/2025 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024421391 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support, the 60-year-old male patient had massive access site bleeding requiring transfusion. the bleeding occurred after 4 larger introducers were implanted. the sheath was removed and bleeding was managed. the patient ultimately expired after 3 days on support. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of access site bleeding was determined to be use issue because it was clearly mentioned that 4 large introducer was placed that was causing the bleed. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
29948 | 1220648-2024-07678 | 18873555 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873555&pc=OZD | 02/18/2024 | 09/24/2025 |  | 03/11/2024 |  |  | 02/18/2024 | 03/11/2024 |  |  | 10/31/2025 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025452657 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that during impella 5.5 support for 56-year-old male patient, the impella position was noted to be 3.7cm. therefore, repositioning to 5cm was recommended. there was no reported patient harm. | The investigation uncovered placement signal issues. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported positioning issue and placement signal issue has been completed. the device was not returned for evaluation. data logs were not returned. however, clinical information provided is sufficient to determine the root cause. according to the clinical, on the 9th day of impella 5.5 support ps and lv curves were lost when increasing the p-level from p8 to p9. a placement signal not reliable alarm was recorded and no attempts were made to resolve the issue. it was also noted that the impella was positioned at 3.7cm. a recommendation was made to move the pump to 5cm followed by the decision to leave the pump in the patient. the cause of the optical signal issue could not be determined because no product or logs were returned and not enough information was given. the most likely cause of the positioning issue was use related. | B5 updated to include optical signal issue. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-07678. b7 other relevant history, including preexisting medical conditions updated. d6a/d6b updated with additional information. d10 concomitant medical products and therapy dates updated. g1 reporting contact email and g1 manufacturing site fax number updated.
29949 | 1220648-2024-07685 | 18873556 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873556&pc=OZD | 02/18/2024 | 04/08/2024 |  | 03/11/2024 |  |  | 02/18/2024 | 03/11/2024 |  |  | 11/30/2025 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024427470 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | No | Yes | Inaccurate Delivery | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 59-year-old female undergoing cardiac surgery was to be implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 was attempted to be delivered via the axillary site through a 10mm graft. the team used all troubleshooting measures but the impella 5.5 would not deliver, and was replaced with an impella cp pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into low delivery issues has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the delivery issue was most likely due to patient condition since the pump would not advance in subclavian artery dove behind clavicle.
29950 | 1220648-2024-07680 | 18873557 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873557&pc=OZD | 05/18/2022 | 03/07/2024 |  | 03/11/2024 |  |  | 12/08/2023 | 03/11/2024 |  |  | 02/29/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023195961 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/08/2023 |  |  |  |  | The user facility reported a 43-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the medical team attempting to implant an impella 5.5 was having difficulty delivering the pump due to the patient's vascular anatomy. after several attempts they "banged on the device" and determined that it would have been better to use a different pump. they proceeded with a new impella 5.5. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition. it was reported that due to the tortuous anatomy of the patient, the device was manipulated roughly during the attempt to implant it. this report is being filed as part of a retrospective review of historical records.
29951 | 1220648-2024-07674 | 18873558 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873558&pc=OZD | 05/02/2023 | 03/07/2024 |  | 03/11/2024 |  |  | 05/02/2023 | 03/11/2024 |  |  |  |  | 03/11/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1198725 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 05/02/2023 |  |  |  |  | The user facility reported a 29-year-old female with low cardiac output syndrome and post cardiotomy cardiogenic shock was being transitioned from an impella cp device to an impella 5.5 device for increased mechanical circulatory support. while the patient was on support, the automated impella controller (aic) failed to recognize a purge cassette. the pump was moved to a new automated impella controller and these issues resolved. | [5 empty] | The automated impella controller was returned to abiomed and evaluated. the reported issues of ¿purge disk not detected¿ was confirmed in the device logs. an inspection of the automated impella controller revealed a broken (cracked) purge disk retainer, which would have contributed to the issue in detecting the disk. the root cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
29952 | 1220648-2024-07692 | 18873559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873559&pc=OZD | 02/21/2024 | 04/09/2024 |  | 03/11/2024 |  |  | 02/21/2024 | 03/11/2024 |  |  | 11/30/2025 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024433278 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/21/2024 |  |  |  |  | The user facility reported a 61-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had cp support that was ongoing for a patient with multiple cardiac comorbidities. the pump was placed but the limb became ischemic on day 2 of the support. the team placed perfusion sheaths/catheter and the team continued the cp support. no explant or further intervention. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
29953 | 1220648-2024-07688 | 18873560 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873560&pc=OZD | 02/20/2024 | 04/23/2024 |  | 03/11/2024 |  |  | 02/20/2024 | 03/11/2024 |  |  | 11/30/2025 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024435208 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient of unknown age and gender undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that following the impella 5.5 being implanted, the automated impella controller (aic) displayed high purge pressure alarms. the impella 5.5 was explanted and another device was used. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation into the high purge pressure issue has been completed since the original report was submitted. the device was not returned for investigation. data log analysis confirmed a gradual uptick in purge pressure with two periods of relief corresponding to an increase in purge flow. however, it could not be determined whether this was due to a kinked purge line or biomaterial. the cause of the high purge pressure could not be determined because no product was returned, and not enough clinical information was given.
29954 | 1220648-2024-07690 | 18873561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873561&pc=OZD | 02/21/2024 | 02/21/2025 |  | 03/11/2024 |  |  | 02/21/2024 | 03/11/2024 |  |  | 08/31/2025 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025443708 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Deformation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/21/2024 |  |  |  |  | The user facility reported a 52-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that a .018 guidewire was kinked prior to use. the guidewire was discarded, and a new guidewire was utilized for the planned impella cp implant. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the product damage has been completed. no product was returned for evaluation. clinical details noted that 0.0180" guidewire was kinked prior to using. a new 0.0180" guidewire was taken from a separate impella kit and was used to successfully insert impella cp. the root cause of the product damage was unable to be determined as no product was returned for analysis.
29955 | 1220648-2024-07682 | 18873562 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873562&pc=OZD | 06/27/2023 | 03/07/2024 |  | 03/11/2024 |  |  | 06/27/2023 | 03/11/2024 |  |  | 02/28/2025 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024327039 |  | 1000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Difficult to Insert | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/27/2023 |  |  |  |  | The user facility reported a 54-year-old male undergoing cardiac surgery was attempted to be implanted with an impella 5.5 for mechanical circulatory support. the device did not pass from the right subclavian to the bifurcation of the brachiocephalic artery, with possibility of a catheter kink. it was reported that the implantation aborted as a result of the patient vasculature anatomy. there was no patient harm reported. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. as per clinical review, impella 5.5 could not be placed due to anatomical issues at brachiocephalic artery with sandwich-like calcification observed and catheter kinked during insertion. the root cause of the delivery issue was patient condition related with catheter kink occurring due to resistance at brachiocephalic artery with calcified vessel conditions. this report is being filed as part of a retrospective review of historical records.
29956 | 1220648-2024-07689 | 18873563 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873563&pc=OZD | 02/20/2024 | 04/22/2024 |  | 03/11/2024 |  |  | 02/20/2024 | 03/11/2024 |  |  | 10/31/2025 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024414358 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that impella cp was explanted as 48-year-old male patient was diagnosed with cerebral hemorrhage. heparin was stopped. | [5 empty] | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. added- h6 impact code 4641. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
29957 | 1220648-2024-07681 | 18873564 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873564&pc=OZD | 11/11/2022 | 03/13/2025 |  | 03/11/2024 |  |  | 11/11/2022 | 03/11/2024 |  |  | 07/31/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023258761 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Insert | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/11/2022 |  |  |  |  | The user facility reported a 76-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that the medical team experienced difficulty implanting an impella cp and the right side insertion was abandoned due to complex body habitus. the impella cp device was successfully placed through the left femoral artery. a hematoma was noted and compressed, and the physician decided to explant due to groin site instability. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29958 | 1220648-2024-07701 | 18873565 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873565&pc=OZD | 02/23/2024 | 04/10/2024 |  | 03/11/2024 |  |  | 02/23/2024 | 03/11/2024 |  |  | 11/30/2025 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025438846 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/23/2024 |  |  |  |  | The user facility reported a 81-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that complications were encountered during attempted delivery of an impella cp pump. it was noted that both femoral sites were accessed during the attempted implant, but both common iliac arteries were heavily calcified and stenosed. as the impella was unable to advance through either site, the decision was made to abort the implant. there was no patient harm. | [5 empty] | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue was most likely patient condition related, both femoral sites were accessed during the attempted implant but both common iliac arteries were heavily calcified and stenosed as per the clinical description provided. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29959 | 1220648-2024-07694 | 18873566 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873566&pc=OZD | 02/15/2024 | 04/23/2024 |  | 03/11/2024 |  |  | 02/21/2024 | 03/11/2024 |  |  |  |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025437996 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/21/2024 |  |  |  |  | The user facility reported a 58-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the guidewire was damaged. the floppy part of the guidewire broke. when the placement wire was shaped into a j-curve as usual and when the impella was placed, the coil in the floppy part, near the boundary between the stiff part and the floppy part, stretched. there was no effect on the patient as a spare placement wire was used. | [5 empty] | The investigation into the product damage (guidewire damage - peripheral vessels) has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the issue was not determined since product was not returned for investigation. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29960 | 1220648-2024-07691 | 18873567 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873567&pc=OZD | 09/11/2022 | 03/13/2025 |  | 03/11/2024 |  |  | 09/11/2022 | 03/11/2024 |  |  | 09/30/2023 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022137555 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 91 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Activation, Positioning or Separation Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/11/2022 |  |  |  |  | The user facility reported a 91-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed a hematoma and bleeding at the implantation site. the patient was not on systemic heparin. manual pressure was held which reduced the size of the hematoma, but the groin was still oozing. the site was redressed, rechecked, and it was dry and intact. additionally, p levels were reduced due to continuous suction. an echocardiogram was performed for an aorta alarm and the pump was successfully repositioned per waveforms. subsequently, the device was explanted. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29961 | 1220648-2024-07693 | 18873568 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873568&pc=OZD | 08/07/2022 | 03/13/2025 |  | 03/11/2024 |  |  | 08/07/2022 | 03/11/2024 |  |  | 01/31/2023 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | IMPELLA CP |  |  | US |  | 2022037660 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/07/2022 |  |  |  |  | The user facility reported a 40-year-old male with cardiomyopathy was implanted with an impella cp to rest the myocardium and to wean off of pressers. the patient developed bleeding around the repositioning sheath. several sutures were placed and the bleeding was stopped. femstop was also placed to prevent any hematoma from developing. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29962 | 1220648-2024-07687 | 18873569 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873569&pc=OZD | 02/18/2024 | 04/19/2024 |  | 03/11/2024 |  |  | 02/19/2024 | 03/11/2024 |  |  | 01/31/2025 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024322616 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 39 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support for 39-year-old patient, the pump was attempted to be repositioned due to suction. however, while repositioning the pump was pulled above the aortic valve, leading to pump having to be explanted. there was no reported patient harm. | [5 empty] | The investigation into the positioning issues has been completed since the original report was submitted. the device was not returned for investigation. the cause of the positioning issue most likely was use related due to improper technique as the pump pulled above aortic valve and was unable to repositioned. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29963 | 1220648-2024-07702 | 18873570 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873570&pc=OZD | 12/15/2022 | 03/13/2025 |  | 03/11/2024 |  |  | 12/15/2022 | 03/11/2024 |  |  | 07/31/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023255562 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/15/2022 |  |  |  |  | The user facility reported a 60-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. an angiogram revealed that there was an occlusion in the left anterior descending in the prox-mid portion. penumbra was used for several passes. a drug-eluting stent was placed prox-mid lad. thrombolysis in myocardial infarction 3 flow established. patient was sent to unit for further unloading per protocol. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29964 | 1220648-2024-07704 | 18873571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873571&pc=OZD | 02/21/2024 | 11/20/2025 |  | 03/11/2024 |  |  | 02/21/2024 | 03/11/2024 |  |  |  |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1275228 |  | 0042-0000-EU-L |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Power Problem | Unintended Application Program Shut Down | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/21/2024 |  |  |  |  | The user facility reported a 54-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that the automated impella controller (aic) shut off during support. it then powered back on with controller error alarm (yellow). the aic was replaced. | [5 empty] | Data analysis: imc logs confirmed the reported failure mode given there was a controller error alarms (opm comm crc error ¿ event set # 1045) triggered during the clinical procedure. there is no evidence of the console unexpectedly shutting down on the day of occurrence. device analysis: console (b)(4) was sent back fs for further investigation. conclusion: this previously investigated failure mode is characterized by values of all signals received from optical bench (placement, snr, contrast, lamp, gain) being shown in rtlog data as -16384. it lasts between few seconds to several minutes, after which all signals go back to normal. the failure has been reproduced using a an optical bench simulator and the root cause investigated to be a deficiency of the communication protocol of the optical bench. the controller error and loss of placement signal is due to an optical bench fw communication protocol deficiency, resulting in misalignment of data communication packets received by epc. console sw operated as designed. placement monitoring algorithms are suspended if placement signal is lost or considered unreliable which can lead to the misinterpretation of the console having shut off. d4: corrected the catalog number. g1: corrected the manufacturer email address. h3: added information regarding the investigation status. h6: added codes per the results of the investigation. h10: added the results of the investigation. | The impella device was received from the customer and¿an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed.
29965 | 1220648-2024-07698 | 18873572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873572&pc=OZD | 02/21/2024 | 04/09/2024 |  | 03/11/2024 |  |  | 02/21/2024 | 03/11/2024 |  |  | 06/30/2025 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024428504 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Activation, Positioning or Separation Problem | Difficult to Advance |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | Followup |  |  |  | 02/21/2024 |  |  |  |  | The user facility reported a (b)(6)-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the physician had significant resistance at the iliac bifurcation when placing the impella cp. impella could not advance, and the case was aborted. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely due to patient condition since physician met significant resistance at the iliac bifurcation and significant calcification was noted on angio. added-h6 device code 2682 d4-updated model number | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29966 | 1220648-2024-07700 | 18873573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873573&pc=OZD | 08/07/2022 | 05/20/2025 |  | 03/11/2024 |  |  | 08/07/2022 | 03/11/2024 |  |  | 04/30/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023216539 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/07/2022 |  |  |  |  | The user facility reported a 60-year-old male undergoing coronary artery bypass grafting was implanted with an impella cp device for mechanical circulatory support. while on support, the heart pump catheter in the patient had migrated back into the aorta with the pigtail across the aortic valve. after failed attempts to advance the catheter back into left ventricle the impella cp was explanted and the patient was successfully implanted with another impella cp device. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29967 | 1220648-2024-07695 | 18873574 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873574&pc=OZD | 02/21/2024 | 03/25/2024 |  | 03/11/2024 |  |  | 02/21/2024 | 03/11/2024 |  |  |  |  | 03/11/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1102528 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/21/2024 |  |  |  |  | The user facility reported a 76-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a console with a battery failure. the battery failure was a communication one and low battery. the rn stated the team would replace the console per ifu recommendations. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console (aic) yellow battery alarm has been completed. the aic was returned for evaluation. visual inspection found contamination and minor damage to the optical pump connector. functional testing could not reproduce the reported alarm. it was likely that the alarms were due to a momentary breakdown of communication between the batteries and pbm. logs were reviewed which confirm the battery alarms on the console. because the alarms could not be reproduced, the root cause of the alarms could not be determined.
29968 | 1220648-2024-07696 | 18873575 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873575&pc=OZD | 11/17/2022 | 03/13/2025 |  | 03/11/2024 |  |  | 11/17/2022 | 03/11/2024 |  |  | 07/31/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023249659 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/17/2022 |  |  |  |  | The user facility reported a 77-year-old female with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding at the impella cp access site. patient received one unit of platelets and the site was redressed. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
29969 | 1220648-2024-07703 | 18873577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873577&pc=OZD | 02/24/2024 | 04/12/2024 |  | 03/11/2024 |  |  | 02/24/2024 | 03/11/2024 |  |  | 05/31/2025 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024351868 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/24/2024 |  |  |  |  | The user facility reported a 74-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a 5.5 placed via the axillary access graft for a patient admitted in cgs (cardiogenic shock). the pump was supporting but with an access site bleed/ooze. the team had to troubleshoot with pressure, redressing, and new sutures x3. the pump remains on for support while team decided upon surgery or tavr (transcatheter aortic valve replacement). the patient instead had care withdrawn the patient expired. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical details provided was sufficient to determine the root cause. the root cause of the access site bleeding was determined to be use issue because the bleed was controlled upon adding additional sutures around the graft surroundings. d4-updated model number.
29970 | 1220648-2024-07697 | 18873578 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873578&pc=OZD | 08/01/2022 | 03/13/2025 |  | 03/11/2024 |  |  | 08/01/2022 | 03/11/2024 |  |  | 04/30/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023223972 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/01/2022 |  |  |  |  | The user facility reported a 56-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was very agitated, moving and thrashing in bed, and developed a hematoma at the impella cp insertion site (groin). the hematoma was expressed by the nurse and femstop was added. the device remained in place. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
29971 | 1220648-2024-07699 | 18873579 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873579&pc=OZD | 11/15/2022 | 03/27/2024 |  | 03/11/2024 |  |  | 11/15/2022 | 03/11/2024 |  |  | 08/31/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023264018 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/15/2022 |  |  |  |  | The user facility reported a 67-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after repositioning the sheath, the patient experienced bleeding despite proper angle match. a forward tension suture was applied and femstop was applied to the access site. the patient later expired while on the impella cp support. no allegations were made towards the impella cp for cause of death. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely due to use issue since the hemostasis was achieved by femstop along with forward tension suture. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29972 | 1220648-2024-07706 | 18873581 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873581&pc=OZD | 02/21/2024 | 04/10/2024 |  | 03/11/2024 |  |  | 02/21/2024 | 03/11/2024 |  |  | 12/31/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024315888 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Undesired Nerve Stimulation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/21/2024 |  |  |  |  | The user facility reported a 44-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient experienced some moderate pain in their right shoulder following impella 5.5 implant via the right axillary access. with scant oozing at the site, the patient was taken to the operating room where a pocket washout was performed on a notable hematoma. heparin was temporarily put on hold in response to the condition. after the washout, the hematoma was no longer putting pressure on the nerve and the pain subsided. support continued with the implanted pump. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation into the access site bleeding ¿ major and the nerve compression issue has been completed since the original report was submitted. the cause of the access site bleeding issue was not determined since product was not returned and due to insufficient clinical information. the cause of the nerve compression issue was not determined due to insufficient clinical information.
29973 | 1220648-2024-07768 | 18873609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873609&pc=OZD | 03/02/2024 | 04/18/2024 |  | 03/11/2024 |  |  | 03/02/2024 | 03/11/2024 |  |  |  |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025441338 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/02/2024 |  |  |  |  | The user facility reported a 44-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that after one day of support, the limb accessed became ischemic and the impella cp was explanted, and support was upgraded to the impella 5.5 pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
29974 | 1220648-2024-07769 | 18873692 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873692&pc=PYX | 03/03/2024 | 03/11/2024 |  | 03/11/2024 |  |  | 03/03/2024 | 03/11/2024 |  |  | 08/31/2025 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024397374 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/03/2024 |  |  |  |  | The user facility reported a 41-year-old male undergoing cardiac surgery was implanted with an impella rp flex device for mechanical circulatory support. it was reported that the impella rp flex flows have decreased after a spike in motor current with return to baseline. the impella rp flex was explanted due to suspected clot entrainment. | [5 empty] | The investigation was completed. the logs showed a high motor current pump stop on the second day of support which restarted two seconds later. the pump ran successfully for 66 hours after the stop. there were multiple instances of motor current spikes with a slight placement signal shift and speed deviation, which indicated biomaterial ingestion. on the last day of support while on p-7, there was a motor current spike and shift in placement signal followed by an immediate drop in flows of over one liter. upon investigation of the returned pump, a large clot was observed wrapped around the impeller. the root cause of the low pump flow was determined to be biomaterial.
29975 | 1220648-2024-07770 | 18873740 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873740&pc=OZD | 03/03/2024 | 04/19/2024 |  | 03/11/2024 |  |  | 03/03/2024 | 03/11/2024 |  |  |  |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025439519 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/03/2024 |  |  |  |  | The user facility reported a 74-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the impella cp was explanted due to thrombosis. there was no harm or injury from the thrombus burden. the pump was replaced. | [5 empty] | The investigation into the thrombus issues has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29976 | 1220648-2024-07771 | 18873858 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18873858&pc=OZD | 03/05/2024 | 10/13/2025 |  | 03/11/2024 |  |  | 03/05/2024 | 03/11/2024 |  |  |  |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1687316 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/05/2024 |  |  |  |  | The user facility reported a 56-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that the automated impella controller (aic) had a frozen screen, and the soft keys were not working. troubleshooting was unsuccessful. the aic was exchanged. there was no harm to the patient. | [5 empty] | The impella device was received from the customer and therefore,¿an evaluation of the device is underway. upon investigation closure, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) ¿ kbd/rotary knob issues has been completed. the impella device was returned for evaluation. data analysis: the logs were reviewed, and no relevant information was found in the logs. the console would not boot up so no logs were available from the reported complaint date/time, only from prior. device analysis: the console was returned for further investigation. the reported complaint was reproduced. the reported complaint was determined to be the corruption of the compact flash cards installed in the epc. during testing, the compact flash cards were replaced with known-good compact flash cards, and the issue was resolved. before returning the console to the customer, field service was instructed to replace the compact flash cards and to perform a software update. the reported complaint was determined to be the corruption of the compact flash cards installed in the epc. d2 revised common device name since the initial manufacturer device report number 1220648-2024-07771 was submitted in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2024-07771 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-07771 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2024-07771.
29977 | 1220648-2024-07772 | 18874519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18874519&pc=OZD | 02/22/2024 | 04/10/2024 |  | 03/11/2024 |  |  | 02/23/2024 | 03/11/2024 |  |  | 09/30/2025 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025453966 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 24 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/23/2024 |  |  |  |  | The user facility reported a 24-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient was on impella 5.5 support due to needing a bridge to a heart transplant. while on support, the patient had oozing at the access site. a wound vacuum was performed. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding issue was not determined since product was not returned and limited information was provided in the clinical description. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
29978 | 1220648-2024-07773 | 18874619 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18874619&pc=OZD | 02/25/2024 | 04/18/2024 |  | 03/11/2024 |  |  | 02/25/2024 | 03/11/2024 |  |  | 11/30/2025 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024427356 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | Pressure Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/25/2024 |  |  |  |  | The user facility reported a 48-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that there were low flows and high purge pressure. the purge cassette was replaced. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the high purge pressure and the low pump flow issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the high purge pressure and the low flow was not determined since product and data logs were not returned.
29979 | 1220648-2024-07774 | 18874943 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18874943&pc=OZD | 02/25/2024 | 04/16/2024 |  | 03/11/2024 |  |  | 02/25/2024 | 03/11/2024 |  |  | 11/30/2025 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024436335 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 23 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Pressure Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/25/2024 |  |  |  |  | The user facility reported a 23-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the implanted impella cp pump began to experience decreased purge flows during patient support. the issue coincided with a purge system blocked alarm and was followed by the pump unexpectedly stopping. the device could not be restarted and the decision was made to replace the pump. there was no patient harm. | [5 empty] | The investigation into the pump stop and the high purge pressure issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the pump stop was not determined since product was not returned and data logs are insufficient. the root cause of the high purge pressure was most likely a kinked purge line, but the exact location of the kink could not be determined because product was not returned. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
29980 | 1220648-2024-07775 | 18875048 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18875048&pc=OZD | 02/23/2024 | 11/21/2025 |  | 03/11/2024 |  |  | 02/23/2024 | 03/11/2024 |  |  |  |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1653274 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Inaccurate Flow Rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/23/2024 |  |  |  |  | The user facility reported the automated impella controller (aic) had possible inaccurate flow readings. the aic was replaced. no patient harm or involvement was reported. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
29981 | 1220648-2024-07776 | 18875153 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18875153&pc=OZD | 02/25/2024 | 04/30/2024 |  | 03/11/2024 |  |  | 02/25/2024 | 03/11/2024 |  |  | 10/31/2025 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024407166 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | Improper Flow or Infusion |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/25/2024 |  |  |  |  | The user facility reported a 48-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported impella 5.5 had to be explanted due high purge pressure experienced during support for 48-year-old male patient. a new impella 5.5 was placed. there was no reported patient harm. | [5 empty] | The investigation into the high purge pressure has been completed since the original report was submitted. the device was not returned for investigation. according to the clinical, shortly after beginning impella 5.5 support high purge pressure and purge system blocked alarms were recorded. after tissue plasminogen activator was added to the purge the alarms were resolved. the most likely cause of the high purge pressure was biomaterial. impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ d.4 revised model number as previously entered incorrectly. d.6a and d.6b revised from the initial submission in accordance with updated procedures. d.8 no was inadvertently left off the previously submitted report and was added. e.1 and g.1 revised facility name and manufacturing site fax number from the initial submission in accordance with updated procedures. h.6 codes 2199, 2954 and 1248 were reported incorrectly on the previously submission report. new code has been added to medical device problem codes. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
29982 | 1220648-2024-07777 | 18875204 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18875204&pc=OZD | 02/22/2024 | 07/16/2024 |  | 03/11/2024 |  |  | 02/27/2024 | 03/11/2024 |  |  | 11/30/2025 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024427951 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [9 empty] | Life Threatening | Disability | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/27/2024 |  |  |  |  | The user facility reported a n unspecified study patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant from the study reported that patient on impella cp support had acute kidney injury stage 3 that was moderate in severity. the event was determined to be possibly related to the study device and the procedure. the patient also had respiratory failure that was moderate in severity. the event was determined to possibly be related to the study device and the procedure. the patient experienced reperfusion arrhythmias that were moderate in severity. the event was determined to possibly be related to the study device and the procedure. the patient also had av block grade 3 that required prolonged existing hospitalization. the event was determined to possibly be related to the study device and the procedure. the patient reportedly had pneumonia that was moderate in severity and required prolonged existing hospitalization. the event was determined to possibly be related to the study device and the procedure. the patient was reported to have hematuria that was moderate in severity and required prolonged existing hospitalization. the event was determined to not be related to the study device or the procedure. the patient had pulmonary hypertension that was mild in severity. the event was determined to not be related to the impella device or the procedure. the patient had a hematoma at the access site that was moderate in severity. however, the event was determined to not be related to the study device or the impella procedure. the event was related to the percutaneous coronary intervention (pci). the patient from the study had left ventricular thrombus that was moderate in severity and required prolonged hospitalization. the event was determined to not be related to the study device or impella procedure. the patient had a positive fecal occult blood test (fobt) that was mild in severity. the event was neither related to the study device or the impella procedure. the event occurred 8 days post impella removal. | [5 empty] | The investigation for the reported renal failure, arrhythmia, hemolysis, and access site bleed has been completed. neither the device nor data logs were not returned for review. the root cause of renal failure was unable to be determined as limited clinical details were provided and no product was returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. the root cause of access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. the root cause of the arrhythmia could not be determined without additional clinical details. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
29983 | 1220648-2024-07778 | 18875209 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18875209&pc=OZD | 02/27/2024 | 04/17/2024 |  | 03/11/2024 |  |  | 02/27/2024 | 03/11/2024 |  |  | 11/30/2025 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024430122 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/27/2024 |  |  |  |  | The user facility reported a 58-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed ischemia in their impella leg during pump support. the patient's leg was found to be cold and pulseless as a result of the condition. due to the ischemia, the decision was made to escalate to impella 5.5 support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
29984 | 1220648-2024-07779 | 18875274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18875274&pc=OZD | 02/27/2024 | 04/17/2024 |  | 03/11/2024 |  |  | 02/27/2024 | 03/11/2024 |  |  |  |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025437995 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/27/2024 |  |  |  |  | The user facility reported a 58-year-old female was to be implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the physician attempted several times to implant the impella 5.5 but would not advance into the ascending aorta from the axillary, it was known that the patient had small vessels. the implant was aborted and impella cp was successfully placed. | [5 empty] | The investigation into the delivery issues has been completed since the original report was filed. the pump was returned, tested, and evaluated. the root cause of delivery issue is determined to be patient condition because of the smaller vessel size of the patient. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
29985 | 1220648-2024-07780 | 18875281 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18875281&pc=OZD | 11/17/2022 | 03/13/2025 |  | 03/11/2024 |  |  | 11/17/2022 | 03/11/2024 |  |  | 07/31/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023251457 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/17/2022 |  |  |  |  | The user facility reported a 66-year-old female noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. the patient had bleeding (oozing) at the impella cp insertion site. the peel away was removed and reposition sheath was advanced into place. the sheath was sutured with forward tension sutures. upon dressing the site, oozing was noted. multiple attempts to adjust angle, tighten perclose sutures and redress were performed but oozing persisted. bicarb was added to the purge. subsequently, impella cp was weaned and explanted. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29986 | 1220648-2024-07782 | 18876150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18876150&pc=OZD | 09/09/2022 | 03/13/2025 |  | 03/11/2024 |  |  | 09/09/2022 | 03/11/2024 |  |  | 07/31/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231486 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/09/2022 |  |  |  |  | The user facility reported a 73-year-old female in sinus tachycardia (st) elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient developed right lower extremity (rle) ischemia. a rle fasciotomy was performed to release pressure with sterile dressing applied post procedure. | [5 empty] | Limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
29987 | 1220648-2024-07783 | 18876241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18876241&pc=OZD | 02/29/2024 | 04/18/2024 |  | 03/11/2024 |  |  | 02/29/2024 | 03/11/2024 |  |  |  |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025438147 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/29/2024 |  |  |  |  | The user facility reported a 53-year-old male with cardiomyopathy was implanted with an impella for mechanical circulatory support. the patient was on impella 5.5 support and had a hematoma as access site. the patient had a washout at the access site and heparin was withheld. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
29988 | 1220648-2024-07784 | 18876345 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18876345&pc=OZD | 02/28/2024 | 03/22/2024 |  | 03/11/2024 |  |  | 02/28/2024 | 03/11/2024 |  |  |  |  | 03/11/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1481995 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/28/2024 |  |  |  |  | It was reported that the automated impella controller had its safety disk holder forcibly removed. the aic was removed from service. no patient was involved. | [5 empty] | The impella device was received from the customer and¿an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were not relevant to the reported issue. the console was evaluated and it was confirmed that the purge disc retainer had broken off on the console. the root cause of this issue was a broken purge retainer.
29989 | 1220648-2024-07785 | 18876360 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18876360&pc=OZD | 09/09/2022 | 03/13/2025 |  | 03/11/2024 |  |  | 09/09/2022 | 03/11/2024 |  |  | 04/30/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023214845 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/09/2022 |  |  |  |  | The user facility reported a 72-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while on impella cp support, the patient experienced bleeding and systemic heparin was discontinued. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29990 | 1220648-2024-07786 | 18876428 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18876428&pc=OZD | 08/09/2022 | 03/13/2025 |  | 03/11/2024 |  |  | 08/09/2022 | 03/11/2024 |  |  | 04/30/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023220668 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/09/2022 |  |  |  |  | The user facility reported a 71-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support the pump dove forward resulting in placement signal low alarm and runs of ventricular tachycardia (vt). the doctor pulled pump back 2 centimeters (cm) which resolved ectopy and the alarm. | [5 empty] | Cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease ¿timing of the arrythmia event in relation to impella support (during or post-implant) ¿electrocardiogram (ekg) recordings of the arrhythmia episodes ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities ¿response to any antiarrhythmic treatments or interventions during the event ¿any documented device-related issues or complications during impella support ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand ¿review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29991 | 1220648-2024-07787 | 18876434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18876434&pc=OZD | 03/01/2024 | 04/18/2024 |  | 03/11/2024 |  |  | 03/01/2024 | 03/11/2024 |  |  | 09/30/2025 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024403877 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/01/2024 |  |  |  |  | The user facility reported a 79-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. the cp support ongoing for a patient via the 14fr sheath. the patient had curvature of the aortic arch. the pump kinked and had low flows. the kinked pump was not replaced. the support remained til explant at 46minutes after hrpci. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the product damage (cannula kink) issue has been completed since the original report was submitted. the device was not returned for investigation. however, data log review and clinical information provided determined there were low pump flow and saturated flow. the root cause of the low pump flow issue was patient condition related since cannula kinking due to patient anatomy. the cause of the saturated flow issue was most likely biomaterial per log review.
29992 | 1220648-2024-07788 | 18876526 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18876526&pc=OZD | 08/10/2022 | 03/13/2025 |  | 03/11/2024 |  |  | 08/10/2022 | 03/11/2024 |  |  | 04/30/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2023214981 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/10/2022 |  |  |  |  | The user facility reported a 66-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced persistent oozing around the peel away sheath. this was treated by applying a pressure dressing and by removing heparin from the purge to manage the condition. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
29993 | 1220648-2024-07789 | 18876559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18876559&pc=OZD | 03/02/2024 | 04/08/2024 |  | 03/11/2024 |  |  | 03/02/2024 | 03/11/2024 |  |  | 10/31/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2023284013 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/02/2024 |  |  |  |  | The user facility reported a 65-year-old male patient of unknown race and ethnicity in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had a hematoma and bleeding at the impella site. the patient had a right femoral artery cut down and the impella was removed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided.
29994 | 1220648-2024-07791 | 18876583 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18876583&pc=OZD | 02/28/2024 | 04/04/2024 |  | 03/11/2024 |  |  | 02/28/2024 | 03/11/2024 |  |  | 09/30/2025 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024398846 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 35 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/28/2024 |  |  |  |  | The user facility reported a 35-year-old female in st-segment elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support as part of a clinical study. it was reported that the patient developed bruising around the access site. a computed tomography (ct) angiogram was performed and showed a right groin subcutaneous hematoma that required a blood transfusion. the condition was monitored via ct imaging. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of access site bleeding was most likely due to patient movement since the patient continued trying move and sit-up during procedure and once in icu. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
29995 | 1220648-2024-07790 | 18876585 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18876585&pc=OZD | 08/07/2022 | 03/13/2025 |  | 03/11/2024 |  |  | 08/07/2022 | 03/11/2024 |  |  | 08/31/2023 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022127060 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/07/2022 |  |  |  |  | The user facility reported a 76-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient had a large hematoma noted to the right groin with femstop in place. it was advised to stop heparin. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. this report is being filed as part of a retrospective review of historical records.
29996 | 1220648-2024-07792 | 18876621 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18876621&pc=OZD | 08/10/2022 | 03/13/2025 |  | 03/11/2024 |  |  | 08/10/2022 | 03/11/2024 |  |  | 04/30/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213585 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/10/2022 |  |  |  |  | The user facility reported a 69-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient experienced red urine. after decreasing p level it cleared. the impella device was weaned and explanted. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
29997 | 1220648-2024-07793 | 18876706 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18876706&pc=OZD | 09/11/2022 | 03/13/2025 |  | 03/11/2024 |  |  | 09/11/2022 | 03/11/2024 |  |  | 03/31/2023 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022056592 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/11/2022 |  |  |  |  | The user facility reported a 61-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support and the patient experienced bleeding (oozing) at the impella cp groin access site. the nurse re-dressed the groin, angle matched and discontinued aggrastat to control the bleeding. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
29998 | 1220648-2024-07794 | 18876838 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18876838&pc=OZD | 08/22/2022 | 03/13/2025 |  | 03/11/2024 |  |  | 08/22/2022 | 03/11/2024 |  |  | 02/29/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023199404 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/22/2022 |  |  |  |  | The user facility reported a 46-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. despite no visual signs of infection or clotting on the device, echo imaging did show a thrombus remaining on the left ventricle. the impella cp was successfully weaned and removed, and the patient remains on heparin. | [5 empty] | The failure mode of "thrombus" is to be used when a thrombus is noted but cannot be connected to any other adverse events or decrease in pump performance. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions. description of the location and characteristics of the thrombus (e.g., size, shape, consistency). relevant laboratory test results, such as coagulation profiles and platelet counts. imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). medications or treatments that the patient was receiving at the time of thrombus formation. surgical or procedural details if applicable (e.g, cardiac catheterization). any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ this report is being filed as part of a retrospective review of historical records.
29999 | 1220648-2024-07795 | 18876840 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18876840&pc=OZD | 02/29/2024 | 06/26/2025 |  | 03/11/2024 |  |  | 02/29/2024 | 03/11/2024 |  |  | 11/30/2025 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024436168 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | Followup | 02/29/2024 |  |  |  |  | The user facility reported a 55-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. there was bleeding/leaking from the 7 french companion sheath. the bleeding was resolved with the removal of the sheath. percutaneous coronary intervention access was obtained from another femoral artery. the patient is stable. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of access site bleeding was unable to be determined as the product was not returned and insufficient clinical details were provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation for the access site bleed has been updated. the introducer has been returned for evaluation. upon inspection, there were no abnormalities found. therefore, the root cause of the access site bleeding was unable to be determined as clinical details were insufficient. in accordance with updated procedures, sections d6, g3, and h10 have been revised from the initial submission. | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | F6/f8-should have been left blank on the initial manufacturer device report number 1220648-2024-07795. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures 1220648-2024-07795. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2024-07795.
30000 | 1220648-2024-07796 | 18876866 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18876866&pc=OZD | 08/10/2022 | 09/25/2024 |  | 03/11/2024 |  |  | 08/10/2022 | 03/11/2024 |  |  | 04/30/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023226496 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | American Indian Or Alaskan Native | Death | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/10/2022 |  |  |  |  | It was reported that the pump was successfully placed and within expected performance parameters. the patient remained relatively stable throughout the removal and reinsertion of the cp, however, prior to leaving the cardiac catheterization laboratory the patient coded an additional 3 times requiring cardiopulmonary resuscitation. return of spontaneous circulation was achieved each time. the patient was then transported to the unit where she coded an additional two times. during the second and final code the patient's family asked that all attempts at resuscitation be halted, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | The user facility reported a 71-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support.it was reported that there were positioning issues with the impella pump. post-angioplasty, it was noted that the impella flows, and map declined significantly. the pump was noted to be against the free wall of the left ventricle and repositioning was attempted. although initial troubleshooting resolved the performance issues, the pump again began to experience complications and further attempts to reposition were required. when repositioning proved difficult, the staff opted to completely remove the pump and re-insert the device using wire assisted re-access. the pump was successfully placed, and support continued. there was no reported patient harm. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07796 was provided in sections b1, b2, b5, d6a, d6b, h1, and h6. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
30001 | 1220648-2024-07797 | 18876927 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18876927&pc=OZD | 02/29/2024 | 04/16/2024 |  | 03/11/2024 |  |  | 02/29/2024 | 03/11/2024 |  |  | 10/31/2025 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024409568 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Separation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/29/2024 |  |  |  |  | The user facility reported a 63-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the sterile sleeve became disconnected. the distal tuohy on the sterile sleeve became disconnected and then reattached. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the product damage has been completed. the impella device was not returned for evaluation. therefore, without sufficient clinical details, the root cause of the reported issues product damage could not be determined.
30002 | 1220648-2024-07798 | 18876928 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18876928&pc=OZD | 08/08/2022 | 03/14/2025 |  | 03/11/2024 |  |  | 08/08/2022 | 03/11/2024 |  |  | 04/30/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023219930 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/08/2022 |  |  |  |  | The user facility reported a 60-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient was placed on impella cp support after the patient went into ventricular tachycardia arrest. electrocardiogram (ekg) showed congenital heart block with heart rate in the 40s and inferior st-segment elevation myocardial infarction and cardiac output response to stress had indicated the dominant proximal left circumflex artery was blocked. the patient developed bleeding at the impella cp insertion site. oozing from all femoral access devices had started overnight with decreasing observed over the morning. the patient had bicarbonate in purge started and systemic heparin was being withheld waiting on activated partial thromboplastin time. when the peel-away sheath was removed, there was still sluggish flow down the superficial femoral artery, so an external fem-fem bypass was done using 6f 5f system. doppler pulses in either leg was still unable to be obtained. the impella cp was successfully explanted, however, it was not explanted early. the patient remained on cp support for two additional days. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
30003 | 1220648-2024-07799 | 18877081 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18877081&pc=OZD | 08/23/2022 | 03/11/2024 |  | 03/11/2024 |  |  | 08/23/2022 | 03/11/2024 |  |  | 02/29/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023198540 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/23/2022 |  |  |  |  | The user facility reported a 58-year-old female with congestive heart failure was implanted with an impella cp device for mechanical circulatory support. the patient developed oozing (bleeding) at the impella cp insertion site after the sheath was peeled. the bleeding was managed with pressure/dressing changes and tightening the pro-glide/sutures. the impella cp was successfully weaned and explanted from the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records.
30004 | 1220648-2024-07800 | 18877125 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18877125&pc=OZD | 03/01/2024 | 04/18/2024 |  | 03/11/2024 |  |  | 03/01/2024 | 03/11/2024 |  |  | 06/30/2025 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024427262 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/01/2024 |  |  |  |  | The user facility reported a 80-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the impella cp was placed but removed due to low flows from a kinked pump. no patient harm was reported due to the issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow has been completed since the original report was filed. the pump was returned, tested, and evaluated. data logs confirmed the low flow. visual inspection found a kink on the cannula. the root cause of low flow was kinked cannula.
30005 | 1220648-2024-07801 | 18877307 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18877307&pc=OZD | 03/02/2024 | 04/19/2024 |  | 03/11/2024 |  |  | 03/02/2024 | 03/11/2024 |  |  |  |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025443346 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/02/2024 |  |  |  |  | The user facility reported an 84-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had runs of ventricular tachycardia (vt) requiring cardiopulmonary resuscitation (cpr). the impella was repositioned, and the arrhythmias resolved. | [5 empty] | The investigation into the access site bleeding ¿ minor and the vf/vt/af/at issues has been completed since the original report was submitted. the device was not returned for investigation the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30006 | 1220648-2024-07803 | 18877338 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18877338&pc=OZD | 11/07/2022 | 03/13/2025 |  | 03/11/2024 |  |  | 11/07/2022 | 03/11/2024 |  |  | 07/31/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023259235 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Great Vessel Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/07/2022 |  |  |  |  | The user facility reported a 42-year-old male was implanted with an impella cp for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced profuse bleeding at the right femoral artery groin access site and was coagulopathic. manual pressure was applied to the groin, and the patient was transfused with 2 units of platelets, 2 units of fresh frozen plasma, and cryoprecipitated antihemophilic factor with no improvement to bleeding. anticoagulant therapy was discontinued, and vascular surgery was performed to repair the right femoral artery after it was discovered it had been dissected all the way up to a side branch. during the surgery to repair the right femoral artery, the impella cp was removed and the patient was escalated to an impella 5.5. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. g1 reporting contact first name and reporting contact last name. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
30007 | 1220648-2024-07802 | 18877341 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18877341&pc=OZD | 08/23/2022 | 05/14/2025 |  | 03/11/2024 |  |  | 08/23/2022 | 03/11/2024 |  |  | 03/31/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204018 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/23/2022 |  |  |  |  | The user facility reported a 76-year-old male after experiencing pulmonary edema was implanted with an impella cp device for mechanical circulatory support. the patient developed bleeding at the impella cp insertion site and a low platelet count. the bleeding was managed with angle matching and light fem-stop, and the thrombocytopenia was managed with a purge change from heparin to argatroban. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30008 | 1220648-2024-07804 | 18877542 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18877542&pc=OZD | 08/24/2022 | 03/13/2025 |  | 03/11/2024 |  |  | 08/24/2022 | 03/11/2024 |  |  | 05/31/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, IND. |  |  | US |  | 2023231220 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/24/2022 |  |  |  |  | The user facility reported a 70-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the device was removed after possible clot ingestion and lower flows than expected and was unable to be placed after cleaning. the device was discarded and a second impella cp device was used successfully. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
30009 | 1220648-2024-07805 | 18877867 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18877867&pc=OZD | 08/28/2022 | 03/13/2025 |  | 03/11/2024 |  |  | 08/28/2022 | 03/11/2024 |  |  | 05/31/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023230086 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/28/2022 |  |  |  |  | The user facility reported a 83-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient developed bleeding at the impella cp access site after the peel away sheath was exchanged to the reposition sheath. cangalor and systemic heparin were discontinued and a fem-stop was applied to control the bleeding. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
30010 | 1220648-2024-07806 | 18878257 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18878257&pc=OZD | 08/24/2022 | 03/27/2024 |  | 03/11/2024 |  |  | 08/24/2022 | 03/11/2024 |  |  | 09/30/2023 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022146593 |  | 0048-0044 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 39 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Cardiac Perforation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/24/2022 |  |  |  |  | The user facility reported a 39-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. during placement of the impella the sheath could not clear two tight bends of the patient's anatomy, and bleeding was noted around the peel away sheath. when the peel away sheath was removed there was a major bleed in the artery. vascular repair was performed and the impella was successfully placed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | The investigation has been completed. the root cause of the access site bleeding could not be determined since the product was not returned and insufficient clinical details was provided. h6 code 2513 cardiac perforation was selected in error on the initial report.
30011 | 1220648-2024-07807 | 18879210 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18879210&pc=OZD | 08/26/2022 | 03/11/2024 |  | 03/11/2024 |  |  | 08/26/2022 | 03/11/2024 |  |  | 09/30/2023 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022147135 |  | 0048-0045 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/26/2022 |  |  |  |  | The user facility reported a 58-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed persistent bleeding around the impella sheath despite redressing. a hematoma was present as well. the patient received two units of packed red blood cells. hemolysis and hematuria was also noted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
30012 | 1220648-2024-07808 | 18879653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18879653&pc=OZD | 08/08/2022 | 03/13/2025 |  | 03/11/2024 |  |  | 08/08/2022 | 03/11/2024 |  |  | 04/30/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218100 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/08/2022 |  |  |  |  | The user facility reported a 77-year-old male patient was being implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient was on impella cp support due to non-disclosed cardiovascular medical history. while on support, the patient required repositioning of the sheath, implanted hematoma was observed, and pressure was held. two units of packed red blood cells were administered. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
30013 | 1220648-2024-07809 | 18879840 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18879840&pc=OZD | 08/28/2022 | 10/20/2025 |  | 03/11/2024 |  |  | 08/28/2022 | 03/11/2024 |  |  | 07/31/2023 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022116290 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/28/2022 |  |  |  |  | The user facility reported a 66-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while on impella cp support, the patient experienced hematuria and developed hemolysis. the medical team confirmed impella cp position with echocardiography and turned the impella cp pump down to p2. the medical team decided to explant the impella cp device and transition the patient to venoarterial extracorporeal membrane oxygenation (va ecmo). | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
30014 | 1220648-2024-07810 | 18880171 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18880171&pc=OZD | 08/26/2022 | 03/11/2024 |  | 03/11/2024 |  |  | 08/26/2022 | 03/11/2024 |  |  | 05/31/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023229476 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/26/2022 |  |  |  |  | The user facility reported a 45-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient¿s right foot was cold to the touch and didn¿t have a palpable pulse. an echo was performed, and the pump was noted to be angled down to the mitral valve. no suction alarms and no hemolysis were noted so the decision was made to not reposition the pump. three units of packed red blood cells were ordered. after the pump was explanted, a lower extremity angiogram revealed an occluded popliteal artery. penumbra was used to extract the clot and flow was restored to the foot. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
30015 | 1220648-2024-07812 | 18880593 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18880593&pc=OZD | 08/29/2022 | 02/18/2025 |  | 03/11/2024 |  |  | 08/29/2022 | 03/11/2024 |  |  | 02/29/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023193097 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Cardiac Arrest | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 08/29/2022 |  |  |  |  | The user facility reported a 63-year-old male was implanted with an impella cp device for mechanical circulatory support. after the impella cp device was placed, it was discovered the device was under a papillary muscle. the device was repositioned unsuccessfully, and the medical team decided to remove the impella cp device, rewire, and reinsert the device. impella cp was successfully placed. | [5 empty] | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation into the reported issue has been completed. device history record review confirmed that the passed all post-sterile inspection checks. no device was received for evaluation. due to the limited clinical information and lack of available data/product the root cause of this investigation was not determined.
30016 | 1220648-2024-07811 | 18880628 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18880628&pc=OZD | 08/08/2022 | 03/13/2025 |  | 03/11/2024 |  |  | 08/08/2022 | 03/11/2024 |  |  | 04/30/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023224720 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/08/2022 | [10 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined. | This mdr related to the puerto rico manufacturing site has been assigned a medwatch number from the medtronic minimed northridge site, per variance 5. select patient information cannot be provided due to regional privacy regulations. currently it is unknown whether or not the device may have caused or contributed to the event as no product has been returned. no conclusion can be drawn at this time. we therefore consider this report complete to the best of our knowledge. medtronic, inc. (medtronic) is submitting this report to comply with 21 c.f.r. part 803, the medical device reporting regulation. this report is based upon information obtained by medtronic, which the company may not have been able to fully investigate or verify prior to the date the report was required by the fda. medtronic has made reasonable efforts to obtain more complete information in the time allotted and has provided as much information as is available to the company as of the submission date this report. this report does not constitute an admission or a conclusion by fda, medtronic, or its employees that the device, medtronic, or its employees caused or contributed to the event described in the report. in particular, this report does not constitute an admission by anyone that the product described in this report has any "defects" or has "malfunctioned". these words are included in the fda 3500a form and are fixed items for selection created by the fda, to categorize the type of event solely for the purpose of reporting pursuant to part 803. medtronic objects to the use of these words and others like it because of the lack of definition and the connotations implied by these terms. this statement should be included with any information or report disclosed to the public under the freedom of information act.
30017 | 1220648-2024-07813 | 18880972 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18880972&pc=OZD | 08/27/2022 | 04/17/2025 |  | 03/11/2024 |  |  | 08/27/2022 | 03/11/2024 |  |  | 05/31/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023230846 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/27/2022 |  |  |  |  | The user facility reported a 62-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient developed oozing (bleeding) at the impella cp insertion site. the patient also exhibited bloody output from the nasogastric tube. because the patient's hemoglobin was decreased and platelets were low, the patient received packed red blood cells and platelets. the family decided to withdraw care. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
30018 | 1220648-2024-07814 | 18881174 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18881174&pc=OZD | 09/09/2022 | 03/13/2025 |  | 03/11/2024 |  |  | 09/09/2022 | 03/11/2024 |  |  | 05/31/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023229088 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/09/2022 |  |  |  |  | The user facility reported a 55-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient developed slight oozing (bleeding) from the impella insertion site. hemostasis was achieved by redressing the site and applying additional sutures to secure the device. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
30019 | 1220648-2024-07816 | 18881599 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18881599&pc=OZD | 08/11/2022 | 03/13/2025 |  | 03/11/2024 |  |  | 08/11/2022 | 03/11/2024 |  |  | 04/30/2022 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023224858 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  |  |  | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/11/2022 |  |  |  |  | The user facility reported a 55-year-old female in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella cp device for mechanical circulatory support. insertion of the device was aborted due to patient's small right femoral artery. due to this difficulty, the patient was then scheduled for axillary implant the next morning. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
30020 | 1220648-2024-07817 | 18881676 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18881676&pc=OZD | 08/27/2022 | 03/11/2024 |  | 03/11/2024 |  |  | 08/27/2022 | 03/11/2024 |  |  | 04/30/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023226430 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/27/2022 |  |  |  |  | The user facility reported a 57-year-old female in segment elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient developed oozing at the groin and insertion site and was given one unit of packed red blood cells. hemostasis was achieved with manual pressure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
30021 | 1220648-2024-07818 | 18881797 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18881797&pc=OZD | 09/14/2022 | 03/14/2025 |  | 03/11/2024 |  |  | 09/14/2022 | 03/11/2024 |  |  | 04/30/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213607 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/09/2022 |  |  |  |  | The user facility reported a 48-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, a computed tomography (ct) scan was performed on the patient's head and thorax, which revealed plural effusions and a few brain infarctions. the patient was successfully escalated to the impella 5.5 and the impella cp was explanted. | [5 empty] | Strokes/cvas can occur during or after impella support, and can be either ischemic or hemorrhagic. to make a determination as to the cause of the stroke, the following clinical information must be considered: patient's medical history, including any previous strokes, cardiovascular disease, or risk factors such as hypertension, diabetes, or atrial fibrillation. timing of the stroke in relation to impella support (during or post-implant). detailed description of the symptoms experienced by the patient. neurological examination findings and any focal deficits observed. diagnostic imaging results, such as ct scan or mri of the brain, to identify the type and location of the stroke (ischemic or hemorrhagic). laboratory test results, including coagulation profiles and cardiac markers. any relevant procedural or device-related factors, such as duration and settings of impella support, presence of embolic protection devices, or any documented procedural complications. response to any previous anticoagulation or antiplatelet therapies. evaluation of potential alternative causes of stroke, such as arrhythmias, embolic sources, or underlying vascular pathology. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
30022 | 1220648-2024-07820 | 18881805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18881805&pc=OZD | 09/09/2022 | 03/14/2025 |  | 03/11/2024 |  |  | 09/09/2022 | 03/11/2024 |  |  | 04/30/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023226195 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/09/2022 |  |  |  |  | The user facility reported a 68-year-old female was placed on impella cp support for impella assisted percutaneous coronary intervention (pci). the impella was in the left femoral artery, however, the physician was unable to engage the right coronary artery with the impella in place. the decision was made to remove the impella and proceed with the pci of the right coronary artery. the impella was successfully explanted. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
30023 | 1220648-2024-07821 | 18881822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18881822&pc=OZD | 08/27/2022 | 03/11/2024 |  | 03/11/2024 |  |  | 08/27/2022 | 03/11/2024 |  |  | 03/31/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2023210399 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/27/2022 |  |  |  |  | The user facility reported a 73-year-old male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support hemolysis was suspected. after an echo was performed, pseudomonas was verified. the impella was successfully weaned and removed from the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
30024 | 1220648-2024-07819 | 18881827 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18881827&pc=OZD | 08/11/2022 | 03/14/2025 |  | 03/11/2024 |  |  | 08/11/2022 | 03/11/2024 |  |  | 02/29/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202936 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/11/2022 |  |  |  |  | The user facility reported an 81-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced oozing at the access site and developed a hematoma. the impella device was repositioned via echo. the issue resolved with manual pressure and dressing change. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
30025 | 1220648-2024-07822 | 18881868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18881868&pc=OZD | 08/11/2022 | 03/11/2024 |  | 03/11/2024 |  |  | 08/11/2022 | 03/11/2024 |  |  | 04/30/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023215708 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/11/2022 |  |  |  |  | The user facility reported that a 78-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding at the impella sheath site. manual pressure was held, femostop was placed. the patient was administered two plasma, two fresh frozen plasma and four packed red blood cells. the device was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
30026 | 1220648-2024-07823 | 18882012 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18882012&pc=OZD | 08/11/2022 | 03/14/2025 |  | 03/11/2024 |  |  | 08/11/2022 | 03/11/2024 |  |  | 05/31/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023226932 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/11/2022 |  |  |  |  | The user facility reported that a 70-year-old male in acute myocardial infarction/cardiogenic shock with hemodynamic compromise was implanted with an impella cp device for mechanical circulatory support. after placement the patient experienced bleeding at the site. heparin was put on hold. an echocardiogram was done and showed that the impella cp was shallow and interacting with mitral valve and pap muscle. the cardiologist came to the bedside and attempted to free the impella cp from the pap muscle but was unsuccessful. the impella cp was explanted and the surgeon performed arterial repair. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records.
30027 | 1220648-2024-07824 | 18882075 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18882075&pc=OZD | 08/11/2022 | 03/11/2024 |  | 03/11/2024 |  |  | 08/11/2022 | 03/11/2024 |  |  | 04/30/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2023217214 |  | 0048-0047 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/11/2022 |  |  |  |  | The user facility reported an 84-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support there were intermittent suction alarms and the patient experienced hemolysis. the impella was weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
30028 | 1220648-2024-07825 | 18882626 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18882626&pc=OZD | 08/12/2022 | 03/14/2025 |  | 03/11/2024 |  |  |  | 03/11/2024 |  |  | 11/30/2023 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2023169890 |  | 0048-0050 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/12/2022 |  |  |  |  | The user facility reported an 70-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after receiving percutaneous coronary intervention (pci), the patient went into cardiac arrest. after a repeat pci, the impella cp was inserted. a stat echocardiogram indicated very low volume in the left and right ventricles. the pump was deep and pulled back at bedside. during support the patient was profoundly hypovolemic so addressed volume status with blood and albumin. all extremities were cold and pale. the impella cp was explanted. | [5 empty] | Limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition ¿ any concomitant medication ¿ vascular size or variations thereof ¿ detailed description of the symptoms experienced by the patient ¿ physical examination findings, including pulse assessment and visual examination of the affected limb ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension ¿ patient's response to previous treatments or interventions for vascular conditions with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
30029 | 1220648-2024-07826 | 18882735 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18882735&pc=OZD | 08/16/2022 | 03/14/2025 |  | 03/11/2024 |  |  |  | 03/11/2024 |  |  | 02/29/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023199242 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/16/2022 |  |  |  |  | The user facility reported a 75-year-old male patient on an impella cp for mechanical circulatory support. the patient developed hematomas at all insertion sites (rij swan, impella site, and l fem interventional site) which resulted in the decision to not perform the percutaneous coronary intervention. the plan was to continue bicarb purge with the addition of systemic heparin. the impella was then weaned and explanted. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
30030 | 1220648-2024-07827 | 18882750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18882750&pc=OZD | 08/17/2022 | 03/14/2025 |  | 03/11/2024 |  |  | 08/17/2022 | 03/11/2024 |  |  | 03/31/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023208334 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Deliver | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/17/2022 |  |  |  |  | The user facility reported a 68-year-old male, race and ethnicity unknown, in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that there was a delivery issue as they were unable to get the device to pass through the patient's ilio-femoral stent. multiple attempts proved unsuccessful, and insertion was aborted as the patient had bilateral ilio-femoral stents. the decision was made to place an intra-aortic balloon pump instead. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
30031 | 1220648-2024-07828 | 18882789 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18882789&pc=OZD | 08/17/2022 | 03/14/2025 |  | 03/11/2024 |  |  | 08/17/2022 | 03/11/2024 |  |  | 05/31/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2023228754 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Improper Flow or Infusion | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/17/2022 |  |  |  |  | The user facility reported a 74-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that there was access site bleeding. the bleeding was stopped with forward sutures and tightening of perclose sutures. it was also noted that the patient had suffered a stroke during support. plans for an embolectomy of the right carotid were put in place to treat the condition. it was additionally noted that the purge flow dropped over the course of a few hours. a tissue plasminogen activator (tpa) protocol was introduced to counteract the issue. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
30032 | 1220648-2024-07829 | 18882873 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18882873&pc=OZD | 08/17/2022 | 03/14/2025 |  | 03/11/2024 |  |  | 08/17/2022 | 03/11/2024 |  |  | 11/30/2023 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2023160524 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/17/2022 |  |  |  |  | The user facility reported a 61-year-old male, noted as high risk percutaneous coronary intervention, was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient was on impella support due to hemolysis. while on support, the patient experienced oozing from the impella site, resolved with placement of forward tension sutures and angle match dressing. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma. remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site. potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it. obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine. hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis. suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure. patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
30033 | 1220648-2024-07830 | 18882928 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18882928&pc=OZD | 08/17/2022 | 03/14/2025 |  | 03/11/2024 |  |  | 08/17/2022 | 03/11/2024 |  |  | 04/30/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218260 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/17/2022 |  |  |  |  | The user facility reported a 75-year-old male was implanted with an impella cp device for mechanical circulatory support. the complainant reported that upon arrival, the patient was shocked twice by life flight for ventricular tachycardia (vt) and experienced multiple complications during impella cp support. patient was placed on impella after an unsuccessful stent. the pump was placed, repositioned, and vt resolved. impella sterile sleeve was torn during placement, slight oozing at impella site. tegaderm was placed on the tear to resolve the issue. the following day, additional oozing was noted at the site. gauze was added to support the repo sheath coming out of peel away. support was maintained with the implanted pump throughout the event. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
30034 | 1220648-2024-07831 | 18882979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18882979&pc=OZD | 08/17/2022 | 03/14/2025 |  | 03/11/2024 |  |  | 08/17/2022 | 03/11/2024 |  |  | 03/31/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204058 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/17/2022 |  |  |  |  | The user facility reported a 48-year-old male was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient experienced limb ischemia while on impella cp support. a small groin hematoma was observed following removal of peel away sheath. a fem fem bypass was completed due to concern about the ischemia, resolving the issue. | [5 empty] | Limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition. any concomitant medication. vascular size or variations thereof. detailed description of the symptoms experienced by the patient. physical examination findings, including pulse assessment and visual examination of the affected limb. diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. laboratory test results, including blood tests for markers of inflammation or clotting disorders. any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
30035 | 1220648-2024-07832 | 18883019 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18883019&pc=OZD | 06/10/2022 | 03/07/2025 |  | 03/11/2024 |  |  | 06/10/2022 | 03/11/2024 |  |  | 12/31/2023 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023180813 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 06/10/2022 |  |  |  |  | The user facility reported a 60-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient was placed on impella cp support after experiencing respiratory arrest. impella placed along with fem bypass. slack removed via echo, pump pulled back 2.5 cm and art line was placed in the unit. labs were drawn after 4 hours, and it indicated that all of the patient¿s labs were hemolyzing. additionally, the patient¿s legs were grey and mottled despite external fem-fem bypass. the patient required increased pressors support. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported hemolysis and limb ischemia ¿ irreversible issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
30036 | 1220648-2024-07833 | 18883042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18883042&pc=OZD | 08/17/2022 | 03/14/2025 |  | 03/11/2024 |  |  | 08/17/2022 | 03/11/2024 |  |  | 04/30/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023226999 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/17/2022 |  |  |  |  | The user facility reported a 69-year-old male was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient developed a moderate-sized hematoma while on the impella support. pressure was held over 15 mins with resolution. sutures were applied and a femostop device was put in place. the hematoma worsened and the site was redressed and the femostop removed. the patient was weaned from the impella and it was explanted. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
30037 | 1220648-2024-07834 | 18883058 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18883058&pc=OZD | 07/31/2022 | 03/14/2025 |  | 03/11/2024 |  |  | 07/30/2022 | 03/11/2024 |  |  | 12/31/2023 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023173116 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 14 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/30/2022 |  |  |  |  | The user facility reported a 14-year-old male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the complainant reported lactate dehydrogenase (ldh) was higher at 900 but the percent fetal hemoglobin remained less than 40. the urine output was also amber. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
30038 | 1220648-2024-07835 | 18883066 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18883066&pc=OZD | 07/30/2022 | 10/13/2025 |  | 03/11/2024 |  |  | 07/30/2022 | 03/11/2024 |  |  | 04/30/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023216065 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Unspecified Infection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/30/2022 |  |  |  |  | The user facility reported a 48-year-old male patient on an impella cp for mechanical circulatory support. while on support, the patient developed hemolysis and an infection at the insertion site. the hemolysis issue was addressed by fine-tuning the position of the device using echocardiogram. the infection was observed after removal of the device and was noted to be non-serious. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
30039 | 1220648-2024-07836 | 18883072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18883072&pc=OZD | 08/01/2022 | 03/10/2025 |  | 03/11/2024 |  |  | 08/01/2022 | 03/11/2024 |  |  | 04/30/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023224486 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/01/2022 |  |  |  |  | The user facility reported a 63-year-old male with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. the impella cp was placed pending coronary artery bypass graft. the patient developed oozing/bleeding at the impella cp insertion site. to control the bleeding, femstop was applied. the patient was not on purge or systemic heparin at this time. | [5 empty] | The investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
30040 | 1220648-2024-07837 | 18883096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18883096&pc=OZD | 08/01/2022 | 03/10/2025 |  | 03/11/2024 |  |  | 08/01/2022 | 03/11/2024 |  |  | 04/30/2023 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP. |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022081024 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/01/2022 |  |  |  |  | The user facility reported a 74-year-old male with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient had an occluded left anterior descending artery and right coronary artery with an end diastolic pressure of 36. it was noted that the patient had hematuria (hemolysis) that had not approved despite repositioning. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
30041 | 1220648-2024-07838 | 18883115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18883115&pc=OZD | 08/03/2022 | 03/11/2025 |  | 03/11/2024 |  |  | 08/03/2022 | 03/11/2024 |  |  | 04/30/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023225104 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/03/2022 |  |  |  |  | The user facility reported a 44-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient¿s skin appeared mottled and pulses were lost at the limb. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
30042 | 1220648-2024-07839 | 18883122 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18883122&pc=OZD | 08/03/2022 | 03/10/2025 |  | 03/11/2024 |  |  | 08/03/2022 | 03/11/2024 |  |  | 03/31/2024 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023206103 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/03/2022 |  |  |  |  | The user facility reported a 52-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after the impella cp was implanted the patient experienced hemolysis due to high level of pump support. this was subsequently treated by patient escalation of care to an impella 5.5 pump for additional support. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
30043 | 1220648-2024-07840 | 18883135 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18883135&pc=OZD | 08/03/2022 | 03/10/2025 |  | 03/11/2024 |  |  | 08/03/2022 | 03/11/2024 |  |  | 09/30/2023 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022146948 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/03/2022 |  |  |  |  | The user facility reported a 88-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. after the impella cp was removed both perclose sutures broke while tying, with an additional one also breaking. consequently, this was treated through manual compression. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
30044 | 1220648-2024-07841 | 18883136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18883136&pc=OZD | 08/03/2022 | 03/11/2025 |  | 03/11/2024 |  |  | 08/03/2022 | 03/11/2024 |  |  | 11/30/2023 |  | 03/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023167609 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/03/2022 |  |  |  |  | The user facility reported a 82-year-old male post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. the complainant reported that the patient developed a groin hematoma during impella cp support. a mattress suture and fem-stop were placed to manage the condition. coinciding with the hematoma, it was also noted that the implanted pump experienced persistent positioning issues throughout support. repositioning was ineffective as imaging was inadequate. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
30045 | 1220648-2024-07761 | 18872209 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18872209&pc=OZD | 01/23/2024 | 03/26/2024 |  | 03/10/2024 |  |  |  | 03/10/2024 |  |  | 10/31/2024 |  | 03/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024420019 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/08/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry indicating a 65 year old male patient presenting with acute myocardial infarction and cardiogenic shock for impella support to have endured thrombocytopenia with a "possible" relationship to the impella device. the condition improved after removal of the impella. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the thrombocytopenia issue has been completed since the initial report was submitted. the product was not returned for investigation the root cause of the thrombocytopenia issue was not determined since product was not returned and insufficient clinical information was provided.
30046 | 1220648-2024-07762 | 18872210 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18872210&pc=OZD | 11/19/2023 | 04/10/2024 |  | 03/10/2024 |  |  |  | 03/10/2024 |  |  | 09/30/2025 |  | 03/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024398730 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/21/2023 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry indicating a patient had endured hemolysis and thrombocytopenia with a "definite" relationship to the impella device. unspecified intervention was required. | A 70 year old male patient presented with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient endured an access site bleed prompting sutures and 4 units of prbc. according to an additional notification, via abiomed¿s long-term outcome and quality indicator impella registry, a quick clot intervention with applied pressure was also administered. |  |  |  |  | Additional information received via loqi study added to b(5), with additional coding added to h(6). upon completion of our investigation, a supplemental mdr will be filed. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major, the hemolysis, and the thrombocytopenia issues has been completed since the original report was submitted. the device was not returned for investigation. based on clinical description & data analysis, the root cause of hemolysis was most likely due to pump was not in the proper position. the root cause of access site bleeding was most likely use related as the access site required additional suturing to stop bleeding. the root cause of thombocytopenia was unable to be determined as limited clinical details were provided and no product was returned for review.
30047 | 1220648-2024-07731 | 18864453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18864453&pc=OZD | 02/28/2024 | 03/21/2024 |  | 03/08/2024 |  |  | 02/28/2024 | 03/08/2024 |  |  |  |  | 03/08/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1684985 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Break | Pressure Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/28/2024 |  |  |  |  | The user facility reported a 59-year-old male in st-segment elevation myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the automated impella controller (aic) had a low purge pressure alarm. purge fluid was found on the floor. the purge door was opened and noticed gray purge disc cover was broken and not attached to the aic. the aic was replaced successfully. | [5 empty] | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were reviewed which confirmed purge pressure low alarm issue. the console was evaluated and it was confirmed that the purge disc retainer was broke off on the console. the root cause of this issue was a broken purge disc retainer. d4-updated model number. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30048 | 1220648-2024-07732 | 18864506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18864506&pc=OZD | 02/28/2024 | 08/05/2024 |  | 03/08/2024 |  |  | 02/28/2024 | 03/08/2024 |  |  | 09/30/2025 |  | 03/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024394968 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 60 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hemolysis | Low Blood Pressure/ Hypotension | No Clinical Signs, Symptoms or Conditions | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | Followup |  | 02/28/2024 |  |  |  |  | We received a notification via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured hemolysis with a "possible" relationship to the impella device used. patient remained hemodynamically stable with no symptoms. patient had gross hematuria with evidence of hemolysis. additionally, patient endured hypotension requiring medication with a "possible" relationship to the impella device used. the patient recovered with no sequelae. | The user facility reported a 60-year-old female with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that during support there were positioning issues. the doctor attempted to reposition the impella, however it kept flipping back into the mitral valve. the decision was made to remove the impella. |  |  |  |  | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the positioning issue was use related. | Section b1, b2, b5, g2 h1 and h6 were updated as per additional information received. in accordance with updated procedures, section d6 has been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported hypotension and hemolysis has been completed. the device was not returned for evaluation. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. the root cause of the hypotension was unable to be determined as limited clinical details were provided and no product was returned for review. additionally, for the positioning issue that was also reported in the initial report, the data logs were reviewed and showed no positioning alarm was triggered during the case. no other issues were found on the logs. based on clinical description, the root cause of positioning issue was most likely use related. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock in the following sections: section: potential adverse events (united states) states: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ added-h6 impact code 4644 in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8 should have been left blank in the initial report.
30049 | 1220648-2024-07733 | 18864570 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18864570&pc=OZD | 02/29/2024 | 05/02/2024 |  | 03/08/2024 |  |  | 02/29/2024 | 03/08/2024 |  |  | 10/31/2025 |  | 03/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024412822 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 91 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Fracture | Difficult to Insert | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/29/2024 |  |  |  |  | The user facility reported a 91-year-old male in acute myocardial infarction/cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. it was reported that the implant of the impella cp failed due to a malfunction/interaction of the pump's easy guide lumen and the guidewire. the lumen fractured and the pump was replaced. no harm was reported. a new pump was successfully placed for patient support. | [5 empty] | The investigation for the reported delivery issues has been completed. the pump was returned for evaluation. visual inspection found the red lumen was folded and came out of outflow cage. red lumen was removed from cannula and a kink was found on the 3 different position of the red lumen. the root cause of delivery issue was most likely a deflective red lumen. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ to conform to updated procedures, sections d6 & d10 were revised with updated information. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30050 | 1220648-2024-07734 | 18864659 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18864659&pc=PYX | 02/29/2024 | 04/16/2024 |  | 03/08/2024 |  |  | 02/29/2024 | 03/08/2024 |  |  | 02/28/2025 |  | 03/08/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024333391 |  | 1000323 |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/29/2024 |  |  |  |  | The user facility reported a 74-year-old female undergoing cardiac surgery was implanted with an impella rp flex device for mechanical circulatory support. an rp flex was placed to support a patient on extracorporeal membrane oxygenation (ecmo) already. the rp flex was placed and ecmo was removed. after 21 hours the pump was explanted with signs of hemolysis. flows were lower than expected. one unit of packed red blood cells was transfused. when the pump was explanted, the team again placed a venoarterial ecmo. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | G.4 revised pma/510(k) number as previously entered incorrectly. | The investigation for the reported hemolysis has been completed. the device was not returned but data logs were investigated. there was a suction event and a motor current spike seen after 90 minutes of pump insertion. hemolysis was reported after that. the pump flow was also seen to be very low right after the motor current spike. flows seen were 1.5l/min at p9. as per the data log investigation, the root cause of the hemolysis issue was most likely biomaterial ingestion. the instructions for use states ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected low s or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿
30051 | 1220648-2024-07735 | 18864854 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18864854&pc=OZD | 03/03/2024 | 04/19/2024 |  | 03/08/2024 |  |  | 03/03/2024 | 03/08/2024 |  |  | 07/31/2025 |  | 03/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025437239 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/03/2024 |  |  |  |  | The user facility reported a 58-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the surgical team was consulted for multivessel coronary artery disease (cad). after 91 hours of support, the team observed signs of hemolysis and the decision was made to remove the pump and not escalate care further. the care was withdrawn and the patient expired. | [5 empty] | The investigation into the hemolysis issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of hemolysis was not able to be determined since product was not returned and there was limited information from data log & clinical review. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ . ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿.
30052 | 1220648-2024-07736 | 18864977 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18864977&pc=OZD | 03/04/2024 | 03/08/2024 |  | 03/08/2024 |  |  | 03/04/2024 | 03/08/2024 |  |  |  |  | 03/08/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1184602 |  | 0042-0000-US-L |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 03/04/2024 |  |  |  |  | The user facility reported a 64-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the automated impella controller (aic) had a controller error alarm. the aic was exchanged. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30053 | 1220648-2024-07737 | 18866136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18866136&pc=OZD | 02/19/2024 | 04/04/2024 |  | 03/08/2024 |  |  | 02/19/2024 | 03/08/2024 |  |  |  |  | 03/08/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1152990 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/19/2024 |  |  |  |  | The user facility's biomedical engineer reported the automated impella controller (aic) had a battery red failure alarm while a patient was on support. the aic was swapped with another aic with no further issues. support proceeded without harm or injury. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the aic - battery failure (red alarm) has been completed since the initial report was submitted. the console was returned, visually inspected, and tested. data logs confirmed that both a battery communication failure alarm as well as a battery failure alarm were issued by the console. the root cause was a defective pbm whose charging circuitry for battery not working as expected.
30054 | 1220648-2024-07738 | 18866272 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18866272&pc=OZD | 02/20/2024 | 03/27/2024 |  | 03/08/2024 |  |  | 02/20/2024 | 03/08/2024 |  |  | 10/31/2025 |  | 03/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024424216 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/20/2024 |  |  |  |  | The user facility reported a 76-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that after looking at the impella cp under echocardiogram, the impella was facing the posterior wall impeding on the mitral valve leaflet. the patient was brought back to the catheterization laboratory to reevaluate placement of device under fluoroscopy. after multiple failed attempts, the impella was unable to be properly placed in the mid left ventricle cavity. the decision was made to remove the impella and reattempt to place a new cp utilizing the reaccess port to maintain the arterial site. the patient is stable. | [5 empty] | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the positioning issue was use related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30055 | 1220648-2024-07739 | 18866368 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18866368&pc=OZD | 02/21/2024 | 07/02/2025 |  | 03/08/2024 |  |  | 02/21/2024 | 03/08/2024 |  |  |  |  | 03/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1143273 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support for male patient of unknown age, the automated impella controller (aic) did not display placement signal. the type of issue was described as a controller error (yellow alarm). the pump is noted to generate flow and was supporting the patient. there was no reported patient harm. | [5 empty] | The automated impella controller (aic) was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. data analysis: the logs were reviewed, and an unreported controller error alarm (yellow grade, event 1043) was identified. the controller error alarm triggered due to a stoppage of the communication between the 4-in-1 and the optical bench. real time (rt) logs confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) are lastly reported as 16384 values and then no more samples were recorded. the absence of signal not reliable (snr) samples impacted the console¿s ability to recognize the pump disconnection and required the user to perform a hard shutdown. preceding the controller error, there was instance of an invalid placement signal the console identifying the snr as being below the minimum expected value. there was also an instance of an optical sensor system error during case start procedure. further, rt logs revealed that at the time of the reported complaint, the snr was below specification. device analysis: console was sent back to field service for further investigation. the 5s parameters were tested with an optical test cavity. the 5s parameters were within specification. the optical system was confirmed to be within specification, and there were no issues running an optical pump and purge. the log entry processsamplingdatafromopticalbench: sentstopacquisitioncmd ack indicates the sentstopacquisitioncmdflag was set when entering ossi_sampling_stopped_state which eventually led to a false communication stopped condition. the root cause of the controller error alarm was due to an aic software issue. this was entered into jama with defect/ bug. the reported optical signal issue was most likely caused by incomplete insertion of the pump into the pump receptacle (air gap). d2 common device name revised on manufacturer device report 1220648-2024-07739 in accordance with updated procedures g1 reporting contact email revised on manufacturer device report 1220648-2024-07739 in accordance with updated procedures.
30056 | 1220648-2024-07740 | 18866405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18866405&pc=OZD | 02/19/2024 | 07/03/2024 |  | 03/08/2024 |  |  | 02/19/2024 | 03/08/2024 |  |  | 11/30/2025 |  | 03/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024433900 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Myalgia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 02/19/2024 |  |  |  |  | The user facility reported a 78-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that following impella cp implant, no flow was found down the impella leg. a reperfusion sheath was placed to treat the condition and support with the pump was maintained. | Additional information received via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured compartment syndrome with a "definite" relationship to the impella device used on (b)(6) 2024. vascular surgery is consulted in the setting of concern for compartment syndrome¿. the patient outcome is unknown at this time. |  |  |  |  | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. | The investigation into the limb ischemia ¿ reversible issue has been completed since the initial report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. | The investigation into the patient's limb ischemia has been completed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ d4-corrected model number. added- d6a/d6b. f6/f8 should have been blank. g1 reporting contact email address corrected. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | Section b5, g2 and h6 were updated as per additional information received.
30057 | 1220648-2024-07741 | 18866478 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18866478&pc=OZD | 12/12/2022 | 03/13/2025 |  | 03/08/2024 |  |  | 12/12/2022 | 03/08/2024 |  |  | 06/30/2024 |  | 03/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023238755 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Use of Device Problem | Priming Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/12/2022 |  |  |  |  | The user facility reported a 56 year-old male in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported during preparation of device, the user broke the side arm of the purge in half, the impella device was inserted without difficulty. the patient later expired while on impella cp support, however there is no allegation of device contributing to death. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30058 | 1220648-2024-07742 | 18866481 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18866481&pc=OZD | 02/21/2024 | 04/23/2024 |  | 03/08/2024 |  |  | 02/21/2024 | 03/08/2024 |  |  |  |  | 03/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025443612 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | An 80-year-old patient received hemodynamic support with an impella cp smartassist heart pump. the patient developed leg ischemia distal to the impella femoral access site. two external bypasses were placed and local lysis therapy was performed. as a result, the blood circulation in the leg improved. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
30059 | 1220648-2024-07743 | 18866579 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18866579&pc=OZD | 02/21/2024 | 04/23/2024 |  | 03/08/2024 |  |  | 02/21/2024 | 03/08/2024 |  |  |  |  | 03/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025443612 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported 80-year-old patient received hemodynamic support from an impella cp smartassist heart pump due to acute myocardial infarction/cardiogenic shock. one day after insertion of the impella cp, deterioration of blood flow in the leg occurred distal to the femoral insertion site. the user decided to place an external femoral-femoral bypass. a day later the situation had not improved significantly. the patient was transferred to another hospital and the impella cp was exchanged for an axillary impella 5.5 smartassist. eight days after inserting the impella cp and five days after replacing it with the impella 5.5, the patient's leg had to be amputated. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
30060 | 1220648-2024-07744 | 18866588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18866588&pc=PYX | 02/21/2024 | 04/05/2024 |  | 03/08/2024 |  |  | 02/21/2024 | 03/08/2024 |  |  | 10/31/2025 |  | 03/08/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | RP FLEX W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2024413940 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Advance | Material Twisted/Bent | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/21/2024 |  |  |  |  | The user facility reported a 65-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was having an impella rp implant and there was an issue with the sheath. the initial sheath kinked so they used a second sheath while maintaining pack pressure on the wire to insert the second peel away. additionally, the patient had an ooze at the access site and there was unknown intervention performed. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the delivery issues and the access site bleeding ¿ major issues has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the access site bleeding issue was most likely use related, precloses were cinched and manual pressure was held to obtain hemostasis as per the clinical description provided. the root cause of the delivery issue was most likely patient condition related, due to patient body habitus, the initial sheath kinked as per the clinical description provided.
30061 | 1220648-2024-07745 | 18866680 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18866680&pc=OZD | 02/21/2024 | 04/09/2024 |  | 03/08/2024 |  |  | 02/21/2024 | 03/08/2024 |  |  | 11/30/2025 |  | 03/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024427574 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/21/2024 |  |  |  |  | The user facility reported a 65-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient experienced ventricular tachycardia (vt). the impella was repositioned, and vt was resolved. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined. the instructions for use states ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ added-h6 impact code 4628.
30062 | 1220648-2024-07746 | 18866700 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18866700&pc=OZD | 02/22/2024 | 04/23/2024 |  | 03/08/2024 |  |  | 02/22/2024 | 03/08/2024 |  |  | 08/31/2025 |  | 03/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024395268 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  |  | White | Malfunction | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support for 54-year-old male patient, the impella had to be explanted due to ischemic leg. following explant, the leg was noted to be better perfused and the lactate had dropped. a proglide closure system was used to achieve hemostasis and the patient was stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
30063 | 1220648-2024-07747 | 18866750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18866750&pc=OZD | 02/22/2024 | 04/10/2024 |  | 03/08/2024 |  |  | 02/22/2024 | 03/08/2024 |  |  | 07/31/2025 |  | 03/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025443414 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/22/2024 |  |  |  |  | The user facility reported a 44-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that following impella cp implant, the staff was unable to identify a dopplerable pulse in the lower extremity. due to the noted condition, the decision was made to remove the pump and escalate to impella 5.5 support. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
30064 | 1220648-2024-07748 | 18866791 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18866791&pc=OZD | 02/22/2024 | 04/08/2024 |  | 03/08/2024 |  |  | 02/22/2024 | 03/08/2024 |  |  | 08/31/2025 |  | 03/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024394909 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic |  | Injury | No | Yes | Restricted Flow rate | Pressure Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/22/2024 |  |  |  |  | The user facility reported a 70-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella had low purge flow and high purge pressure alarms. the impella was running for less than 10 hours, was weaned, and removed with minor concerns of the purge issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation into the access site bleeding ¿ minor and the high purge pressure issues has been completed since the original report was filed. the pump was returned, tested, and evaluated. the root cause of the access site bleeding issue was not determined since there were no abnormalities found on the repo sheath and blue butterfly and other clinical information during the time of bleeding was not provided. the root cause of the high purge issue was not determined since the failure could not be reproduced and the data logs were inconclusive.
30065 | 1220648-2024-07749 | 18866957 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18866957&pc=OZD | 02/24/2024 | 05/02/2024 |  | 03/08/2024 |  |  | 02/24/2024 | 03/08/2024 |  |  | 11/30/2025 |  | 03/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024425504 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user reports a patient with hemodynamic support from an impella cp smartassist heart pump who experienced massive bleeding at the access point of the heart pump at a ptt of 180 seconds after temperature control measures and ventilation had ended. the systemic administration of heparin was then paused until the ptt decreased significantly. however, the bleeding continued despite pressure dressing, so the pump was removed and hemostasis was achieved using a femstop. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The impella device was received from the customer and an investigation regarding the returned device has been completed. no abnormalities were found. the root cause of the access site bleeding issue was not determined as insufficient information is provided in the clinical description. impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d.6a and d.6b revised from the initial submission in accordance with updated procedures. g.1manufacturing site fax number was revised from the initial submission in accordance with updated procedures. h.6 health effect - impact code was revised as code was inadvertently omitted from the previously submitted report. h.10 additional manufacturer narrative has been revised as the previous report was incorrect and stated that the device was not returned. h.10 additional manufacturer narrative instructions for use were revised from the initial submission in accordance with updated procedures.
30066 | 1220648-2024-07750 | 18867004 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18867004&pc=OZD | 02/22/2024 | 04/10/2024 |  | 03/08/2024 |  |  | 02/22/2024 | 03/08/2024 |  |  | 06/30/2024 |  | 03/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023237246 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/22/2024 |  |  |  |  | The user facility reported a 65-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was during support a right ventricle thrombus was noted. heparin was given to assist with the thrombus. | [5 empty] | The investigation into the thrombus issue has been completed since the original report was submitted. the product was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30067 | 1220648-2024-07751 | 18867086 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18867086&pc=OZD | 02/22/2024 | 04/23/2024 |  | 03/08/2024 |  |  | 02/22/2024 | 03/08/2024 |  |  | 09/30/2025 |  | 03/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024406966 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported during impella 5.5 support for 65-year-old male patient, the pump was malpositioned in aorta. the pump was unable to be repositioned and was successfully explanted. there was no reported patient harm. | [5 empty] | The investigation into the positioning issues has been completed since the original report was submitted. the device was not returned for investigation. based on clinical description, the root cause of positioning issue was most likely use related. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30068 | 1220648-2024-07752 | 18867359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18867359&pc=OZD | 02/29/2024 | 04/17/2024 |  | 03/08/2024 |  |  | 02/29/2024 | 03/08/2024 |  |  |  |  | 03/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025438005 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/29/2024 |  |  |  |  | The user facility reported a 70-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the implanted impella 5.5 pump began to trigger high purge pressure alarms during patient support. the purge cassette, purge bag, and automated impella controller (aic) were all exchanged to troubleshoot the issue, but the alarms persisted. due to the ongoing issue, the decision was made to replace the pump. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the high purge pressure (hpp) issue has been completed since the original report was submitted. the device was not returned for investigation. based on the data log, the root cause of the hpp was determined to be use related as the pump was not primed before implant.
30069 | 1220648-2024-07753 | 18867453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18867453&pc=OZD | 02/28/2024 | 04/17/2024 |  | 03/08/2024 |  |  | 02/29/2024 | 03/08/2024 |  |  | 11/30/2025 |  | 03/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024422151 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/29/2024 |  |  |  |  | The user facility reported a 58-year-old male of an unknown race with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the patient developed a small hematoma at the at the access site. the patient was taken to the operating room for a washout and a drain was placed to manage the condition. heparin was temporarily stopped for the procedure but was planned to be restarted later that afternoon. support with the implanted pump was maintained. the patient was reported as stable. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding issue was not determined since product was not returned and not much information is provided in the clinical description. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
30070 | 1220648-2024-07754 | 18867494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18867494&pc=OZD | 03/01/2024 | 03/21/2024 |  | 03/08/2024 |  |  | 03/01/2024 | 03/08/2024 |  |  |  |  | 03/08/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1692374 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/01/2024 |  |  |  |  | The user facility reported a 33-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the complainant reported that the automated impella controller (aic) displayed a "purge not detected¿ alarm despite a purge cassette being installed. the aic was swapped to resolve the noted issue. there was no patient harm. | [5 empty] | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were reviewed and there was no evidence of the reported ¿purge not detected¿ alarms found. however, there were multiple instances of purge system open. the console was evaluated and it was confirmed that the blue purge disc retainer was found to be broken. the root cause of this issue was a broken purge disc retainer. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30071 | 1220648-2024-07755 | 18867534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18867534&pc=PYX | 03/01/2024 | 04/18/2024 |  | 03/08/2024 |  |  | 03/01/2024 | 03/08/2024 |  |  | 11/30/2025 |  | 03/08/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC |  |  | US |  | 2024431758 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Insufficient Flow or Under Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/01/2024 |  |  |  |  | The user facility reported a 66-year-old female was implanted with an impella rp for mechanical circulatory support. it was reported that the implanted impella rp flex pump began to experience a drop in flow rates during patient support. despite being set to a performance level of p6, flows were measured at 1.1 l/m and continued to decrease. due to the minimal flows and lack of support for the patient, the decision was made to explant the pump. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow issue has been completed since the original report was filed. the pump was returned, tested, and evaluated. visually inspected the pump and biomaterial observed around impeller and inside cannula. the cause of the low pump flow issue was biomaterial ingestion.
30072 | 1220648-2024-07756 | 18867645 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18867645&pc=OZD | 03/02/2024 | 04/18/2024 |  | 03/08/2024 |  |  | 03/02/2024 | 03/08/2024 |  |  | 11/30/2025 |  | 03/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025439273 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/02/2024 |  |  |  |  | The user facility reported a 46-year-old male with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella placement was slightly shallow in the cardiac catheterization lab, but the physician was satisfied with placement. the impella came back during turn and inlet on aortic valve (av). the physician advanced at bedside, but impella flipped and was catching on the anterior leaflet of mitral valve (mv). the decision was made to go back to cardiac catheterization lab to reposition under fluoroscopy. the pigtail was caught on the pap muscle and got pulled back across av. the physician attempted to advance back across valve but was unsuccessful and the impella was removed. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed since the original report was submitted. the device was not returned for investigation. data logs confirmed pump was in improper position corresponded with clinical description that triggered alarms impella position in ventricle/aorta. based on clinical description & data analysis, the root cause of positioning issue was most likely use related.
30073 | 1220648-2024-07757 | 18867676 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18867676&pc=OZD | 03/02/2024 | 04/18/2024 |  | 03/08/2024 |  |  | 03/02/2024 | 03/08/2024 |  |  | 07/31/2025 |  | 03/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025437225 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/02/2024 |  |  |  |  | The user facility reported a 56-year-old female in acute myocardial infarction cardiogenic shock was to be implanted with an impella cp device for mechanical circulatory support. it was reported that impella cp placement was aborted due to the patient having the iliac too small. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the delivery issue was most likely patient condition related. impella was unable to pass through illiac due to vessel size as per the clinical description.
30074 | 1220648-2024-07758 | 18867692 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18867692&pc=OZD | 03/03/2024 | 04/18/2024 |  | 03/08/2024 |  |  | 03/03/2024 | 03/08/2024 |  |  | 09/30/2025 |  | 03/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2025452529 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Separation | Unraveled Material | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/03/2024 |  |  |  |  | The user facility reported a 54-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the connection between the red plug and impella driveline had broken down to the point that the internal wiring was visible. it was noted that the patient had been moving a lot during support. due to the noted damage, the staff was advised to replace the pump. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the product damage issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the product damage issue was not determined since product was not returned for investigation.
30075 | 1220648-2024-07759 | 18867716 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18867716&pc=OZD | 03/03/2024 | 04/22/2024 |  | 03/08/2024 |  |  | 03/03/2024 | 03/08/2024 |  |  | 07/31/2025 |  | 03/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025437259 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/03/2024 |  |  |  |  | The user facility reported a 66-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that after peel away sheath was removed, large hematoma developed despite angle match. angiogram confirmed the bleed around repo sheath. the pump was removed, and a new pump placed. leaving the introducer sheath in place. bleed was no longer noted on angiogram. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. d4-updated model number
30076 | 1220648-2024-07760 | 18870808 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18870808&pc=OZD | 07/01/2023 | 09/26/2024 |  | 03/08/2024 |  |  | 02/08/2024 | 03/08/2024 |  |  | 03/31/2025 |  | 03/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024345182 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 47 YR |  | Female |  | Non Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/08/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry indicating this patient to have endured a subarachnoid and intraventricular hemorrhage with a "possible" relationship to the impella device. the patient's family ultimately chose to withdraw care and the patient expired. | [5 empty] | The investigation into the stroke/cva issue has been completed since the initial report was submitted. the product was not returned for investigation. as the clinical information was not provided, the root cause of the stroke/cva could not be determined. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information was provided in sections h1 and h6.
30077 | 1220648-2024-07657 | 18855161 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18855161&pc=OZD | 03/03/2023 | 03/07/2024 |  | 03/07/2024 |  |  | 03/03/2023 | 03/07/2024 |  |  |  |  | 03/07/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1666616 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a 72-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that while the patient was on impella 5.5 support, the automated impella controller failed to recognize a purge cassette. the impella 5.5 device was successfully transferred to a back up automated impella controller. | [5 empty] | The automated impella controller console was received from the customer and an investigation regarding the returned device has been completed. device analysis found that the console had broken purge retainer stand and broken purge flag. the root cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records
30078 | 1220648-2024-07658 | 18855517 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18855517&pc=OZD | 11/03/2022 | 03/07/2024 |  | 03/07/2024 |  |  | 11/03/2022 | 03/07/2024 |  |  | 03/31/2024 |  | 03/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023209682 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 23 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/03/2022 |  |  |  |  | The user facility reported a 23-year-old female undergoing cardiac surgery, and was being implanted with an impella 5.5 device for mechanical circulatory support and the impella 5.5 pump was unable to pass through the innominate artery during attempted delivery. multiple techniques were utilized during attempts to implant the pump, but the device was still unable to advance into the left ventricle. as a result of the noted issue, the decision was made to abort the implant, and an impella cp was subsequently placed for patient support. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issue - anatomy was most likely due to the patient condition. this report is being filed as part of a retrospective review of historical records.
30079 | 1220648-2024-07659 | 18855521 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18855521&pc=OZD | 02/10/2024 | 05/06/2024 |  | 03/07/2024 |  |  | 03/05/2024 | 03/07/2024 |  |  | 11/30/2025 |  | 03/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2024428943 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/05/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant in japan reported that the impella cp placement was aborted as the patient¿s vessel had tortuosity from the femoral artery to the aortic arch. there was no patient harm. | [5 empty] | The investigation into the delivery issues has been completed since the original report was submitted. the device was not returned by the user facility for investigation. the root cause of the delivery issue is determined to be patient condition because of tortuous anatomy from the femoral artery to the aortic arch of the patient. in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30080 | 1220648-2024-07660 | 18855797 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18855797&pc=OZD | 11/04/2022 | 10/28/2025 |  | 03/07/2024 |  |  | 11/04/2022 | 03/07/2024 |  |  | 07/31/2024 |  | 03/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023250324 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Improper Flow or Infusion | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/04/2022 |  |  |  |  | The user facility reported a 51-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the patient experienced multiple complications during impella cp support. it was noted that the impella flows were lower than expected up until explant. multiple support levels were utilized, however the flows remained below expected results. it was further documented that the patient had oozing around their access site at the start of support. manual pressure and femstop were applied to treat the oozing. heparin and aggrastat were temporarily halted in response to condition. the following day, the patient's impella leg had no pulse and the skin was mottled. the pump was explanted and the patient was escalated to impella 5.5 support. a fasciotomy was performed and clot was removed from the femoral artery. flow was restored to the level of the calf, but tissue death was present. heparin was reintroduced due to clot burden in distal right leg. support continued with the newly implanted 5.5 pump following the intervention. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
30081 | 1220648-2024-07661 | 18855996 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18855996&pc=OZD | 12/18/2022 | 03/07/2024 |  | 03/07/2024 |  |  | 12/18/2022 | 03/07/2024 |  |  |  |  | 03/07/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1301047 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Alarm System | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 12/18/2022 |  |  |  |  | The user facility reported that the automated impella controller (aic) had a controller error. the aic was swapped out. | [5 empty] | The investigation into the controller error (yellow alarm) has been completed. product was returned and investigated. data logs were reviewed, there were impella position in aorta and intermittent suction alarms. neither of these alarms would have advised the console swap. there was no issue found during the case. the device functioned operated to specifications. this report is being filed as part of a retrospective review of historical records.
30082 | 1220648-2024-07662 | 18856066 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18856066&pc=OZD | 11/19/2022 | 09/15/2024 |  | 03/07/2024 |  |  | 11/19/2022 | 03/07/2024 |  |  | 01/31/2024 |  | 03/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023186110 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/19/2022 |  |  |  |  | The user facility reported a female patient was being implanted with an impella 5.5 device for mechanical circulatory support, and that the impella 5.5 pump could not be placed. the doctor performed a cutdown for the right axillary placement, but was unable to find the artery. direct insertion was suggested by the onsite representative, however the doctor proceeded to attempt to place the pump without a graft. although the pump was able to advance to the aorta, it would not cross the aortic valve. the implant was subsequently aborted. | It was noted that the patient expired in the operating room. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of delivery issue: anatomy was most likely due to the patient condition. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07662 was provided.
30083 | 1220648-2024-07663 | 18856346 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18856346&pc=OZD | 12/21/2022 | 03/07/2024 |  | 03/07/2024 |  |  | 12/21/2022 | 03/07/2024 |  |  | 08/31/2024 |  | 03/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023269658 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/21/2022 |  |  |  |  | The user facility reported a 54-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, there was a large right atrium (ra) clot (thrombus) visualized in the pericardium obstructing preload. patient continued to show improvement and the impella was explanted. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
30084 | 1220648-2024-07664 | 18856444 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18856444&pc=OZD | 12/11/2022 | 09/12/2024 |  | 03/07/2024 |  |  | 12/11/2022 | 03/07/2024 |  |  | 08/31/2024 |  | 03/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023261428 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/11/2022 |  |  |  |  | The user facility reported a 50-year-old female being placed on impella cp required mechanical circulatory support. the patient experienced significant bleeding post peel away sheath removal. two units of blood products were given for intervention of access site bleeding. the patient repeatedly experienced ventricular fibrillation (vf) arrest throughout the procedure. the impella cp was placed on p2 and cpr was performed multiple times for over 30 mins. the patient was on impella cp support for 1.12 hours before the patient expired. no allegations were made towards the impella cp for cause of death. | It was noted that the patient expired in the operating room. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | B2: date of patient death is unknown. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07664 was provided in sections b2, b5, h1, and h6.
30085 | 1220648-2024-07665 | 18856498 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18856498&pc=OZD | 05/31/2023 | 03/12/2025 |  | 03/07/2024 |  |  | 05/31/2023 | 03/07/2024 |  |  | 01/31/2024 |  | 03/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023184657 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/31/2023 |  |  |  |  | The user facility reported a 86-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the complainant reported that upon explant of the device, a clot was noted in the right iliac. a penumbra catheter was utilized to remove the thrombus and flow was restored to the area. support continued with the subsequently implanted impella 5.5 pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | "the failure mode of ¿thrombus¿ is to be used when a thrombus is noted but cannot be connected to any other adverse events or decrease in pump performance. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions description of the location and characteristics of the thrombus (e.g., size, shape, consistency). relevant laboratory test results, such as coagulation profiles and platelet counts imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). medications or treatments that the patient was receiving at the time of thrombus formation surgical or procedural details if applicable (e.g, cardiac catheterization). any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined." this report is being filed as part of a retrospective review of historical records.
30086 | 1220648-2024-07666 | 18856515 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18856515&pc=OZD | 07/07/2023 | 09/22/2025 |  | 03/07/2024 |  |  | 07/07/2023 | 03/07/2024 |  |  | 05/31/2025 |  | 03/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024343805 |  | 1000211 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | Device Handling Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/07/2023 |  |  |  |  | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during the procedure to insert the impella 5.5, the pump did not advance from the right subclavian artery to the vicinity of the brachiocephalic artery bifurcation. a placement signal 'not reliable' alarm appeared. the impella had to be removed and upon removal, a kink was observed near the motor. it was decided to discontinue insertion. there was no harm to the patient. | [5 empty] | The investigation for the reported delivery issues has been completed. the returned complaint 5.5 pump (b)(6) was visually inspected. catheter kink near motor housing observed. no other abnormality found. the cause of the delivery issue was the patient's condition since their vessel diameter was below the recommended range and the anatomy was reported to be challenging. pump discarded prior to being evaluated for the optical signal issue. the cause of the optical signal issue most likely was damaged optical fiber line due to excessive force that kinked the catheter. this report is being filed as part of a retrospective review of historical records. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the cause of the optical signal issue most likely was damaged optical fiber line due to excessive force that kinked the catheter. all pertinent information available to abiomed has been submitted at this time. d4 model number updated. d6a/d6b updated with additional information. f6 and f8 should have been left blank from the initial manufacturer device report number 1220648-2024-07666. h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2024-07666.
30087 | 1220648-2024-07667 | 18856686 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18856686&pc=OZD | 02/23/2023 | 03/07/2024 |  | 03/07/2024 |  |  | 02/23/2023 | 03/07/2024 |  |  |  |  | 03/07/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1103168 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Alarm System | Unexpected Shutdown | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a 38-year-old male was implanted with an impella 5.5 device for mechanical circulatory support. during the positioning of the impella 5.5 device, the physician noted numerous twists in the catheter. the catheter was unplugged from the automated impella controller (aic) to remove coils and twists from the line. upon re-insertion and attempted re-start of catheter, a yellow error message appeared on the aic, "impella catheter defective". the aic was restarted multiple times and the same error message occurred. the aic was successfully replaced with a different unit. there was no reported patient harm reported. | [5 empty] | The automated impella controller console was received from the customer and an investigation regarding the returned device has been completed. the console was extensively tested, using multiple pumps and pump simulators, but the issues were not reproduced. during data analysis, the console logs from the reported day of event were consistent with the complaint. the root cause of the issue was not determined since the issue could not be reproduced. this report is being filed as part of a retrospective review of historical records.
30088 | 1220648-2024-07669 | 18856851 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18856851&pc=OZD | 07/16/2023 | 03/07/2024 |  | 03/07/2024 |  |  | 07/16/2023 | 03/07/2024 |  |  | 04/30/2025 |  | 03/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024337451 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/16/2023 |  |  |  |  | The user facility reported a 56-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. during support, the device had reoccurring suction alarms. echo showed placement incorrect, the doctor attempted to torque and was unable to move device. the doctor decided to exchange device as it had been in 40 days and had also lost placement signal. | [5 empty] | The investigation into optical signal issue and positioning issues has been completed. the product and data logs were returned for evaluation. the placement signal is low throughout the whole case. during this case there are suction alarms, and position in aorta alarms present. when the placement signal spikes to 3000 in the second lt log, there are multiple suction alarms present. additionally, there is sensor drift. when the sensor began to drift, they were counteracting is with lv offsets. when they began to step up the lv offset, the placement signal drifts, offset is added, the placement signal adjusts and then it drifts again. the pump was received and run in the lab. the place signal not reliable (psnr) did not reproduce but the placement signal was drifting. the root cause of the optical signal issue is most likely due to sensor wear that is within the indication of use. additionally, it was stated that the pump was in the wrong position and the physician struggled to reposition it. the root cause of the positioning issues is most likely due to use. this report is being filed as part of a retrospective review of historical records.
30089 | 1220648-2024-07670 | 18856853 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18856853&pc=OZD | 12/11/2022 | 03/07/2024 |  | 03/07/2024 |  |  | 12/11/2022 | 03/07/2024 |  |  | 08/31/2024 |  | 03/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023262830 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/11/2022 |  |  |  |  | The user facility reported a 64-year-old male had a cp support and developed a hemorrhagic stroke. the impella was removed. | [5 empty] | The investigation for the reported hemorrhagic stroke has been completed. the device was not returned for evaluation nor was there sufficient clinical information provided. therefore, the root cause of the access site bleed could not be determined. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
30090 | 1220648-2024-07668 | 18856949 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18856949&pc=OZD | 06/23/2023 | 03/07/2024 |  | 03/07/2024 |  |  | 06/23/2023 | 03/07/2024 |  |  | 03/31/2025 |  | 03/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024326955 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/23/2023 |  |  |  |  | The user facility reported a 75-year-old male patient was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the medical team experienced difficulty implanting the device due to patient's anatomy. ultimately the medical team aborted the procedure. there was no patient harm reported. | [5 empty] | The investigation into the reported delivery issue has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine a root cause. the most likely cause of the delivery issue was due to patient condition. this report is being filed as part of a retrospective review of historical records.
30091 | 1220648-2024-07671 | 18856977 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18856977&pc=OZD | 04/10/2023 | 03/07/2024 |  | 03/07/2024 |  |  | 04/10/2023 | 03/07/2024 |  |  | 03/31/2025 |  | 03/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024320024 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/10/2023 |  |  |  |  | The user facility reported a 68-year-old female with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. the user facility reported that a "placement signal unreliable" alarm had occurred. switching to a new console did not resolve the issue. the user was instructed to turn the placement signal alarm off and monitor the device for motor current and flows. the patient was successfully weaned and the impella 5.5 was explanted. there was no harm to the patient as a result of this incident. | [5 empty] | The impella device was not returned, and the investigation has been completed. the logs were returned and investigated. the placement signal not reliable alarms were confirmed. at the end of support, an electrical pump stop was observed because an alarm 119 was found. the root cause of the pump stop issue was not determined since product was not returned and log analysis is inconclusive of the root cause. the root cause of the optical signal issue was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records.
30092 | 1220648-2024-07672 | 18857151 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18857151&pc=OZD | 12/11/2022 | 09/12/2024 |  | 03/07/2024 |  |  | 12/11/2022 | 03/07/2024 |  |  | 04/30/2024 |  | 03/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218783 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Infusion or Flow Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/11/2022 |  |  |  |  | The user facility reported a 64-year-old female had a cp support and developed limb ischemia and lost pulses in either foot, as confirmed via doppler. during this time the patient also experienced low flow/suction events that did not subside, cpr was performed but eventually the patient expired after 6.08 hours of impella support. no allegations were made towards the impella as the cause of the death. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07672 was provided in sections b2, b5, h1, and h6.
30093 | 1220648-2024-07707 | 18858415 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18858415&pc=OZD | 08/07/2022 | 03/12/2025 |  | 03/07/2024 |  |  | 08/07/2022 | 03/07/2024 |  |  | 03/31/2024 |  | 03/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023210583 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/07/2022 |  |  |  |  | The user facility reported a 60-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. after being positioned on operating room table, it was found that the catheter had migrated into the aorta with only the pigtail going across the aortic valve. after failed attempts to advance by the physician, the decision was made to explant and reimplant a new cp device. the exchange was successful, and later the patient was weaned off the pump with no issues. | [5 empty] | The impella device was discarded by the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30094 | 1220648-2024-07708 | 18858559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18858559&pc=OZD | 02/26/2024 | 03/19/2024 |  | 03/07/2024 |  |  | 02/26/2024 | 03/07/2024 |  |  |  |  | 03/07/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1165512 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/26/2024 |  |  |  |  | Us complainant reported the automated impella controller (aic) had a battery failure (red alarm). the aic was sent in for investigation. no patient is involved. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the reported console (aic) red alarm issue has been completed. the aic was returned for evaluation. upon functional evaluation, battery failure issue was able to be reproduced. lcd charge level indicator on battery 2 was completely empty. the data logs were reviewed and confirmed the battery failure on the occurrence date. the root cause of this issue is battery internal cell imbalance.
30095 | 1220648-2024-07709 | 18858724 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18858724&pc=OZD | 02/27/2024 | 04/12/2024 |  | 03/07/2024 |  |  | 02/27/2024 | 03/07/2024 |  |  |  |  | 03/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025444072 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/27/2024 |  |  |  |  | The user facility reported a 71-year-old male with cardiomyopathy was implanted with an impella for mechanical circulatory support. the patient was on impella 5.5 support and while on support the patient had access site bleeding. the impella 5.5 was removed and the impella cp was implanted. patient is stable post intervention. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
30096 | 1220648-2024-07710 | 18858813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18858813&pc=OZD | 12/27/2022 | 03/07/2024 |  | 03/07/2024 |  |  | 12/27/2022 | 03/07/2024 |  |  | 06/30/2024 |  | 03/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231315 |  | 0048-0371 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/27/2022 |  |  |  |  | The user facility reported a 66-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while the patient was on support, the staff observed a small leak in purge fluid at check valve/yellow hub. it is unknown if the purge cassette was replaced. there was no patient harm related to this event. | [5 empty] | The impella device was not received from the customer, therefore investigation of the device was not possible. should additional information be provided in the future, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
30097 | 1220648-2024-07711 | 18858908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18858908&pc=PYX | 02/27/2024 | 09/12/2024 |  | 03/07/2024 |  |  | 02/27/2024 | 03/07/2024 |  |  |  |  | 03/07/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025442092 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | Hemothorax | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/27/2024 |  |  |  |  | The user facility reported an 81-year-old male with cardiomyopathy was implanted with an impella rp for mechanical circulatory support. it was reported that upon implant of the impella rp, the physician was believed to have nicked something resulting in a hemothorax. the patient subsequently experienced bradycardia and suffered a cardiac arrest. three rounds of cardiopulmonary resuscitation (cpr) were performed before a return of spontaneous circulation was achieved. support continued with the implanted pump. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: avoid manual compression of the inlet and outlet areas of the cannula assembly.¿ impella rp: ¿handle with care. the impella rp with smartassist system catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07711 was provided in sections b2, b5, h1, and h6. | The investigation into the vf/vt/af/at (bradycardia) and vascular damage - great vessel issues has been completed since the original report was submitted. the device was not returned for investigation. as the necessary information was not provided, the root cause of the vt/vf was not determined. the cause of the vascular damage issue was not determined since product was not returned and due to insufficient clinical information.
30098 | 1220648-2024-07712 | 18859020 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18859020&pc=OZD | 11/15/2022 | 03/07/2025 |  | 03/07/2024 |  |  | 11/15/2022 | 03/07/2024 |  |  | 03/31/2024 |  | 03/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023206837 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/15/2022 |  |  |  |  | The user facility reported a patient was implanted with an impella cp device for mechanical circulatory support. the patient had bleeding (oozing) at the impella cp insertion site, and worsening cardiogenic and distributive shock with multiple organ dysfunction syndrome. the patient required frequent packed red blood cell transfusions, totaling 6 units since admission due to a decrease in platelets and oozing at the impella site. the decision was made to remove the impella device. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30099 | 1220648-2024-07715 | 18859216 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18859216&pc=OZD | 02/28/2024 | 04/17/2024 |  | 03/07/2024 |  |  | 02/28/2024 | 03/07/2024 |  |  | 09/30/2025 |  | 03/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024408235 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Insufficient Flow or Under Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/28/2024 |  |  |  |  | The user facility reported a 35-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on support the pump had low flow. attempted repositioning the pump, however flows remained low. it was suspected there was a clot, but there were no clots noted upon explant. the pump was replaced with a new device and no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow issue has been completed since the original report was filed. the pump was returned, tested, and evaluated. the failure did not reproduce. no motor current spikes or any positioning issues shown in the data logs. the cause of the low pump flow issue cannot be determined as no abnormalities were found on the device and it was able to flow within specification.
30100 | 1220648-2024-07716 | 18859251 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18859251&pc=OZD | 02/29/2024 | 04/17/2024 |  | 03/07/2024 |  |  | 02/29/2024 | 03/07/2024 |  |  | 11/30/2025 |  | 03/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024423666 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/29/2024 |  |  |  |  | The user facility reported a 70-year-old male post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the after 30 minutes of support the pump stopped and was unable to be re-started despite troubleshooting measures inclusive of and automated impella controller (aic) swap. the pump was explanted, placed on different device for support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the pump stop issue has been completed since the original report was filed. the pump was returned, tested, and evaluated. product evaluation confirmed the presence of a blue tubing material likely from an interaction with another device inside the cannula and wrapped around the impeller. the cause of the pump stop was an interaction with another device resulting in the pump ingesting a blue tubing material.
30101 | 1220648-2024-07713 | 18859275 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18859275&pc=OZD | 12/14/2022 | 05/22/2025 |  | 03/07/2024 |  |  | 12/14/2022 | 03/07/2024 |  |  | 08/31/2024 |  | 03/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023262333 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/14/2022 |  |  |  |  | The user facility reported a white 51-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support, and the patient experienced a spontaneous coronary artery dissection (scad). cardiopulmonary resuscitation (cpr) was performed while in the cardiac catch lab (ccl) before being sent emergently to the or; where the physician performed a coronary artery bypass graft surgery (cabg) x2, left internal mammary artery (lima) to left anterior descending (lad) and saphenous vein graft (svg) to obtuse marginal (om). it was determined that the patient would need to be escalated to extracorporeal membrane oxygenation (ecmo) as a result of low oxygenation, small left ventricle (lv), and poor right heart (rh) function. upon removal of the impella cp, the patient remained stable and on the ecmo. the impella site was reportedly stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30102 | 1220648-2024-07714 | 18859292 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18859292&pc=OZD | 02/28/2024 | 04/15/2024 |  | 03/07/2024 |  |  | 02/28/2024 | 03/07/2024 |  |  |  |  | 03/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC., |  |  | US |  | 2025438627 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/28/2024 |  |  |  |  | The user facility reported a 78-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that an implanted impella cp pump began to trigger high purge pressure alarms during patient support. pump flows remained appropriate and the patient was stable at the time of the noted event. due to noted issue, the decision was made to replace the pump. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported high purge pressure has been completed. product was not returned for investigation. after around 20 minutes of pump running there was motor current spike followed by a gradual purge pressure increase and then steading out around 1050mmhg. low flow was seen once the purge pressure hit 1050mmhg gradually decreasing till the end of the case. per the data log review, the root cause of the high purge pressure issue was most likely biomaterial. the instructions for use states ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿
30103 | 1220648-2024-07717 | 18859419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18859419&pc=OZD | 02/21/2024 | 04/10/2024 |  | 03/07/2024 |  |  | 02/21/2024 | 03/07/2024 |  |  | 11/30/2025 |  | 03/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024422869 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/21/2024 |  |  |  |  | The user facility reported a 61-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that upon impella 5.5 insertion, the 5.5 was backloaded onto the wire, advanced into the left ventricle (lv), and the patient experienced ventricle tachycardia. cardiopulmonary resuscitation was started by the physician and the patient was shocked once, returning to a normal rhythm. the initial measurement was 6.8cm in lv. the impella was pulled back using transesophageal echocardiography¿(tee) to the depth of 5.6cm. the impella was initiated at p-2. the patient was stable. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the vf/vt (ventricular tachycardia) issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined.
30104 | 1220648-2024-07718 | 18859470 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18859470&pc=OZD | 02/22/2024 | 04/08/2024 |  | 03/07/2024 |  |  | 02/22/2024 | 03/07/2024 |  |  | 11/30/2025 |  | 03/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024432594 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/22/2024 |  |  |  |  | The user facility reported a 65-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that a patient experienced persistent bleeding from their access site following implant. the patient was given five units packed red blood cells, one unit of fresh frozen plasma, one unit of cryoprecipitate, and one unit of platelets to counteract the blood loss. the pump was subsequently explanted one day after implant. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation into low access site bleeding - major has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the access site bleeding issue was not determined since product was not returned, and not enough information is provided in the clinical description provided.
30105 | 1220648-2024-07719 | 18859591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18859591&pc=OZD | 02/23/2024 | 08/19/2024 |  | 03/07/2024 |  |  | 02/23/2024 | 03/07/2024 |  |  | 03/31/2024 |  | 03/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023214381 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/23/2024 |  |  |  |  | The user facility reported a 60-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient had no distal pulses on support. a reperfusion catheter was placed with positive distal perfusion. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The following have been added to manufacturer device report 1220648-2024-07719. d6a and d6b: h6 component code added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-07719. g1 reporting contact fax number missing.
30106 | 1220648-2024-07720 | 18859743 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18859743&pc=OZD | 12/12/2022 | 03/13/2025 |  | 03/07/2024 |  |  | 12/12/2022 | 03/07/2024 |  |  | 08/31/2024 |  | 03/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023270061 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/12/2022 |  |  |  |  | The user facility reported a 74-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the flows were lower than expected, around 2l/min on p-auto, so medical staff looked on fluororoscopy and impella cannula was visibly kinked. the doctor attempted to adjust the position, but the kink partially remained with the flows improving to around 2.8l/min. no further intervention was required. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30107 | 1220648-2024-07721 | 18859777 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18859777&pc=OZD | 02/24/2024 | 04/03/2024 |  | 03/07/2024 |  |  | 02/24/2024 | 03/07/2024 |  |  | 09/30/2025 |  | 03/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024399694 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/24/2024 |  |  |  |  | The user facility reported a 61-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the patient had a hematoma while on support. heparin was stopped and three units of blood products were provided. the patient was stable post implant. | [5 empty] | The investigation into the access site bleeding major issue has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was most likely due to anticoagulation management since the hemostasis was achieved and bleed resolved by turning off the heparin. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
30108 | 1220648-2024-07722 | 18860027 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18860027&pc=PYX | 03/03/2024 | 05/23/2024 |  | 03/07/2024 |  |  | 03/03/2024 | 03/07/2024 |  |  | 10/31/2025 |  | 03/07/2024 |  | RP FLEX W SMART ASSIST SET, US |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | RP FLEX W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2024413655 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/03/2024 |  |  |  |  | The user facility reported a 53-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an implanted impella rp flex pump experienced multiple complications during patient support. two days into support, low purge pressure alarms began to appear. troubleshooting identified no notable leaks and purge cassette replacement failed to resolve the issue. that same day, pump flows began to drop and the pump stopped unexpectedly. the pump was able to restart, but the decision was subsequently made to replace the pump. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ ¿ do not clean with or expose any part of the clear sidearm of the impella catheter (eg, infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir.¿ | The investigation into the temporary pump stop and the purge leak ¿ pump issues has been completed since the original report was filed. the pump was returned for investigation. biomaterial was found wrapped around the impeller blades. upon removal of the biomaterial the pump ran at optimal flows and without any issues. the root cause of the temporary pump stop issue was biomaterial found on the impeller blades. the root cause of the purge leak - cassette issue was not determined since the affected purge cassette was not returned for investigation. in accordance with updated procedures, section d6 has been revised from the initial submission.
30109 | 1220648-2024-07723 | 18860265 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18860265&pc=OZD | 02/27/2024 | 04/12/2024 |  | 03/07/2024 |  |  | 02/27/2024 | 03/07/2024 |  |  | 11/30/2025 |  | 03/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024432818 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/27/2024 |  |  |  |  | The user facility reported a 66-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that following implant, the device displayed a flat motor current and began to trigger continuous suction alarms. the device was removed, cleaned, and re-inserted, but the issue persisted with flows as low as .4 l/m at support level p-9. due to the persistent issue, the decision was made to replace the pump. there was no patient harm. | [5 empty] | The investigation into the low pump flow issue has been completed since the original report was filed. the pump was returned, tested, and evaluated. visually inspected the returned pump and a big chunk of biomaterial observed around the impeller. the cause of the low pump flow issue was biomaterial ingestion. | The impella device was received from the customer and¿an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed.
30110 | 1220648-2024-07725 | 18860382 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18860382&pc=OZD | 11/15/2022 | 03/07/2024 |  | 03/07/2024 |  |  | 11/15/2022 | 03/07/2024 |  |  | 07/31/2024 |  | 03/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023255887 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/15/2022 |  |  |  |  | The user facility reported a 62-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after repositioning the sheath, the patient experienced bleeding despite proper angle match. a forward tension suture was applied, and thirty (30) minutes of manual pressure was held above and below site to maintain angle of entry. the patient was transferred to icu with a femstop in place. the impella cp was successfully explanted, and the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ impella cp: for access site bleeding, ¿assess access site for bleeding and hematoma.¿ introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
30111 | 1220648-2024-07724 | 18860513 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18860513&pc=OZD | 03/04/2024 | 04/19/2024 |  | 03/07/2024 |  |  | 03/04/2024 | 03/07/2024 |  |  |  |  | 03/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025443322 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Pressure Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 03/04/2024 |  |  |  |  | The user facility reported a 36-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported a high purge pressure alarm that did not resolve with a cassette change. tpa (tissue plasminogen activator) will be administered. | [5 empty] | The investigation into the high purge pressure issue has been completed since the original report was submitted. the device was not returned for investigation. per the data logs, the root cause of the high purge pressure issue was most likely biomaterial. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
30112 | 1220648-2024-07726 | 18860656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18860656&pc=OZD | 02/27/2024 | 04/01/2024 |  | 03/07/2024 |  |  | 02/27/2024 | 03/07/2024 |  |  | 03/31/2025 |  | 03/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024340377 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/27/2024 |  |  |  |  | The user facility reported a 56-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed hemolysis during support. troubleshooting steps were provided for treating the condition, but the decision was ultimately made to explant the pump the following day. | [5 empty] | The investigation for the reported hemolysis has been completed. the device nor sufficient clinical information was returned for evaluation. data logs reviewed showed there were no issues with the pump the entire case. therefore, the root cause of the hemolysis could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
30113 | 1220648-2024-07727 | 18861407 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18861407&pc=OZD | 07/31/2022 | 03/13/2025 |  | 03/07/2024 |  |  | 07/31/2022 | 03/07/2024 |  |  | 01/31/2024 |  | 03/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | CP IMPELLA |  |  | ABIOMED, INC. |  |  | US |  | 2023186650 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/31/2022 |  |  |  |  | The user facility reported a 53-year-male with cardiomyopathy experienced shortness of breath and began to decompensate. he was implanted with an impella cp device for mechanical circulatory support. the patient reportedly had an ischemic leg. the impella cp was explanted and the plan was to replace it with an intra-aortic balloon pump therapy (iabp) due to the smaller size. | [5 empty] | The impella cp device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | Limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
30114 | 1220648-2024-07728 | 18861572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18861572&pc=OZD | 02/21/2023 | 03/07/2024 |  | 03/07/2024 |  |  | 02/21/2023 | 03/07/2024 |  |  |  |  | 03/07/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1631552 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Mechanical Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 02/21/2023 |  |  |  |  | The user facility reported a 64-year-old male in acute myocardial infarction cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. during initiation of support, the staff observed that the automated impella controller (aic) failed to detect an installed purge cassette. the purge cassette was replaced, however the issue remained. the automated impella controller was subsequently swapped to resolve the reported failure. upon further inspection of the replaced automated impella controller, it was noted that the cassette door would not properly close. the patient was successfully weaned, and the impella 5.5 was explanted. there was no harm to the patient. | [5 empty] | The investigation for the controller (aic) has been completed. the aic was returned for evaluation. functional testing confirmed the inability to detect the purge cassette. visual inspection found that the purge disk retainer was broken. additionally, the purge lid door was found to be stuck. data logs were reviewed but no apparent issues were found with detecting the purge cassette were present. the purge door mechanism does not have relevant data logs. the cause of the purge fluid not detected issue was a broken purge disc retainer. the cause of the inability to close the purge door was most likely due to the contamination of the latch assembly. this report is being filed as part of a retrospective review of historical records.
30115 | 1220648-2024-07729 | 18861755 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18861755&pc=OZD | 11/14/2022 | 03/13/2025 |  | 03/07/2024 |  |  | 11/14/2022 | 03/07/2024 |  |  | 07/31/2024 |  | 03/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023251626 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/14/2022 |  |  |  |  | The user facility reported a 69-year-old female with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported that the patient experienced mild continuous bleeding at the impella access site. the impella cp was repositioned to resolve the bleeding. the impella cp was successfully explanted. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
30116 | 1220648-2024-07730 | 18861850 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18861850&pc=OZD | 11/18/2022 | 03/13/2025 |  | 03/07/2024 |  |  | 11/18/2022 | 03/07/2024 |  |  | 03/31/2024 |  | 03/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023208045 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/18/2022 |  |  |  |  | The user facility reported a 54 year-old-male with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. while on the impella cp suppor, the patient experienced bleeding from the groin access site. argatroban was discontinued. cardiology redressed site and secured the angle with success. | [5 empty] | The impella cp device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | The root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided. this report is being filed as part of a retrospective review of historical records.
30117 | 1220648-2024-07527 | 18844549 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18844549&pc=OZD | 12/02/2022 | 03/12/2025 |  | 03/06/2024 |  |  | 12/02/2022 | 03/06/2024 |  |  | 08/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023266393 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/02/2022 |  |  |  |  | The user facility reported a 54-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced a small hematoma at the insertion site. the patient was successfully weaned and the impella device was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
30118 | 1220648-2024-07528 | 18844563 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18844563&pc=OZD | 12/03/2022 | 03/06/2024 |  | 03/06/2024 |  |  | 12/03/2022 | 03/06/2024 |  |  | 07/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023266086 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/03/2022 |  |  |  |  | The user facility reported a 73-year-old female implanted with an impella 5.5 requiring mechanical circulatory support for heart failure/cardiogenic shock. it was reported that the impella 5.5 pump was unable to advance through the patient's vasculature during attempted delivery. due to the resistance caused by the patient's anatomy, the decision was made to abort the implant. there was no patient harm reported. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition as the patient's vasculature would not let the pump advance. this report is being filed as part of a retrospective review of historical records.
30119 | 1220648-2024-07529 | 18844578 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18844578&pc=OZD | 12/08/2022 | 09/12/2024 |  | 03/06/2024 |  |  | 12/08/2022 | 03/06/2024 |  |  | 07/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023255862 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/08/2022 |  |  |  |  | The user facility reported a 43-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that there was oozing from the impella site. units of packed red blood cells (prbc) were given as a result. subsequently, a pump stop was experienced. the pump was explanted. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07529 was provided in sections b2, b5, h1, and h6.
30120 | 1220648-2024-07531 | 18844651 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18844651&pc=OZD | 12/05/2022 | 03/12/2025 |  | 03/06/2024 |  |  | 12/05/2022 | 03/06/2024 |  |  | 07/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023247154 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/05/2022 |  |  |  |  | The user facility reported a 76-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced bleeding and a hematoma at the access site. the patient required a blood transfusion. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. the root cause of the access site bleeding cannot be determined since the product was not returned and enough clinical info was not provided.
30121 | 1220648-2024-07530 | 18844652 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18844652&pc=OZD | 12/08/2022 | 03/12/2025 |  | 03/06/2024 |  |  | 12/08/2022 | 03/06/2024 |  |  | 04/30/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023223269 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/08/2022 |  |  |  |  | The user facility reported a 77-year-old female noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported during extended inflation, there was pulsatility loss on impella and patient had bradycardia. the patient immediately recovered. impella was weaned and ultimately pulled across valve. the impella was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30122 | 1220648-2024-07532 | 18844798 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18844798&pc=OZD | 12/08/2022 | 03/12/2025 |  | 03/06/2024 |  |  | 12/08/2022 | 03/06/2024 |  |  | 08/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023271718 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Gastrointestinal Hemorrhage | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/08/2022 |  |  |  |  | The user facility reported a 58-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient started to decompensate as they were taken back to the cath lab, intra-aortic balloon pump (iabp) was placed and patient was intubated. the patient continued to decompensate, and the decision was made for the impella to be implanted. there was oozing from all stick sites. there were some intermittent suctioning, and the nurse was provided with troubleshooting steps if it continued. additionally, the patient had a gi bleed. medical staff discussed anticoagulation recommendations once able to add it back on. it was decided that no systemic anticoagulation due to upper gi bleed and oozing. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30123 | 1220648-2024-07533 | 18844879 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18844879&pc=OZD | 12/05/2022 | 09/12/2024 |  | 03/06/2024 |  |  | 12/05/2022 | 03/06/2024 |  |  | 02/29/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200625 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tricuspid Valve Insufficiency/ Regurgitation | Muscle/Tendon Damage | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/05/2022 |  |  |  |  | The user facility reported a 75-year-old female noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that after closure of a ventricular septal defect (vsd), the catheter became caught in a papillary muscle while being removed, and resulted in a ruptured papillary. this caused severe tricuspid regurgitation (tr). blood products and further intervention were declined by patient's family and care was withdrawn. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿impella catheter in papillary muscle if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ this report is being filed as part of a retrospective review of historical records. | B2: date of death is unknown. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07533 was provided in sections b2, b5, h1, and h6.
30124 | 1220648-2024-07534 | 18844919 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18844919&pc=OZD | 12/03/2022 | 03/06/2024 |  | 03/06/2024 |  |  | 12/03/2022 | 03/06/2024 |  |  | 08/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023271123 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/03/2022 |  |  |  |  | The user facility reported a 78-year-old female requiring an impella cp device for mechanical circulatory support. it was reported that the patient had intermittent pulses of the right lower leg. the patient had a cold leg overnight and reperfusion sheath was placed from left to right to restore flow. no patient harm noted. | [5 empty] | The investigation for the reported limb ischemia has been completed. the impella cp device has not been returned for evaluation. clinical details were not sufficient to determine a root cause. therefore, the root cause of the issue could not be determined. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
30125 | 1220648-2024-07535 | 18844968 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18844968&pc=OZD | 12/05/2022 | 03/12/2025 |  | 03/06/2024 |  |  | 12/05/2022 | 03/06/2024 |  |  | 08/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023269952 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/05/2022 |  |  |  |  | The user facility reported a 79-year-old female noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that during implantation, the impella cannula developed a kink upon insertion, however the valve was crossed with minimal resistance allowing approximately 2.3 liters of flow, and all three vessels had orbital atherectomy followed by stenting. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿avoid manual compression of the inlet and outlet areas of the cannula assembly.¿ ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30126 | 1220648-2024-07536 | 18845042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18845042&pc=OZD | 12/03/2022 | 03/06/2024 |  | 03/06/2024 |  |  | 12/03/2022 | 03/06/2024 |  |  | 07/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023252267 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/03/2022 |  |  |  |  | The user facility reported a 53-year-old male on impella cp for mechanical circulatory support. it was reported that the patient's urine changed color from yellow to red. despite attempts at the bedside to reposition, the urine did not clear. an echo was performed, and the pigtail was hung in the papillary muscle. the impella was repositioned in the cath lab and the urine started to change back to yellow. additionally, the patient was placed on a swan with dobutamine but soon became tachycardic. as a result, the patient was eventually escalated to the impella 5.5. no further harm was reported. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
30127 | 1220648-2024-07537 | 18845531 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18845531&pc=OZD | 12/05/2022 | 03/12/2025 |  | 03/06/2024 |  |  | 12/05/2022 | 03/06/2024 |  |  | 07/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023254094 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/05/2022 |  |  |  |  | The user facility reported a 55-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experience bleeding in the groin access site and the hemoglobin level had dropped, which required a blood transfusion. however, due to religious views the family refused the transfusion and the patient was administered an experimental drug hemopure. the groin bleeding resolved with no further intervention. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
30128 | 1220648-2024-07538 | 18845600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18845600&pc=OZD | 12/12/2022 | 03/12/2025 |  | 03/06/2024 |  |  | 12/12/2022 | 03/06/2024 |  |  | 02/29/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023196734 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Difficult to Insert | Material Twisted/Bent | [5 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/12/2022 |  |  |  |  | The user facility reported a 64-year-old male undergoing cardiac surgery after a brief cardiac arrest, was implanted with an impella cp device for mechanical circulatory support. it was reported upon implanting the impella cp, the impella was backloaded over the 0.027" wire with difficulty. staff had to clip the end of the wire due to a kink. the wire was lubricated with ultrasound jelly and it backloaded without issue and the impella insertion was successful. additionally, staff was unable to achieve doppler pulses in patient's right leg (impella side). heparin was given. patient later expired while on support, however there are no allegations that the impella device contributed to death. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30129 | 1220648-2024-07539 | 18845611 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18845611&pc=OZD | 12/12/2022 | 03/28/2025 |  | 03/06/2024 |  |  | 12/12/2022 | 03/06/2024 |  |  | 09/30/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023273589 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/12/2022 |  |  |  |  | The user facility reported a 59-year-old male in myocardial infarction was implanted with an impella cp device for mechanical circulatory support. the patient experienced atrial fibrillation while on impella cp support and was treated using intravenous drips and cardiovert. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. ¿timing of the arrythmia event in relation to impella support (during or post-implant) ¿electrocardiogram (ekg) recordings of the arrhythmia episodes. ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. ¿response to any antiarrhythmic treatments or interventions during the event. ¿any documented device-related issues or complications during impella support. ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. ¿review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined." positioning issue- the root cause of positioning issue was determined to be use issue because of the extra slack which is due to incorrect placement and was adjusted by pulling the extra slack. hence the root cause is determined to use issue. this report is being filed as part of a retrospective review of historical records.
30130 | 1220648-2024-07540 | 18845743 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18845743&pc=OZD | 08/31/2022 | 03/06/2024 |  | 03/06/2024 |  |  | 08/31/2022 | 03/06/2024 |  |  | 05/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023223627 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/31/2022 |  |  |  |  | The user facility reported a 65-year-old male in acute myocardial infarction/cardiogenic shock was to be implanted with an impella 5.5 device for mechanical circulatory support. the medical staff attempting to deliver an impella pump 5.5 via axillary artery, had difficulty due to the calcification of the artery. after failed attempts, the medical staff aborted the procedure. the patient was implanted with a new impella 5.5 into the aorta via a thoracotomy. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition of calcified, diseased vessels. this report is being filed as part of a retrospective review of historical records.
30131 | 1220648-2024-07542 | 18845804 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18845804&pc=OZD | 12/04/2022 | 03/06/2024 |  | 03/06/2024 |  |  | 12/04/2022 | 03/06/2024 |  |  | 08/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023256721 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Restricted Flow rate | Pumping Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/04/2022 |  |  |  |  | The user facility reported a 69-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. while the patient was supported with the impella cp device, there were frequent suction alarms due to low pump flow. due to the frequent alarms, the patient was weaned and the device was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
30132 | 1220648-2024-07541 | 18845810 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18845810&pc=OZD | 12/12/2022 | 03/12/2025 |  | 03/06/2024 |  |  | 12/12/2022 | 03/06/2024 |  |  | 02/29/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023196015 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/12/2022 |  |  |  |  | The user facility reported a 41-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed hemolysis and a large clot in the proximal left anterior descending coronary artery. penumbra atherectomy was attempted, however the patient lost pulsatility and experienced pulseless electrical activity. the patient received an epinephrine push twice and contractility regained. the pump was subsequently explanted the following day and the patient was started on dialysis. after the impella was explanted, the patient was hemodynamically stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30133 | 1220648-2024-07543 | 18845855 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18845855&pc=OZD | 12/08/2022 | 03/06/2024 |  | 03/06/2024 |  |  | 12/08/2022 | 03/06/2024 |  |  | 07/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023257031 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/08/2022 |  |  |  |  | The user facility reported a 79-year-old female noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient went into ventricular tachycardia requiring 2 electric shocks. return of spontaneous circulation (rosc) was achieved. upon check-in to intensive care unit (icu), the patient pulseless electrical activity (pea) arrested and hemoglobin (hgb) dropped from 10 to 6.6. patient was given 2 rounds of cardiopulmonary resuscitation (cpr) and eliquis. impella was explanted. | [5 empty] | The investigation for the reported ventricular tachycardia has been completed. the device was not returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined. this report is being filed as part of a retrospective review of historical records.
30134 | 1220648-2024-07545 | 18845922 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18845922&pc=OZD | 12/08/2022 | 03/12/2025 |  | 03/06/2024 |  |  | 12/08/2022 | 03/06/2024 |  |  | 06/30/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023248752 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/08/2022 |  |  |  |  | The user facility reported a 67-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that post percutaneous coronary intervention (pci), there was a small hematoma with oozing. the decision was made to explant the impella. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
30135 | 1220648-2024-07544 | 18845966 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18845966&pc=OZD | 12/12/2022 | 03/06/2024 |  | 03/06/2024 |  |  | 12/12/2022 | 03/06/2024 |  |  | 08/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023271671 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Heart Block | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/12/2022 |  |  |  |  | The user facility reported an 87-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. the patient experienced a brief complete heart block immediately following the pump insertion. the pacer initiated 60 heart beats per minute (bpm) with quick recovery. the patient reportedly had an excellent result with complete re-vascularization. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not returned for evaluation. as clinical information was not provided and the device was not returned, the root cause of the reported issue was not determined. this report is being filed as part of a retrospective review of historical records.
30136 | 1220648-2024-07546 | 18845970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18845970&pc=OZD | 09/02/2022 | 03/06/2024 |  | 03/06/2024 |  |  | 09/02/2022 | 03/06/2024 |  |  | 04/30/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023214597 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/02/2022 |  |  |  |  | The user facility reported a 58-year-old male in acute myocardial infarction/cardiogenic shock was to be implanted with an impella 5.5 device for mechanical circulatory support. the medical team was having difficulty in delivering the impella pump due to the patient's tortuous vasculature. the procedure was aborted. the patient will remain on venoarterial extracorporeal membrane oxygenation (va-ecmo). | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition of diseased vessels. this report is being filed as part of a retrospective review of historical records.
30137 | 1220648-2024-07547 | 18846017 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18846017&pc=OZD | 12/12/2022 | 09/13/2024 |  | 03/06/2024 |  |  | 12/12/2022 | 03/06/2024 |  |  | 09/30/2023 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022135762 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Insufficient Flow or Under Infusion | Patient-Device Incompatibility | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/12/2022 |  |  |  |  | The user facility reported a 67-year-old male in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after reviewing the low pulsatility waveforms, the physician became concerned that due to the patient's size, they may not be fully supported with the impella cp. the impella cp was removed and the patient was subsequently escalated to an impella 5.5. | It was noted after the patient was escalated to the impella 5.5 that the patient was made a do not resuscitate and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). the user facility reported a 67-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after reviewing the low pulsatility waveforms, the physician became concerned that due to the patient's size, they may not be fully supported with the impella cp. the impella cp was removed and the patient was subsequently escalated to an impella 5.5. (cp was alleged to not adequately serving patient needs so escalated to 5.5). |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type.
30138 | 1220648-2024-07548 | 18846037 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18846037&pc=OZD | 12/09/2022 | 03/06/2024 |  | 03/06/2024 |  |  | 12/09/2022 | 03/06/2024 |  |  | 06/30/2023 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022096129 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 17 YR |  | Female |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/09/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 17-year-old female patient for mechanical circulatory support. it was reported that the medical team experienced difficulty implanting impella due to vessels spasms and clamping down. ultimately the medical team aborted the procedure. there was no patient harm reported. | [5 empty] | The investigation into the reported delivery issue has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine a root cause. the most likely cause of the delivery issue was due to patient condition. this report is being filed as part of a retrospective review of historical records.
30139 | 1220648-2024-07549 | 18846038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18846038&pc=OZD | 05/16/2022 | 03/12/2025 |  | 03/06/2024 |  |  | 05/16/2022 | 03/06/2024 |  |  | 01/01/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2023192398 |  | 0048-0003 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Activation, Positioning or Separation Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/16/2022 |  |  |  |  | The user facility reported a 63-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the complainant reported a patient developing limb ischemia in the impella inserted extremity. as a result, a fem-fem bypass was completed. additionally, the pump was found to be malpositioned under echocardiogram imagery. the pump was removed, rewired, and re-inserted, however, optimal positioning could not be achieved. the pump was removed and replaced. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30140 | 1220648-2024-07550 | 18846054 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18846054&pc=OZD | 02/18/2024 | 04/08/2024 |  | 03/06/2024 |  |  | 02/18/2024 | 03/06/2024 |  |  |  |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024433465 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/18/2024 |  |  |  |  | The user facility reported a 70 -year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. as soon as the impella cp started there was a quick notice of flows diminishing, suction alarm, and flat motor current. the impella was replaced due to suspected clot ingestion. | [5 empty] | The investigation into the low pump flow issues has been completed since the original report was filed. the pump was returned, tested, and evaluated. visual inspection found no issues or damages on the pump. no biomaterial was found in pump housing & cannula. clot was likely flushed or removed before complaint pump was returned. the failure mode could not be reproduced. based on clinical description, data review & product analysis, the root cause of low flow was most likely due to biomaterial ingestion. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30141 | 1220648-2024-07551 | 18846061 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18846061&pc=OZD | 12/12/2022 | 03/12/2025 |  | 03/06/2024 |  |  | 12/12/2022 | 03/06/2024 |  |  | 08/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023265747 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/12/2022 |  |  |  |  | The user facility reported an 81-year-old male in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. while on impella support, the patient experienced gastrointestinal (gi) bleeding and the decision was made to remove the impella device due to a worsening gastrointestinal (gi) bleed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30142 | 1220648-2024-07552 | 18846066 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18846066&pc=OZD | 12/12/2022 | 03/12/2025 |  | 03/06/2024 |  |  | 12/13/2022 | 03/06/2024 |  |  | 08/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023270367 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/13/2022 |  |  |  |  | The user facility reported a 79-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced some bleeding from the groin insertion site, requiring 1 unit of packed red blood cells (prbcs). the staff switched the patient to bicarbonate and withheld the systemic heparin drip. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. the root cause of the access site bleeding cannot be determined since the product was not returned and enough clinical info was not provided.
30143 | 1220648-2024-07553 | 18846078 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18846078&pc=OZD | 12/04/2022 | 03/06/2024 |  | 03/06/2024 |  |  | 12/04/2022 | 03/06/2024 |  |  | 02/29/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023199899 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/04/2022 |  |  |  |  | The user facility reported a 47-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. while the patient was on impella cp support, the patient was noted to not be moving his right side. the patient was taken for ct head scan and results revealed a left anterior myocardial infarction and petechial hemorrhage. heparin was discontinued due to petechial hemorrhage to minimize risk for hemorrhagic conversion. | [5 empty] | The investigation for the reported stroke/cva has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. as the necessary clinical information was not provided, the root cause of the stroke/cva was not determined. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
30144 | 1220648-2024-07554 | 18846084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18846084&pc=OZD | 12/09/2022 | 03/12/2025 |  | 03/06/2024 |  |  | 12/09/2022 | 03/06/2024 |  |  | 07/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2023246587 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 88 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/09/2022 |  |  |  |  | The user facility reported an 88-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after implantation, the patient experienced access site bleeding and a hematoma. the patient required a blood transfusion. the device was explanted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. the root cause of the access site bleeding cannot be determined since the product was not returned and enough clinical info was not provided.
30145 | 1220648-2024-07555 | 18846101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18846101&pc=OZD | 11/09/2022 | 03/06/2024 |  | 03/06/2024 |  |  | 11/09/2022 | 03/06/2024 |  |  | 07/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023249695 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/09/2022 |  |  |  |  | The user facility reported a 29-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an impella 5.5 pump was unable to advance through the patient's vasculature due to the vessel size and anatomy. after multiple unsuccessful attempts, the implant was aborted and an impella cp pump was placed through the axillary graft for patient support. there was no patient harm. | [5 empty] | The investigation into the reported delivery issue has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue - anatomy was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
30146 | 1220648-2024-07556 | 18846149 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18846149&pc=OZD | 12/13/2022 | 03/12/2025 |  | 03/06/2024 |  |  | 12/13/2022 | 03/06/2024 |  |  | 08/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023267664 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/13/2022 |  |  |  |  | The user facility reported a patient that a 28-year-old female with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the impella had to be replaced with another impella due to continuous suction even with the extracorporeal membrane oxygenation flows adjusted. when ramp up was attempted with the replacement impella, the same situation occurred. the echocardiogram showed a dilated left ventricle that would collapse within seconds of the impella being turned on. the extra-corporeal membrane oxygenation flows adjusted with zero influence on impella. the replacement impella was explanted and the 14fr sheath was replaced with the 16fr cook sheath, along with an antegrade reperfusion sheath. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30147 | 1220648-2024-07557 | 18846165 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18846165&pc=OZD | 02/21/2024 | 03/28/2024 |  | 03/06/2024 |  |  | 02/21/2024 | 03/06/2024 |  |  | 11/30/2025 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024435065 |  | 100100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/21/2024 |  |  |  |  | The user facility reported a 48-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella 5.5 support due to cardiogenic shock. while on support, the patient had hematoma. washout was performed to resolve the hematoma. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ . ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ . ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿. | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
30148 | 1220648-2024-07558 | 18846242 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18846242&pc=OZD | 02/17/2024 | 05/08/2024 |  | 03/06/2024 |  |  | 02/17/2024 | 03/06/2024 |  |  | 11/30/2025 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024422035 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Optical Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/17/2024 |  |  |  |  | The user facility reported a 58-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the patient was on impella 5.5 support ongoing for cardiac and systemic comorbidities and was awaiting a transplant. after 39 days, the pump was explanted with a transplant. during the support the optical signal had been lost. no harm was reported. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the optical signal issue has been completed since the original report was submitted. the device was not returned for investigation. upon review of the data logs, it is apparent that the console is the potential cause of the issue. investigation into the console is done (b)(4). in accordance with updated procedures, section d6 has been revised from the initial submission.
30149 | 1220648-2024-07560 | 18846246 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18846246&pc=OZD | 02/20/2024 | 04/09/2024 |  | 03/06/2024 |  |  | 02/20/2024 | 03/06/2024 |  |  | 11/30/2025 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024433210 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/20/2024 |  |  |  |  | The user facility reported a 54-year-old female/male noted as high risk percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. the implanted impella cp pump experienced low pump flows coinciding with no lv signal during patient support. repositioning was attempted, but the flow remained low and the motor current remained flat. due to the noted failure, the decision was made to replace the pump. there was no patient harm. | [5 empty] | The investigation into the low pump flow issue has been completed since the original report was filed. the pump was returned, tested, and evaluated. visual inspection found no issues on the pump. the failure mode could not be reproduced as pump ran without issue under bench test. the root cause of the low flow was most likely a patient condition as no issue was found on the pump. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30150 | 1220648-2024-07559 | 18846248 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18846248&pc=OZD | 12/13/2022 | 03/12/2025 |  | 03/06/2024 |  |  | 12/13/2022 | 03/06/2024 |  |  | 04/30/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023220583 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/13/2022 |  |  |  |  | The user facility reported a 76-year-old male in post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding at the access site. sutures were placed and repositioning occurred, and the bleeding was resolved. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. the root cause of the access site issue was not determined since product was not returned.
30151 | 1220648-2024-07561 | 18846279 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18846279&pc=OZD | 12/04/2022 | 03/11/2025 |  | 03/06/2024 |  |  | 12/04/2022 | 03/06/2024 |  |  | 01/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023183587 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/04/2022 |  |  |  |  | The user facility reported a 76-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the staff observed a thrombus on the end of the pigtail. a sentinel device was placed to entrain clot in the catheter. upon removal, no clot was found in the sentinel device or the impella device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
30152 | 1220648-2024-07562 | 18846304 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18846304&pc=OZD | 12/06/2022 | 04/10/2024 |  | 03/06/2024 |  |  | 02/22/2024 | 03/06/2024 |  |  | 06/30/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023247977 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Hospitalization | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/06/2022 |  |  |  |  | The study team of protect iv reported a 75-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. the patient had a cp for support in (b)(6) 2022. the patient was enrolled in the protect iv. the study team forwarded that there was an sepsis/osteomyelitis with a "possible relationship" to the use of impella. the severity on the study documentation was listed as severe. this prompted the patient to be re-admitted after hospital discharge. patient was treated with medication and is stable. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ | The investigation for the reported sepsis/infection has been completed. the device was not returned for evaluation. not enough clinical information was provided to determine the root cause of the reported issue. therefore, the root cause of the sepsis could not be determined. the instructions for use states ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ added-h6 impact code 4644 d4-updated model number.
30153 | 1220648-2024-07563 | 18846351 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18846351&pc=OZD | 12/14/2022 | 03/11/2025 |  | 03/06/2024 |  |  | 12/14/2022 | 03/06/2024 |  |  | 09/30/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023273948 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/14/2022 |  |  |  |  | The user facility reported a 71-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced two episodes of ventricular tachycardia/ventricular fibrillation it was necessary to shock the patient twice. the patient maintained perfusion pressure with both episodes and never lost consciousness. the decision was made to leave the impella in place overnight to allow heart to rest. impella sheath was removed. during the explant of the impella cp, impella was weaned and removed without issue and hemodynamics remained stable immediately after removal. the perclose sutures (two) that were placed with the initial insertion had failed. post closure was attempted with dual wire technique/smaller sheath. a total of four more perclose sutures were deployed without hemostasis. left femoral groin site access was established for balloon tamponade of opposite iliac artery. when balloon was inflated, patient continued to have bleeding at the insertion site. initial stick was found to be very low at the level of the bifurcation. bleeding subsided with 10 french sheath in right groin site. vascular surgery was consulted and decided to take patient to operating room for surgical closure of site. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
30154 | 1220648-2024-07565 | 18846358 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18846358&pc=OZD | 12/09/2022 | 03/06/2024 |  | 03/06/2024 |  |  | 12/09/2022 | 03/06/2024 |  |  | 09/30/2023 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022136572 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/09/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 67-year-old male patient for mechanical circulatory support. it was reported that there was oozing at the insertion site. the patient was redressed, no new oozing was noted. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
30155 | 1220648-2024-07564 | 18846369 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18846369&pc=OZD | 02/24/2024 | 04/12/2024 |  | 03/06/2024 |  |  | 02/24/2024 | 03/06/2024 |  |  | 11/30/2025 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025438252 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/24/2024 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had cp support ongoing for a patient along with ecmo (extracorporeal membrane oxygenation) support. the patient experienced limb ischemia which prompted the team to place bilateral reperfusion sheaths. the pump remains on for support and has not yet been explanted. pump running for 238 hours to date. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
30156 | 1220648-2024-07566 | 18846427 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18846427&pc=OZD | 02/25/2024 | 04/15/2024 |  | 03/06/2024 |  |  | 02/25/2024 | 03/06/2024 |  |  | 07/31/2025 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024389534 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/25/2024 |  |  |  |  | The user facility reported a 76-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient experienced significant bleeding at the impella access site. the md put forward tension sutures on impella and albumin given for volume status and site bleeding was stabilized. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was use related, forward tension sutures on impella stopped the bleeding as per the clinical description provided.
30157 | 1220648-2024-07567 | 18846499 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18846499&pc=OZD | 02/26/2024 | 04/16/2024 |  | 03/06/2024 |  |  | 02/26/2024 | 03/06/2024 |  |  |  |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2025441936 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/26/2024 |  |  |  |  | The user facility reported a 33-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an implanted impella 5.5 pump came back across the aortic valve during patient support. after multiple unsuccessful attempts to re-advance the pump, the decision was made to explant the pump. an iabp (intra-aortic balloon pump) was subsequently placed for ongoing support. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ | The investigation into the positioning issues has been completed since the original report was submitted. the device was not returned for investigation. the cause of the positioning issue most likely was use related issue due to improper technique.
30158 | 1220648-2024-07568 | 18846511 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18846511&pc=OZD | 11/10/2022 | 03/11/2025 |  | 03/06/2024 |  |  | 11/10/2022 | 03/06/2024 |  |  | 07/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023258905 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/10/2022 |  |  |  |  | The user facility reported a 76-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported that after the impella cp was successfully implanted, the patient developed a small hematoma post procedure at the groin impella insertion site. manual pressure was held to compress the area, but after an additional assessment by the physician, the groin site bleeding had increased and the decision was made to take the patient to operating room for surgical cutdown and repair of the insertion site. the impella cp position was confirmed and repositioned using transesophageal echocardiography (tee) guidance. patient was transfused with blood products. post surgical repair, some oozing was noted at the access site. access site was re-dressed and the oozing subsided. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
30159 | 1220648-2024-07569 | 18846518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18846518&pc=OZD | 12/10/2022 | 05/23/2025 |  | 03/06/2024 |  |  | 12/10/2022 | 03/06/2024 |  |  | 08/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023267041 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/10/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 81-year-old male patient for mechanical circulatory support. it was reported while on impella cp support, the patient was experiencing a hematoma at the access site. after the administration of heparin the hematoma subsequently improved after several days. the impella cp pump was explanted and replaced with an impella 5.5. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The root cause of the access site bleeding cannot be determined since the product was not returned and enough clinical info was not provided. this report is being filed as part of a retrospective review of historical records.
30160 | 1220648-2024-07570 | 18846564 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18846564&pc=OZD | 02/26/2024 | 04/11/2024 |  | 03/06/2024 |  |  | 02/26/2024 | 03/06/2024 |  |  | 11/30/2025 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024434897 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/26/2024 |  |  |  |  | The user facility reported a 53-year-old female of unknown race was implanted with an impella 5.5 as a bridge to transplant. it was reported on the day of implant the sterile sleeve became damaged. to troubleshoot the issue the team placed tegaderm, there was no infection noted. the patient remained on support for the next 3 days and was bridged to transplant. | Additional information provided states that upon removal of the 5.5 for the heart transplant, biomaterial was found on pump inlet. there was no harm reported. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported product damage and biomaterial has been completed. no product was returned. data logs were not returned for evaluation. the cause of the product damage (sterile sleeve) issue was not determined since product was not returned. as the necessary information was not provided, the root cause of the thrombus was not determined. the instructions for use states ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿
30161 | 1220648-2024-07571 | 18846588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18846588&pc=OZD | 11/12/2022 | 03/11/2025 |  | 03/06/2024 |  |  | 11/12/2022 | 03/06/2024 |  |  | 08/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023263735 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/12/2022 |  |  |  |  | The user facility reported a 67-year-old female in st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. after being placed on impella cp support, it was reported that the distal pulses were unable to be palpated and there was continuous bleeding from the groin, without any evidence of hematoma. the dressing was changed at the bedside and the physician made the decision to explant the impella cp device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
30162 | 1220648-2024-07573 | 18846628 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18846628&pc=OZD | 12/14/2022 | 03/06/2024 |  | 03/06/2024 |  |  | 12/14/2022 | 03/06/2024 |  |  | 09/30/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023273968 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/14/2022 |  |  |  |  | The user facility reported a 62-year-old female in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patients limb pulses were absent but capillary refill was slow. nitro was started to lower blood pressure. the impella was weaned and explanted. | [5 empty] | The investigation for limb ischemia has been completed. the impella device was not returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
30163 | 1220648-2024-07574 | 18846631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18846631&pc=OZD | 12/06/2022 | 03/06/2024 |  | 03/06/2024 |  |  | 12/06/2022 | 03/06/2024 |  |  | 08/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023270614 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/06/2022 |  |  |  |  | The user facility reported a 68-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient experienced access site bleeding and was treated and secured with placement of purse string suture, which resolved the issue. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
30164 | 1220648-2024-07575 | 18846665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18846665&pc=OZD | 12/10/2022 | 03/06/2024 |  | 03/06/2024 |  |  | 12/10/2022 | 03/06/2024 |  |  | 07/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023252243 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | Difficult to Advance | [5 empty] | Arrhythmia | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/10/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 74-year-old female patient for mechanical circulatory support. it was reported after insertion, there were issues with the placement of the impella. when advancing, the inflow kept directing toward the mitral valve. the color doppler showed that it had caused mitral regurgitation (mr) that had not been there previously. with repositioning attempts, the pigtail was pulled all the way out of the left ventricle and the pump was unable to readvance across the valve. impella was pulled into the descending aorta and left on p1 for case. the decision was made to remove the impella cp, but due to the patient¿s low platelets and dropping blood pressure, the removal was postponed. heparin was resumed (stopped previously due to the potential removal). patient is also experienced atrial fibrillation (af). | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
30165 | 1220648-2024-07576 | 18846676 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18846676&pc=OZD | 02/26/2024 | 04/03/2024 |  | 03/06/2024 |  |  | 02/27/2024 | 03/06/2024 |  |  | 12/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024304711 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/27/2024 |  |  |  |  | The user facility reported a 62-year-old female was implanted with an impella 5.5 for mechanical circulatory support pre-operatively for coronary artery bypass grafting. it was reported that left ventricle thrombus was found on trans-thoracic echo (tte). the patient was reported as stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | Abiomed, inc. submitted manufacturer report number 1220648-2024-07576 on 07mar2024. after the initial report was submitted, additional information was received on 11mar2024 stating the user facility did not find a left ventricle thrombus. this supplemental report is being filed to retract the initial report.
30166 | 1220648-2024-07577 | 18846743 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18846743&pc=PYX | 02/27/2024 | 04/15/2024 |  | 03/06/2024 |  |  | 02/27/2024 | 03/06/2024 |  |  | 11/30/2025 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC |  |  | US |  | 2024434372 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Remove | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/27/2024 |  |  |  |  | The user facility reported a 36-year-old male with extensive medical comorbidities was implanted with an impella rp for mechanical circulatory support. it was reported that after just 27 hours of the support the rp was explanted due to concerns of hemolysis. at the time of explant, the rp and peripherally inserted central catheter picc line interacted. there was no lasting harm but, the picc was additionally removed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported hemolysis and removal issues has been completed. the pump was returned for evaluation. pump was inspected and biomaterial was found on impeller blade. clinical details noted that pump was in good position and patient had adequate volume when urine was observed to be dark red indicating hematuria. pump was being run at p-6 with no alarms throughout the night, however lab showed blood still be hemolyzing. data logs were reviewed and a motor current spike occurred with a drop in minimum p-level flow and drop in motor speed at the same time. upon explant of the pump, patient's peripherally inserted central catheter was pulled out of right internal jugular vein with the pump. the picc line was removed fully after pump was explanted. the root cause of the hemolysis was due to biomaterial on the impeller blade. the root cause of the removal issues was most likely due to interaction with other devices as the picc line was partially pulled out with pump during explant. the instructions for use states ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected low s or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿
30167 | 1220648-2024-07578 | 18846809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18846809&pc=OZD | 03/01/2024 | 05/01/2024 |  | 03/06/2024 |  |  | 03/01/2024 | 03/06/2024 |  |  |  |  | 03/06/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1692374 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 03/01/2024 |  |  |  |  | The user facility reported a 56-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a console replaced due to purge alarms. the team had attempted to troubleshoot the purge alarms by purge bag exchange. the troubleshooting did not resolve and so the aic was swapped. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported console (aic) purge issue has been completed. the console was returned for evaluation. upon visual inspection, the purge unit area did not reveal any abnormalities. a known good pump and purge line was attached to the console and run. the purge system open alarm was manually induced and the change purge fluid bag procedure initiated. the console was able to complete the procedure with no observable abnormality. data logs were reviewed which confirmed that the purge system open alarm was issued. real time logs from the console confirm that the average purge pressure was < 100 mmhg for 20 seconds or more, leading to the issuance of the aforementioned alarm. the data logs confirm that the user then proceeded to initiate change of purge fluid bag as recommended to clear the alarm. although the procedure is initiated, it is never completed by the user before the pump is unplugged and console shutdown. the root cause for the purge not completing as reported during purge fluid change procedure could not be determined because the event could not be reproduced as reported. d4-updated model number. to conform with updated procedures, section d6 has been revised with updated information.
30168 | 1220648-2024-07579 | 18846873 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18846873&pc=OZD | 02/20/2024 | 04/03/2024 |  | 03/06/2024 |  |  | 02/20/2024 | 03/06/2024 |  |  | 07/31/2025 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025440510 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/20/2024 |  |  |  |  | The user facility reported a 73-year-old male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the pump was pulled by the patient on day three of support. the pump was then turned off while in the aorta, and upon re-advancement to the left ventricle the pump failed to start back. a new pump was placed. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the positioning issues and the pump stop issue has been completed since the initial report was submitted. the device was not returned for investigation. per the clinical data provided, the root cause of the positioning issues was most likely patient movement related as the patient was combative and pulled on the pump out of position. the root cause of the pump stop issue was most likely an open electrical line due to excessive force when the patient pulled the pump.
30169 | 1220648-2024-07580 | 18846906 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18846906&pc=OZD | 12/14/2022 | 03/06/2024 |  | 03/06/2024 |  |  | 12/14/2022 | 03/06/2024 |  |  | 08/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023267009 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/14/2022 |  |  |  |  | The user facility reported a 66-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the patient developed ventricular tachycardia and had to be shocked, but tolerated well. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the reported issue was not determined. this report is being filed as part of a retrospective review of historical records.
30170 | 1220648-2024-07581 | 18846993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18846993&pc=OZD | 12/06/2022 | 05/19/2025 |  | 03/06/2024 |  |  | 12/06/2022 | 03/06/2024 |  |  | 11/30/2023 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023159818 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | Required Intervention | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/06/2022 |  |  |  |  | The user facility reported a 73-year-old male was implanted with an impella cp device for mechanical circulatory support. the patient developed no peripheral pulse in the right leg and later expired while on the impella cp support. no allegations were made towards the impella cp for cause of death. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation for the reported limb ischemia has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07581 was provided. | Limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition. any concomitant medication. vascular size or variations thereof detailed description of the symptoms experienced by the patient. physical examination findings, including pulse assessment and visual examination of the affected limb. diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. laboratory test results, including blood tests for markers of inflammation or clotting disorders. any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
30171 | 1220648-2024-07593 | 18847584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18847584&pc=OZD | 02/20/2024 | 04/09/2024 |  | 03/06/2024 |  |  | 02/20/2024 | 03/06/2024 |  |  |  |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025438984 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/20/2024 |  |  |  |  | The user facility reported a 56-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient developed a hematoma in their right groin during impella cp support. due to the noted condition, the decision was made to explant the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. the instructions for use states ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
30172 | 1220648-2024-07594 | 18847606 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18847606&pc=OZD | 12/12/2022 | 03/06/2024 |  | 03/06/2024 |  |  | 12/12/2022 | 03/06/2024 |  |  | 07/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  |  |  | 2023255090 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Failure to Deliver | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/12/2022 |  |  |  |  | The user facility reported a 69-year-old male noted as high risk percutaneous cardiac insertion was unsuccessfully implanted with an impella cp device for mechanical circulatory support. it was reported that the medical team attempting the impella implant had difficulty with delivery due to the patient's tortuous and calcified anatomy of left femoral artery. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the root cause of the delivery issue was most likely patient condition related. this report is being filed as part of a retrospective review of historical records.
30173 | 1220648-2024-07595 | 18847641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18847641&pc=OZD | 11/12/2022 | 03/11/2025 |  | 03/06/2024 |  |  | 11/12/2022 | 03/06/2024 |  |  | 04/30/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023220611 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/12/2022 |  |  |  |  | The user facility reported a 69-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported after the impella cp was successfully implanted, the patient experienced bleeding from insertion site that leaked into the sterile sleeve. the patient was redressed, and position of the sheath was adjusted, heparin was withheld, and no new oozing was noted. . | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
30174 | 1220648-2024-07596 | 18847701 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18847701&pc=OZD | 09/05/2022 | 03/06/2024 |  | 03/06/2024 |  |  | 09/05/2022 | 03/06/2024 |  |  | 06/30/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023233425 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/05/2022 |  |  |  |  | The user facility reported a 69-year-old female in acute myocardial infarction/cardiogenic shock was to be implanted with an impella 5.5 device for mechanical circulatory support. the medical team was having difficulty delivering the impella pump due to the patient's anatomy. after several attempts, they decided to abort the impella implant and deliver venoarterial extracorporeal membrane oxygenation (va ecmo) instead. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
30175 | 1220648-2024-07597 | 18847747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18847747&pc=OZD | 11/14/2022 | 03/11/2025 |  | 03/06/2024 |  |  | 12/14/2022 | 03/06/2024 |  |  | 01/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191712 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/14/2022 |  |  |  |  | The user facility reported an 80-year-old patient noted as high risk percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support and. the patient experienced significant bleeding. a balloon tamponade was attempted with successful access to radiofrequency ablation. angiograms revealed bleeding at the 8 fr sheath site. the sheaths were upsized and the impella sheath was placed. hemostasis was noted on fluoroscopy. the impella cp was weaned down and removed. surgical closure and sheath removal were performed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. b5 date of event was corrected.
30176 | 1220648-2024-07598 | 18847822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18847822&pc=OZD | 12/14/2022 | 03/11/2025 |  | 03/06/2024 |  |  | 12/14/2022 | 03/06/2024 |  |  | 08/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023272887 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/14/2022 |  |  |  |  | The user facility reported a 84-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was on impella support for 127.87 hours before the patient was successfully weaned and the device was explanted. numerous clots were removed after explant and a stent to the femoral artery was placed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ assess access site for bleeding and hematoma. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
30177 | 1220648-2024-07599 | 18847836 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18847836&pc=OZD | 12/07/2022 | 03/06/2024 |  | 03/06/2024 |  |  | 12/07/2022 | 03/06/2024 |  |  | 06/30/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023248101 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/07/2022 |  |  |  |  | The user facility reported a 73-year-old male was attempted to be implanted with a impella cp device. prior to the securement of the impella cp, the patient moved their leg aggressively, which caused the sheath to be pulled out about four inches and the impella came back across the aortic valve. the impella was unable to be repositioned from pap prior to weaning from the bypass. during the attempt, the impella bounced back across aortic valve (ao). the decision was made to remove the impella and surgically repair the right femoral artery (rfa). it was noted that the high stick and angle of entry to be greater than 45 degrees. patient required 3 rounds of epinephrine and 3 rounds of levophed upon arrival to cardiovascular intensive care unit (cvicu). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
30178 | 1220648-2024-07600 | 18847841 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18847841&pc=OZD | 02/23/2024 | 09/16/2024 |  | 03/06/2024 |  |  | 02/23/2024 | 03/06/2024 |  |  | 11/30/2025 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024430945 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Air Embolism | Tachycardia | [11 empty] | Required Intervention | Death | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 02/23/2024 |  |  |  |  | The user facility reported a 66-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 pump was inserted via the axillary. the patient was intubated on support with impella 5.5 and extracorporeal membrane oxygenation (ecmo). the team had observed what looked to be an air emboli in the left ventricle and so the team troubleshot by replacing the purge. the thought was maybe there was a ventricular septal defect (vsd). additionally, there was a run of ventricular tachycardia (vt) that prompted treatment with repositioning and lidocaine. the pump remained on for support until care was withdrawn after eight days. | Additional information noted care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the embolism and the ventricular tachycardia (vt) issues has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the embolism was not determined since product was not returned. as the necessary information was not provided, the root cause of the vt/vf was not determined. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr was provided in sections b2, b5, h1, and h6. | The following have been reported incorrectly or missing from manufacturer device report 1220648-2024-07600: d6a and d6b revised from the initial submission in accordance with updated procedures. g1 reporting contact fax number missing. f6/f8- should have been left blank.
30179 | 1220648-2024-07601 | 18847955 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18847955&pc=OZD | 11/13/2022 | 03/10/2025 |  | 03/06/2024 |  |  | 11/13/2022 | 03/06/2024 |  |  | 06/30/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023238046 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/13/2022 |  |  |  |  | The user facility reported a 64-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported a large hematoma was on the insertion site of the pump and the fem stopped which was applied with minimal pressure. the patient was bleeding from all orifices and sites. hematoma was resolved and the groin was stable. systemic heparin was withheld. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleeding issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
30180 | 1220648-2024-07602 | 18847971 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18847971&pc=OZD | 12/14/2022 | 03/11/2025 |  | 03/06/2024 |  |  | 12/14/2022 | 03/06/2024 |  |  | 08/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023268374 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/14/2022 |  |  |  |  | The user facility reported a 61-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced access site bleeding. blood products were administered to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
30181 | 1220648-2024-07603 | 18847992 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18847992&pc=OZD | 12/10/2022 | 03/11/2025 |  | 03/06/2024 |  |  | 12/10/2022 | 03/06/2024 |  |  | 08/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023272398 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Insert | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/10/2022 |  |  |  |  | The user facility reported a 43-year-old male being placed on impella cp required mechanical circulatory support. it was reported the patient was on impella cp support and experienced delivery issues, each attempt resulted in the patient developing ectopy and ventricular fibrillation when trying to implant the device. the patient also experienced bleeding at the insertion site. ectopy eventually resolved and the impella cp was advanced over wire and started to function properly. the patient was on impella cp support for 4.35 hours before patient was successfully weaned off the device. the event required removal of the pump as the intervention. no patient harm was reported. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
30182 | 1220648-2024-07604 | 18848044 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18848044&pc=OZD | 11/17/2022 | 05/23/2025 |  | 03/06/2024 |  |  | 11/17/2022 | 03/06/2024 |  |  | 07/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023256466 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/17/2022 |  |  |  |  | The user facility reported a 73-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported after the impella cp was successfully implanted, the patient experienced bleeding from the right femoral access site. the sheath was placed at optimal angle match to decrease bleeding and surgicel was applied. | Additional information noted the patient expired while on support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07604 was provided in sections b2, b5, h1, and h6. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30183 | 1220648-2024-07605 | 18848118 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18848118&pc=OZD | 12/07/2022 | 03/06/2024 |  | 03/06/2024 |  |  | 12/07/2022 | 03/06/2024 |  |  | 03/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023225897 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/07/2022 |  |  |  |  | The user facility reported a 69-year-old female was implanted with an impella cp device for mechanical circulatory support. during support, the patient¿s bilateral distal pulses were absent. a round of cardiopulmonary resuscitation was required and epinephrine bolus was given before obtaining return of spontaneous circulation. the patient remained on impella cp support. | [5 empty] | The investigation for the reported limb ischemia has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
30184 | 1220648-2024-07606 | 18848217 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18848217&pc=OZD | 09/06/2022 | 03/06/2024 |  | 03/06/2024 |  |  | 09/06/2022 | 03/06/2024 |  |  | 01/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023188771 |  | P140003 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/06/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 69-year-old female patient for mechanical circulatory support. it was reported upon placement of impella, the patient went into ventricular tachycardia (vt) and was shocked five (5) times. patient was stabilized and was sent to unit for continuation of care. | [5 empty] | The investigation for the reported ventricular tachycardia has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
30185 | 1220648-2024-07607 | 18848223 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18848223&pc=OZD | 02/25/2024 | 04/16/2024 |  | 03/06/2024 |  |  | 02/25/2024 | 03/06/2024 |  |  | 11/30/2025 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025438234 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/25/2024 |  |  |  |  | The user facility reported a 78-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed ischemia during impella cp support. it was noted that there was a decrease in palpable pulse in the lower right extremity coinciding with pain in the right leg. due to the noted condition, the decision was made to explant the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
30186 | 1220648-2024-07608 | 18848268 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18848268&pc=OZD | 09/07/2022 | 03/06/2024 |  | 03/06/2024 |  |  | 09/07/2022 | 03/06/2024 |  |  | 04/30/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023217233 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/07/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 74-year-old male patient for mechanical circulatory support. it was reported while being placed on impella cp support, the patient developed an ischemic leg on the impella side. the impella cp pump was explanted and replaced with an impella 5.5. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. this report is being filed as part of a retrospective review of historical records.
30187 | 1220648-2024-07609 | 18848409 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18848409&pc=OZD | 11/18/2022 | 03/11/2025 |  | 03/06/2024 |  |  | 11/18/2022 | 03/06/2024 |  |  | 02/29/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023199132 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/18/2022 |  |  |  |  | The user facility reported a 50-year-old female was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient developed a clot from the 14f x 13 cm abiomed sheath in the leg and the patient underwent surgical closure. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
30188 | 1220648-2024-07610 | 18848454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18848454&pc=OZD | 02/26/2024 | 04/16/2024 |  | 03/06/2024 |  |  | 02/26/2024 | 03/06/2024 |  |  |  |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024436376 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Activation, Positioning or Separation Problem | Difficult to Advance |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/26/2024 |  |  |  |  | The user facility reported a 55-year-old female undergoing cardiac surgery was to be implanted with an impella 5.5 device for mechanical circulatory support. it was reported that complications were encountered during attempted delivery of an impella 5.5 pump. it was documented that multiple attempts were made to advance the pump, but the outflow cage was unable to traverse through the artery. the decision was subsequently made to abort the implant due to the patient's anatomy. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition because of lesser vessel diameter as noted to be 6mm which is less than minimum vessel diameter of 7mm needed for impella 5.5 insertion. h6 impact code added 4627.
30189 | 1220648-2024-07611 | 18848511 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18848511&pc=OZD | 12/10/2022 | 03/11/2025 |  | 03/06/2024 |  |  | 12/10/2022 | 03/06/2024 |  |  | 07/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023259045 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/10/2022 |  |  |  |  | The user facility reported a 40-year-old male being placed on impella cp required mechanical circulatory support. it was reported that the patient quickly decompensated just after the impella cp crossed the valve. five (5) rounds of epinephrine were used to support the patient's blood pressure. the patient was taken to the or on impella cp for ecmo support. the physician left the impella cp in place as a ventricular vent. the impella cp was weaned and explanted from the patient, with plans to convert to vv ecmo. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
30190 | 1220648-2024-07612 | 18848532 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18848532&pc=OZD | 02/26/2024 | 04/17/2024 |  | 03/06/2024 |  |  | 02/26/2024 | 03/06/2024 |  |  | 11/30/2025 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025438238 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/26/2024 |  |  |  |  | The user facility reported a 78-year-old female in st-segment elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced ventricular tachycardia during impella cp support. the pump was repositioned in a response to the condition and no further issues were encountered. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the vf/vt/af/at (ventricular tachycardia) has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined.
30191 | 1220648-2024-07613 | 18848760 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18848760&pc=OZD | 09/13/2022 | 03/06/2024 |  | 03/06/2024 |  |  | 09/13/2022 | 03/06/2024 |  |  | 02/29/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023193409 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/13/2022 |  |  |  |  | The user facility reported a 62-year-old male was implanted with an impella cp device for mechanical circulatory support. while explanting the impella cp device, a clot was noted and was unable to be aspirated. best practices were reviewed regarding removal of peel away sheath at time of the procedure and/or hooking the sidearm to pressure to keep lumen patent. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ this report is being filed as part of a retrospective review of historical records.
30192 | 1220648-2024-07614 | 18849139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18849139&pc=OZD | 02/26/2024 | 04/16/2024 |  | 03/06/2024 |  |  | 02/26/2024 | 03/06/2024 |  |  | 03/31/2025 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024319970 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/26/2024 |  |  |  |  | The user facility reported a 72-year-old female noted as high-risk percutaneous coronary intervention was to be implanted with an impella 5.5 device for mechanical circulatory support. it was reported that when inserting impella 5.5 the surgeon noted resistance and buckling of the impella catheter once past the anastomosis. several unsuccessful attempts were made. the decision was made to abort impella 5.5 and place impella cp. the patient was stable and no harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition since surgeon noted resistance and buckling of the impella catheter once past the anastomosis.
30193 | 1220648-2024-07615 | 18849257 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18849257&pc=OZD | 09/07/2022 | 03/10/2025 |  | 03/06/2024 |  |  | 09/07/2022 | 03/06/2024 |  |  | 04/30/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023223888 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/07/2022 |  |  |  |  | The user facility reported a 83-year-old female noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. while the patient was receiving support from the pump, there were difficulties deploying the second preclose and the impella was removed with fluoroscopy guidance. the initial images did not show any abnormalities in the femoral artery; however, once there was bleeding manual pressure was required. iliac artery imaging was repeated and dissection was noted. to resolve the access complication, a balloon and a 8f angioseal was deployed while being held with manual pressure. once 1 perclose was deployed and a dissection was noted on imaging, hemostasis was achieved shortly after. eventually, the impella device was explanted. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30194 | 1220648-2024-07616 | 18849306 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18849306&pc=OZD | 09/14/2022 | 03/06/2024 |  | 03/06/2024 |  |  | 09/14/2022 | 03/06/2024 |  |  | 03/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213178 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/14/2022 |  |  |  |  | The user facility reported a 80-year-old male was implanted with an impella cp device for mechanical circulatory support and developed a hematoma at the impella cp insertion site. the patient was given packed red blood cells, nabicarb, and systemic heparin. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
30195 | 1220648-2024-07617 | 18849373 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18849373&pc=OZD | 02/27/2024 | 04/18/2024 |  | 03/06/2024 |  |  | 02/27/2024 | 03/06/2024 |  |  | 08/31/2025 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024380589 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/27/2024 |  |  |  |  | The user facility reported a 75-year-old female post cardiotomy cardiogenic shock/low cardiac output syndrome was to be implanted with an impella 5.5 for mechanical circulatory support. it was reported that an impella 5.5 was prepped and primed and inserted through the graft via the axillary graft, could not advance wirelessly. the device was removed and re-prepped. wire and catheter advanced through the graft via fluoroscopy. the wire could not advance cross the aortic valve after multiple attempts and multiple wire and catheters utilized. the decision was made to abort and placed impella cp instead via left femoral artery. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition since aortic was innominate due lack of space on the aorta.
30196 | 1220648-2024-07618 | 18849435 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18849435&pc=OZD | 09/07/2022 | 05/28/2025 |  | 03/06/2024 |  |  | 09/07/2022 | 03/06/2024 |  |  | 08/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023207195 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 09/07/2022 |  |  |  |  | The user facility reported a 54-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the physician was concerned with the lack of distal pulses and considered sending the patient to the lab for distal angio as the patient lost pulses at the limb. the patient's feet and extremities were slightly dusky. unfortunately, the patient experienced hemolysis, due to internal bleeding and the results from the electroencephalogram (eeg) indicated the patient had limited/minimal brain function. the family withdrew care and the patient eventually expired. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07618 was provided in sections b2, b5, h1, and h6. | Section d6 : added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30197 | 1220648-2024-07619 | 18849523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18849523&pc=OZD | 02/27/2024 | 04/18/2024 |  | 03/06/2024 |  |  | 02/27/2024 | 03/06/2024 |  |  | 10/31/2025 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024421555 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/27/2024 |  |  |  |  | The user facility reported a 46-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the patient had runs with ventricular tachycardia overnight while on support. the impella was repositioned, and no more ectopy noted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported ventricular fibrillation/ventricular tachycardia (vf/vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf/vt could not be determined. the instructions for use ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
30198 | 1220648-2024-07620 | 18849620 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18849620&pc=OZD | 12/16/2022 | 05/20/2025 |  | 03/06/2024 |  |  | 12/16/2022 | 03/06/2024 |  |  | 08/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023272842 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/16/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 80-year-old male patient for mechanical circulatory support. it was reported that while the patient was under impella cp support, the impella started to have suction alarms, and patient went into a pulseless electrical activity rhythm. advance cardiac life support protocol was initiated, but they were unable to achieve return of spontaneous circulation. no allegations were made towards the impella device for cause of death. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07620 was provided. | Analysis summary: "cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. timing of the arrythmia event in relation to impella support (during or post-implant) electrocardiogram (ekg) recordings of the arrhythmia episodes. evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. response to any antiarrhythmic treatments or interventions during the event. any documented device-related issues or complications during impella support. evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. review of any concomitant medications that may influence cardiac electrophysiology. with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined."
30199 | 1220648-2024-07621 | 18849779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18849779&pc=OZD | 02/29/2024 | 11/20/2025 |  | 03/06/2024 |  |  | 02/29/2024 | 03/06/2024 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1147694 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | Followup |  | 02/29/2024 |  |  |  |  | The complainant reported that an automated impella controller (aic) displayed a controller failure alarm following power up. the power was cycled to troubleshoot the issue, but the alarm remained. the device was swapped as a result of the failure. there was no patient involved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | H.6 type of investigation code 4114 device not returned was report on the original manufacture device report (mdr) inadvertently. the device was returned. therefore, code 4114 should not of been reported. h.10 the additional manufacturer narrative on the original mdr stated that the device was not received from the customer and it was. the additional manufacturer narrative should of read as: : the impella device was received from the customer and an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed.
30200 | 1220648-2024-07622 | 18849820 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18849820&pc=OZD | 10/20/2022 | 03/19/2024 |  | 03/06/2024 |  |  | 10/20/2022 | 03/06/2024 |  |  | 01/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023179603 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  |  | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/20/2022 |  |  |  |  | The user facility reported a 70-year-old male was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that unspecified complications were encountered during impella 5.5 implant due to the patient's anatomy. the implant was initially aborted; however, it was later noted that the doctor was eventually able to implant the device successfully. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition.
30201 | 1220648-2024-07623 | 18849845 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18849845&pc=OZD | 12/16/2022 | 03/11/2025 |  | 03/06/2024 |  |  | 12/16/2022 | 03/06/2024 |  |  | 07/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023252376 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/16/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 63-year-old female patient for mechanical circulatory support. it was reported that during impella support, an alarm of the impella in the aorta occurred. the patient become disoriented and flailed limbs which resulted in the impella cable getting caught around foot and the pump became pulled back to 65 centimeters. it was previously at 100 centimeters. the patient was taken to the cardiac catheterization laboratory to advance the pump back across the valve. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
30202 | 1220648-2024-07624 | 18849907 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18849907&pc=OZD | 03/02/2024 | 04/19/2024 |  | 03/06/2024 |  |  | 03/02/2024 | 03/06/2024 |  |  |  |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025438170 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Improper Flow or Infusion | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 03/02/2024 |  |  |  |  | The user facility reported a 79-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that an implanted impella 5.5 pump ingested a clot and began to experience flow saturation during patient support. due to the noted issue, the decision was made to explant the pump and implant an intra-aortic balloon pump (iabp) for ongoing support. upon removal, a large clot was seen on the inlet and within the outlet of the pump. | [5 empty] | The investigation into the saturated flow and the thrombus issues has been completed since the original report was filed. the pump was returned, tested, and evaluated. there was heavy biomaterial seen around the impeller blades. data logs confirmed motor current spike followed by saturated flows further confirming biomaterial as the issue. the root cause of the saturated flow and the thrombus issues was heavy biomaterial wrapped around the impeller blades. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental mdr will be filed.
30203 | 1220648-2024-07625 | 18849921 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18849921&pc=OZD | 12/10/2022 | 10/27/2025 |  | 03/06/2024 |  |  | 12/10/2022 | 03/06/2024 |  |  | 04/30/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023219205 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Tachycardia | [11 empty] | Death | Required Intervention |  |  |  |  | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/10/2022 |  |  |  |  | The user facility reported a 78-year-old female being placed on impella cp required mechanical circulatory support. it was reported the patient was on impella cp support and showed some signs of hemolysis. just a bit later, the patient went into pulseless ventricular tachycardia. the family decided to make do not resuscitate order. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07625 was provided in sections b2, b5, h1, and h6. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated. f6 and f8 should have been left blank on the initial report. g1 reporting contact email and reporting contact fax number updated.
30204 | 1220648-2024-07626 | 18849926 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18849926&pc=OZD | 12/12/2022 | 03/06/2024 |  | 03/06/2024 |  |  | 12/12/2022 | 03/06/2024 |  |  | 08/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023270574 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/12/2022 |  |  |  |  | The user facility reported a 65-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient exhibited poor perfusion to operative leg, peel away was left in place and side port connected to antegrade sheath for distal perfusion. it was reported the reperfusion sheath to the right leg was clotted off in the morning but there was a distal pulse. the patient was successfully weaned from the extracorporeal membrane oxygenation (ecmo) and impella in the operating room. the patient remained stable post explant per the perfusionist. | [5 empty] | The impella device was not received from the customer, therefore investigation of the device was not possible. should additional information be provided in the future, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records.
30205 | 1220648-2024-07627 | 18849961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18849961&pc=OZD | 12/15/2022 | 03/06/2024 |  | 03/06/2024 |  |  | 12/15/2022 | 03/06/2024 |  |  |  |  | 03/06/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1181077 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 12/15/2022 |  |  |  |  | The user facility reported a 65-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the console alarmed and the purge fluid was no longer running. the issue was resolved after replacing the console, the cassette and the purge fluid. there was no patient harm reported. | [5 empty] | The investigation for the reported issue has been completed. the impella device was returned for evaluation. logs are consistent with the complaint. the occurrences of ¿controller failure¿ were due to purge pressure sensor voltage below specification. and according to the logs, it was confirmed that staff attempted to enter the wizard of purge cassette change (but failed). during testing with pump and purge simulator, the controller failure red alarm was not reproduced. additionally, it was confirmed that the pc change wizard option grey out was due to the mechanism (not malfunction) of aic console software¿ ineligible to change pc during controller alarms. the cause of the issue was a defective component. no corrective action is recommended at this time because the failure was determined to have a risk as low as practical (alap).this report is being filed as part of a retrospective review of historical records.
30206 | 1220648-2024-07629 | 18850003 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18850003&pc=OZD | 12/10/2022 | 04/15/2025 |  | 03/06/2024 |  |  | 12/10/2022 | 03/06/2024 |  |  | 07/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023257296 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/10/2022 |  |  |  |  | The user facility reported a 71-year-old female being placed on impella cp required mechanical circulatory support. it was reported the patient was on impella cp support and the device exhibited a purge pressure high alarm. the high purge pressure alarm was resolved, but pressure remained elevated. the impella cp was removed post impella 5.5 placement. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. the initial mfr 1220648-2024-07629 was incorrectly sent for the incorrect pump. the report for the impella cp pump with serial number (b)(6) and filed with mfr 1220648-2024-07629 does not have any reportable malfunction nor any adverse event to report. the correct information has been submitted for the impella 5.5 pump with serial number (b)(6) in mfr 1220648-2024-09728.
30207 | 1220648-2024-07628 | 18850017 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18850017&pc=OZD | 09/08/2022 | 03/10/2025 |  | 03/06/2024 |  |  | 09/08/2022 | 03/06/2024 |  |  | 03/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023204013 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/08/2022 |  |  |  |  | The user facility reported a 65-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella was repositioned by the physician, due to the insertion measurements were too deep. once repositioning was completed, the patient¿s pulses were unable to be found in either foot and the feet were dusky in color. activated clotting time was greater than 1000s and purge heparin was being held until activated clotting time floated down some. patient had remarkable turn around 24 hours post implant. eventually, the device would be explanted. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported limb ischemia issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
30208 | 1220648-2024-07630 | 18850188 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18850188&pc=OZD | 12/16/2022 | 10/28/2025 |  | 03/06/2024 |  |  | 12/16/2022 | 03/06/2024 |  |  | 08/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023270968 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Improper Flow or Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/16/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 60-year-old female patient for mechanical circulatory support. it was reported, while moving the patient, the impella alarmed continuous suction alarms. as a result of the noted issue, p-levels were reduced to p-2 to resolve the alarms. additionally, echocardiogram results found the impella was in the papillary muscle. multiple attempts were made to manipulate the impella back to the mid-ventricular space without success. the patient's condition had improved significantly, so the physician decided for patient to remain on support at the max p level that would not cause suction alarms until successfully weaned. | Us complainant reported that prior to being placed on impella cp support, the patient was admitted for orthopnea. during an xray, while moving the patient, the impella alarmed continuous suction alarms. p-levels were reduced to p-2 to finally resolve the alarms. echocardiogram results found the impella was in the papillary muscle. multiple attempts were made to manipulate the impella back to the mid-ventricular space without success. because the patient's condition had improved significantly; the physician made the decision to keep patient on max impella support without causing suction and to continue to wean from the pump, pressors and lasix as tolerated. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the low pump flow issue has been completed. no product was returned. upon reviewing the data logs, multiple suction alarms were noticed which were resolved with the reduction in the p - level. also, during the start of the case impella flow reduced alarms were noted even though the p - level was 9. there were no spikes in the motor current and the motor current was in good agreement with the p - level. the root cause of the low pump flow could not be determined due to insufficient clinical details.
30209 | 1220648-2024-07631 | 18850661 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18850661&pc=OZD | 12/17/2022 | 03/11/2025 |  | 03/06/2024 |  |  | 12/17/2022 | 03/06/2024 |  |  | 08/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023267005 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/17/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 65-year-old female patient for mechanical circulatory support. it was reported that during impella support, the patient developed hematoma at the impella site. as a result of the noted issue, 5 units of packed red blood cells (prbcs), 2 units of platelets (plts), and 500ml of albumin were given. the site was no longer bleeding. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
30210 | 1220648-2024-07632 | 18850756 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18850756&pc=OZD | 10/21/2022 | 03/06/2024 |  | 03/06/2024 |  |  | 10/21/2022 | 03/06/2024 |  |  | 04/30/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218216 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/21/2022 |  |  |  |  | The user facility reported for a 66-year-old male, that resistance was met during the attempted implant of the impella cp pump. as a result of the patient's tortuous anatomy, the decision was made to abort the implant and an impella 5.5 pump was subsequently implanted for patient support. | [5 empty] | The investigation for the reported delivery anatomy issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery anatomy issue was patient condition related, as the impella was unable to pass through the vascular anatomy as per the clinical description provided. this report is being filed as part of a retrospective review of historical records.
30211 | 1220648-2024-07633 | 18850927 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18850927&pc=OZD | 12/15/2022 | 04/16/2024 |  | 03/06/2024 |  |  | 12/15/2022 | 03/06/2024 |  |  | 07/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023238128 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 70-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient¿s pedal pulses were reported to not be palpable. the patient was noted to be on brilique. the impella was weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation into the access site bleeding ¿ minor and the limb ischemia ¿ reversible has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia and the access site bleeding issues was not determined.
30212 | 1220648-2024-07634 | 18851386 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18851386&pc=OZD | 10/23/2022 | 03/06/2024 |  | 03/06/2024 |  |  | 10/23/2022 | 03/06/2024 |  |  | 07/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023248904 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/23/2022 |  |  |  |  | The user facility reported a 54-year-old female was to be implanted with an impella 5.5 for mechanical circulatory support. the medical team had difficulty delivering an impella pump due to the patient's torturous anatomy. after several failed attempts they decided to abort the procedure. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issue: anatomy was most likely due to the patient¿s condition. this report is being filed as part of a retrospective review of historical records.
30213 | 1220648-2024-07635 | 18851570 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18851570&pc=OZD | 12/15/2022 | 03/06/2024 |  | 03/06/2024 |  |  | 12/15/2022 | 03/06/2024 |  |  | 03/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200352 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/15/2022 |  |  |  |  | The user facility reported an 82-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the impella cp was unable to be inserted due to stenosis in the iliac artery. a stent was placed, and reinsertion was attempted without success. after extra-corporeal membrane oxygenation was introduced, the impella cp was successfully implanted. | [5 empty] | The investigation into the delivery issues has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related (diseased vessels). this report is being filed as part of a retrospective review of historical records.
30214 | 1220648-2024-07636 | 18851688 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18851688&pc=OZD | 12/15/2022 | 03/06/2024 |  | 03/06/2024 |  |  | 12/15/2022 | 03/06/2024 |  |  | 09/30/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023273970 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/15/2022 |  |  |  |  | The user facility reported a 67-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the patient had no palpable or dopplerable distal pulses. the patient was maxed on pressors and unable to tolerate continuous renal replacement therapy (crrt). placement signal alarms were noted. crrt was discontinued. | [5 empty] | The investigation into the reported limb ischemia has been completed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
30215 | 1220648-2024-07637 | 18851786 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18851786&pc=OZD | 09/08/2022 | 03/07/2025 |  | 03/06/2024 |  |  | 09/08/2022 | 03/06/2024 |  |  | 03/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023205947 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/08/2022 |  |  |  |  | The user facility reported a male in acute myocardial infarction cardiogenic shock was unable to be implanted with an impella cp device for mechanical circulatory support. it was reported after the patient was transferred with a perforated left anterior descending (lad) artery, the physician was unable to pass the impella beyond the common femoral artery (cfa) after multiple attempts. no patient harm. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30216 | 1220648-2024-07638 | 18851931 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18851931&pc=OZD | 10/27/2022 | 03/06/2024 |  | 03/06/2024 |  |  | 10/27/2022 | 03/06/2024 |  |  | 07/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023254983 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/27/2022 |  |  |  |  | The user facility reported an unspecified patient was being implanted with an impella 5.5 device for mechanical circulatory support, and the medical team had difficulty with delivery due to the patient's vasculature. after several attempts they were unable to deliver the impella 5.5 and they placed a iabp (intra-aortic balloon pump) instead. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issue - anatomy was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
30217 | 1220648-2024-07639 | 18851940 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18851940&pc=OZD | 12/15/2022 | 03/11/2025 |  | 03/06/2024 |  |  | 12/15/2022 | 03/06/2024 |  |  | 06/30/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023248988 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/15/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 76-year-old male patient for mechanical circulatory support. it was reported that during impella support, the patient developed bleeding at the impella site. as a result of the noted issue, sutures and manual pressure were applied. additionally, the impella cp was explanted. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
30218 | 1220648-2024-07640 | 18851985 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18851985&pc=OZD | 12/15/2022 | 04/04/2025 |  | 03/06/2024 |  |  | 12/15/2022 | 03/06/2024 |  |  | 08/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023267598 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/15/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 62-year-old male patient for mechanical circulatory support. it was reported that during impella support, the patient experienced an impella cp alarm after rolling while in the icu. the physician was able to get the pump back in the correct position with echo guidance and j-wire technique. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
30219 | 1220648-2024-07641 | 18852024 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18852024&pc=OZD | 12/17/2022 | 03/11/2025 |  | 03/06/2024 |  |  | 12/17/2022 | 03/06/2024 |  |  | 09/30/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023274438 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/17/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 68-year-old male patient for mechanical circulatory support. the patient was experiencing oozing at the access site. the purge was switched to sodium bicarbonate and heparin was discontinued. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury". ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
30220 | 1220648-2024-07642 | 18852046 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18852046&pc=OZD | 10/29/2022 | 03/06/2024 |  | 03/06/2024 |  |  | 10/29/2022 | 03/06/2024 |  |  | 05/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023228273 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/29/2022 |  |  |  |  | The user facility reported a 70-year-old male was to be implanted with an impella 5.5 for mechanical circulatory support. the team had difficulty with delivery due to an innominate occlusion in the patient's vessels. therefore, the procedure was aborted. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the anatomy was most likely due to the patient condition. this report is being filed as part of a retrospective review of historical records.
30221 | 1220648-2024-07643 | 18852057 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18852057&pc=OZD | 09/08/2022 | 03/11/2025 |  | 03/06/2024 |  |  | 09/08/2022 | 03/06/2024 |  |  | 04/30/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2023221096 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/08/2022 |  |  |  |  | The user facility reported a 66-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient was placed on impella cp support after experiencing dyspnea on exertion for two months. the patient had elevated troponin, non-specific sinus tachycardia (st) changes on electrocardiogram (ekg), and the transthoracic echocardiogram (tte) revealed ejection fraction (ef) 25-30% with wall motion abnormalities. left heart catheterization revealed mitral valve disease (left anterior descending artery, left circumflex artery, and right coronary artery). after a successful implant, it was reported that labs were unable to be obtained due to hemolyzed samples. physician was informed to recheck positioning as well as reducing p-levels if tolerated. the physician reduced p level to p 6. the patient¿s urine output was minimal and was unresponsive to lasix and fluid challenges. additionally, there were absent doppler signals to right lower extremity. the patient had poor pulses and the leg was cool. surgical resident felt the leg had worsened. the patient remained on support for three additional days. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
30222 | 1220648-2024-07644 | 18852134 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18852134&pc=OZD | 09/08/2022 | 05/29/2025 |  | 03/06/2024 |  |  |  | 03/06/2024 |  |  | 01/31/2023 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022035963 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/08/2022 |  |  |  |  | The user facility reported a 77-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient was on impella cp support and the patient experienced bleeding at the impella cp groin insertion site while on a flight. manual pressure was held and systemic heparin was discontinued. after the patient arrived at the intensive care unit, the patient was switched to bicarb purge. bleeding was stopped with manual compression and ensuring the position of reposition sheath was at right angle. the pump was explanted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "assess access site for bleeding and hematoma.¿ "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30223 | 1220648-2024-07645 | 18852196 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18852196&pc=OZD | 09/09/2022 | 03/07/2025 |  | 03/06/2024 |  |  |  | 03/06/2024 |  |  | 04/30/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023213979 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/09/2022 |  |  |  |  | The user facility reported a 62-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the patient had a notable hematoma at the pump sheath site which continue to grow. the physician was notified of the hematoma and weaned the support for possible explant, pending patient¿s cardiac function. heparin was not added. eventually the device was explanted. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30224 | 1220648-2024-07646 | 18852216 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18852216&pc=OZD | 11/02/2022 | 03/06/2024 |  | 03/06/2024 |  |  |  | 03/06/2024 |  |  | 07/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023253293 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/02/2022 |  |  |  |  | The user facility reported a 63-year-old female was implanted with an impella cp device for mechanical circulatory support. the patient was bleeding excessively at the insertions site and pressure was applied to control the bleeding. two units of packed red blood cells (prbc) were given to the patient, and the heparin dosage was higher than recommended at 50 units/ml. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
30225 | 1220648-2024-07647 | 18852271 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18852271&pc=OZD | 09/09/2022 | 03/07/2025 |  | 03/06/2024 |  |  |  | 03/06/2024 |  |  | 03/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023217641 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/09/2022 |  |  |  |  | The user facility reported a 61-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the patient developed an ooze (bleeding) at the insertion site and precloses two times with additional 8fr angioseal. hemostasis was achieved. the device was eventually explanted. . | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: "potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site¿. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30226 | 1220648-2024-07648 | 18852314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18852314&pc=OZD | 11/02/2022 | 03/06/2024 |  | 03/06/2024 |  |  |  | 03/06/2024 |  |  | 08/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023260846 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/02/2022 |  |  |  |  | The user facility reported a 77-year-old female was implanted with an impella cp device for mechanical circulatory support. however, distal limb ischemia was a concern for the patient and the family decided to withdraw care and the device was explanted. | [5 empty] | The investigation for the reported limb ischemia has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. instructions for use for the related event are as follows: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
30227 | 1220648-2024-07649 | 18852386 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18852386&pc=OZD | 09/10/2022 | 03/11/2025 |  | 03/06/2024 |  |  |  | 03/06/2024 |  |  | 06/30/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023236396 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/10/2022 |  |  |  |  | The user facility reported a 60-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support after she had chronic heart failure exacerbation. the patient had a left and right heart catheter for a transcatheter aortic valve replacement work up for severe aortic stenosis. the patient went into fluid overload, went into respiratory arrest and required intubation. an external fem-fem by-pass was noted. extremities were cold and gray. the right leg still felt cold and faint pulses were present with doppler. the device was explanted. | [5 empty] | The impella cp device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
30228 | 1220648-2024-07650 | 18852574 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18852574&pc=OZD | 09/09/2022 | 03/06/2024 |  | 03/06/2024 |  |  |  | 03/06/2024 |  |  | 03/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023205528 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/10/2022 |  |  |  |  | The user facility reported a (b)(6) male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. while on the impella support, the patient had an external fem fem bypass performed for an occlusive right femoral. pulses/distal flow was dopplerable. additionally, the groin site was soft and oozing a little (from the impella sheath). after the patient was successfully weaned, no complications were noted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
30229 | 1220648-2024-07651 | 18852664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18852664&pc=OZD | 09/12/2022 | 03/11/2025 |  | 03/06/2024 |  |  |  | 03/06/2024 |  |  | 06/30/2024 |  | 03/06/2024 |  | CORONARY ARTERY DISEASE AND DIABETES MELLITUS |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235928 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/12/2022 |  |  |  |  | The user facility reported a (b)(6) male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. while on the impella support, the patient had an external fem fem bypass performed for an occlusive right femoral. pulses/ distal flow was dopplerable. additionally, the groin site was soft and oozing a little (from the impella sheath). after the patient was successfully weaned, no complications were noted. | [5 empty] | The impella cp device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
30230 | 1220648-2024-07652 | 18852732 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18852732&pc=OZD | 09/13/2022 | 03/06/2024 |  | 03/06/2024 |  |  |  | 03/06/2024 |  |  | 02/28/2023 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022056402 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/13/2022 |  |  |  |  | The user facility reported a 70-year-old female was implanted with an impella cp device for mechanical circulatory support. it was reported while being placed on impella cp support, the patient developed ischemic leg on the impella side. an external femorofemoral bypass was performed and the impella cp pump was explanted. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
30231 | 1220648-2024-07653 | 18852785 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18852785&pc=OZD | 09/13/2022 | 05/20/2025 |  | 03/06/2024 |  |  |  | 03/06/2024 |  |  | 05/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023227931 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/13/2022 |  |  |  |  | The user facility reported a 49-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported upon removing 14x25cm sheath, the impella kinked. the impella cp was explanted and a replacement impella was used. no patient harm reported. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30232 | 1220648-2024-07654 | 18852809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18852809&pc=OZD | 09/13/2022 | 05/22/2025 |  | 03/06/2024 |  |  |  | 03/06/2024 |  |  | 03/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204973 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/13/2022 |  |  |  |  | The user facility reported a (b)(6)female was implanted with an impella cp device for mechanical circulatory support. it was reported that while being placed on impella cp support, the patient began to show signs of tamponade. covered stent was placed, which allowed the percutaneous coronary intervention (pci) to be completed. the patient then underwent a pericardiocentesis. hemodynamics was normalized after that. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30233 | 1220648-2024-07655 | 18852964 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18852964&pc=OZD | 08/13/2023 | 05/13/2024 |  | 03/06/2024 |  |  | 02/05/2024 | 03/06/2024 |  |  | 11/30/2023 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2023167193 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 57 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Low Blood Pressure/ Hypotension | Sepsis | Thrombocytopenia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/05/2024 |  |  |  |  | Additional information was received via new loqi notifications. on (b)(4) 2024, the patient endured hypotension with a possible relationship to the device that prompted "7.5 ivf and one unit of rbc." additionally, the patient was "started on levophed and was discontinued on (b)(4) 2023. while patient was recovering from rv failure, patient was also treated with dopamine which helped with the maps throughout the remainder of this admission." further information received indicates the patient to have endured sepsis with a possible relationship to the device. as treatment, a wound vac was performed. | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry regarding a 57 year old patient the presented with acute myocardial infarction and cardiogenic shock for impella support. during support, the patient endured worsening thrombocytopenia and hemolysis with a "possible" relationship to the impella device. platelets were delivered in response to worsening thrombocytopenia, and a red blood cell infusion was performed in response to hemolysis. |  |  |  |  | Upon completion of our investigation, a supplemental mdr will be filed. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the reported issues has been completed since the original report was submitted. the device was not returned for investigation. the cause of the thrombocytopenia and the minor access site bleeding was not determined because no product was returned and limited clinical information was provided. the cause of the hemolysis was most likely patient condition since the patient had a history of anemia. because no associated information was provided, the root cause of the infection cannot be determined. potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
30234 | 1220648-2024-07656 | 18853041 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853041&pc=OZD | 08/21/2023 | 07/26/2024 |  | 03/06/2024 |  |  | 02/05/2024 | 03/06/2024 |  |  | 05/31/2025 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024341842 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiomyopathy | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/05/2024 |  |  |  |  | Additional information received via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured worsening of heart failure with reduced ejection fraction and worsening of non-ischemic cardiomyopathy with a "possible" relationship to the impella device used. impella waveform and settings observed throughout admission. | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry regarding a 66 year old male patient presenting with cardiomyopathy for impella support. during support, signs of hemolysis were endured with a "possible" relationship to the impella device. unspecified intervention was performed and patient was monitored. the patient expired after being made "comfort measures only." |  |  |  |  | Section b5 and h6 were updated as per additional information received. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | Additional information was received from the coordinator for abiomed¿s long-term outcome and quality indicator impella registry regarding an updated relatedness for the adverse event: heart failure with reduced ejection fraction to "not related" to the impella device. the investigation for the reported hemolysis has been completed. the cause of the hemolysis was unable to be determined since the product was not returned for review. relation to impella is unknown since the hemolysis reported begin prior to impella placement. the data logs were reviewed and no positioning issues were observed and not spikes in motor current. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ b.7 revised as this information was inadvertently omitted from the initial submission of manufacturer device report number 1220648-2024-07656. d.4 revised model number from the initial submission of manufacturer device report number 1220648-2024-07656 in accordance with updated procedures. d.6a and d.6b revised from the initial submission of manufacturer device report number 1220648-2024-07656 in accordance with updated procedures. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-07656 and should not have been. g.1 revised facility name and manufacturing site fax number from the initial submission of manufacturer device report number 1220648-2024-07656 in accordance with updated procedures. h.6 added code 925 since initial submission of manufacturer device report number 1220648-2024-07656 in accordance with updated procedures. h.6 added code 4115 as it was inadvertently omitted from the initial submitted manufacturer device report number 1220648-2024-07656. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2024-07656 in accordance with updated procedures. h.6 and h.10 investigation results were inadvertently omitted on supplemental manufacturer device report follow-up 1 and have been added to this report. information to section b.5 reporting heart failure with reduced ejection fraction reporting, h.6 health effect-clinical code: 1764 cardiomyopathy and its corresponding health effect¿impact code: unexpected 4641 medical intervention captured under supplemental manufacturer device report follow-up 1 acknowledged/received on may 16, 2024, should be disregarded.
30235 | 1220648-2024-07065 | 18853088 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853088&pc=OZD | 04/18/2022 | 10/28/2025 |  | 03/06/2024 |  |  | 04/18/2022 | 03/06/2024 |  |  | 06/30/2023 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022100437 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/18/2022 |  |  |  |  | The user facility reported a 69-year-old male patient was implanted with an impella cp device for mechanical circulatory support. it was reported that while on support, the patient developed a right groin hematoma. the nurse decompressed the site. 250 milliliters of normal saline bolus was started. the heparin drip was stopped due to activated clotting time of 240 seconds. the patient continued on support. | Us complainant reported the patient was on impella cp support and the patient developed a right groin hematoma. the nurse decompressed the site. 250 normal saline bolus was started. heparin drip was stopped due to activated clotting time of 240. the nurse rechecked the patient in one hour. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the access site bleeding - major issue has been completed. the root cause of the access site bleeding - major could not be determined as insufficient clinical details were provided; however, according to information received from the healthcare professional, access site bleeding was most likely due to anticoagulation management.
30236 | 1220648-2024-07066 | 18853089 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853089&pc=OZD | 10/23/2022 | 05/20/2025 |  | 03/06/2024 |  |  | 10/23/2022 | 03/06/2024 |  |  | 06/30/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235260 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/23/2022 |  |  |  |  | The user facility reported a 70-year-old male patient needing mechanical circulatory support was implanted with an impella cp device. the patient experienced hemolysis as the urine was red tinged. transesophageal echocardiogram showed the pump diving into mitral valve apparatus. the pump was removed prior to the patient going on bypass. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30237 | 1220648-2024-07006 | 18853091 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853091&pc=OZD | 02/12/2024 | 04/01/2024 |  | 03/06/2024 |  |  | 02/12/2024 | 03/06/2024 |  |  | 11/30/2025 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024430095 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/12/2024 |  |  |  |  | The user facility reported a 60-year-old female patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that it was planned to remove an impella cp and escalate care to the impella 5.5. the team attempted to deliver the impella 5.5 via the surgical graft anastomosis but, delivery failed. the wire-pump failed to deliver. the pump was removed and at the time of explant, biomaterial was adhered and the pulse was lost at the graft and the bypass was placed to the axillary artery. the impella 5.5 aborted delivery made the team keep the patient on the impella cp. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation into the delivery issues, the thrombus issue, and the limb ischemia - irreversible has been completed since the initial report was submitted. the product was not returned for investigation. the root cause of delivery issue was determined to be patient condition since it was mentioned that the patient has small vessel size. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus issue, and the limb ischemia - irreversible was not determined.
30238 | 1220648-2024-07132 | 18853092 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853092&pc=PYX | 02/14/2024 | 03/21/2024 |  | 03/06/2024 |  |  | 02/14/2024 | 03/06/2024 |  |  | 07/31/2025 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC |  |  | US |  | 2024392299 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Female |  |  |  | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/14/2024 |  |  |  |  | The user facility reported a 32-year-old female patient of unknown race and ethnicity was in acute myocardial infarction cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. it was reported that following placement of an impella rp pump, the doctor was placing deep mattress sutures when a suture snapped. the doctor's hand subsequently jerked and pulled the repositioning sheath and pump out of position. after multiple unsuccessful attempts to reposition the device, the decision was made to explant the device. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the positioning issue was use related.
30239 | 1220648-2024-07063 | 18853093 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853093&pc=OZD | 10/18/2022 | 06/02/2025 |  | 03/06/2024 |  |  | 10/18/2022 | 03/06/2024 |  |  | 01/31/2023 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022048401 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  |  | Death | Yes | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 10/18/2022 |  |  |  |  | The user facility reported a 60-year-old male patient was reported to be implanted with an impella cp device for mechanical circulatory support. during the placement of the impella cp in the right femoral artery, a recently placed femoral stent prevented placement. the procedure was aborted. there was no patient harm reported. | It was noted that the patient expired in the operating room. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely due to patient condition as a recent femoral stent was preventing placement. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07063 was provided in sections b2, b5, h1, and h6. | The root cause of the delivery issue was most likely due to patient condition since patient had recent femoral stent. this report is being filed as part of a retrospective review of historical records.
30240 | 1220648-2024-07064 | 18853094 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853094&pc=OZD | 10/20/2022 | 05/30/2025 |  | 03/06/2024 |  |  | 10/20/2022 | 03/06/2024 |  |  | 06/30/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023236889 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/20/2022 |  |  |  |  | The user facility reported an (b)(6) female patient needing mechanical circulatory support. it was reported while being on the impella cp and the patient developed a large hematoma, and no distal pulses were found. manual pressure was held by the staff. the impella cp and the sheath were explanted. the user facility reported an 82-year-old female patient was needing mechanical circulatory support. it was reported while being on the impella cp no distal pulses were found. the impella cp and the sheath were explanted. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported limb ischemia issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
30241 | 1220648-2024-07317 | 18853095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853095&pc=OZD | 11/14/2022 | 03/04/2024 |  | 03/06/2024 |  |  | 11/14/2022 | 03/06/2024 |  |  | 09/30/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023259983 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/14/2022 |  |  |  |  | The user facility reported that the medical team experienced difficulty implanting an impella 5.5 in a 66-year-old male patient for mechanical circulatory support due to the patient's anatomy. it was reported the impella 5.5 device was unable to advance through the subclavian artery during attempted delivery, and the impella 5.5 implant was aborted due to patient anatomy. | [5 empty] | The investigation into the reported delivery issue has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue: anatomy was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records
30242 | 1220648-2024-07431 | 18853096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853096&pc=OZD | 02/17/2024 | 04/01/2024 |  | 03/06/2024 |  |  | 02/17/2024 | 03/06/2024 |  |  |  |  | 03/06/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1461597 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  | Asian | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/17/2024 |  |  |  |  | The user facility reported a unknown age, gender and ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the complainant reported that an aic failed to recognize an installed purge cassette during preparation for impella implant. the purge cassette was replaced to troubleshoot the issue, but the purge disc not detected alarm persisted. priming was unable to proceed and the decision was made to abort the implant due to the absence of a replacement console. there was no patient harm. | [5 empty] | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were reviewed which confirmed the inability to detect the purge disc. the console was evaluated and the issue with detecting the purge disc was reproduced, and the purge flag was confirmed to be broken . the root cause of this issue was a broken purge flag. d4-updated model number | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
30243 | 1220648-2024-07428 | 18853097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853097&pc=OZD | 02/09/2024 | 04/17/2024 |  | 03/06/2024 |  |  | 02/16/2024 | 03/06/2024 |  |  | 09/30/2025 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2024411411 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/16/2024 |  |  |  |  | The user facility reported a 35-year-old male with cardiomyopathy was implanted with an impella for mechanical circulatory support. during repositioning, the displayed impella flow dropped to 0l/min. the issue persisted for roughly one to two minutes until the impella was successfully restarted. patient support was maintained at the time of the noted issue with the assistance of ecmo. upon return to functionality, no further issues were noted with the pump for the remainder of support. there was no patient harm. | [5 empty] | The investigation for the reported pump stop has been completed. the impella device was returned for evaluation. product evaluation confirmed the presence of blood in purge gap and purge line. however, the temporary pump stop could not be reproduced. data log analysis did not include the temporary pump stop as the logs did not cover the early hours of support. according to the clinical, during the first day of impella cp support a temporary pump stop was recorded. when the temporary pump stop occurred the patient was being repositioned. no decision was made to remove the device so it remained in the patient for a duration of support. the cause of the temporary pump stop could not be determined because logs were not returned and not enough information was given. the instructions for use states ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
30244 | 1220648-2024-07427 | 18853098 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853098&pc=OZD | 12/01/2022 | 03/05/2024 |  | 03/06/2024 |  |  | 12/01/2022 | 03/06/2024 |  |  | 07/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023257996 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/01/2022 |  |  |  |  | The user facility reported a 67-year-old male patient on impella cp for support during urgent coronary artery bypass graft. while on the impella device, the staff was unable to find pulses in the right foot and had weak left pulses. the ultrasound showed thrombus in right groin at insertion site. vascular cut down was performed and thrombectomy of large clot was successful. good bilateral flow was confirmed on angiogram following thrombectomy. the patient continued on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ this report is being filed as part of a retrospective review of historical records.
30245 | 1220648-2024-07429 | 18853099 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853099&pc=OZD | 11/23/2022 | 03/12/2025 |  | 03/06/2024 |  |  | 11/23/2022 | 03/06/2024 |  |  | 03/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023214141 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/23/2022 |  |  |  |  | The user facility reported a (b)(6) male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during this time expanding groin was noted at impella access site. the introducer sheath was peeled away, and bleeding increased. the physician elected to wean and explant the impella cp due to the bleeding. manual pressure was held for 40 minutes, and the medical team gained control of the bleeding. blood products were administered due to the blood loss. the patient maintained stable mean arterial pressure post explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. the root cause of the access site bleeding issue was not determined since product was not returned.
30246 | 1220648-2024-07430 | 18853100 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853100&pc=OZD | 02/16/2024 | 04/03/2024 |  | 03/06/2024 |  |  | 02/16/2024 | 03/06/2024 |  |  | 09/30/2025 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024399650 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/16/2024 |  |  |  |  | The user facility reported a 30-year-old male with cardiomyopathy was implanted with an impella for mechanical circulatory support. during attempted repositioning, the pump got caught in the papillary muscle and was accidentally pulled out of the left ventricle. as a result of the noted issue, the decision was made to explant the device. the patient remained on ecmo for ongoing support. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed since the initial report was submitted. the device was not returned for investigation. per the clinical data provided, the cause of the positioning issue was use related due to improper technique.
30247 | 1220648-2024-07133 | 18853101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853101&pc=OZD | 09/14/2022 | 03/12/2025 |  | 03/06/2024 |  |  | 09/14/2022 | 03/06/2024 |  |  | 03/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023210971 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/14/2022 |  |  |  |  | The user facility reported a 71-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was on impella cp support and experienced bleeding from all access sites. an echocardiogram revealed the impella cp had pushed into the left ventricle. no further information has been provided. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30248 | 1220648-2024-07433 | 18853102 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853102&pc=OZD | 12/01/2022 | 05/28/2025 |  | 03/06/2024 |  |  | 12/01/2022 | 03/06/2024 |  |  | 08/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023262600 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/01/2022 |  |  |  |  | The user facility reported a 33-year-old female patient on impella cp for mechanical circulatory support. the patient was experiencing hemolysis. after weaning p-levels, hemolysis cleared up. additionally, there was a purge pressure alarm for the duration of the case. the staff checked for kinks in the line and couldn¿t find any. they attempted to change the purge cassette but was unable to disconnect the yellow connector from the side arm. since the pump was being explanted the next day, the decision was made not to change it, to prevent cracking the luer. . | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30249 | 1220648-2024-07439 | 18853103 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853103&pc=OZD | 12/15/2021 | 03/12/2025 |  | 03/06/2024 |  |  | 12/15/2021 | 03/06/2024 |  |  | 07/31/2023 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022120072 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/15/2021 |  |  |  |  | The user facility reported a 63-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the medical team determined the patient needed to be escalated to an impella 5.5 for additional mechanical support. it was reported that after the impella cp was removed, there was a loss of distal flow to the femoral artery. a vascular surgeon removed 3 large clots from the artery and flow returned. the access site was surgically closed. the patient was implanted with an impella 5.5 device successfully. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | "the failure mode of ¿thrombus¿ is to be used when a thrombus is noted but cannot be connected to any other adverse events or decrease in pump performance. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions description of the location and characteristics of the thrombus (e.g., size, shape, consistency). relevant laboratory test results, such as coagulation profiles and platelet counts imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). medications or treatments that the patient was receiving at the time of thrombus formation. surgical or procedural details if applicable (e.g, cardiac catheterization). any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined". this report is being filed as part of a retrospective review of historical records.
30250 | 1220648-2024-07436 | 18853104 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853104&pc=OZD | 02/23/2024 | 04/10/2024 |  | 03/06/2024 |  |  | 02/23/2024 | 03/06/2024 |  |  | 11/30/2025 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025438701 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/23/2024 |  |  |  |  | The user facility reported a 78-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and there were limb ischemia issues. the patient lost distal pulses and the vascular team was evaluating peripheral arteries for external fem/fem by-pass possible axillary access if fem/fem by-pass could not be established. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
30251 | 1220648-2024-07437 | 18853105 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853105&pc=OZD | 09/08/2022 | 03/05/2024 |  | 03/06/2024 |  |  | 09/08/2022 | 03/06/2024 |  |  | 04/30/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023219267 |  | P140003 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/08/2022 |  |  |  |  | The user facility reported a 43-year-old female was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient had signs of hemolysis with dark red urine and hemolyzed labs. the impella cp pump was replaced with an impella 5.5. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
30252 | 1220648-2024-07435 | 18853106 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853106&pc=OZD | 11/23/2022 | 03/12/2025 |  | 03/06/2024 |  |  | 11/23/2022 | 03/06/2024 |  |  | 08/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023269746 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/23/2022 |  |  |  |  | The user facility reported a 76-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced oozing at the insertion site. an additional suture was placed, the site was redressed, and heparin was discontinued due to the bleeding. the patient continued on support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
30253 | 1220648-2024-07432 | 18853107 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853107&pc=OZD | 02/16/2024 | 09/13/2024 |  | 03/06/2024 |  |  | 02/16/2024 | 03/06/2024 |  |  | 11/30/2025 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024431190 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | No | Yes | Pumping Stopped | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/16/2024 |  |  |  |  | The user facility reported a 71-year-old male post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. the implanted impella 5.5 pump experienced saturated flow and subsequently stopped during patient support. the pump was explanted as a result of the noted issue. there was no patient harm. | It was noted that the patient expired in the operating room. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the pump stop issue has been completed since the initial report was submitted. the device was not returned for investigation. the cause of the pump stop issue was most likely biomaterial per log and clinical review. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrected information for the initial mfr 1220648-2024-07432 was provided.
30254 | 1220648-2024-07434 | 18853108 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853108&pc=PYX | 02/22/2024 | 03/25/2024 |  | 03/06/2024 |  |  | 02/22/2024 | 03/06/2024 |  |  | 10/31/2025 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | RP FLEX W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2024413665 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/22/2024 |  |  |  |  | The user facility reported a 55-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella rp support. while on support, the patient had hemolyzed lab results and hematuria. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the hemolysis issue has been completed since the initial report was submitted. the product was not returned for investigation. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review.
30255 | 1220648-2024-07443 | 18853109 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853109&pc=OZD | 02/21/2022 | 03/05/2024 |  | 03/06/2024 |  |  | 02/21/2022 | 03/06/2024 |  |  | 06/30/2023 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022102546 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 02/21/2022 |  |  |  |  | The user facility reported a male patient was being implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the physician was having difficulty getting the impella 5.5 to pass through the graft. once it passed, the pump was unable to cross the valve. therefore, placement was of impella 5.5 was aborted. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not received from the customer. however, clinical details were sufficient to determine the root cause. the root cause of the delivery issue was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
30256 | 1220648-2024-07590 | 18853110 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853110&pc=OZD | 11/11/2022 | 04/07/2025 |  | 03/06/2024 |  |  | 11/11/2022 | 03/06/2024 |  |  | 07/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023254937 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/11/2022 |  |  |  |  | The user facility reported a 64-year-old female was implanted with an impella cp device for mechanical circulatory support. it was reported after the impella cp was successfully implanted, the patient experienced bleeding from insertion site. as a result of the noted issue, position of the repo sheath was adjusted, 2 units of packed red blood cells (prbcs), and extra sutures were placed. no new oozing was noted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: implant and explant dates added. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
30257 | 1220648-2024-07442 | 18853111 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853111&pc=OZD | 02/16/2024 | 04/17/2024 |  | 03/06/2024 |  |  | 02/16/2024 | 03/06/2024 |  |  |  |  | 03/06/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1798913 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/16/2024 |  |  |  |  | The user facility reported a 69-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the automated impella controller (aic) failed to detect an installed purge cassette during set up. the aic was replaced as a result of the noted failure. there was no patient harm. | [5 empty] | The investigation into the console (aic) purge disc/flag issue has been completed. the device was not returned for evaluation. data logs were reviewed finding no instances of the reported "cassette not detected" alarms found in the logs from the reported occurrence date. however, there were 2 instances of purge disc not detected alarms (event set #402) found from the reported occurrence date. both instances of purge disc not detected alarms were resolved quickly after the purge disc was inserted. the root cause of the purge disc/flag issue was unable to be determined because the product was not returned. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30258 | 1220648-2024-07446 | 18853112 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853112&pc=OZD | 02/16/2024 | 04/04/2024 |  | 03/06/2024 |  |  | 02/16/2024 | 03/06/2024 |  |  | 03/31/2025 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024321649 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  |  |  | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/16/2024 |  |  |  |  | The user facility reported a 82-year-old male of unknown race and ethnicity undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the purge flow of an implanted impella 5.5 pump dropped to below 2ml/hr resulting in high purge pressure alarms. plans were made to start a higher strength heparin concentration, but a purge blocked alarm appeared and the decision was made to explant the device. the patient hemodynamics were improved and there was no subsequent patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into high purge pressure issue has been completed since the initial report was submitted. the device was not returned for investigation. a review of the data logs suggests that there was a possible biomaterial build up leading to a pump stop. per the data log review, the root cause of the high purge pressure issue was most likely biomaterial.
30259 | 1220648-2024-07440 | 18853113 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853113&pc=OZD | 02/25/2024 | 04/03/2024 |  | 03/06/2024 |  |  | 02/25/2024 | 03/06/2024 |  |  | 10/31/2025 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024423479 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/25/2024 |  |  |  |  | The user facility reported a 84-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella had an aorta alarm, and it was unable to be resolved at bedside using echo. the pigtail fell back across the aorta valve. the patient on impella support remained hemodynamically stable. the decision was made to explant impella at bedside. | [5 empty] | The investigation into the positioning issues has been completed since the initial report was submitted. the device was not returned for investigation. per the clinical description and data analysis provided, the root cause of positioning issue was most likely use related. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿
30260 | 1220648-2024-07438 | 18853114 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853114&pc=OZD | 02/23/2024 | 04/10/2024 |  | 03/06/2024 |  |  | 02/23/2024 | 03/06/2024 |  |  |  |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025443186 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/23/2024 |  |  |  |  | The user facility reported a 60-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant was attempting delivery of the 5.5 pump for support of a patient admitted with a stemi and ischemic cardiomyopathy. the team was unable to deliver the pump/wire to the lv (left ventricle) and support was aborted. the ct (computed tomography) imaging revealed the issue was native anatomy with an underlying stenosis at the subclavian-innominate artery. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition. added- h6 impact code 4627. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30261 | 1220648-2024-07445 | 18853115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853115&pc=OZD | 09/07/2022 | 03/11/2025 |  | 03/06/2024 |  |  | 09/07/2022 | 03/06/2024 |  |  | 11/30/2023 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023161931 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/07/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 58-year-old male patient for mechanical circulatory support. it was reported while being placed on impella cp support the patient was observed to have a thrombus/biomaterial. the decision was made to explant the femoral impella cp along with the embolectomy and the patient was escalated to the impella 5.5. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. no device was received for evaluation. the device failed post sterile inspection due to ppm test failure. this issue is unrelated to the failure mode. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
30262 | 1220648-2024-07588 | 18853116 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853116&pc=OZD | 12/10/2022 | 09/12/2024 |  | 03/06/2024 |  |  | 12/10/2022 | 03/06/2024 |  |  | 08/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023263098 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Positioning Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/10/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 85-year-old female patient for mechanical circulatory support. it was reported that the impella appeared to be in the papillary muscle, so it was repositioned. there was no patient harm reported. | Additional information noted the patient care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07588 was provided in sections b2, b5, h1, and h6.
30263 | 1220648-2024-07586 | 18853117 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853117&pc=OZD | 12/07/2022 | 05/13/2025 |  | 03/06/2024 |  |  | 12/07/2022 | 03/06/2024 |  |  | 04/30/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023223209 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/07/2022 |  |  |  |  | The user facility reported a 71-year-old male was implanted with an impella cp device for mechanical circulatory support. at the end of the impella cp implant procedure, flows and positioning were suspected. as a result, the impella cp pump was removed and the left ventricle via 14 french sheath was rewired. the pump was reinserted with better positioning and flows. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30264 | 1220648-2024-07447 | 18853118 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853118&pc=OZD | 09/08/2022 | 03/05/2024 |  | 03/06/2024 |  |  | 09/08/2022 | 03/06/2024 |  |  | 05/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231037 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/08/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 79-year-old female patient for mechanical circulatory support. it was reported while being placed on impella cp support, the patient developed ischemic leg on impella side. an external femorofemoral bypass was performed. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
30265 | 1220648-2024-07444 | 18853119 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853119&pc=OZD | 10/10/2021 | 03/05/2024 |  | 03/06/2024 |  |  | 10/10/2021 | 03/06/2024 |  |  | 07/31/2023 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2022123860 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Disability | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/10/2021 |  |  |  |  | The user facility reported a 58-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the complainant reported that a patient developed ischemia in their impella leg during support. after withstanding the condition for sixteen hours, the patient's leg was cold and blue with rising lactate levels. the pump was subsequently removed and the patient was escalated to impella 5.5 support. once removed, distal pulses were restored, however, the patient was taken to the operating room three days later for an amputation above the knee. the amputation was attributed to the ischemia. support continued with the implanted 5.5 pump. | [5 empty] | The investigation for the reported limb ischemia has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
30266 | 1220648-2024-07441 | 18853120 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853120&pc=OZD | 11/23/2022 | 03/12/2025 |  | 03/06/2024 |  |  | 11/23/2022 | 03/06/2024 |  |  | 06/30/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235515 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/23/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 52-year-old female patient for mechanical circulatory support. the impella was placed via the patient¿s left femoral artery and upon insertion into the left ventricle, it was noted that it had caught on the papillary. the impella was removed and reinserted with successful placement into the left ventricle. procedure was completed without further issue. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30267 | 1220648-2024-07582 | 18853121 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853121&pc=OZD | 11/10/2022 | 03/12/2025 |  | 03/06/2024 |  |  | 11/10/2022 | 03/06/2024 |  |  | 08/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023261412 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/10/2022 |  |  |  |  | The user facility reported a 57-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced hematuria and hemolyzed labs indicating that the patient had developed hemolysis. the impella cp pump was explanted and was escalated to an impella 5.5. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30268 | 1220648-2024-07583 | 18853122 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853122&pc=OZD | 02/18/2024 | 03/06/2024 |  | 03/06/2024 |  |  | 02/20/2024 | 03/06/2024 |  |  |  |  | 03/06/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1204376 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 02/20/2024 |  |  |  |  | The user facility reported a 63-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that a ¿purge disc not detected¿ alarm occurred. the automated impella controller (aic) was exchanged. there was no harm to the patient. | [5 empty] | The investigation into the reported aic - purge disc/flag issues has been completed. after reviewing the logs, the issues with automated impella controller's inability to detect the purge disc was confirmed. multiple instances of purge disc not detected were found in the logs on the reported occurrence date. the console was tested after arriving. the issue with detecting the purge disc was reproduced, and the purge flag was confirmed to be broken. the cause of the issue was a broken purge flag.
30269 | 1220648-2024-07584 | 18853123 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853123&pc=OZD | 12/06/2022 | 03/12/2025 |  | 03/06/2024 |  |  | 12/06/2022 | 03/06/2024 |  |  | 08/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023263783 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility |  |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/06/2022 |  |  |  |  | The user facility reported a 63-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. it was reported during impella cp support, the motor current spiked and a pump stop occurred. the pump was restarted but there was low flow with minimally pulsatile motor current, persistent suction, and a negative left ventricle signal throughout the cardiac cycle. a second pump was placed and started without issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. correction: g1 reporting contact first name and reporting contact last name.
30270 | 1220648-2024-07585 | 18853124 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853124&pc=OZD | 02/20/2024 | 04/01/2024 |  | 03/06/2024 |  |  | 02/20/2024 | 03/06/2024 |  |  | 11/30/2025 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024433023 |  | 0048-0045 |  | No |  |  | No |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/20/2024 |  |  |  |  | The user facility reported a 78-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed a hematoma at the access site. the team applied femostop and stopped heparin. the pump was explanted and shortly afterwards was made a do-not-resuscitate with care/comfort only and expired. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation into the access site bleeding ¿ major issue has been completed since the initial report was submitted. the product was not returned for investigation. the root cause of the access site bleeding was most likely due to anticoagulation management since ptt was greater than 400 which is way higher than recommended safe number.
30271 | 1220648-2024-07589 | 18853125 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853125&pc=OZD | 12/12/2022 | 03/06/2024 |  | 03/06/2024 |  |  | 12/12/2022 | 03/06/2024 |  |  | 09/30/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023273540 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/12/2022 |  |  |  |  | The user facility reported an 85-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after being placed on impella cp support, it was reported the staff observed lower extremity on impella cp side mottled with no flow identified. the sheath was exchanged, but the patient developed a loss of pedal pulses. the decision was made to wean and remove the impella cp pump to resolve limb ischemia. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
30272 | 1220648-2024-07587 | 18853126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853126&pc=OZD | 11/11/2022 | 03/12/2025 |  | 03/06/2024 |  |  | 11/11/2022 | 03/06/2024 |  |  | 08/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023265694 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 91 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/11/2022 |  |  |  |  | The user facility reported a (b)(6)male was implanted with an impella cp device for mechanical circulatory support. it was reported that after the impella cp was successfully implanted, the patient experienced persistent bleeding at the impella cp insertion site following a failed perclose. an 8fr angioseal was deployed and hemostasis was achieved. no new oozing was noted. additionally, it was reported when the pump was crossing the aortic valve, the wire was pulled back and the pump was inadvertently pulled back behind the value, preventing it from being advanced. the pump was then removed, reinserted and started without issue after repeating the insertion process. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30273 | 1220648-2024-07591 | 18853129 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853129&pc=OZD | 09/02/2022 | 03/06/2024 |  | 03/06/2024 |  |  | 09/02/2022 | 03/06/2024 |  |  | 05/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023222323 |  | 0550-0008 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/02/2022 |  |  |  |  | The user facility reported a 65-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the medical team was having difficulty delivering the impella pump due to the patient's anatomy and after failing to do so they aborted the procedure. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
30274 | 1220648-2024-07592 | 18853130 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18853130&pc=OZD | 12/07/2022 | 05/13/2025 |  | 03/06/2024 |  |  | 12/07/2022 | 03/06/2024 |  |  | 07/31/2024 |  | 03/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023256729 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/07/2022 |  |  |  |  | The user facility reported a 74-year-old male was implanted with an impella device for mechanical circulatory support. while the patient was on impella cp support, the impella cp pulled back across the aorta, as a result, the device had to be reinserted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30275 | 1220648-2024-07407 | 18835627 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18835627&pc=OZD | 10/11/2021 | 09/15/2024 |  | 03/05/2024 |  |  | 10/11/2021 | 03/05/2024 |  |  | 01/31/2023 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022032557 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/11/2021 |  |  |  |  | The user facility reported a 74-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the medical team attempting to implant an impella 5.5 had difficulty delivering the device due to patient's anatomy. the implant was aborted and an impella cp was placed. | The procedural outcome is patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the reported delivery issue has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue - anatomy was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07407 was provided in sections b2, b5, h1, and h6.
30276 | 1220648-2024-07408 | 18835679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18835679&pc=OZD | 10/29/2021 | 03/05/2024 |  | 03/05/2024 |  |  | 10/29/2021 | 03/05/2024 |  |  | 07/31/2023 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022115164 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/29/2021 |  |  |  |  | The user facility reported a 72-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that resistance was met when attempting to advance an impella 5.5 pump during implant. it was noted that the device would not move forward through the innominate artery due to an existing calcium shelf. the pump was removed and an abbott 8mmx40mm angioplasty balloon was inserted and inflated at the calcium shelf. the impella 5.5 pump was inserted again and implant went smoothly without further issue. there was no patient harm resulting from the resistance caused by the patient's anatomy. | [5 empty] | The investigation into the reported delivery issue has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue: anatomy was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
30277 | 1220648-2024-07409 | 18835758 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18835758&pc=OZD | 10/29/2021 | 03/05/2024 |  | 03/05/2024 |  |  | 10/29/2021 | 03/05/2024 |  |  | 07/31/2023 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022119302 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/29/2021 |  |  |  |  | The user facility reported a 77-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that an impella 5.5 pump was unable to advance beyond the graft anastomosis and axillary junction during attempted implant via the right axillary artery. after resistance was met due to the heavy calcification and tortuous anatomy, an attempt was made to place the impella via the left axillary artery. once again, the pump was unable to advance through the patient's tortuous anatomy and the implant was aborted. the patient was placed on extracorporeal membrane oxygenation (ecmo) for support. there was no patient harm. | [5 empty] | The investigation into the reported delivery issue has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue - anatomy was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records
30278 | 1220648-2024-07410 | 18835827 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18835827&pc=OZD | 11/28/2022 | 06/02/2025 |  | 03/05/2024 |  |  | 11/28/2022 | 03/05/2024 |  |  | 07/31/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023258738 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Suction Problem | Insufficient Flow or Under Infusion | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 11/28/2022 |  |  |  |  | The user facility reported a 71-year-old male patient needing mechanical circulatory support. it was reported that the patient was on impella cp support and the impella cp appeared ¿kinked¿ near the motor housing on the distal side of the aortic valve. the staff observed reduced flows and inconsistent suctioning. the device was successfully removed and replaced. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. b1: corrected from ¿adverse event¿ to ¿product problem¿ h1: corrected from ¿serious injury¿ to ¿malfunction¿.
30279 | 1220648-2024-07411 | 18835857 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18835857&pc=OZD | 12/01/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 12/01/2022 | 03/05/2024 |  |  | 08/31/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023265602 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 78 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/01/2022 |  |  |  |  | The user facility reported a 78-year-old male patient was being implanted with an impella cp device for mechanical circulatory support. the site reported that while attempting impella cp support the staff was unable to obtain axillary access. therefore, the case was not completed. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not received from the customer. however, clinical details were sufficient to determine the root cause. the root cause of the delivery issue was determined to be patient condition related. this report is being filed as part of a retrospective review of historical records.
30280 | 1220648-2024-07412 | 18835858 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18835858&pc=OZD | 11/11/2021 | 03/05/2024 |  | 03/05/2024 |  |  | 11/11/2021 | 03/05/2024 |  |  | 08/31/2023 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022128818 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/11/2021 |  |  |  |  | The user facility reported that the medical team experienced difficulty implanting an impella 5.5 in a 59-year-old male patient for mechanical circulatory support due to the patient's anatomy. after multiple failed attempts, the impella 5.5 implant was aborted | [5 empty] | The investigation into the reported delivery issue has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue: anatomy was patient condition since patient had coarctation of the aorta. this report is being filed as part of a retrospective review of historical records.
30281 | 1220648-2024-07413 | 18835935 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18835935&pc=OZD | 11/21/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 11/21/2022 | 03/05/2024 |  |  | 08/31/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023263343 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/21/2022 |  |  |  |  | The user facility reported a 78-year-old male noted as post cardiotomy cardiogenic shock¿/low cardiac output syndrome was implanted with an impella cp device for mechanical circulatory support. it was reported that during initial implantation, the 0.18 wire curled up on itself and became tangled in the impella. the device and sheath were removed and the artery was reassessed. the pump was again prepped and placed per protocol. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
30282 | 1220648-2024-07414 | 18835948 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18835948&pc=OZD | 11/29/2021 | 03/05/2024 |  | 03/05/2024 |  |  | 11/29/2021 | 03/05/2024 |  |  | 06/30/2023 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022104877 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/29/2021 |  |  |  |  | The user facility reported a 65-year-old female in acute myocardial infarction cardiogenic shock was to be implanted with an impella 5.5 device for mechanical circulatory support. the physician had difficulty delivering the 5.5 device due to the patient's anatomy. after failed attempts, the procedure was aborted. an impella cp was placed through axillary graft for continued hemodynamic support. there was no patient harm. | [5 empty] | The investigation into the delivery issues has been completed. the impella products were not returned for evaluation. the root cause of the delivery issue: anatomy was patient condition because of the patient's complex vascular anatomy. this report is being filed as part of a retrospective review of historical records.
30283 | 1220648-2024-07415 | 18836168 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18836168&pc=OZD | 11/21/2022 | 03/12/2025 |  | 03/05/2024 |  |  | 11/21/2022 | 03/05/2024 |  |  | 09/30/2023 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022136202 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/21/2022 |  |  |  |  | The user facility reported a 52-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the patient experienced a slow ooze at the insertion site. manual pressure was held for 30 minutes. the site was then redressed and propped up with a kerlix ball. systemic heparin was stopped due to oozing from impella site. after the oozing was resolved, heparin was restarted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30284 | 1220648-2024-07416 | 18836211 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18836211&pc=OZD | 11/30/2021 | 11/12/2025 |  | 03/05/2024 |  |  | 11/30/2021 | 03/05/2024 |  |  |  |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1591091 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  |  |  | Malfunction | No | Yes | Computer Operating System Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 11/30/2021 |  |  |  |  | The user facility reported a 62-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that at the end of the case the automated impella controller (aic) needed to be replaced due to an alarm with the message "controller error" when the console cable was disconnected. the aic was replaced and the new aic operated as expected. | The complainant reported that the pump in a patient on impella support had tendency to rotate so the device needed to be adjusted several times before position was correct. at the end of the case the automated impella controller (aic) needed to be replaced and said "controller error" when they disconnected the cable. the aic was changed out and the av and lv waveform returned. there was no adverse event reported. |  |  |  |  | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Investigation has been completed. console logs from the reported day showed several occurrences of the placement signal low (event #439) and multiple positioning alarms following the adjustment of the pump position. no controller error alarms were recorded as noted in the clinical description. lt log data indicates that the snr was low for throughout the case, with varying intensity in the placement signal. the root cause for the placement signal low alarms, could not be determined as the automated impella controller (aic)was not returned. | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. the impella device was returned for evaluation. aic - controller error (yellow alarm): no controller error was observed in the data logs, and no issues were identified within the aic d9 device returned to manufacturer date added d2 common device name revised on manufacturer device report 1220648-2024-07416 in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2024-07416 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-07416 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2024-07416.
30285 | 1220648-2024-07417 | 18836255 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18836255&pc=OZD | 11/28/2022 | 03/12/2025 |  | 03/05/2024 |  |  | 11/28/2022 | 03/05/2024 |  |  | 06/30/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023237269 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/28/2022 |  |  |  |  | The user facility reported a 42-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed a thrombus in the left ventricle. physician elected to keep the impella cp support implanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Analysis summary the failure mode of ¿thrombus¿ is to be used when a thrombus is noted but cannot be connected to any other adverse events or decrease in pump performance. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions. ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency). ¿relevant laboratory test results, such as coagulation profiles and platelet counts. ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states) ¿medications or treatments that the patient was receiving at the time of thrombus formation ¿surgical or procedural details if applicable (e.g, cardiac catheterization) ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events) ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. this report is being filed as part of a retrospective review of historical records. cardiac arrythmia analysis summary: cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. ¿timing of the arrythmia event in relation to impella support (during or post-implant) ¿electrocardiogram (ekg) recordings of the arrhythmia episodes. ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. ¿response to any antiarrhythmic treatments or interventions during the event. ¿any documented device-related issues or complications during impella support. ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. ¿review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined.
30286 | 1220648-2024-07418 | 18836405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18836405&pc=OZD | 11/22/2022 | 03/12/2025 |  | 03/05/2024 |  |  | 11/22/2022 | 03/05/2024 |  |  | 08/31/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023259685 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/22/2022 |  |  |  |  | The user facility reported that a 80-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient lost complete pulsatility and pressure. the physician also discovered a perforation from initial fellow stick on left side. the patient was administered two units of packed red blood cells (prbcs). a balloon tamponade and covered stent were also placed. after treatment and revascularization, the physician elected to wean and explant the impella device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury" ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. the root cause of access site bleeding was not determined because of the limited clinical information provided and the product was not returned. it was mentioned that the there was perforation to lv, but this does not explain the cause of access site bleeding. hence the root cause is not determined.
30287 | 1220648-2024-07419 | 18836430 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18836430&pc=OZD | 12/01/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 12/01/2022 | 03/05/2024 |  |  | 07/31/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231438 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/01/2022 |  |  |  |  | The user facility reported a 72-year-old male was implanted with an impella cp support and the staff noticed a thrombus formation in the common femoral artery after the device removal. a six french (6fr) sheath was inserted in the ipsilateral, groin site and the thrombus was treated with angioplasty. the sheath was left in place with single preclose at site, to be closed later when physician is satisfied that flow is maintained to the leg. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu, these conditions are known potential adverse events associated with impella support: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ this report is being filed as part of a retrospective review of historical records.
30288 | 1220648-2024-07420 | 18836464 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18836464&pc=OZD | 09/27/2021 | 03/05/2024 |  | 03/05/2024 |  |  | 09/27/2021 | 03/05/2024 |  |  | 02/28/2023 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022046357 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/27/2021 |  |  |  |  | The user facility reported a 60-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 pump was unable to advance through the patient's small vasculature during attempted delivery. despite multiple attempts, the pump could not be placed and the decision was made to abort the implant. there was no patient harm. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition of small vessels. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
30289 | 1220648-2024-07421 | 18836497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18836497&pc=OZD | 09/07/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 09/07/2022 | 03/05/2024 |  |  | 04/30/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023219149 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/07/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 83-year-old male patient for mechanical circulatory support. it was reported while on impella cp support, the patient was experiencing a hematoma at the access site. the hematoma does not seem to be progressing, and some drainage is present at site as well. no further information has been provided. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury" this report is being filed as part of a retrospective review of historical records.
30290 | 1220648-2024-07422 | 18836540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18836540&pc=OZD | 12/10/2021 | 03/05/2024 |  | 03/05/2024 |  |  | 12/10/2021 | 03/05/2024 |  |  | 09/30/2023 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022143137 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/10/2021 |  |  |  |  | The user facility reported that the medical team experienced difficulty implanting an impella 5.5 in a 57-year-old male patient for mechanical circulatory support due to the patient's anatomy. after multiple failed attempts, the impella 5.5 implant was aborted. | [5 empty] | The investigation into the reported delivery issue has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue - anatomy was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records
30291 | 1220648-2024-07423 | 18836558 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18836558&pc=OZD | 11/22/2022 | 03/12/2025 |  | 03/05/2024 |  |  | 11/22/2022 | 03/05/2024 |  |  | 03/31/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023214346 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/22/2022 |  |  |  |  | The user facility reported a 67 year-old female in non-st elevated acute myocardial infarction/cardiogenic was implanted with an impella cp device for mechanical circulatory support. the patient experienced left ventricle thrombus. the impella cp was weaned, explanted, and exchanged for an intra-aortic balloon pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
30292 | 1220648-2024-07424 | 18836623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18836623&pc=OZD | 02/14/2024 | 03/19/2024 |  | 03/05/2024 |  |  | 02/14/2024 | 03/05/2024 |  |  |  |  | 03/05/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1669290 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/14/2024 |  |  |  |  | The user facility reported a 55-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an aic (automated impella controller) displayed a controller error alarm during ground transport above 9000 ft elevation. there was no loss of aic function or impella function at the time of the noted issue. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the console was returned and tested after arriving in fs. the optical system was tested and confirmed to be working correctly. there was no evidence of a device malfunction. the device functioned/operated to specification. even though the data logs confirmed the reported issue, the root cause is undetermined.
30293 | 1220648-2024-07425 | 18836638 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18836638&pc=OZD | 11/23/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 11/23/2022 | 03/05/2024 |  |  | 06/30/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023248983 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/23/2022 |  |  |  |  | The user facility reported an impella cp was attempted to be implanted in a male patient for mechanical circulatory support. it was reported that the medical team experienced difficulty implanting impella due to patient's anatomy. ultimately the medical team aborted the procedure. there was no patient harm reported. | [5 empty] | The investigation into the reported delivery issue has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine a root cause. the most likely cause of the delivery issue was due to patient condition. this report is being filed as part of a retrospective review of historical records.
30294 | 1220648-2024-07426 | 18836677 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18836677&pc=OZD | 02/16/2024 | 04/04/2024 |  | 03/05/2024 |  |  | 02/16/2024 | 03/05/2024 |  |  | 11/30/2025 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025436887 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/16/2024 |  |  |  |  | The user facility reported a 69-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. the implanted impella cp pump experienced saturated flows at start up. the max/min flows of the pump were both 4.3 and the ao/lv placement signals were inverted. due to the noted issue, the decision was made to replace the pump prior to the scheduled pci. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into saturated flow and access site bleeding ¿ minor issues has been completed since the initial report was submitted. the device was not returned for investigation; only the data stick was returned. based on the clinical details and data log revision, the root cause of the saturated flow cannot be determined because there was no evidence in the data logs as to what caused the motor current to produce a saturated pump flow and no product was returned for examination. the root cause of the minor access site bleeding was also unknown.
30295 | 1220648-2024-07449 | 18837527 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18837527&pc=OZD | 09/08/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 09/08/2022 | 03/05/2024 |  |  | 05/31/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231240 |  | P140003 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/08/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 58-year-old male patient for mechanical circulatory support. it was reported that there was bleeding at the insertion site. as a result of the noted issue, peel away sheath was removed, pressure was held to the site, and anticoagulation was stopped. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury" this report is being filed as part of a retrospective review of historical records.
30296 | 1220648-2024-07450 | 18837545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18837545&pc=OZD | 09/08/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 09/08/2022 | 03/05/2024 |  |  | 05/31/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023226475 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/08/2022 |  |  |  |  | The user facility reported a 70-year-old male (race unknown) in acute myocardial infarction/cardiogenic shock was to be implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 pump could not be delivered successfully due to the patient's anatomy. the implant was aborted and an impella cp pump was placed for patient support. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issue - anatomy was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
30297 | 1220648-2024-07448 | 18837563 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18837563&pc=OZD | 02/17/2024 | 04/01/2024 |  | 03/05/2024 |  |  | 02/17/2024 | 03/05/2024 |  |  | 07/31/2025 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025442178 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 27 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Alarm System | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/17/2024 |  |  |  |  | The user facility reported a 27-year-old female with cardiomyopathy was to be implanted with an impella 5.5 for mechanical circulatory support. when the impella 5.5 was plugged into the automated impella controller, an alarm occurred stating "impella defective. do not use". the impella 5.5 was replaced and support proceeded. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | H.6 type of investigation code 4114 and investigation findings code 3233 were inadvertently left off the original manufacture device report that was submitted.
30298 | 1220648-2024-07451 | 18837637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18837637&pc=OZD | 12/20/2021 | 03/05/2024 |  | 03/05/2024 |  |  | 12/20/2021 | 03/05/2024 |  |  | 09/30/2023 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022144170 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/20/2021 |  |  |  |  | The user facility reported that the medical team experienced difficulty implanting an impella 5.5 in a male patient of unknown age for mechanical circulatory support due to the patient's anatomy. after multiple failed attempts, the impella 5.5 implant was aborted. | [5 empty] | The investigation into the reported delivery issue has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue - anatomy was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
30299 | 1220648-2024-07453 | 18837644 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18837644&pc=OZD | 03/31/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 03/31/2022 | 03/05/2024 |  |  | 12/31/2023 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023176918 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 03/31/2022 |  |  |  |  | The user facility reported a 61-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported an inability to pass the impella pump through the graft to deliver the device. insertion was aborted with no harm to the patient. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not received from the customer. however, clinical details were sufficient to determine the root cause. the root cause of the delivery issue was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
30300 | 1220648-2024-07452 | 18837645 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18837645&pc=OZD | 09/08/2022 | 03/26/2024 |  | 03/05/2024 |  |  | 09/08/2022 | 03/05/2024 |  |  | 04/30/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218646 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/08/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 47-year-old male patient for mechanical circulatory support. it was reported upon placement of impella, the patient went into ventricular fibrillation and was shocked twice. patient was stabilized and impella cp was explanted. | [5 empty] | The investigation for the reported issue has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the reported issue could not be determined. | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
30301 | 1220648-2024-07455 | 18837837 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18837837&pc=OZD | 01/14/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 01/14/2022 | 03/05/2024 |  |  | 09/30/2023 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022144191 |  | 0550-0008 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/14/2022 |  |  |  |  | The user facility reported a (b)(6) female was being placed on an impella 5.5 for mechanical circulatory support. the complainant reported that an impella 5.5 pump could not be placed during attempted delivery due to the patient's anatomy. after failed attempts the medical staff decided to abort the procedure. | [5 empty] | The investigation into the reported delivery issue has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue - anatomy was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records
30302 | 1220648-2024-07454 | 18837838 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18837838&pc=OZD | 09/10/2022 | 03/12/2025 |  | 03/05/2024 |  |  | 09/10/2022 | 03/05/2024 |  |  | 06/30/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235768 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/10/2022 |  |  |  |  | The user facility reported a 65-year-old male with post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. the complainant reported the patient was on impella cp support and the patient had bleeding (oozing) at the impella cp insertion site. sheath was sutured and dressed with 4x4 and tegaderm. the impella access site was cut down and repaired with sutures to control the bleeding. the patient remained on cp support for two additional days. the pump was explanted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30303 | 1220648-2024-07456 | 18837921 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18837921&pc=OZD | 09/28/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 09/28/2022 | 03/05/2024 |  |  | 08/31/2023 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022127924 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/28/2022 |  |  |  |  | The user facility reported an impella cp on a male patient requiring mechanical circulatory support. it was reported that the patient developed ischemia on the impella side. there were additionally 2 sheaths in the arterial/impella side. a crossover bypass was not possible and therefore the patient was escalated to an impella 5.0 via axillary access. | [5 empty] | The investigation for the reported limb ischemia has been completed. the impella device was not returned for evaluation nor was sufficient clinical details provided to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
30304 | 1220648-2024-07457 | 18838052 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18838052&pc=OZD | 11/22/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 11/22/2022 | 03/05/2024 |  |  | 07/31/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023251476 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/22/2022 |  |  |  |  | The user facility reported a 57-year-old female was being placed on an impella 5.5 for mechanical circulatory support. the impella 5.5 was unable to advance through the patient¿s vasculature due to the size of their axillary artery. the implant was aborted and an impella cp was placed. | [5 empty] | The investigation into the reported delivery issue has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue - anatomy was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records
30305 | 1220648-2024-07458 | 18838067 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18838067&pc=OZD | 09/09/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 09/09/2022 | 03/05/2024 |  |  | 11/30/2023 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023161740 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | NA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient of unknown gender undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the patient developed groin complications due to low blood flow during impella cp support. the pump was surgically removed as a result of the hypoperfusion and the patient was escalated to impella 5.5 support. | [5 empty] | The investigation for the reported limb ischemia has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
30306 | 1220648-2024-07459 | 18838095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18838095&pc=OZD | 09/29/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 09/29/2022 | 03/05/2024 |  |  | 06/30/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023234660 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/29/2022 |  |  |  |  | The user facility reported a 76-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. it was reported that the impella 5.5 pump was unable to advance into the aortic arch during attempted delivery. it was noted that there was relatively severe stenosis in that portion of the artery. after multiple failed attempts, the implant was aborted and an impella cp was placed for patient support. there was no patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the delivery issue - anatomy was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
30307 | 1220648-2024-07460 | 18838118 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18838118&pc=OZD | 02/17/2024 | 04/03/2024 |  | 03/05/2024 |  |  | 02/17/2024 | 03/05/2024 |  |  | 07/31/2025 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024378410 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/17/2024 |  |  |  |  | The user facility reported a 61-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support as escalation of care from an impella cp device. while on impella 5.5 device, the patient experienced oozing from the graft site and from midline incision. additional sutures were placed and one unit of packed red blood cells was given. the patient remained on support for two additional days. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation into the access site bleeding ¿ major issue has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was determined to be use issue because after applying the additional sutures hemostasis was achieved.
30308 | 1220648-2024-07461 | 18838120 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18838120&pc=OZD | 04/19/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 04/19/2022 | 03/05/2024 |  |  | 10/31/2023 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022154341 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/19/2022 |  |  |  |  | The user facility reported a 52-year-old male with heart failure was to be implanted with an impella 5.5 for mechanical circulatory support. the medical team had difficulty delivering the pump due to the patient's anatomy. after several failed attempts, the procedure was aborted. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
30309 | 1220648-2024-07462 | 18838166 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18838166&pc=OZD | 09/29/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 09/29/2022 | 03/05/2024 |  |  | 04/30/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023216598 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  |  |  | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/29/2022 |  |  |  |  | The user facility reported an impella cp was attempted to be implanted in a 79 year-old male in acute myocardial infarction/cardiogenic shock. it was reported that the impella pump was being inserted during resuscitation, but the physician determined placement was not possible due to severe aortic stenosis. as a result, placement was aborted. the patient subsequently expired with no allegation of harm caused by device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
30310 | 1220648-2024-07463 | 18838228 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18838228&pc=OZD | 02/19/2024 | 04/09/2024 |  | 03/05/2024 |  |  | 02/19/2024 | 03/05/2024 |  |  | 07/31/2025 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025437214 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/19/2024 |  |  |  |  | The user facility reported a 48-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the implanted impella cp pump experienced high purge pressure and purge blocked alarms shortly following implant. both the purge cassette and purge bag were swapped to troubleshoot the issue, but the failure remained. as a result of the ongoing issue, the decision was made to replace the pump. upon explant, it was noted that the catheter on the impella was kinked near the motor housing. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the high purge pressure issue has been completed since the original report was filed. the pump was returned, tested, and evaluated. according to the clinical, after a couple hours of impella cp support high purge pressure and purge system blocked alarms were recorded. product evaluation confirmed evidence of a catheter kink and purge line kink. no biomaterial was found. the cause of the high purge pressure was a kink in the purge line.
30311 | 1220648-2024-07464 | 18838260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18838260&pc=OZD | 05/01/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 05/01/2022 | 03/05/2024 |  |  | 09/30/2023 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022151058 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/01/2022 |  |  |  |  | The user facility reported a 72-year-old female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the medical staff was having difficulty delivering the impella 5.5 pump due to the patient's vascular anatomy. after multiple failed attempts they aborted the impella 5.5 implant and decided to use impella cp instead, which they were able to implant successfully. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
30312 | 1220648-2024-07465 | 18838272 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18838272&pc=OZD | 10/01/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 10/01/2022 | 03/05/2024 |  |  | 06/30/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023230342 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  |  | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/01/2022 |  |  |  |  | The user facility reported a 72-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the medical staff had difficulty delivering the impella pump due to the patient's anatomy. the device was removed and flushed. impella 5.5 was inserted over the 0.18 and advanced into place without complications. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to the patient condition. this report is being filed as part of a retrospective review of historical records.
30313 | 1220648-2024-07466 | 18838281 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18838281&pc=OZD | 11/23/2022 | 03/12/2025 |  | 03/05/2024 |  |  | 11/23/2022 | 03/05/2024 |  |  | 06/30/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023249032 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/23/2022 |  |  |  |  | The user facility reported a 76-year-old male in post cardiotomy cardiogenic shock¿and with low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. during explant of the device, bleeding complications were noted at the femoral access site. vascular surgery was performed to repair the site and treat the bleed. the patient remained supported with an implanted 5.5 pump following the cp explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
30314 | 1220648-2024-07467 | 18838368 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18838368&pc=OZD | 02/26/2024 | 04/30/2024 |  | 03/05/2024 |  |  | 02/26/2024 | 03/05/2024 |  |  |  |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMEPLLA CP |  |  | ABIOMED, INC |  |  | US |  | UNKNOWN |  | UNKONWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Literature | Health Professional | User facility | Company representation |  | 62 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/26/2024 |  |  |  |  | The user facility reported information from a literature article regarding a 62-year-old male of unknown race and ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella device was removed after 3 days of support and it was noted a thrombus formation, which was thought to be induced by impella placement. a bilateral femoral artery thrombectomy was performed as treatment and the patient was reported as stable. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the thrombus issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ to conform with updated procedures, section h10 was revised with updated information.
30315 | 1220648-2024-07468 | 18838372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18838372&pc=OZD | 11/24/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 11/24/2022 | 03/05/2024 |  |  | 07/31/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023252292 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/24/2022 |  |  |  |  | The user facility reported a 78-year-old female with cardiomyopathy was to be implanted with an impella 5.5 for mechanical circulatory support. the impella 5.5 pump was unable to advance through the patient's vasculature during attempted delivery. as a result of the noted issue, the decision was made to abort the implant. there was no patient harm. | [5 empty] | The investigation into the reported delivery issue has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue - anatomy was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
30316 | 1220648-2024-07469 | 18838436 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18838436&pc=OZD | 02/22/2024 | 04/09/2024 |  | 03/05/2024 |  |  | 02/22/2024 | 03/05/2024 |  |  | 09/30/2025 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024408893 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/22/2024 |  |  |  |  | The user facility reported a 66-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient developed hemolysis during impella cp support. an echocardiogram verified positioning of the pump and discussions were initiated on escalating care to impella 5.5 support. the decision was ultimately made to withdraw care. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the reported hemolysis issue has been completed. data logs revealed pump was started and run without issue for entire case. there were some intermittent suction alarms triggered during support. product was not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review.
30317 | 1220648-2024-07470 | 18838441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18838441&pc=OZD | 02/26/2024 | 04/03/2024 |  | 03/05/2024 |  |  | 02/26/2024 | 03/05/2024 |  |  |  |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025442083 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/26/2024 |  |  |  |  | The user facility a 64-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the access site became very oozy, no hematoma was noted. systemic anticoagulation was held du to the bleeding and the patient received 1 unit of packed red blood cells. the bleeding stopped after pressure dressing applied. systemic anticoagulation resumed the next day with and no bleeding issues were noted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided.
30318 | 1220648-2024-07471 | 18838447 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18838447&pc=OZD | 09/14/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 09/14/2022 | 03/05/2024 |  |  | 03/31/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023205520 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/14/2022 |  |  |  |  | The user facility reported a 51-year-old male undergoing cardiac surgery to be implanted with an impella 5.5 device for mechanical circulatory support. during attempted implant, the guidewire met resistance and was unable to cross the left ventricle. several failed attempts were made to deliver the impella 5.5 over the steel core .18 wire and v-18. the team decided to place 0.035lunderquist in the left ventricle dipped in rota glide over v-18, which was successful. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery anatomy issue was most likely patient condition related, as a .035 lunderquist guidewire was used to insert a impella instead of 0.18 guidewire and no product was damaged as per the clinical description provided. this report is being filed as part of a retrospective review of historical records.
30319 | 1220648-2024-07472 | 18838482 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18838482&pc=OZD | 02/27/2024 | 04/17/2024 |  | 03/05/2024 |  |  | 02/27/2024 | 03/05/2024 |  |  | 10/31/2025 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024414347 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/27/2024 |  |  |  |  | The user facility reported a 76-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during impella cp placement, the cp catheter was unable to advance past iliac bifurcation. cp and wire removed with pigtail advanced to distal aorta for iliac angio. tight calcified stenosis revealed in the left iliac artery (lia). the decision was made to abort the procedure and reschedule for axillary access approach. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the clinical details states that the patient had calcified and tortuosity vessels condition. the cause of the delivery issue was patient condition related due to calcified and tortuosity vessels condition.
30320 | 1220648-2024-07473 | 18838536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18838536&pc=OZD | 02/27/2024 | 05/01/2024 |  | 03/05/2024 |  |  | 02/27/2024 | 03/05/2024 |  |  | 09/30/2025 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024398495 |  | 1000211 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/27/2024 |  |  |  |  | The user facility a 66-year-old male patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported impella 5.5 had a purge leak from around the side arm connection. the catheter was replaced. after the purge housing leaked, the motor current consumption increased, and the pump stopped. the impella was explanted. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup. | The investigation into the purge sidearm leak and pump stop has been completed. the data logs confirmed high motor current and low purge pressure as reported. followed by an increase in motor current and a high motor current pump stop . the product was inspected and purge leak issues were replicated where a crack along the yellow luer was found. inspection also showed blood ingress in the purge line confirming reverse flow. there was also biomaterial in the purge gap. the root cause of the pump stop was determined to be a damaged sidearm yellow luer. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ to conform with updated procedures, sections d6 & h10 have been revised with updated information.
30321 | 1220648-2024-07474 | 18838573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18838573&pc=OZD | 02/29/2024 | 04/01/2024 |  | 03/05/2024 |  |  | 02/29/2024 | 03/05/2024 |  |  | 11/30/2025 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025438878 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/29/2024 |  |  |  |  | The user facility reported a 65-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had placed a cp to support a patient admitted in cgs (cardiogenic shock). the patient had cad (coronary artery disease) and was in shock, and after 1 day of support the cp was explanted. the cp had signs of hemolysis. the cp was explanted and replaced by a 5.5 pump/escalation of care. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the hemolysis has been completed since the initial report was submitted. the product was not returned for investigation. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review.
30322 | 1220648-2024-07475 | 18838583 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18838583&pc=OZD | 02/23/2024 | 04/03/2024 |  | 03/05/2024 |  |  | 02/23/2024 | 03/05/2024 |  |  | 11/30/2025 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024425041 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/23/2024 |  |  |  |  | The user facility reported a 47-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported a hematoma at the impella insertion site. hematoma is resolving with sandbag and heparin was on hold for a day. the patient is stable. | [5 empty] | The investigation into the access site bleeding ¿ major issue has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was most likely due to anticoagulation management since the hemostasis was achieved and bleed resolved by turning off the heparin. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
30323 | 1220648-2024-07476 | 18838625 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18838625&pc=OZD | 09/16/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 09/16/2022 | 03/05/2024 |  |  | 03/31/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023208296 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 12 YR |  | Female |  |  |  | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/16/2022 |  |  |  |  | The user facility reported a 12-year-old female with cardiomyopathy was to be implanted with an impella 5.5 for mechanical circulatory support. the medical team was having difficulty delivering an impella 5.5 pump due to the patient's left axillary artery being too small. the decision was made to go with impella cp instead. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issue - anatomy was most likely due to the patient¿s condition. this report is being filed as part of a retrospective review of historical records.
30324 | 1220648-2024-07477 | 18838666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18838666&pc=OZD | 04/05/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 04/05/2022 | 03/05/2024 |  |  | 01/31/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023181399 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/05/2022 |  |  |  |  | The user facility reported a (b)(6)male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that a patient on impella cp was being transitioned to impella 5.5. the medical staff had difficulty inserting the impella 5.5 in the right axillary artery due to the graft being anastomosed. after multiple failed attempts to insert the pump through the right axillary they decided to use the left axillary artery and they were able to deliver the pump successfully. there was no harm to the patient as a result of this incident. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not received from the customer. however, clinical details were sufficient to determine the root cause. the root cause of the delivery issue - anatomy, was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
30325 | 1220648-2024-07478 | 18838718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18838718&pc=OZD | 12/02/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 12/02/2022 | 03/05/2024 |  |  | 04/30/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213479 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/02/2022 |  |  |  |  | The user facility reported a 59-year-old-male patient on impella cp support. it was reported that the patient experienced bleeding from the venous sheath with slight oozing from repositioning sheath on impella. systemic heparin was turned off due to bleeding and hematoma the site. activated clotting time was greater than 375. the pump stopped and was removed. the patient continued to decline so a new pump was used. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ this report is being filed as part of a retrospective review of historical records.
30326 | 1220648-2024-07479 | 18838809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18838809&pc=OZD | 12/02/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 12/02/2022 | 03/05/2024 |  |  | 04/30/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023223902 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 12/02/2022 |  |  |  |  | The user facility reported that after the patient was placed on impella cp support, it was removed due to limb ischemia. | [5 empty] | The investigation into the limb ischemia has been completed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
30327 | 1220648-2024-07480 | 18838825 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18838825&pc=OZD | 09/22/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 09/22/2022 | 03/05/2024 |  |  | 06/30/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023237564 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/22/2022 |  |  |  |  | The user facility reported a 61-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support, and the patient experienced excessive bleeding at the impella site. bicarb was added to the purge. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury" this report is being filed as part of a retrospective review of historical records.
30328 | 1220648-2024-07481 | 18838878 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18838878&pc=OZD | 02/23/2024 | 09/12/2024 |  | 03/05/2024 |  |  | 02/23/2024 | 03/05/2024 |  |  | 11/30/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024303544 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/23/2024 |  |  |  |  | The user facility reported a(b)(6) male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the us complainant had placed an impella cp to support a patient with multiple cardiac and systemic comorbidities. the patient was deemed too critical for surgery. the patient had extensive thoracic and aortic dissection history. after the pump was delivered, the pump positioning was not optimal. the medical doctor was able to cross the valve and deliver. the patient repositioned multiple times but, the patient decompensated, coded, and expired. there is no allegation made that the use of impella contributed to the death. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed since the original report was submitted. the device was not returned for investigation. the cause of the positioning issue most likely was use related due to improper technique, per the data log review. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr1220648-2024-07481 was provided in sections b2, b5, h1, and h6.
30329 | 1220648-2024-07482 | 18838887 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18838887&pc=OZD | 04/18/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 04/18/2022 | 03/05/2024 |  |  | 12/31/2023 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023180819 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/18/2022 |  |  |  |  | The user facility reported a 65-year-old male patient on impella cp for mechanical circulatory support. it was reported that the patient experienced a bleed out of the impella site at the area of repo sheath insertion. the physician repositioned the repo sheath. the bleeding was resolved, and dressing was applied. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿ potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records.
30330 | 1220648-2024-07483 | 18838938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18838938&pc=OZD | 12/02/2022 | 05/19/2025 |  | 03/05/2024 |  |  | 12/02/2022 | 03/05/2024 |  |  | 06/30/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023253011 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Other | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/02/2022 |  |  |  |  | The user facility reported a 72-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during suturing the pump alerted to suction alarms and pulsatility was lost. right ventricle function rapidly decompensated on echo. heart failure occurred which was consulted and no further support was suggested. the patient coded and the family agreed to no further support, therefore, the impella was disconnected. unfortunately, the patient expired 1.47 hours of impella support. no allegations were made towards the impella for cause of death. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation into the limb ischemia has been completed. the device was not returned nor was sufficient clinical information available. therefore, the root cause of the limb ischemia could not be determined. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07483 was provided. | Limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition any concomitant medication vascular size or variations thereof detailed description of the symptoms experienced by the patient. physical examination findings, including pulse assessment and visual examination of the affected limb. diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. laboratory test results, including blood tests for markers of inflammation or clotting disorders. any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
30331 | 1220648-2024-07484 | 18838969 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18838969&pc=OZD | 11/29/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 11/29/2022 | 03/05/2024 |  |  | 06/30/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023247477 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/29/2022 |  |  |  |  | The user facility reported a 76-year-old female was being placed on an impella 5.5 for mechanical circulatory support. the complainant reported that the impella 5.5 pump was unable to advance through the patient's vasculature during attempted implant. after multiple unsuccessful attempts utilizing different techniques and wires, the decision was made to abort the implant. there was no patient harm. | [5 empty] | The investigation into the reported delivery issue has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue - anatomy was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
30332 | 1220648-2024-07485 | 18839096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18839096&pc=OZD | 02/24/2024 | 04/12/2024 |  | 03/05/2024 |  |  | 02/24/2024 | 03/05/2024 |  |  |  |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025442091 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/24/2024 |  |  |  |  | The user facility reported a 59-year-old female undergoing cardiac surgery was being implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that complications were encountered during the attempted delivery of an impella 5.5 pump. despite multiple attempts, it was documented that the impella was unable to advance into the left ventricle due to the tortuous axillary/subclavian artery. the decision was subsequently made to abort the implant. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, impella was unable to advance into the left ventricle due to the tortuous axillary/subclavian artery as per the clinical description provided. added- e1 facility name, h6 impact code 4627.
30333 | 1220648-2024-07486 | 18839134 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18839134&pc=OZD | 12/02/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 12/02/2022 | 03/05/2024 |  |  | 12/31/2023 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023174932 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/02/2022 |  |  |  |  | The user facility reported a 75-year-old male in high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp could not be inserted as intended into the femoral access site. however, the impella was able to be inserted into the right axillary/subclavian artery. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not received from the customer. however, clinical details were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, the impella was unable to pass through the vasculature due to patient's anatomy. this report is being filed as part of a retrospective review of historical records.
30334 | 1220648-2024-07487 | 18839301 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18839301&pc=OZD | 05/09/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 05/09/2022 | 03/05/2024 |  |  | 02/29/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023195764 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/09/2022 |  |  |  |  | The user facility reported a 72-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the medical staff was having difficulty delivering the impella 5.5 pump due to the patient's vascular anatomy. after failed attempts, the procedure was aborted. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
30335 | 1220648-2024-07488 | 18839392 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18839392&pc=OZD | 12/02/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 12/02/2022 | 03/05/2024 |  |  | 05/31/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023227104 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/02/2022 |  |  |  |  | The user facility reported a 79- year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella pump was unable to be advanced through the anterior iliac due to high degree of stenosis. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not received from the customer. however, clinical details were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, the impella was unable to pass through the vasculature due to patient's anatomy. this report is being filed as part of a retrospective review of historical records.
30336 | 1220648-2024-07490 | 18839492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18839492&pc=OZD | 12/02/2022 | 05/16/2025 |  | 03/05/2024 |  |  | 12/02/2022 | 03/05/2024 |  |  | 03/31/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023205539 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/02/2022 |  |  |  |  | The user facility reported a 68-year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced bleeding at the access site. access site required additional sutures to resolve bleeding. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | The cause of the major access site bleed was most likely due to lack of vessel recoil as the bleeding occurred after insertion of the gwru. this report is being filed as part of a retrospective review of historical records.
30337 | 1220648-2024-07489 | 18839494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18839494&pc=OZD | 02/26/2024 | 04/12/2024 |  | 03/05/2024 |  |  | 02/26/2024 | 03/05/2024 |  |  | 04/30/2025 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024330381 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Material Deformation | Material Twisted/Bent | Material Split, Cut or Torn |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/26/2024 |  |  |  |  | The user facility implanted a 33-year-old female patient undergoing coronary artery bypass grafting with an impella 5.5 device for mechanical circulatory support. it was reported during the implant the guidewire got caught on the pigtail of the catheter. the physician was able to get the wire and catheter separated and used a v-18 wire to advance the impella 5.5 over and into the left ventricle and the impella 5.5 was successfully placed. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed since the original report was filed. the pump was returned, tested, and evaluated. visually inspected the guidewire and several guidewire kinks and partial fracture observed due to excessive force during insertion. as per clinical, guidewire was getting stuck within pump and likely damaged when trying to advance at aortic valve. the cause of the product damage (guidewire damage- great vessel) issue was use related due to excessive force being applied during support.
30338 | 1220648-2024-07491 | 18839506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18839506&pc=OZD | 09/23/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 09/23/2022 | 03/05/2024 |  |  | 01/31/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023188104 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/23/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a female patient for mechanical circulatory support and the patient¿s leg mottled and got pale. the patient lost complete pulsatility when the physician was engaging as well as when stenting. the impella was removed with 1 perclose, and the leg quickly regained color. | [5 empty] | The investigation for the reported limb ischemia has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
30339 | 1220648-2024-07492 | 18839553 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18839553&pc=OZD | 12/02/2022 | 03/07/2025 |  | 03/05/2024 |  |  | 12/02/2022 | 03/05/2024 |  |  | 07/31/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023249617 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/02/2022 |  |  |  |  | The user facility reported a 73-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. upon repositioning of the pump, plasma-free hemoglobin (pfhgb) followed causing the urinary output(uop) to become red. the decision was to reposition the impella a second time; however, plasma-free hemoglobin (pfhgb) appeared again. it was advised to reposition the pump once more if plasma-free hemoglobin (pfhgb) are still elevated. the patient urinary output(uop) decreased with continue pink urine . | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30340 | 1220648-2024-07494 | 18839713 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18839713&pc=OZD | 05/13/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 05/13/2022 | 03/05/2024 |  |  | 01/31/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023186165 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/13/2022 |  |  |  |  | The user facility reported a (b)(6) male was to be implanted with an impella 5.5 for mechanical circulatory support as a bridge to left ventricle assist device. pre-op transesophageal echocardiography showed heavily calcified acute aortic occlusion and arch. several attempts at changing the cannula angulation and altering the direction of graft were attempted. a decision was made to switch to an impella cp. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
30341 | 1220648-2024-07495 | 18839762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18839762&pc=OZD | 12/04/2022 | 05/30/2025 |  | 03/05/2024 |  |  | 12/04/2022 | 03/05/2024 |  |  | 07/31/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023254933 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Tachycardia | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/04/2022 |  |  |  |  | The user facility reported a (b)(6) male in st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. after being placed on impella cp support, the patient developed a large hematoma that was actively growing in the intensive care unit. the physician pushed the hematoma out, resutured the blue hub, and redressed the site, to improve the ability to maintain the insertion angle. the hematoma resolved. it was also reported while on impella cp support, the patient experienced ventricular tachycardia (vt) arrest multiple times and the patient was placed on a do not resuscitate order. the patient had another vt event and the patient expired. no allegations were made towards the impella cp device for cause of death. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07495 was provided in sections b2, b5, h1, and h6. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined.
30342 | 1220648-2024-07496 | 18839846 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18839846&pc=OZD | 02/26/2024 | 04/16/2024 |  | 03/05/2024 |  |  | 02/26/2024 | 03/05/2024 |  |  |  |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025442038 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/26/2024 |  |  |  |  | The user facility a 84-year-old female patient was to be implanted with an impella 5.5 device for mechanical circulatory support. it was reported that placement of impella 5.5 was attempted but the impella would not cross into the aorta. the impella 5.5 was removed and an impella cp was placed via axillary approach so the patient could continue to be hemodynamic supported. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition since the pump could not cross the aorta as vessels were pad/pvd.
30343 | 1220648-2024-07497 | 18839892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18839892&pc=OZD | 12/02/2022 | 03/07/2025 |  | 03/05/2024 |  |  | 12/02/2022 | 03/05/2024 |  |  | 08/31/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023270553 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/02/2022 |  |  |  |  | The user facility reported a 77-year-old male in post cardiotomy cardiogenic shock /low cardiac output syndrome was implanted with an impella cp device for mechanical circulatory support. once the pump was implanted, the patient started to have runs of ventricular tachycardia (vt) after the wire accessed into the left ventricle; which caused a rapid decreased in pressure. the patient was subsequently stabilized with pressors and the impella. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30344 | 1220648-2024-07498 | 18839925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18839925&pc=OZD | 09/23/2022 | 05/20/2025 |  | 03/05/2024 |  |  | 09/23/2022 | 03/05/2024 |  |  | 03/31/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023214762 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Disability | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 09/23/2022 |  |  |  |  | The user facility reported a 75-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. while on impella cp support, there was no pedal pulse on the right foot. arterial doppler was done, which showed monophasic flow, but nurse states provider reported no clots or obstruction. the patient was on multiple pressers and bicarb drops with pressures labile. the legs are both cool and mild mottling from the knee down. the patient continued on support. | Additional information noted care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation for the reported limb ischemia has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07498 was provided in sections b2, b5, h1, and h6. | Limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition. any concomitant medication. vascular size or variations thereof. detailed description of the symptoms experienced by the patient. physical examination findings, including pulse assessment and visual examination of the affected limb. diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. laboratory test results, including blood tests for markers of inflammation or clotting disorders. any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
30345 | 1220648-2024-07499 | 18839964 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18839964&pc=OZD | 10/06/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 10/06/2022 | 03/05/2024 |  |  | 06/30/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023229769 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/06/2022 |  |  |  |  | The user facility reported a 63-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient experienced ischemia. the complication was managed with reperfusion line from venoarterial extracorporeal membrane oxygenation. | [5 empty] | The investigation for the reported limb ischemia has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
30346 | 1220648-2024-07500 | 18840070 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18840070&pc=OZD | 12/02/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 12/02/2022 | 03/05/2024 |  |  | 09/30/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023261829 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/02/2022 |  |  |  |  | The user facility reported a 62-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. an attempt was made to escalate to the impella 5.5; however, the procedure was aborted due to vessel diameter being too small. the physician switched to the impella cp axillary via axillary access. there was no reported patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was due to patient condition. this report is being filed as part of a retrospective review of historical records.
30347 | 1220648-2024-07501 | 18840316 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18840316&pc=OZD | 02/27/2024 | 04/17/2024 |  | 03/05/2024 |  |  | 02/27/2024 | 03/05/2024 |  |  | 08/31/2025 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024391694 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | Material Twisted/Bent |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/27/2024 |  |  |  |  | The user facility reported a 73-year-old female undergoing cardiac surgery was to be implanted with an impella 5.5 device for mechanical circulatory support. it was reported that while attempting to implant the patient there were delivery issues. the impella 5.5 would not pass through axillary artery due to patient's anatomy. | [5 empty] | The impella device was not received from the customer and therefore.an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the delivery issues has been completed. the pump was returned and a cannula kink was observed. the root cause of the delivery issue was most likely patient condition related, the impella was unable to pass through due to patients anatomy as per the clinical description.
30348 | 1220648-2024-07503 | 18840415 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18840415&pc=OZD | 12/03/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 12/03/2022 | 03/05/2024 |  |  | 06/30/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023247257 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/03/2022 |  |  |  |  | The user facility reported an (b)(6)male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced worsening symptoms of cold feet and absent pulses. the impella was successfully weaned and explanted in the operating room by vascular surgery. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ "assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records.
30349 | 1220648-2024-07504 | 18840723 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18840723&pc=OZD | 02/27/2024 | 01/08/2025 |  | 03/05/2024 |  |  | 02/27/2024 | 03/05/2024 |  |  |  |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1163567 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/27/2024 |  |  |  |  | The biomedical department of the user facility observed the automated impella controller had a red alarm/battery failure. there was no noted patient use at the time of the failure. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - battery failure (red alarm) has been completed. the aic logs confirmed the reported failure. the failure was reproduced on start up in the lab. analysis of the battery through a battttest revealed that cell 4 of battery 1 was at a significantly lower voltage that the rest of the cells in battery. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-07504.
30350 | 1220648-2024-07506 | 18840876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18840876&pc=OZD | 12/05/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 12/05/2022 | 03/05/2024 |  |  | 08/31/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023269645 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/05/2022 |  |  |  |  | The user facility reported a 63-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. it was reported when first attempting to place the impella cp, the right femoral artery was totally occluded and the impella cp could not be placed. the perclose sutures and impella sheath and dilator had already been placed and had to be removed. the impella cp was then successfully placed through the left femoral artery. | [5 empty] | The investigation into the reported delivery issue has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related. this report is being filed as part of a retrospective review of historical records
30351 | 1220648-2024-07507 | 18840968 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18840968&pc=OZD | 02/29/2024 | 04/18/2024 |  | 03/05/2024 |  |  | 02/29/2024 | 03/05/2024 |  |  |  |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025437459 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/29/2024 |  |  |  |  | The user facility reported a (b)(6) male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the staff noticed dripping from purge sidearm, apparently coming from reservoir. it appeared around 2/3 of the purge volume was leaking out. the impella cp was exchanged for the impella 5.5. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation for the reported pump leak has been completed. no product was returned for evaluation. clinical details noted that a leak was observed from purge reservoir with purge fluid leaking out. purge flow and pressure were maintained at normal values. the root cause of the purge leak was most likely a damaged sidearm reservoir as clinical details noted a leak was coming from the sidearm reservoir. the reason for the leak is unknown as no product was returned.
30352 | 1220648-2024-07505 | 18841200 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18841200&pc=OZD | 09/23/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 09/23/2022 | 03/05/2024 |  |  | 06/30/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023246491 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/23/2022 |  |  |  |  | The user facility reported a (b)(6) male implanted with an impella cp for percutaneous coronary intervention (pci). after being placed on impella cp support, the patient went into ventricular fibrillation requiring defibrillation. when preparing to exchange the peel-away for the repo sheath, a hematoma was noted following the exchange as well as a hematoma to the left femoral artery (lfa) access site. manual pressure was applied to both with little improvement in the sites. the sheaths were secured and dressed, and the patient was transferred urgently to the or for a coronary artery bypass grafting (cabg). the patient did require blood transfusions during the surgery. at the end of the case the groins were assessed and were thought to be stable and no longer bleeding or growing, so no exploration was performed. systemic heparin was withheld due to bleeding. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
30353 | 1220648-2024-07510 | 18841331 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18841331&pc=OZD | 10/11/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 10/11/2022 | 03/05/2024 |  |  | 07/31/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023250319 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/11/2022 |  |  |  |  | The user facility reported a 47-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced a large hematoma at the access site during pump explant. a surgical cut-down, repair, and closure were performed on the site as a result of the condition. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. the cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical information was provided. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
30354 | 1220648-2024-07508 | 18841351 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18841351&pc=OZD | 10/06/2022 | 10/17/2025 |  | 03/05/2024 |  |  | 10/09/2022 | 03/05/2024 |  |  | 08/31/2023 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022128037 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/09/2022 |  |  |  |  | The user facility reported a 46-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding at the access site. the dressing was changed and the angle of the repositioning sheath was elevated to troubleshoot the issue. femostop was also placed and the bleeding resolved. subsequently, during explanation/escalation of the patient from impella cp to impella 5.5 support, the patient experienced more bleeding at the access site. attempts to close the groin were unsuccessful and slight hemostasis was only achieved when a 16 fr gore sheath was placed. the decision was made to perform a surgical cutdown to effectively close the femoral artery. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | B1 product problem was erroneously selected on the initial manufacturer device report number 1220648-2024-07508. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2024-07508. d6a and d6b revised from the initial manufacturer device report 1220648-2024-07508 in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2024-07508 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-07508 in accordance with updated procedures. g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2024-07508. h5 was inadvertently omitted on the initial manufacturer device report number 1220648-2024-07508. h6 component code revised from the initial submission in accordance with updated procedures. h6 investigation conclusions code 11 was erroneously entered on the initial manufacturer device report number 1220648-2024-07508.
30355 | 1220648-2024-07509 | 18841405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18841405&pc=OZD | 09/24/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 09/24/2022 | 03/05/2024 |  |  | 06/30/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023247916 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Disability | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/24/2022 |  |  |  |  | The user facility reported a 59-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. while on impella cp support, the patient¿s right leg was pulseless and cold. distal flow management was performed. the patient continued on support. | [5 empty] | The investigation for the reported limb ischemia issues has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
30356 | 1220648-2024-07511 | 18841445 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18841445&pc=OZD | 05/19/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 05/19/2022 | 03/05/2024 |  |  | 11/30/2023 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023162174 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 18 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/19/2022 |  |  |  |  | The user facility reported a 18-year-old male with congenital heart disease, undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 insertion was aborted due to the inadequate size of the subclavian and the tortuosity and unique anatomy of the patient. a decision was made that although very close to a successful implant, it was in the best interest of the patient to abort the case, remove the 5.5 from the body and leave the existing impella cp in place for now. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
30357 | 1220648-2024-07513 | 18841486 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18841486&pc=OZD | 09/26/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 09/26/2022 | 03/05/2024 |  |  | 05/31/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023230100 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/26/2022 |  |  |  |  | The user facility reported a 74-year-old female patient was implanted with an impella cp. the patient developed groin site bleeding. pressure was held for forty-five (45) minutes, but the site continued to bleed. the decision was made to cut the sutures and redo. this controlled the bleed and no issues were noted afterwards. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
30358 | 1220648-2024-07502 | 18841544 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18841544&pc=OZD | 12/03/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 12/03/2022 | 03/05/2024 |  |  | 06/30/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023236838 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 72 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/03/2022 |  |  |  |  | The user facility reported a 72-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced episodes of ventricular fibrillation with rapid loss of hemodynamics. the patient was shocked and administered epinephrine and an amiodarone bolus. the patient stabilized. subsequently, the patient developed an infection and possible sepsis. the physician elected to explant the impella. | [5 empty] | The impella device was not received from the customer, therefore investigation of the device was not possible. should additional information be provided in the future, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿infusion filter the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ this report is being filed as part of a retrospective review of historical records.
30359 | 1220648-2024-07514 | 18841643 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18841643&pc=OZD | 10/11/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 10/11/2022 | 03/05/2024 |  |  | 03/31/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218726 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Female |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/11/2022 |  |  |  |  | The user facility reported a 71-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient experienced significant bleeding at the access site. the patient was in operating room at the time of explant / pump removal. no further information was provided. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
30360 | 1220648-2024-07515 | 18841789 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18841789&pc=OZD | 05/21/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 05/21/2022 | 03/05/2024 |  |  | 02/29/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023190579 |  | 0550-0008 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/21/2022 |  |  |  |  | The user facility reported a 60-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an impella 5.5 could not be placed during attempted delivery due to the patient's anatomy. there was no patient harm and support continued with the already implanted impella cp pump. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
30361 | 1220648-2024-07516 | 18841911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18841911&pc=OZD | 05/25/2022 | 09/11/2024 |  | 03/05/2024 |  |  | 05/25/2022 | 03/05/2024 |  |  | 02/29/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2023195174 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 05/25/2022 |  |  |  |  | The user facility reported a 59-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that while on the impella cp, the patient experienced oozing from both the extracorporeal membrane oxygenation (ecmo) and impella cp site and was required to receive 1 unit of blood. | It was reported that the patient survived impella explant, however the procedural outcome is expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The investigation for the reported limb ischemia has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07516 was provided in sections b5, h1, and h6.
30362 | 1220648-2024-07517 | 18841925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18841925&pc=OZD | 12/04/2022 | 09/21/2024 |  | 03/05/2024 |  |  | 12/04/2022 | 03/05/2024 |  |  | 08/31/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023263831 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/04/2022 |  |  |  |  | The user facility reported a (b)(6) male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that the medical team attempting the impella implant had difficulty with delivery due the staff accidentally pulling wire access port and access was lost. | The pump was removed for bleeding issues and swapped to the right femoral. during placement in right femoral, access was lost when they pulled the perclose by accident. insertion into right femoral was then aborted because patient was stable. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted the use of the device cannot conclusively be associated with the outcome. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrected information for the initial mfr 1220648-2024-07517 was provided in sections b1, b2, b5, d6a, d6b, h1, and h6.
30363 | 1220648-2024-07512 | 18841933 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18841933&pc=OZD | 12/03/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 12/03/2022 | 03/05/2024 |  |  | 07/31/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023253687 |  | 0048-0045 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/03/2022 |  |  |  |  | The user facility reported an (b)(6) male requiring an impella cp device for mechanical circulatory support as a bridge to left ventricular assist device (lvad). it was reported that the patient developed hemolysis after having suction events due to the deep position of the impella. the impella was repositioned, and the hemolysis started to clear. | [5 empty] | The impella device was not received from the customer, therefore investigation of the device was not possible. should additional information be provided in the future, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
30364 | 1220648-2024-07518 | 18842042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18842042&pc=OZD | 10/21/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 10/21/2022 | 03/05/2024 |  |  | 06/30/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023233382 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/21/2022 |  |  |  |  | The user facility reported a 60-year-old male with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. the medical staff had difficulty delivering an impella pump due to the patient's anatomy. after several attempts they were able to deliver the pump. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition. this report is being filed as part of a retrospective review of historical records.
30365 | 1220648-2024-07519 | 18842195 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18842195&pc=OZD | 12/07/2022 | 03/12/2025 |  | 03/05/2024 |  |  | 12/07/2022 | 03/05/2024 |  |  | 08/31/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023266428 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 12/07/2022 |  |  |  |  | The user facility reported a 74-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding from the access site. femo stop and manual pressure did not resolve the bleeding. cp support was weaned and the device was explanted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. the cause of the minor access site bleed cannot be determined due to insufficient clinical information provided and lack of product returned.
30366 | 1220648-2024-07520 | 18842242 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18842242&pc=OZD | 12/08/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 12/08/2022 | 03/05/2024 |  |  | 09/30/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023273568 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/08/2022 |  |  |  |  | The user facility reported a 65-year-old male with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the impella was unable to be advanced via the left axillary site; therefore, it was placed without complications in the left femoral site where the intra aortic balloon pump was placed. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the delivery issue was due to patient condition. this report is being filed as part of a retrospective review of historical records.
30367 | 1220648-2024-07521 | 18842245 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18842245&pc=OZD | 07/06/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 07/06/2022 | 03/05/2024 |  |  | 05/31/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023217661 |  | 0550-0008 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/06/2022 |  |  |  |  | The user facility reported a (b)(6) female undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the medical team had difficulty delivering the pump due to the patient's vessels being too small. after failed attempts the procedure was aborted. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
30368 | 1220648-2024-07522 | 18842341 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18842341&pc=OZD | 08/06/2022 | 05/29/2025 |  | 03/05/2024 |  |  | 08/06/2022 | 03/05/2024 |  |  | 10/31/2023 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022152710 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  |  |  | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 48-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient developed hemolysis during impella cp support. repositioning was attempted multiple times, however the condition remained. due to the hemolysis and overall condition of the patient, the decision was made to escalate to an impella 5.5 pump for ongoing support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. b1 corrected from product problem to adverse event b2 changed required intervention to prevent permanent impairment/damage (devices) from no to yes. d6: added implant/explant dates.
30369 | 1220648-2024-07523 | 18842464 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18842464&pc=OZD | 08/06/2022 | 03/12/2025 |  | 03/05/2024 |  |  |  | 03/05/2024 |  |  | 12/31/2023 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2023215523 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 08/06/2022 |  |  |  |  | The user facility reported a 59-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported a hematoma developing after advancing the impella's repositioning sheath. manual pressure was held, however, the pump was ultimately explanted and a new 14fr was inserted with plans to place a new impella cp. imagery found ischemia and a vascular tear. vascular surgery was performed, in response. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30370 | 1220648-2024-07524 | 18842599 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18842599&pc=OZD | 08/16/2022 | 03/05/2024 |  | 03/05/2024 |  |  |  | 03/05/2024 |  |  | 05/31/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023217658 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/16/2022 |  |  |  |  | The user facility reported a 69-year-old male noted as high risk percutaneous coronary intervention was implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 pump was unable to advanced through the anastomosis. the implant was aborted due to the patient's anatomy and the decision was made to implant an impella cp for patient support. there was no patient harm. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
30371 | 1220648-2024-07525 | 18842698 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18842698&pc=OZD | 08/22/2022 | 03/05/2024 |  | 03/05/2024 |  |  |  | 03/05/2024 |  |  | 09/30/2023 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022143587 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/22/2022 |  |  |  |  | The user facility reported a 49-year-old male in cardiac arrest implanted with an impella cp for mechanical circulatory support. the patient was being escalated to the impella 5.5. the medical staff was having difficulty delivering the impella 5.5 pump due to the patient's anatomical limitations. after multiple failed attempts, the procedure was aborted. the impella cp remained as support for the patient. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition of small vessels. this report is being filed as part of a retrospective review of historical records.
30372 | 1220648-2024-07526 | 18842722 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18842722&pc=OZD | 08/29/2022 | 03/05/2024 |  | 03/05/2024 |  |  | 08/29/2022 | 03/05/2024 |  |  | 05/31/2024 |  | 03/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2023223534 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/29/2022 |  |  |  |  | The user facility reported a (b)(6)male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the medical staff was having difficulty delivering the impella pump due to the patient's small axillary artery. after failed attempts they decided to abort the procedure. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition of a small axillary artery. this report is being filed as part of a retrospective review of historical records.
30373 | 1220648-2024-07311 | 18825067 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18825067&pc=OZD | 11/13/2022 | 03/04/2024 |  | 03/04/2024 |  |  | 11/13/2022 | 03/04/2024 |  |  | 06/30/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023237105 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/13/2022 |  |  |  |  | The user facility reported a 73-year-old male noted as post cardiotomy cardiogenic shock/low cardiac output syndrome was attempted to be implanted with an impella 5.5 device for mechanical circulatory support. it was reported that during attempted implant of an impella 5.5 pump, the guidewire met resistance and was unable to cross the aortic valve. the wire was removed and a .035¿ j wire was inserted; however, it bent when resistance was met. finally, a new .035¿ j wire was used in conjunction with a 6 fr angled pigtail to cross into the ventricle. the j wire was then exchanged for a boston scientific .018¿ platinum plus wire and the pump was implanted successfully. the bent .035" j wire is not an abiomed product. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. however, the clinical report has been reviewed. the root cause of the delivery issue was most likely patient condition related, as the guidewire was unable to pass through the vasculature as per the clinical description provided. this report is being filed as part of a retrospective review of historical records.
30374 | 1220648-2024-07312 | 18825187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18825187&pc=OZD | 03/18/2023 | 03/25/2024 |  | 03/04/2024 |  |  | 02/03/2024 | 03/04/2024 |  |  | 04/30/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 2023207390 |  | 005062-RO |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/03/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry indicating a 73 year old male patient presenting with acute myocardial infarction and cardiogenic shock had endured transcaval access site bleeding occurred with a probable relationship to the impella device used on this patient. a blood transfusion was performed and an iv of protonix was delivered. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the access site bleeding issue was not determined because the product was not returned, and the clinical description had limited information.
30375 | 1220648-2024-07313 | 18825436 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18825436&pc=OZD | 11/24/2022 | 05/13/2025 |  | 03/04/2024 |  |  | 11/24/2022 | 03/04/2024 |  |  | 08/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023266112 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/24/2022 |  |  |  |  | The user facility reported a 75-year-old male patient on an impella cp for mechanical circulatory support. while the patient was on impella cp support, the patient was going in and out of ventricular fibrillation requiring defibrillator shocks. the impella was found to be caught in the papillary muscle. while attempting to reposition the device, the impella cp was pulled back across the atrioventricular (av). the impella cp was reinserted. no patient harm reported. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records | Section d6a / d6b: added implant and explant dates. this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30376 | 1220648-2024-07314 | 18825549 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18825549&pc=OZD | 11/24/2022 | 11/19/2024 |  | 03/04/2024 |  |  | 11/24/2022 | 03/04/2024 |  |  | 08/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023267058 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/24/2022 |  |  |  |  | The user facility reported a 81-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. after the pump was explanted, pulses were not found in the distal lower extremity. however, on ultrasound, popliteal pulses were found. limb ischemia could not be confirmed. no additional information was provided related to intervention. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use indicate the following: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident, stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported thrombus has been completed. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
30377 | 1220648-2024-07315 | 18825553 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18825553&pc=OZD | 10/10/2022 | 10/21/2025 |  | 03/04/2024 |  |  | 10/10/2022 | 03/04/2024 |  |  | 06/30/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023248745 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/10/2022 |  |  |  |  | The user facility reported a 56-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the patient experienced a hematoma and oozing at the site. therefore, the physician decided to wean and explant the impella cp. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. the root cause of the access site bleeding cannot be determined since the product was not returned and enough clinical info was not provided.
30378 | 1220648-2024-07316 | 18825585 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18825585&pc=OZD | 10/12/2022 | 03/04/2024 |  | 03/04/2024 |  |  | 10/12/2022 | 03/04/2024 |  |  | 06/30/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023238267 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/12/2022 |  |  |  |  | The user facility reported a 77-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support, and there were concerns with pedal pulses on the implantation side. therefore, the impella cp was explanted. | [5 empty] | The investigation for the reported limb ischemia has been completed. the impella device was not received from the customer. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
30379 | 1220648-2024-07318 | 18825749 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18825749&pc=OZD | 11/24/2022 | 10/07/2025 |  | 03/04/2024 |  |  | 11/24/2022 | 03/04/2024 |  |  | 05/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023228845 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/24/2022 |  |  |  |  | The user facility reported a 31-year-old female being placed on the impella cp for mechanical circulatory support. the first attempt to place the impella cp in the cath lab failed, due to access issues. the patient was sent to the operating room for successful intra-aortic balloon pump (iabp) placement. after iabp placement, the team was successful with placement of the impella cp at the iabp site. there was no patient harm reported. | [5 empty] | The impella device has not been returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. the investigation of delivery issues has been completed. the root cause of the delivery issue was not determined due to lack of clinical details and no product returned. d6a and d6b revised from the initial manufacturer device report 1220648-2024-07318 in accordance with updated procedures. f6/f8-should have been left blank on the initial manufacturer device report number 1220648-2024-07318 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-07318 in accordance with updated procedures. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2024-07318.
30380 | 1220648-2024-07319 | 18825766 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18825766&pc=OZD | 10/11/2022 | 03/04/2024 |  | 03/04/2024 |  |  | 10/11/2022 | 03/04/2024 |  |  | 02/29/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023201436 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/11/2022 |  |  |  |  | The user facility reported a 61-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support with plans to escalate to an impella 5.5 in a day or two. after being placed on impella cp support (via right femoral artery), an external fem bypass was done to maintain flow to the lower extremity. the vessel size was marginal. the site was dressed and doppler flow was maintained. the patient was escalated to the impella 5.5 as originally planned. | [5 empty] | The investigation for the reported limb ischemia has been completed. the impella device was not received from the customer. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
30381 | 1220648-2024-07320 | 18825781 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18825781&pc=OZD | 11/14/2022 | 05/23/2025 |  | 03/04/2024 |  |  | 11/14/2022 | 03/04/2024 |  |  | 07/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023254708 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Required Intervention | Death | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/14/2022 |  |  |  |  | The user facility reported a 84-year-old male was implanted with an impella cp requiring mechanical circulatory support. during support the patient experienced a brief period of loss of preload with suction and ventricular tachycardia (vt). the patient received amiodarone and vt was resolved. | Despite stabilization following the initial episode of ventricular tachycardia, the patient was reported to have passed away on support the following day. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device has not been returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07320 was provided in sections b2, b5, d6, h1, and h6. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30382 | 1220648-2024-07321 | 18826010 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18826010&pc=OZD | 11/14/2022 | 09/17/2025 |  | 03/04/2024 |  |  | 11/14/2022 | 03/04/2024 |  |  | 03/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023197406 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  |  |  | Malfunction | No | Yes | Difficult to Insert | Failure to Deliver | Failure to Advance | Patient-Device Incompatibility |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/14/2022 |  |  |  |  | The user facility reported a 59-year-old male in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella cp device for mechanical circulatory support. it was reported that insertion was aborted due to patients' vascular anatomy. no additional information was provided. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation has been completed. no product was returned. per the clinical investigation, the root cause of delivery issue was determined to be patient condition as the pump delivery was aborted due to the patient's anatomy. this report is being filed as part of a retrospective review of historical records. d3 country was left blank on the initial manufacture report. d6a and d6b revised from the initial submission in accordance with updated procedures. g1 reporting contact first name revised on manufacturer device report in accordance with updated procedures.
30383 | 1220648-2024-07322 | 18826149 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18826149&pc=OZD | 11/25/2022 | 03/04/2024 |  | 03/04/2024 |  |  | 11/25/2022 | 03/04/2024 |  |  | 04/30/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213451 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/25/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a in 56-year-old male patient for mechanical circulatory support. it was reported that during the impella support, the patient experienced oozing at the impella access. the team was able to place sutures and the bleeding resolved. additionally, it was determined under echocardiogram that the impella cp was rotated toward the mitral valve. the physician tried to reposition the impella cp but was unsuccessful. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury." this report is being filed as part of a retrospective review of historical records.
30384 | 1220648-2024-07323 | 18826197 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18826197&pc=OZD | 02/07/2024 | 03/26/2024 |  | 03/04/2024 |  |  | 02/07/2024 | 03/04/2024 |  |  | 11/30/2025 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024422158 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/07/2024 |  |  |  |  | The user facility reported a 68-year-old female with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that during patient transport, the impella was unplugged for 10 seconds to change the console and the patient became unresponsive. the patient recovered was as reported stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the temporary pump stop issue has been completed since the initial report was submitted. the product was not returned for investigation. there was no issue found during the case. the device functioned to specification.
30385 | 1220648-2024-07324 | 18826203 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18826203&pc=OZD | 02/23/2024 | 04/10/2024 |  | 03/04/2024 |  |  | 02/23/2024 | 03/04/2024 |  |  | 06/30/2025 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024373721 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/23/2024 |  |  |  |  | The user facility reported a 83-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient on impella support the right lower extremity became mottled. the decision was made to wean and explant impella. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined.
30386 | 1220648-2024-07325 | 18826238 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18826238&pc=OZD | 02/23/2024 | 04/04/2024 |  | 03/04/2024 |  |  | 02/23/2024 | 03/04/2024 |  |  | 10/31/2025 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024422803 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/23/2024 |  |  |  |  | The user facility reported a 52-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that the decision was made to remove the impella cp due to bleeding significantly from the access site. five units of packed red blood cells was transfused as treatment. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was most likely due to anticoagulation management since the act at the time of bleed and end was 233, which is above the limit range.
30387 | 1220648-2024-07326 | 18826255 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18826255&pc=OZD | 02/24/2024 | 04/03/2024 |  | 03/04/2024 |  |  | 02/24/2024 | 03/04/2024 |  |  | 11/30/2025 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025453996 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/24/2024 |  |  |  |  | The user facility reported a(b)(6) male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported after two days of persistent suction alarms, the imeplla cp was replaced with a new one. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed since the initial report was submitted. the device was not returned for investigation. per the clinical description and data analysis provided, the root cause of the positioning issues was most likely use related.
30388 | 1220648-2024-07328 | 18826271 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18826271&pc=OZD | 11/25/2022 | 03/04/2024 |  | 03/04/2024 |  |  | 11/25/2022 | 03/04/2024 |  |  | 05/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023230937 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/25/2022 |  |  |  |  | The user facility reported a 63-year-old female in post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella cp device for mechanical circulatory support. it was reported that there was no pulses to the impella leg. a doppler ultrasound was performed, however, the results had not been read. | [5 empty] | The investigation for the reported limb ischemia has been completed. the impella device was not received from the customer. however, clinical details were insufficient to determine the root cause. limb ischemia may occur during impella support; unfortunately, determining the exact cause requires clinical information which was not provided. the instructions for use states: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
30389 | 1220648-2024-07327 | 18826272 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18826272&pc=OZD | 02/25/2024 | 04/15/2024 |  | 03/04/2024 |  |  | 02/25/2024 | 03/04/2024 |  |  |  |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025439417 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Material Split, Cut or Torn | [5 empty] | Hemolysis | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/25/2024 |  |  |  |  | The user facility reported a 60-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported after two days of support the impella cp was escalated to impella 5.5 due to hemolysis concerns and a torn sterile sleeve. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the reported hemolysis and product damage has been completed. the impella device was not returned for evaluation. data log analysis confirmed no recorded abnormalities. according to the clinical, during the first day of impella cp support the pump position was too deep. over the next day the pump was repositioned a total of 4cm and the sterile sleeve was damaged. the decision was then made to leave the pump until it was exchanged for a 5.5 the following day. during this time hemolysis was noted due to the improper positioning of the pump. after replacing the pump the hemolysis was resolved. the most likely cause of the sterile sleeve product damage was a use related issue. the most likely cause of the hemolysis was improper positioning. impact code-4627 was coded twice in error. 2nd code should have been 4628.
30390 | 1220648-2024-07329 | 18826304 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18826304&pc=OZD | 02/26/2024 | 04/09/2024 |  | 03/04/2024 |  |  | 02/26/2024 | 03/04/2024 |  |  | 11/30/2025 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024430631 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | Yes | Yes | Difficult to Remove | [6 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/26/2024 |  |  |  |  | The user facility reported a 64-year-old female noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that at the conclusion of the procedure, the impella cp was weaned and removed through the peel away sheath, when the peel away sheath was removed, closure was unsuccessful with the initial preclose devices. the physician achieved hemostasis by deploying covered stent via contralateral femoral artery. the patient was reported as stable. | [5 empty] | The investigation for the reported removal issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the most likely cause of the major access site bleeding was use related. the instructions for use states ¿assess access site for bleeding and hematoma.¿ | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30391 | 1220648-2024-07330 | 18826396 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18826396&pc=OZD | 02/23/2024 | 04/10/2024 |  | 03/04/2024 |  |  | 02/23/2024 | 03/04/2024 |  |  | 10/31/2025 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024409483 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/23/2024 |  |  |  |  | The user facility reported a 69-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that impella 5.5 was explanted from the right axillary artery and physician noted thrombus embedded in the outflow cage. the patient is stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the thrombus issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined.
30392 | 1220648-2024-07331 | 18826439 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18826439&pc=OZD | 11/29/2022 | 03/11/2025 |  | 03/04/2024 |  |  | 11/29/2022 | 03/04/2024 |  |  | 08/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023270223 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/29/2022 |  |  |  |  | The user facility reported a 77-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced some oozing at the access site postoperatively while the patient was not on systemic heparin. heparin was not administered, and manual pressure was applied to address bleeding. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | The root cause of the access site oozing cannot be determined due to insufficient clinical information and lack of product returned. this report is being filed as part of a retrospective review of historical records.
30393 | 1220648-2024-07332 | 18826450 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18826450&pc=OZD | 11/25/2022 | 03/07/2025 |  | 03/04/2024 |  |  | 11/25/2022 | 03/04/2024 |  |  | 01/26/2022 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023166864 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/25/2022 |  |  |  |  | The user facility reported a 61-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported there was a lot of oozing coming from the access site, which was treated by transfusing 3 units of packed red blood cells (prbc) into the patient. it was also noted that the hemoglobin (hgb) dropped to 5.3 and albumin was given. unfortunately, when the staff was moving the patient the impella was pulled as well as the catheter, and the repositioning sheath out of place. upon the arrival of the physician ,the decision was made to place a new impella cp within the patient. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30394 | 1220648-2024-07333 | 18826466 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18826466&pc=OZD | 02/23/2024 | 04/10/2024 |  | 03/04/2024 |  |  | 02/23/2024 | 03/04/2024 |  |  | 11/30/2025 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024434167 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/23/2024 |  |  |  |  | The user facility reported a (b)(6) male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient had a visible hematoma forming around the impella access site of the left groin during the pci (percutaneous coronary intervention). after the peel away sheath was removed, the groin was bleeding and md attempted to place repositioning sheath. in an effort to stop the bleeding, the md inserted the end of the blue suture hub into the patient to ¿plug the hole¿. when that didn¿t work to control the bleeding, the md opted to explant pump. a 0.35 wire was placed in the access port of the repositioning sheath and pump was explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
30395 | 1220648-2024-07334 | 18826512 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18826512&pc=OZD | 02/24/2024 | 09/24/2024 |  | 03/04/2024 |  |  | 02/24/2024 | 03/04/2024 |  |  | 07/31/2025 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025442935 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Break | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Ventricular Fibrillation | No Clinical Signs, Symptoms or Conditions | [9 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/24/2024 |  |  |  |  | The user facility reported a 67-year-old male with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported that the purge pressure reservoir and infusion filter broke off during impella cp procedure. purge bypass was initiated, and the purge system open alarm cleared. the impella catheter operated without difficulty or incident with the purge bypass in place. no patient harm was reported. | It was also reported the patient experienced a lack of pulses to their right arm. the physician performed a percutaneous transluminal angioplasty to restore flow to the patient's right arm. it was also noted that the patient had an episode of ventricular fibrillation that required defibrillation and the patient was administered amiodarone to resolve. there was also oozing noted at the patient's impella access site. the site was treated with a pressure dressing and the patient received multiple units of blood as a result of the blood loss. the patient's family decided to withdraw care, and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. the following have been reported incorrectly or missing from manufacturer device report. b1, b2: corrected to include serious injury. b5 included new information for serious injuries found during retrospective review and the medical safety officer statement. d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact fax number missing. h1: changed to serious injury. h6: added codes for ventricular fibrillation, limb ischemia, bleeding, medication required, blood transfusion, additional surgery. | The investigation into the purge leak ¿ pump issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the purge leak - pump was most likely a leak from the sidearm but the exact location of the leak cannot be confirmed because the product was not returned.
30396 | 1220648-2024-07335 | 18826569 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18826569&pc=OZD | 11/25/2022 | 05/23/2025 |  | 03/04/2024 |  |  | 11/25/2022 | 03/04/2024 |  |  | 09/30/2023 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023265751 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/25/2022 |  |  |  |  | The user facility reported a 71-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient developed bleeding at the impella cp insertion site. the physician managed the bleeding by placing mattress sutures around the site. | The bleed was managed and the patient was returned to the intensive care unit. however, once in the unit, the patient subsequently coded and passed away. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial manufacturer report 1220648-2024-07335 was provided in sections b2, b5, d6, h1, and h6. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30397 | 1220648-2024-07336 | 18826587 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18826587&pc=OZD | 02/25/2024 | 04/16/2024 |  | 03/04/2024 |  |  | 02/25/2024 | 03/04/2024 |  |  | 09/30/2025 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024394945 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Insufficient Flow or Under Infusion | Pressure Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/25/2024 |  |  |  |  | The user facility reported a 51-year-old male in st elevated myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that after 45 days of support the pump began to demonstrate low flows and purge block/high purge pressure alarms. the pump was exchanged for a new 5.5 pump and no patient harm was reported. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed | The investigation into the low pump flow and the high purge pressure issues has been completed since the original report was filed. the pump was returned, tested, and evaluated. heavy biomaterial was found stuffed in the outflow cade and around the impeller and in the pump. data logs showed a gradual increase in the purge pressure and drop in flow further confirming biomaterial buildup. the root cause of the low pump flow and the high purge pressure issues was biomaterial.
30398 | 1220648-2024-07337 | 18826659 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18826659&pc=OZD | 02/26/2024 | 08/16/2024 |  | 03/04/2024 |  |  | 02/25/2024 | 03/04/2024 |  |  |  |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024433386 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 02/25/2024 |  |  |  |  | The user facility reported an (b)(6) female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp was replaced with intra-aortic balloon pump (iabp) due to concerns of left leg ischemia and hemolysis. patient was reported to be stable. | It was later reported via abiomed's long-term outcome and quality indicator impella registry, that the patient endured thrombocytopenia with a probable relationship to the impella device. a consistent downward trend in platelets was noted with oozing from multiple lines. two units of platelets were transfused in response to the downward trend. the patient was reported to have recovered from the condition. | Additional information received via abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured thrombocytopenia with a "probable" relationship to the impella device used: consistent downward trends in platelets. oozing from multiple lines. platelets down to 48 and 1 unit transfused. the patient recovered with sequelae. |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ ischemia: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. | The investigation for the reported thrombocytopenia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombocytopenia could not be determined. the instructions for use referenced in the initial report is for the impella cp with smartassist system in the following sections: section: potential adverse events (united states), which now includes the statement "cardiac or vascular injury (including ventricular perforation.¿ section: use of echocardiography for positioning of the impella catheter. section: hemolysis. section: suction a1-corrected patient identifier. added- d6a/d6b. f6/f8 should have been left blank in the initial report. 4642 added as an impact code in h6 to coincide with the report of an iabp being placed due to concerns of ischemia and hemolysis. | Section b5, g2 and h6 were updated as per additional information received. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
30399 | 1220648-2024-07338 | 18826688 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18826688&pc=OZD | 11/25/2022 | 03/07/2025 |  | 03/04/2024 |  |  | 11/25/2022 | 03/04/2024 |  |  | 08/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023263747 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/25/2022 |  |  |  |  | The user facility reported a 86-year-old female noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that while adjusting the impella the pump moved out of position. the decision was made to remove the pump and reimplant it. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30400 | 1220648-2024-07339 | 18827035 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18827035&pc=OZD | 11/25/2022 | 03/12/2025 |  | 03/04/2024 |  |  | 11/25/2022 | 03/04/2024 |  |  | 06/30/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235952 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 26 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/25/2022 |  |  |  |  | The user facility reported a 26-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced significant bleeding around the sheath. a suture was replaced around the peel away to manage the bleeding. the patient continued on support. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. the instructions for use indicates the following: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
30401 | 1220648-2024-07340 | 18827139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18827139&pc=OZD | 10/12/2022 | 03/04/2024 |  | 03/04/2024 |  |  | 10/12/2022 | 03/04/2024 |  |  | 06/30/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023247460 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/12/2022 |  |  |  |  | The user facility reported a 63-year-old female was attempting to be placed on an impella 5.5 device for mechanical circulatory support. the impella 5.5 pump was unable to pass through the patient's vasculature during attempted delivery. the patient's right side was noted to be tortuous while the left side was occluded. as direct access was also not viable, the decision was made to abort the impella 5.5 implant. an impella cp pump was subsequently placed for patient support. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the root cause of the delivery issue was most likely due to patient anatomy as per the clinical description provided. this report is being filed as part of a retrospective review of historical records.
30402 | 1220648-2024-07341 | 18827168 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18827168&pc=OZD | 10/12/2022 | 03/04/2024 |  | 03/04/2024 |  |  | 10/12/2022 | 03/04/2024 |  |  | 02/29/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023197451 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/12/2022 |  |  |  |  | The user facility reported a 78-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the pump continued to exhibit high purge pressure alarms with no resolution after the second tissue plasminogen activator (tpa) bag after the second day. the patient alternated between tpa and bicarb. additionally, it was reported that there was a hematoma at the insertion site. the patient was improving after given one unit of blood. the pump was removed with no issue. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
30403 | 1220648-2024-07342 | 18827250 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18827250&pc=OZD | 10/28/2022 | 04/16/2024 |  | 03/04/2024 |  |  | 10/28/2022 | 03/04/2024 |  |  | 08/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023249839 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/28/2022 |  |  |  |  | The user facility reported a 65-year-old of unknown gender undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the patient was unable to be weaned off the impella cp support due to a second sepsis with 39 °c (102.2 °f) fever. following another sepsis and ventricular tachycardia, the patient expired on support. no information was provided on intervention used for sepsis or ventricular tachycardia. no allegation that device contributed to patient death. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿infusion filter the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported sepsis and ventricular tachycardia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the sepsis and ventricular tachycardia could not be determined.
30404 | 1220648-2024-07343 | 18827265 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18827265&pc=OZD | 11/24/2022 | 04/07/2025 |  | 03/04/2024 |  |  | 11/24/2022 | 03/04/2024 |  |  | 06/30/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023269535 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/24/2022 |  |  |  |  | The user facility reported a 72-year-old female implanted with an impella cp for ongoing mechanical circulatory support due to the patients need for percutaneous coronary intervention (pci). it was reported that pulses could not be found in the leg. the pump was explanted prematurely after 30 hours due to concerns for leg ischemia. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
30405 | 1220648-2024-07344 | 18827295 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18827295&pc=OZD | 11/26/2022 | 03/07/2025 |  | 03/04/2024 |  |  | 11/26/2022 | 03/04/2024 |  |  | 06/30/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023236934 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemolysis | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/26/2022 |  |  |  |  | The user facility reported a 62-year-old female in cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported that there were multiple episodes of vtach (ventricular tachycardia) as the pump migrated deep in the left ventricle (lv). the patient was intubated and the patient¿s peripheral vascular resistance (svr) was trending upwards and ast and alt were both trending downwards. systematic heparin was paused. patient experienced a stroke. the purged was switched to bicarb. due to low hematocrit and hemoglobin (h/h), the patient received one unit of packed red blood cells (prbc). | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
30406 | 1220648-2024-07345 | 18827298 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18827298&pc=OZD | 10/29/2022 | 03/04/2024 |  | 03/04/2024 |  |  | 10/29/2022 | 03/04/2024 |  |  | 06/30/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023248576 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/29/2022 |  |  |  |  | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support, and the patient experienced continued oozing at insertion site. heparin was discontinued, and due to a drop in hemoglobin/hematocrit, the patient received a blood transfusion and a femstop was placed on the groin. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records.
30407 | 1220648-2024-07346 | 18827311 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18827311&pc=OZD | 11/24/2022 | 03/04/2024 |  | 03/04/2024 |  |  | 11/24/2022 | 03/04/2024 |  |  | 08/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023269651 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/24/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 51-year-old male patient for mechanical circulatory support. it was reported that during impella support, the patient developed bleeding at the impella site. as a result of the noted issue, surgical and manual pressure were applied. the impella cp was explanted. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury" this report is being filed as part of a retrospective review of historical records.
30408 | 1220648-2024-07347 | 18827380 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18827380&pc=OZD | 11/01/2022 | 03/04/2024 |  | 03/04/2024 |  |  | 11/01/2022 | 03/04/2024 |  |  | 07/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023251619 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/01/2022 |  |  |  |  | The user facility reported a 65-year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. during impella cp support, the creatine kinase myocardial band fraction (ckmb) started to rise and the impella leg started getting tight. a fasciotomy was performed for leg ischemia and required blood products for access site bleeding. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records.
30409 | 1220648-2024-07348 | 18827415 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18827415&pc=OZD | 11/24/2022 | 03/04/2024 |  | 03/04/2024 |  |  | 11/24/2022 | 03/04/2024 |  |  | 05/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023217629 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/24/2022 |  |  |  |  | The user facility reported an impella cp in a patient for mechanical circulatory support. it was reported that the medical team was having difficulty delivering the pump due to the patient's anatomy and the procedure was ultimately aborted. there was no patient harm reported. | [5 empty] | The investigation into the reported position issue has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine a root cause. the most likely cause of the delivery issue was patient condition related because the impella was unable to pass through vasculature. this report is being filed as part of a retrospective review of historical records.
30410 | 1220648-2024-07349 | 18827480 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18827480&pc=OZD | 09/29/2022 | 03/12/2025 |  | 03/04/2024 |  |  | 09/29/2022 | 03/04/2024 |  |  | 03/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023230300 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/29/2022 |  |  |  |  | The user facility reported 67-year-old male patient needing mechanical circulatory support. it was reported that after being placed on impella cp support, the femoral artery was bleeding after the reposition sheath was placed. femostop placed. patient care was withdrawn, and patient placed on hospice. impella support discontinued. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30411 | 1220648-2024-07351 | 18827502 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18827502&pc=OZD | 11/01/2022 | 03/04/2024 |  | 03/04/2024 |  |  | 11/01/2022 | 03/04/2024 |  |  | 04/30/2023 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022072637 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/01/2022 |  |  |  |  | The user facility reported a patient of unknown gender, was undergoing cardiac surgery and was unable to be implanted with an impella cp device for mechanical circulatory support due to the patient¿s anatomy. the case was completed using an intra-aortic balloon pump. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not received from the customer. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely due to the patient¿s condition since resistance was met due to anatomy at aorta. this report is being filed as part of a retrospective review of historical records.
30412 | 1220648-2024-07350 | 18827503 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18827503&pc=OZD | 11/24/2022 | 03/04/2024 |  | 03/04/2024 |  |  | 11/24/2022 | 03/04/2024 |  |  | 08/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023260493 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/24/2022 |  |  |  |  | The user facility reported an impella cp in a 70-year-old male patient for mechanical circulatory support. it was reported oozing at the impella access site. as a result of the noted issue, a standard bleeding protocol was followed to treat the bleeding. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury" this report is being filed as part of a retrospective review of historical records.
30413 | 1220648-2024-07352 | 18827535 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18827535&pc=OZD | 10/31/2022 | 03/26/2024 |  | 03/04/2024 |  |  | 10/31/2022 | 03/04/2024 |  |  | 05/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023228081 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/31/2022 |  |  |  |  | The user facility reported a 79-year-old female in st segment elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. after being placed on impella cp support, the patient developed no pedal pulses. the patient¿s family decided to not escalate care and wanted to withdraw support. the impella cp device was weaned and removed, and the patient remained intubated in levo medication and dobutamine. | The user facility reported a 79-year-old female in st segment elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. after being placed on impella cp support, the patient developed no pedal pulses. no medical or surgical intervention was reported. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined.
30414 | 1220648-2024-07353 | 18827631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18827631&pc=OZD | 11/16/2022 | 03/04/2024 |  | 03/04/2024 |  |  | 11/16/2022 | 03/04/2024 |  |  | 07/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023256701 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/16/2022 |  |  |  |  | The user facility reported a 64-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. while being placed under impella cp support, it was determined that the patient had bilateral calcified femoral arteries. as a result, the procedure was aborted due to the patient¿s anatomy. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issue was most likely patient condition related, the impella was unable to pass through due to bilateral calcified femoral arteries as per the clinical description provided. this report is being filed as part of a retrospective review of historical records.
30415 | 1220648-2024-07354 | 18827699 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18827699&pc=OZD | 11/16/2022 | 03/04/2024 |  | 03/04/2024 |  |  | 11/16/2022 | 03/04/2024 |  |  | 07/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023257992 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/16/2022 |  |  |  |  | The user facility reported a 62-year-old male in st segment elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. after being placed on impella cp support, the patient experienced a massive hematoma and the site oozing/bleeding. manual pressure was performed and the hematoma softened. dressing was also applied with angle matching gauze. additionally, bicarbonate was added to the purge and the patient received 4 units of packed red blood cells (prbc¿s), 2 units of fresh frozen plasma (ffp) and 1 platelet. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
30416 | 1220648-2024-07355 | 18827708 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18827708&pc=OZD | 11/25/2022 | 05/23/2025 |  | 03/04/2024 |  |  | 11/25/2022 | 03/04/2024 |  |  | 08/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023270046 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | Cardiovascular Dysfunction/Insufficiency | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/25/2022 |  |  |  |  | The user facility reported an impella cp in a 67-year-old male patient for mechanical circulatory support. it was reported that during impella support, the patient experienced excessive bleeding at the impella site. under fluoroscopy, it was shown the catheter was kinked and there was extravasation at the site. a 8 mm x 10 cm stent was placed in the axillary artery. penumbra device could be used to extract some of the clot that had developed at the end of the catheter, but the doctor wanted to medically manage the thrombus. | Following the initial incidents, support was continued with the implanted pump until planned weaning 8 days later. it was documented that the patient's native heart function had not significantly improved at the time of weaning. the patient subsequently coded and the family made the decision to stop resuscitation efforts. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury" this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial manufacturer report 1220648-2024-07355 was provided in sections b2, b5, d6, h1, and h6. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30417 | 1220648-2024-07356 | 18827740 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18827740&pc=OZD | 11/17/2022 | 05/22/2025 |  | 03/04/2024 |  |  | 11/17/2022 | 03/04/2024 |  |  | 01/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023187125 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/17/2022 |  |  |  |  | The user facility reported a 89-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. after explant, the patient's left femoral artery (lfa) was unable to be closed with existing sutures and the vascular surgeon performed a cut down for closure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30418 | 1220648-2024-07357 | 18827779 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18827779&pc=OZD | 11/17/2022 | 05/23/2025 |  | 03/04/2024 |  |  | 11/17/2022 | 03/04/2024 |  |  | 07/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023259537 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/17/2022 |  |  |  |  | The user facility reported a 44-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. during the implant, there was no distal flow past the repo sheath, so a distal perfusion cannula was placed in the post tibial and into the arterial ecmo. the site position was oozing but intact and the angle was maintained. distal perfusion was improving per skeletal muscle oxygen saturation (sto2). bicarb ordered for the purge. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | The root cause of access site bleed was determined to be patient condition because the distal perfusion was because of the patient condition. this report is being filed as part of a retrospective review of historical records.
30419 | 1220648-2024-07358 | 18827875 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18827875&pc=OZD | 11/17/2022 | 03/04/2024 |  | 03/04/2024 |  |  | 11/17/2022 | 03/04/2024 |  |  | 05/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023227910 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/17/2022 |  |  |  |  | The user facility reported a 65-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while the impella cp support was being removed from the patient, a hematoma developed that required blood product administration. a femstop was used for hemostasis at time of removal. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
30420 | 1220648-2024-07359 | 18827954 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18827954&pc=OZD | 11/25/2022 | 09/11/2024 |  | 03/04/2024 |  |  | 11/25/2022 | 03/04/2024 |  |  | 01/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023186942 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/25/2022 |  |  |  |  | The user facility reported an impella cp in 56-year-old female patient for mechanical circulatory support. it was reported that while on impella cp support, the patient experienced global ischemia. | Following implant, the patient was taken to the intensive care unit (icu) where their mean arterial pressure began trending downward. the patient eventually reached pulseless electrical activity and care was withdrawn. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. reportable due to patient's ischemia while on impella cp support and death. |  |  |  |  | The investigation for the reported ischemia has been completed. the impella device has not been returned for evaluation. however, as the necessary clinical information was not provided and the device was not returned, the root cause of the ischemia could not be determined. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial manufacturing report 1220648-2024-07359 was provided in sections b2, b5, d6, h1, and h6.
30421 | 1220648-2024-07360 | 18827956 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18827956&pc=OZD | 11/17/2022 | 03/04/2024 |  | 03/04/2024 |  |  | 11/17/2022 | 03/04/2024 |  |  | 07/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023250125 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/17/2022 |  |  |  |  | The user facility reported a 74-year-old female on a impella cp for mechanical circulatory support, that during insertion, multiple attempts to cross into the left ventricle were made with the 0.018 wire unsuccessfully. a buddy wire technique also failed. the procedure was aborted and an intraaortic balloon pump was placed instead. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related. upon injection of the aortic root and selective engagement of right and left coronaries there appeared to be a large aneurysmal area in the non-coronary cusp as per the clinical description provided. this report is being filed as part of a retrospective review of historical records.
30422 | 1220648-2024-07361 | 18828025 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18828025&pc=OZD | 11/18/2022 | 03/04/2024 |  | 03/04/2024 |  |  | 11/18/2022 | 03/04/2024 |  |  | 06/30/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023247251 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/18/2022 |  |  |  |  | The user facility reported a (b)(6) male in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella cp device for mechanical circulatory support, and that the impella cp could not be placed due to patient¿s anatomy. an introducer was not used. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not received from the customer. however, clinical details provided were sufficient to determine a root cause. the most likely root cause of the reported delivery issue was use related as no introducer was used as per the clinical description. this report is being filed as part of a retrospective review of historical records.
30423 | 1220648-2024-07362 | 18828036 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18828036&pc=OZD | 11/30/2022 | 03/04/2024 |  | 03/04/2024 |  |  | 11/30/2022 | 03/04/2024 |  |  | 08/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023270285 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Difficult to Insert | Failure to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/30/2022 |  |  |  |  | The user facility reported a 62-year-old female being placed on impella cp required mechanical circulatory support. the complainant reported that the surgical team was unable to cross the valve with pig and j wire and using a straight and al 1 wire caused the wire to cross at a strangle angle. the patient chart revealed that the patient had a mechanical aortic valve placed previously. the implant was aborted with plan to convert to venoarterial extracorporeal membrane oxygenation (va ecmo). | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, because of the patient's prexisiting mechanical aortic valve as per the clinical description provided. this report is being filed as part of a retrospective review of historical records.
30424 | 1220648-2024-07363 | 18828188 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18828188&pc=OZD | 12/01/2022 | 03/04/2024 |  | 03/04/2024 |  |  | 12/01/2022 | 03/04/2024 |  |  | 07/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023255188 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/01/2022 |  |  |  |  | The user facility reported a (b)(6) female patient was being implanted with an impella cp device for mechanical circulatory support. it was reported that the pump was unable to advance due to high calcification within the patient's vasculature. therefore, the impella cp implant was aborted. there was no patient harm reported. | [5 empty] | The investigation into the reported delivery issue has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine a root cause. the most likely cause of the delivery issue was due to patient condition of high calcification. this report is being filed as part of a retrospective review of historical records.
30425 | 1220648-2024-07364 | 18828250 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18828250&pc=OZD | 11/26/2022 | 03/04/2024 |  | 03/04/2024 |  |  | 11/26/2022 | 03/04/2024 |  |  | 08/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023267603 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/26/2022 |  |  |  |  | The user facility reported a 72-year-old male being placed on impella cp required mechanical circulatory support. while on impella cp support, the patient was experiencing distal limb ischemia. no information was provided regarding the treatment of the limb ischemia. however, the family decided to withdraw care due to unrelated causes. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device was not returned for evaluation nor was there sufficient clinical information available. therefore, the root cause of the limb ischemia could not be determined. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
30426 | 1220648-2024-07365 | 18828355 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18828355&pc=OZD | 11/18/2022 | 03/04/2024 |  | 03/04/2024 |  |  | 11/18/2022 | 03/04/2024 |  |  | 07/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023253392 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/18/2022 |  |  |  |  | The user facility reported a 66-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. after transferring the patient to a bed and rolling, oozing noted from the site. the dressing and existing sutures were removed and a new mattress suture was placed with a forward tension suture placed around the sheath to control the bleeding. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
30427 | 1220648-2024-07366 | 18828628 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18828628&pc=OZD | 11/30/2022 | 03/12/2025 |  | 03/04/2024 |  |  | 11/30/2022 | 03/04/2024 |  |  | 12/31/2023 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023175276 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/30/2022 |  |  |  |  | The user facility reported a 80-year-old female in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. after being placed on impella cp support, it was noted that the patient had limb ischemia that progressively became worse. no pulses were able to be found in right foot. patient had arterial doppler done overnight that showed flow to distal popliteal and the decision was made to explant the impella cp device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30428 | 1220648-2024-07367 | 18829545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18829545&pc=OZD | 11/18/2022 | 03/12/2025 |  | 03/04/2024 |  |  | 11/18/2022 | 03/04/2024 |  |  | 03/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023207359 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/18/2022 |  |  |  |  | The user facility reported a 65-year-old male was implanted with an impella cp device for mechanical circulatory support. the patient experienced a marked hematoma at the groin site. the device was successfully weaned and explanted with no change in the hematoma size compared to when it was implanted with no new groin concern. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided.
30429 | 1220648-2024-07369 | 18829626 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18829626&pc=OZD | 11/30/2022 | 03/12/2025 |  | 03/04/2024 |  |  | 11/30/2022 | 03/04/2024 |  |  | 06/30/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023247158 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/30/2022 |  |  |  |  | The user facility reported an impella cp in 53-year-old male patient for mechanical circulatory support. it was reported that during impella support, the patient had experienced oozing at the right groin insertion site. the purge was transitioned from heparin to bicarbonate. additionally, pressure was held, fem-stop was applied, and hemostasis was achieved. the patient was transfused with two units of pack red blood cells. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
30430 | 1220648-2024-07370 | 18829868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18829868&pc=OZD | 11/25/2022 | 03/04/2024 |  | 03/04/2024 |  |  | 11/25/2022 | 03/04/2024 |  |  | 07/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC |  |  | US |  | 2023257982 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/25/2022 |  |  |  |  | The user facility reported a 74-year-old male patient needing an impella cp for mechanical circulatory support. it was reported that the impella cp was attempted to be inserted with a 14f sheath in the left femoral artery. the impella cp could not be advanced through the aorta as it was very heavily diseased and calcified. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
30431 | 1220648-2024-07368 | 18830096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18830096&pc=OZD | 11/26/2022 | 03/10/2025 |  | 03/04/2024 |  |  | 11/26/2022 | 03/04/2024 |  |  | 02/29/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023200026 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/26/2022 |  |  |  |  | The user facility reported a 74-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that after the sheath was placed, the physician was unable to advance the catheter. angiogram results revealed a left femoral dissection. the procedure was completed using a right femoral access site. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30432 | 1220648-2024-07372 | 18830535 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18830535&pc=OZD | 12/01/2022 | 03/04/2024 |  | 03/04/2024 |  |  | 12/01/2022 | 03/04/2024 |  |  | 07/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023251160 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/01/2022 |  |  |  |  | The user facility reported a 69-year-old female patient was being implanted with an impella cp device for mechanical circulatory support. it was reported that during impella support, the patient developed bleeding at the impella access site. as a result of the noted issue, four (4) units of packed red blood cells (prbcs), two (2) units of platelets (plts), and ten (10) units of cryoprecipitate (cryo) were given. the physician placed extra sutures and redressed the site to stop the bleeding. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
30433 | 1220648-2024-07371 | 18830558 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18830558&pc=OZD | 11/27/2022 | 03/04/2024 |  | 03/04/2024 |  |  | 11/27/2022 | 03/04/2024 |  |  | 08/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023261252 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 11/27/2022 |  |  |  |  | The user facility reported a 75-year-old was on impella cp support and the patient had bleeding (oozing) at the impella cp insertion site (groin). the patient required two units of packed red blood cells (2u prbc's). systemic heparin was held temporarily then restarted. hemoglobin was then stable and the access site was patent. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records.
30434 | 1220648-2024-07374 | 18830835 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18830835&pc=OZD | 09/23/2022 | 03/04/2024 |  | 03/04/2024 |  |  | 09/23/2022 | 03/04/2024 |  |  | 02/29/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023199091 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/23/2022 |  |  |  |  | The user facility reported a 68-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding at the access site. manual pressure was applied and femostop was used. the patient was administered blood products. no further information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
30435 | 1220648-2024-07373 | 18830964 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18830964&pc=OZD | 11/30/2022 | 03/12/2025 |  | 03/04/2024 |  |  | 11/30/2022 | 03/04/2024 |  |  | 08/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023268425 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/30/2022 |  |  |  |  | The user facility reported a 55 year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced slight oozing at the groin insertion site and episodes of ventricular tachycardia (vt) to ventricular fibrillation (vf). systemic heparin was discontinued to treat bleeding, and the patient was taken to the operating room for extra-corporeal membrane oxygenation (ecmo) cannulation in response to the rounds of vt/vf. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30436 | 1220648-2024-07375 | 18830993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18830993&pc=OZD | 11/29/2022 | 03/12/2025 |  | 03/04/2024 |  |  | 11/29/2022 | 03/04/2024 |  |  | 08/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023265666 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Malfunction | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/29/2022 |  |  |  |  | The user facility reported a 64-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. a bedside echo was performed and it was discovered that the impella cp was trapped in the mitral valve. the physician attempted to reposition the impella cp in the free left ventricular space but was not successful. the device was removed and the patient was escalated to the impella 5.5 to bridge to coronary artery bypass graft. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ ¿impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30437 | 1220648-2024-07376 | 18831009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18831009&pc=OZD | 11/27/2022 | 03/04/2024 |  | 03/04/2024 |  |  | 11/27/2022 | 03/04/2024 |  |  | 08/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023270413 |  | 0048-0045 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 11/27/2022 |  |  |  |  | The user facility reported a 57-year-old male patient needing mechanical circulatory support. it was reported that while on impella cp cupport, bleeding was noted from the impella sheath site. the dressing was changed and the sheath was elevated. the physician instructed to run the purge heparin at full strength in 2 hours pending bleeding. systemic heparin was then withheld. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records.
30438 | 1220648-2024-07377 | 18831218 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18831218&pc=OZD | 09/26/2022 | 03/04/2024 |  | 03/04/2024 |  |  | 09/26/2022 | 03/04/2024 |  |  | 05/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023220146 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Female |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/26/2022 |  |  |  |  | The user facility reported a 70-year-old female patient was implanted with an impella cp for mechanical circulatory support. it was reported that the impella cp support experienced a pump stop. therefore, the pump was explanted and a new pump was placed. there was no reported patient harm. | [5 empty] | The investigation for the reported pump stop has been completed. the impella device was not received from the customer, therefore, evaluation of the device was not possible. the root cause of the pump stop was not determined since product and data logs were not returned. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
30439 | 1220648-2024-07378 | 18831412 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18831412&pc=OZD | 11/27/2022 | 03/04/2024 |  | 03/04/2024 |  |  | 11/27/2022 | 03/04/2024 |  |  | 08/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023269829 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 11/27/2022 |  |  |  |  | The user facility reported a 63-year-old male implanted with impella cp needing mechanical circulatory support. it was reported that while the patient was on impella cp support the patient experienced purge pressure rising from 400-500 to over 800 and the purge flow dropped to 3.2 from 18ml/hr. no kinks were observed. no further information regarding treatment was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
30440 | 1220648-2024-07379 | 18831968 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18831968&pc=OZD | 11/28/2022 | 03/04/2024 |  | 03/04/2024 |  |  | 11/28/2022 | 03/04/2024 |  |  | 06/30/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023248952 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 11/28/2022 |  |  |  |  | The user facility reported an 80-year-old female patient was implanted with an impella cp needing mechanical circulatory support. it was reported that the staff was unable to find doppler pulse in left foot. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device was not returned for evaulation nor was there sufficient clinical information available. therefore, the root cause of the limb ischemia could not be determined. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
30441 | 1220648-2024-07381 | 18832355 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18832355&pc=OZD | 05/15/2021 | 03/04/2024 |  | 03/04/2024 |  |  | 05/15/2021 | 03/04/2024 |  |  | 10/31/2021 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP C7 W V3 GWRU SET, US |  |  | ABIOMED, INC. |  |  | US |  | 1457877 |  | 0048-0032 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/15/2021 |  |  |  |  | The user facility reported a 52-year-old male was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient developed a bleeding from their groin during the impella support. the repositioning sheath was advanced forward and the site was re-sutured with forward pressure. manual pressure was applied, however femostop was needed as some oozing was still noted from the site. support continued with the implanted pump following the intervention. | [5 empty] | The investigation for the reported access site bleeding issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the access site bleeding issue was user related. the sheath had been pulled back and flattened against the leg as per the clinical description provided. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
30442 | 1220648-2024-07382 | 18832440 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18832440&pc=OZD | 09/26/2022 | 03/12/2025 |  | 03/04/2024 |  |  | 09/26/2022 | 03/04/2024 |  |  | 06/30/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023240342 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/26/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 82-year-old male patient for mechanical circulatory support. it was reported that while stenting the right coronary, flow was lost in the vessel and the patient was non-pulsatile on the impella. flow was restored and the patient regained pulsatility. the impella cp was weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Analysis summary: limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition. ¿ any concomitant medication. ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient. ¿ physical examination findings, including pulse assessment and visual examination of the affected limb. ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders. ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. ¿ patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
30443 | 1220648-2024-07383 | 18832582 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18832582&pc=OZD | 10/13/2022 | 05/13/2025 |  | 03/04/2024 |  |  |  | 03/04/2024 |  |  | 07/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023249584 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/13/2022 |  |  |  |  | The user facility reported that a 57-year-old male was placed on impella cp support. the impella cp became malpositioned after placement. the attempts to reposition the pump were not successful and the pump was explanted. the patient was supported with vasoactives/inotropes until the computed tomography (ct) surgery consult could be completed for impella 5.5 placement. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30444 | 1220648-2024-07384 | 18832602 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18832602&pc=OZD | 11/28/2022 | 03/04/2024 |  | 03/04/2024 |  |  |  | 03/04/2024 |  |  |  |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/28/2022 |  |  |  |  | The user facility reported that an unknown patient had a cp support and the device exhibited a purge pressure high alarm for the duration of the case. no kinks were observed. the patient was successfully weaned at the end of the case. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ this report is being filed as part of a retrospective review of historical records.
30445 | 1220648-2024-07385 | 18832616 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18832616&pc=OZD | 10/13/2022 | 03/04/2024 |  | 03/04/2024 |  |  |  | 03/04/2024 |  |  | 05/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023227166 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Break | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/13/2022 |  |  |  |  | The user facility reported an impella cp was placed in a female patient for mechanical circulatory support. it was reported that the pump had leakage at the yellow luer junction due to the broken purge arm. the pump was replaced. there was no patient harm reported. | [5 empty] | The pump was not returned for investigation. according to the clinical details, the pump was replaced after the observed leak from the sidearm yellow luer. the cause of the purge leak was most likely sidearm yellow luer damage, based on given clinical details, however no product was returned so the exact location of the leak from the luer or how the damage occurred cannot be determined. instructions for use for the related event are as follows: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g., infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
30446 | 1220648-2024-07386 | 18832669 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18832669&pc=OZD | 12/02/2022 | 03/04/2024 |  | 03/04/2024 |  |  |  | 03/04/2024 |  |  | 08/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023269595 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/02/2022 |  |  |  |  | The user facility reported a 58-year-old male patient needing mechanical circulatory support was implanted with an impella cp. the patient experienced a clotting issue (thrombus) while on support. blood products were given to the patient. providers decided to remove the pump and continue medical management. no allegations were made towards the impella for cause of death. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ this report is being filed as part of a retrospective review of historical records.
30447 | 1220648-2024-07387 | 18832694 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18832694&pc=OZD | 12/01/2022 | 03/04/2024 |  | 03/04/2024 |  |  | 12/01/2022 | 03/04/2024 |  |  | 08/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023270290 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Tachycardia | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/01/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a 66-year-old male patient for mechanical circulatory support. the patient had frequent ectopy and progressed to supraventricular tachycardia (svt). the impella cp was explanted and the blood pressure improved with fluids. | [5 empty] | The investigation for the reported ventricular fibrillation, ventricular tachycardia issue has been completed. the impella device was not received from the customer. as the necessary clinical information was not provided and the device was not returned, the root cause of the ventricular fibrillation, ventricular tachycardia was not determined. this report is being filed as part of a retrospective review of historical records.
30448 | 1220648-2024-07388 | 18832838 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18832838&pc=OZD | 12/02/2022 | 03/04/2024 |  | 03/04/2024 |  |  |  | 03/04/2024 |  |  | 01/31/2023 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022040519 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/02/2022 |  |  |  |  | The user facility reported a 69-year-old female patient needing mechanical circulatory support was implanted with an impella cp. it was reported that while on the impella support, the patient experienced limb ischemia and inadequate blood supply to the impella limb. the impella cp was explanted and upgraded to the impella 5.5. the patient was successfully weaned from device, however, did later expire following explant of the impella 5.5 device device. no allegations were made towards the impella for cause of death. | [5 empty] | The investigation for the reported limb ischemia has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records.
30449 | 1220648-2024-07389 | 18832920 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18832920&pc=OZD | 12/01/2022 | 03/04/2024 |  | 03/04/2024 |  |  |  | 03/04/2024 |  |  | 08/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023270287 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 12/01/2022 |  |  |  |  | The user facility reported experiencing difficulty while implanting impella cp in 57-year-old male patient for mechanical circulatory support due to aortic dissection. ultimately the medical team aborted the procedure. | [5 empty] | The investigation into the reported event has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
30450 | 1220648-2024-07390 | 18833022 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18833022&pc=OZD | 11/19/2022 | 03/04/2024 |  | 03/04/2024 |  |  |  | 03/04/2024 |  |  | 08/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023267956 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/19/2022 |  |  |  |  | The user facility reported a 75-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed a left bundle block during impella initiation which caused heart function to decline. there was no pulsatility and very low flow. epinepherine was added and the flows improved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left tachycardia volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
30451 | 1220648-2024-07391 | 18833048 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18833048&pc=OZD | 05/24/2021 | 03/04/2024 |  | 03/04/2024 |  |  |  | 03/04/2024 |  |  | 02/28/2023 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022047444 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 05/24/2021 |  |  |  |  | The user facility reported a 70-year-old male patient in acute myocardial infarction/cardiogenic shock was escalated to an impella 5.5 device for mechanical circulatory support. while on support, the patient experienced bleeding from the injection site which required the patient to receive multiple blood products. patient is stable post intervention. | [5 empty] | The investigation into the reported access site bleeding has been completed. the cause of the bleeding could not be determined as the pump was not returned nor was clinical information provided sufficient to determine the root cause. the failure modes will be monitored and trended. instructions for use for the related event are as follows: assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿.
30452 | 1220648-2024-07392 | 18833061 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18833061&pc=OZD | 11/19/2022 | 03/04/2024 |  | 03/04/2024 |  |  |  | 03/04/2024 |  |  | 06/30/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231352 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/19/2022 |  |  |  |  | The user facility reported a 75-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced oozing around the repo sheath. the impella device was explanted early. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
30453 | 1220648-2024-07393 | 18833142 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18833142&pc=OZD | 09/25/2022 | 05/12/2025 |  | 03/04/2024 |  |  |  | 03/04/2024 |  |  | 06/30/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023238410 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 09/25/2002 |  |  |  |  | The user facility reported a 52-year-old male patient was implanted with an impella cp requiring mechanical circulatory support. it was reported that the patient experienced consistent oozing at the right femoral insertion site. the physician removed a stitch to better angle match the device. the patient improved on cp support, but experienced ongoing issues with placement. subsequently, the patient developed a mild-moderate hematoma at the right femoral site, but remained stable and did not progress. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | Additional information provided in d6, h6 and h10. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrected information was provided in b3, g3, f6, f8 and h10. b3 - corrected date of event to 25-sep-2022. g3 - the correct date received by manufacturer for the initial report is (b)(6) 2022. f6, f8 - these fields should have been left blank in the initial report. h10 - corrected truncated ifu statement.
30454 | 1220648-2024-07394 | 18833147 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18833147&pc=OZD | 09/26/2022 | 03/04/2024 |  | 03/04/2024 |  |  |  | 03/04/2024 |  |  | 06/30/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235415 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Positioning Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/26/2022 |  |  |  |  | The user facility reported a 72-year-old female with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. the patient experienced episodes of oozing due to poor angle matching and a ¿cool¿ leg. the catheter was repostioned with gauze. vascular surgery was consulted to remove the impella cp surgically. removal and surgical closure was successful. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records.
30455 | 1220648-2024-07395 | 18833201 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18833201&pc=OZD | 11/20/2022 | 05/28/2025 |  | 03/04/2024 |  |  |  | 03/04/2024 |  |  | 08/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023267146 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Death | Yes | Yes | Failure to Advance | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 11/20/2022 |  |  |  |  | The user facility reported a male patient in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported while the pump was being inserted into the patient, the patient expired. the procedure was therefore aborted. | Medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07395 was provided. | The root cause of the delivery issue was most likely patient condition related, the impella was unable to pass through the vasculature as per the clinical description provided. this report is being filed as part of a retrospective review of historical records.
30456 | 1220648-2024-07396 | 18833218 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18833218&pc=OZD | 11/20/2022 | 03/04/2024 |  | 03/04/2024 |  |  |  | 03/04/2024 |  |  | 08/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023266544 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/20/2022 |  |  |  |  | The user facility reported a 41-year-old female with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. during insertion of the impella cp the physician decided that the patient's anatomy may not be able to handle the device and decided to implant a balloon pump instead. there was no patient harm reported. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not received from the customer. however, clinical details were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, the impella was unable to pass through the vasculature due to patient's anatomy. this report is being filed as part of a retrospective review of historical records.
30457 | 1220648-2024-07397 | 18833272 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18833272&pc=OZD | 11/20/2022 | 03/04/2024 |  | 03/04/2024 |  |  |  | 03/04/2024 |  |  | 08/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023265311 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/20/2022 |  |  |  |  | The user facility reported a 64-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding and a superficial hematoma at the insertion site. to achieve hemostasis, gauze was placed under the sheath to maintain the angle of entry. the oozing stopped and dressing was then placed. the purge was switched, systemic heparin and aggrastat was stopped. the hematoma became soft and stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records.
30458 | 1220648-2024-07398 | 18833279 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18833279&pc=OZD | 07/18/2021 | 03/04/2024 |  | 03/04/2024 |  |  |  | 03/04/2024 |  |  | 04/30/2023 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 WITH SMARTASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2022072712 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult or Delayed Positioning | Activation, Positioning or Separation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 07/18/2021 |  |  |  |  | The user facility reported a (b)(6) male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an impella 5.5 pump was unable to advance through the point of anastomosis during attempted delivery. the implant was subsequently aborted and the patient was later placed on extracorporeal membrane oxygenation (ecmo) for support. there was no patient harm. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
30459 | 1220648-2024-07399 | 18833283 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18833283&pc=OZD | 11/21/2022 | 03/10/2025 |  | 03/04/2024 |  |  |  | 03/04/2024 |  |  | 08/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023259998 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/21/2022 |  |  |  |  | The user facility reported a 80-year-old female patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the decision was made for the impella to be left within the patient for at least 24 hours. however, the peel away sheath was removed and the repo was placed with some oozing. angle matching and perclose was performed to treat the oozing. the patient received 1 unit of blood transfusion and the heparin was withheld. the pump was eventually explanted. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleeding issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
30460 | 1220648-2024-07400 | 18833287 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18833287&pc=OZD | 11/21/2022 | 03/04/2024 |  | 03/04/2024 |  |  |  | 03/04/2024 |  |  | 03/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213030 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 91 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/21/2022 |  |  |  |  | The user facility reported a (b)(6) male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient had very tortuous iliaces. both the short and long sheath were tried to push impella past the tortuosity. amplatz wire was used as a buddy wire as well. the physician decided to pull impella out after multiple attempts of getting it past the iliacs. there was no reported patient harm. | [5 empty] | The investigation for the reported delivery issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, the impella was unable to pass through as the patient had very tortuous iliaces. this report is being filed as part of a retrospective review of historical records.
30461 | 1220648-2024-07401 | 18833289 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18833289&pc=OZD | 11/21/2022 | 05/20/2025 |  | 03/04/2024 |  |  |  | 03/04/2024 |  |  | 08/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023265289 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Twisted/Bent | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/21/2022 |  |  |  |  | The user facility reported a (b)(6) female noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was being placed on impella cp support, the long sheath kinked. the sheath was exchanged and the impella cp was placed. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30462 | 1220648-2024-07402 | 18833298 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18833298&pc=OZD | 11/21/2022 | 03/12/2025 |  | 03/04/2024 |  |  |  | 03/04/2024 |  |  | 07/30/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023246683 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/21/2022 |  |  |  |  | The user facility reported a 49-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported after being placed on impella cp support, the patient developed bleeding and was unable to lay flat much longer. the impella was explanted early. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
30463 | 1220648-2024-07403 | 18833319 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18833319&pc=OZD | 11/21/2022 | 03/04/2024 |  | 03/04/2024 |  |  |  | 03/04/2024 |  |  | 05/31/2024 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023230193 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/21/2022 |  |  |  |  | The user facility reported a 69-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after sheath swap, the patient developed a large hematoma on the left side of the chest and was rapidly losing pulsatility after being placed on impella cp support. multiple rounds of epinephrine (epi) were administered. manual pressure was held on the area. attempts were made to advance the sheath with no improvement on hemostasis. the impella was weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records.
30464 | 1220648-2024-07404 | 18833360 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18833360&pc=OZD | 07/20/2021 | 09/13/2024 |  | 03/04/2024 |  |  |  | 03/04/2024 |  |  | 03/31/2023 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022069677 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | Difficult to Advance | [5 empty] | Unspecified Vascular Problem | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/20/2021 |  |  |  |  | The user facility reported an impella 5.5 in (b)(6)female patient for mechanical circulatory support. it was reported that the impella 5.5 pump was unable to pass through the anastomoses. the graft was subsequently removed and placed more proximal, following which the impella insertion was successful. there was no patient harm. | Initial implant site had to be changed to an auxillary site due to patients vascular issues. it was reported the patient was withdrawn from care and expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The investigation into the reported delivery issue has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine a root cause. the most likely cause of the delivery issue was due to patient condition as the pump was unable to pass through the anastomosis. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07404 was provided in sections b1, b2, b5, d6a, d6b, h1, and h6.
30465 | 1220648-2024-07405 | 18833533 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18833533&pc=OZD | 08/13/2021 | 03/04/2024 |  | 03/04/2024 |  |  |  | 03/04/2024 |  |  | 02/28/2023 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2022071154 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/13/2021 |  |  |  |  | The user facility reported a 76-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the complainant reported that a patient lost pulse in their impella leg during the impella cp pump support. the leg subsequently became cold and the decision was made to escalate to the impella 5.5 support. upon removal of the impella cp pump, it was noted that pulses returned to the patient's leg. | [5 empty] | The investigation for the reported limb ischemia has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
30466 | 1220648-2024-07406 | 18833931 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18833931&pc=OZD | 09/23/2021 | 09/24/2024 |  | 03/04/2024 |  |  | 09/23/2021 | 03/04/2024 |  |  | 02/28/2023 |  | 03/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022052260 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Difficult to Setup or Prepare | Patient-Device Incompatibility | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/23/2021 |  |  |  |  | The user facility reported an 80-year-old male was implanted with an impella 5.5 for mechanical circulatory support. the md was having trouble with delivering the impella pump due to the patient's anatomy. after a few attempts, the md was finally able to deliver the pump successfully. | It was reported that due to some post-operative bleeding, the dr. did not want systemic heparin; mattress suture was added after the patient was moved and the patient was transfused with 1 unit of red blood cells. the procedure outcome was that patient expired as family withdrew care. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to the patient's condition of diseased and small vessels. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07406 was provided in sections b1, b2, b5, d6, h1, and h6.
30467 | 1220648-2024-07245 | 18815380 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18815380&pc=OZD | 11/01/2022 | 03/11/2025 |  | 03/01/2024 |  |  | 11/01/2022 | 03/01/2024 |  |  | 03/31/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023205655 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/01/2022 |  |  |  |  | The user facility reported a 60-year-old male was implanted with an impella cp as a bridge to heart transplant. it was reported that while on impella cp support, there was oozing at the groin access site. to resolve bleeding at groin, heparin was discontinued and forward tension sutures were placed. it was reported the patient experienced dark urine output and laboratory testing reported hemolyzed samples. the impella cp device was explanted, and the patient was escalated to an impella 5.5 device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury" ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30468 | 1220648-2024-07246 | 18815496 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18815496&pc=OZD | 10/20/2022 | 03/11/2025 |  | 03/01/2024 |  |  | 10/20/2022 | 03/01/2024 |  |  | 02/28/2023 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022052562 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/20/2022 |  |  |  |  | The user facility reported a 76 year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that during insertion, the staff observed a torn sterile sleeve. a different pump was successfully placed. the patient later expired while on the impella cp support. no allegations were made towards the impella cp for cause of death. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30469 | 1220648-2024-07247 | 18815533 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18815533&pc=OZD | 11/07/2022 | 03/11/2025 |  | 03/01/2024 |  |  | 11/07/2022 | 03/01/2024 |  |  | 08/31/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023263177 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/07/2022 |  |  |  |  | The user facility reported a 64-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the patient was successfully weaned and the impella cp device was explanted. a femoral angiography was performed after explant which revealed a vascular dissection from the left femoral artery (lfa) to the abdominal aorta. to treat the dissection, perclose was deployed and femstop was used. and the dissection was allowed to heal. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30470 | 1220648-2024-07248 | 18815818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18815818&pc=OZD | 11/08/2022 | 03/11/2025 |  | 03/01/2024 |  |  | 11/08/2022 | 03/01/2024 |  |  | 06/30/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023235809 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/08/2022 |  |  |  |  | The user facility reported a 60-year-old female noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the patient was successfully weaned and the impella cp device was removed. after explant of the impella cp device, an angiogram of the groin was performed, and a vascular dissection was revealed. the dissection was successfully treated with an angioplasty balloon. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30471 | 1220648-2024-07249 | 18816157 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18816157&pc=OZD | 11/09/2022 | 03/11/2025 |  | 03/01/2024 |  |  | 11/09/2022 | 03/01/2024 |  |  | 07/31/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023256350 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/09/2022 |  |  |  |  | The user facility reported a 51-year-old male was implanted with an impella cp for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced hematoma at the left groin access site and bleeding around the impella sheath. the bleeding was treated unsuccessfully with manual compression, and the patient was transfused with 2 units of packed red blood cells (prbc). during impella cp support it was reported the patient also experienced intermittent loss of pulsatility in left lower extremity. the decision was made to remove the impella cp device earlier than expected to treat the loss of pulsatility. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
30472 | 1220648-2024-07250 | 18816270 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18816270&pc=OZD | 11/01/2022 | 03/01/2024 |  | 03/01/2024 |  |  | 11/01/2022 | 03/01/2024 |  |  | 07/31/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2023254679 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Difficult to Insert | Failure to Advance | [5 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/01/2022 |  |  |  |  | The user facility reported a 59-year-old female patient experienced a loss of pulses (ischemia) while on the impella cp for mechanical circulatory support. multiple sheaths were used, and access was obtained for a successful femoral bypass. the access site was sutured, and dressings were placed. the patient presented with good hemostasis. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
30473 | 1220648-2024-07251 | 18816313 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18816313&pc=OZD | 10/20/2022 | 03/11/2025 |  | 03/01/2024 |  |  | 10/20/2022 | 03/01/2024 |  |  | 04/30/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023208826 |  | 0048-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/20/2022 |  |  |  |  | The user facility reported a 63 year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after being placed on impella cp support, patient developed leg thrombosis. no additional information provided for patient treatment or outcome. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ this report is being filed as part of a retrospective review of historical records. | The failure mode of ¿thrombus¿ is to be used when a thrombus is noted but cannot be connected to any other adverse events or decrease in pump performance. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions. description of the location and characteristics of the thrombus (e.g., size, shape, consistency). relevant laboratory test results, such as coagulation profiles and platelet counts. imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). medications or treatments that the patient was receiving at the time of thrombus formation. surgical or procedural details if applicable (e.g, cardiac catheterization). any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. this report is being filed as part of a retrospective review of historical records.
30474 | 1220648-2024-07252 | 18816456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18816456&pc=OZD | 11/09/2022 | 03/11/2025 |  | 03/01/2024 |  |  | 11/09/2022 | 03/01/2024 |  |  | 08/31/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023260135 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/09/2022 |  |  |  |  | The user facility reported a 74 year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced access site bleeding. the patient received two units of packed red blood cells (prbc) and a decision was made to escalate the patient to an impella 5.5 device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
30475 | 1220648-2024-07253 | 18816483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18816483&pc=OZD | 10/20/2022 | 05/15/2024 |  | 03/01/2024 |  |  | 10/20/2022 | 03/01/2024 |  |  | 08/31/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023250796 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/20/2022 |  |  |  |  | The user facility reported a 83-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. after being placed on impella cp support, the patient had bleeding at the access site that was treated with tranexamic acid and pressure. the patient continued on support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission.
30476 | 1220648-2024-07254 | 18816684 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18816684&pc=OZD | 10/01/2022 | 03/01/2024 |  | 03/01/2024 |  |  | 10/01/2022 | 03/01/2024 |  |  | 05/31/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023224586 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/01/2022 |  |  |  |  | The user facility reported a 76-year-old male with impella 5.5 for mechanical circulatory support. it was reported the user was initially unable to insert the device, the action was taken to remove the existing cp and then reattempt delivery of the impella 5.5. the 5.5 was subsequently able to be inserted successfully on the second attempt. there was no patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not received from the customer, therefore, evaluation of the device was not possible. clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue - anatomy was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
30477 | 1220648-2024-07255 | 18816958 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18816958&pc=OZD | 11/14/2022 | 03/01/2024 |  | 03/01/2024 |  |  | 11/14/2022 | 03/01/2024 |  |  | 05/17/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204914 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  |  |  | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/14/2022 |  |  |  |  | The user facility reported that a procedure was aborted on a 68 year old female patient for placement of impella cp due to extensive aortic disease. no patient injury was reported. | [5 empty] | The investigation for the reported delivery issues was completed. the device was not returned for evaluation. however, the clinical report has been reviewed. the root cause of the delivery issue was most likely related to patient condition; the device was unable to pass through due to extensive aortic disease as per the clinical description provided. this report is being filed as part of a retrospective review of historical records.
30478 | 1220648-2024-07256 | 18817106 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18817106&pc=OZD | 02/13/2024 | 03/29/2024 |  | 03/01/2024 |  |  | 02/13/2024 | 03/01/2024 |  |  | 10/31/2025 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024423300 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/13/2024 |  |  |  |  | The user facility reported a 75-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the touhy-borst valve was not locked and the implanted impella cp pump was accidentally pulled out of the lv (left ventricle) during transfer. wired re-access was attempted but the pump was unable to re-cross the valve. due to the noted issue, the decision was made to explant the pump. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ | The investigation into the positioning issues has been completed since the initial report was submitted. the product was not returned for investigation. as per the clinical details, the root cause of the positioning issue was touhy-borst related to improper technique.
30479 | 1220648-2024-07257 | 18817164 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18817164&pc=OZD | 02/14/2024 | 09/11/2024 |  | 03/01/2024 |  |  | 02/14/2024 | 03/01/2024 |  |  | 09/30/2025 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024401483 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/14/2024 |  |  |  |  | The user facility reported a 62-year-old male in acute myocardial infarction cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. during insertion of the impella 5.5 pump, bleeding was noted from the anastomosis of the graft. upon recommendations from the surgical team, the physician revisited the site following implant to resolve the bleed. the patient was given a total of six units packed red blood cells and one unit of platelets to counteract the blood loss. | It was noted that the patient expired in the operating room. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrected information for the initial mdr 1220648-2024-07257 was provided in sections b2, b5, d6, h1, h6, h10. | The investigation for the reported access site bleed has been completed. the impella device was not returned for evaluation. however, clinical information states the bleed was caused from the anastomosis of the graft and was resolved after re-visiting the graft. this is because of the improper closing around the graft. hence, the root cause is determined to be use issue.
30480 | 1220648-2024-07258 | 18817209 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18817209&pc=OZD | 11/14/2022 | 02/21/2025 |  | 03/01/2024 |  |  | 11/14/2022 | 03/01/2024 |  |  | 08/31/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023261422 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Fluid/Blood Leak | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/14/2022 |  |  |  |  | The user facility reported a 68-year-old female requiring an impella cp device for mechanical circulatory support. after being placed on impella cp support, the patient developed a slow ooze from the valve of the sheath where the impella was inserted in. the patient received a unit of blood. additionally, it was also reported that there was an unspecified leak associated with the pump. the patient's family made the decision to go comfort care. no patient harm was reported. | [5 empty] | The impella cp device has not been returned for evaluation. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation of the reported issue was completed. no data logs or device was not returned. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30481 | 1220648-2024-07259 | 18817229 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18817229&pc=OZD | 02/16/2024 | 03/28/2024 |  | 03/01/2024 |  |  | 02/16/2024 | 03/01/2024 |  |  | 10/31/2025 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024414444 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/16/2024 |  |  |  |  | The user facility reported a 76-year-old female in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support after the patient had coronary artery bypass grafting. while the patient was on support, a hematoma was noted. the cardiologist decided to remove the impella because of possible bleed and for patient transfer for high level of care. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
30482 | 1220648-2024-07260 | 18817273 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18817273&pc=OZD | 11/09/2022 | 03/11/2025 |  | 03/01/2024 |  |  | 11/09/2022 | 03/01/2024 |  |  | 07/31/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023253750 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/09/2022 |  |  |  |  | The user facility reported a 44-year-old male was implanted with an impella cp as a bridge to heart transplant. it was reported that while the patient was on impella cp support, the patient experienced bleeding at the groin access site. the bleed was managed by holding manual pressure, applying fem stop, and redressing site for angle matching and the patient was transfused with 2 units of packed red blood cells. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
30483 | 1220648-2024-07261 | 18817328 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18817328&pc=OZD | 02/19/2024 | 01/29/2025 |  | 03/01/2024 |  |  | 02/19/2024 | 03/01/2024 |  |  | 12/31/2023 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023170755 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  |  | Black Or African American | Malfunction | No | Yes | Expiration Date Error | Activation, Positioning or Separation Problem | Manufacturing, Packaging or Shipping Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/19/2024 |  |  |  |  | The user facility reported a 82-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient had an impella cp implant and the impella was expired. the pump was quickly removed and replaced. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation has been completed. the failure mode was changed from "delivery issues" to "cosmetic defect" as there were no issues with delivery reported. as per clinical information provided, the pump was implanted into the patient and it was later found out that the pump was expired. the pump was explanted after that but the pump should have been checked before implant. therefore, as per the information provided, the root cause of the issue was use related. b.7 revised as information was inadvertently omitted from initial manufacturer device report number 1220648-2024-07261. d.4 revised as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-07261. d.6a and d.6b revised from the initial submission of manufacturer device report number 1220648-2024-07261 in accordance with updated procedures. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-07261 and should not have been. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-07261 was submitted in accordance with updated procedures. h.6 revised medical device problem code since initial manufacturer device report number 1220648-2024-07261 was submitted due to investigation results. | G.1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-07261 was submitted in accordance with updated procedures. h.6 code 2199 was reported incorrectly on manufacturer device report number 1220648-2024-07261. a new code was added to health effect - impact code. h.6 addded code 3221 as it should of been on the initial manufacturer device report number 1220648-2024-07261 and was not.
30484 | 1220648-2024-07262 | 18817376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18817376&pc=OZD | 02/19/2024 | 04/09/2024 |  | 03/01/2024 |  |  | 02/19/2024 | 03/01/2024 |  |  | 09/30/2025 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024413091 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Ventricular Fibrillation | [10 empty] | Required Intervention | [5 empty] | ECMO | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/20/2024 |  |  |  |  | The user facility reported a 54-year-old male patient of unknown ethnicity and race in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced some oozing bleeding from their access site during impella support. the site was initially redressed, and angle matched while pressure was held to manage the condition. the patient experienced a ventricular fibrillation (vfib) arrest requiring shock x2 which caused the access site bleeding to worsen. five units packed red blood cells were infused, quick clot and gauze were utilized and tegaderm was placed over the site. the next day, the patient developed ischemia and a reperfusion catheter was placed. the patient was subsequently escalated to impella 5.5 support because of the condition. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation into the access site bleeding ¿ major and the limb ischemia ¿ reversible issue has been completed since the original report was submitted. per clinical information provided, the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
30485 | 1220648-2024-07263 | 18817447 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18817447&pc=OZD | 02/21/2024 | 04/09/2024 |  | 03/01/2024 |  |  | 02/21/2024 | 03/01/2024 |  |  | 11/30/2025 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024434100 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/21/2024 |  |  |  |  | The user facility reported a 77-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient experienced some bleeding, oozing from their access site following impella cp implant. forward sutures were placed, but the oozing persisted. the sutures were then replaced, and the site was redressed with angle matching technique to resolve the condition. the following day, an echo echocardiogram (ultrasound) was performed and an apical thrombus was discovered. due to the noted thrombus, the decision was made to explant the pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed and thrombus has been completed. the device nor sufficient clinical information was returned for evaluation. the root cause of access site bleeding was determined to use issue because it was mentioned that the bleeding issue was resolved after angle matching. as the necessary clinical information was not provided and the device was not returned, the root cause of the thrombus was not determined. d4-updated model number.
30486 | 1220648-2024-07264 | 18817498 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18817498&pc=OZD | 02/21/2024 | 04/09/2024 |  | 03/01/2024 |  |  | 02/21/2024 | 03/01/2024 |  |  | 09/30/2025 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024401132 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/21/2024 |  |  |  |  | The user facility reported a 53-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support and during support there was an access site bleed. bleeding was noted around the peel away sheath. when the peel away was removed there was a major bleed in the artery. vascular had to repair the bleed. it is unclear if the bleed was caused by a balloon getting stuck. the sheath came out in take. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. therefore, the root cause of the access site bleeding was most likely anticoagulation management as the bleeding was resolved with a pausing the use of aggrastat.
30487 | 1220648-2024-07265 | 18817533 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18817533&pc=OZD | 02/22/2024 | 09/12/2024 |  | 03/01/2024 |  |  | 02/22/2024 | 03/01/2024 |  |  | 08/31/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  |  |  | 2023260442 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/22/2024 |  |  |  |  | The user facility reported a 55-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a bipella cp-rp support ongoing for 4 days when the patient expired when care was withdrawn from the bipella patient. the support had signs of hemolysis throughout and hematuria. the team did not treat the hemolysis. | It was reported by medical staff that the patient expired while on impella cp support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrected information for the initial mfr 1220648-2024-07265 was provided in sections b2, b5, d6, h1, h6 and h10. | The investigation for the reported hemolysis has been completed. the device nor sufficient clinical information was returned for evaluation. the data logs were reviewed finding no motor current spikes observed. no significant positioning issue indicate the hemolysis problem. therefore, the root cause of the hemolysis could not be determined.
30488 | 1220648-2024-07266 | 18817540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18817540&pc=OZD | 11/09/2022 | 03/11/2025 |  | 03/01/2024 |  |  | 11/09/2022 | 03/01/2024 |  |  | 08/31/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023264454 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Insert | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/09/2022 |  |  |  |  | The user facility reported a 69-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that during the placement of the impella cp, delivery of the device was difficult and a forced stick on the upper 3rd of the femoral head created a small dissection of the vessel. the insertion proceeded with a placement of a 14f sheath to seal the damaged vessel while intervening and the impella cp device was inserted without issue. the patient was successfully weaned and the impella cp was explanted. after the impella cp device was explanted, the vascular damage was treated with manual pressure and placement of a stent. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
30489 | 1220648-2024-07267 | 18817597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18817597&pc=OZD | 02/24/2024 | 04/11/2024 |  | 03/01/2024 |  |  | 02/24/2024 | 03/01/2024 |  |  | 11/30/2025 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024424205 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/24/2024 |  |  |  |  | The user facility reported a 65-year-old male undergoing coronary artery bypass grafting was implanted with an impella 5.5 device for mechanical circulatory support. it was reported upon explant there were removal issues with the impella 5.5. the physician reported that the blue tuohy squeeze release would not allow for the catheter to be pulled back during the time of explant. the suture hub and catheter were pulled back and removed as a unit. no harm was done to the patient. the patient had a successful wean and was closed with suture. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported removal issues has been completed. the impella 5.5 pump was returned for investigation. the pump was visually inspected and found the tuohy borst was left unlocked with sterile sleeve left loose, unattached at both ends along with the cath anchor side. no sterile sleeve damage was found. the sterile sleeve was locked at both ends along with cath anchor cap and the catheter was able to easily glide by moving the sterile sleeve with the mechanism of blue button. no other abnormalities were found. the cause of the removal issue was use related per clinical review and field investigation.
30490 | 1220648-2024-07268 | 18817656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18817656&pc=OZD | 02/24/2024 | 04/12/2024 |  | 03/01/2024 |  |  | 02/24/2024 | 03/01/2024 |  |  | 11/30/2025 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024435805 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | FEM-FEM BYPASS | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/24/2024 |  |  |  |  | The user facility reported an 81-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient¿s access site became ischemic and was treated by fem-fem bypass. the limb was perfused, but the team later made the choice to withdraw care. the cp had continued to support the patient until the care was withdrawn. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
30491 | 1220648-2024-07269 | 18817740 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18817740&pc=OZD | 11/14/2022 | 03/01/2024 |  | 03/01/2024 |  |  | 11/14/2022 | 03/01/2024 |  |  | 08/31/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023259994 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/14/2022 |  |  |  |  | The user facility reported a 66-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after being placed on impella cp support, the patient developed leg ischemia. the patient was escalated to the impella 5.5 to bridge to transplant. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. the instructions for use states the following as it relates to limb ischemia: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
30492 | 1220648-2024-07270 | 18817765 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18817765&pc=OZD | 11/10/2022 | 03/01/2024 |  | 03/01/2024 |  |  | 11/10/2022 | 03/01/2024 |  |  | 01/31/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED |  |  | US |  | 2023178598 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/10/2022 |  |  |  |  | The user facility reported a 69-year-old male in post cardiotomy cardiogenic shock/low cardiac output syndromewas implanted with an impella 5.5 device for mechanical circulatory support. during insertion of the impella 5.5 the physicians felt resistance while the pump wire was navigated through the patient's anastomosis. despite using multiple attempts to advance the pump, the impella would not advance and the decision was made to remove the impella 5.5 and use a different wire. unfortunately, various wires and techniques were used and the pump continued to encounter resistance at the innominate ostium/aorta junction. after numerous failed attempts of implanting the pump, the decision was made to remove impella 5.5 from the patient, due to the patient's anatomy. no patient harm reported. | [5 empty] | The investigation into the reported delivery issue has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue - anatomy was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
30493 | 1220648-2024-07271 | 18817932 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18817932&pc=OZD | 11/14/2022 | 03/01/2024 |  | 03/01/2024 |  |  | 11/14/2022 | 03/01/2024 |  |  | 04/30/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218099 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/14/2022 |  |  |  |  | The user facility reported a 69-year-old male requiring an impella cp device for mechanical circulatory support. after being placed on impella cp support, the patient developed right leg ischemia. the impella cp was removed and the patient was escalated to an impella 5.5. no patient harm was reported. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella cp device has not been returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. this report is being filed as part of a retrospective review of historical records.
30494 | 1220648-2024-07272 | 18818033 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18818033&pc=OZD | 11/14/2022 | 03/01/2024 |  | 03/01/2024 |  |  | 11/14/2022 | 03/01/2024 |  |  | 06/30/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023237617 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/14/2022 |  |  |  |  | The user facility reported a 47-year-old male requiring an impella cp device for mechanical circulatory support. after being placed on impella cp support, an magnetic resonance imaging (mri) was performed and it was revealed that the patient had a stroke. it was reported that the patient recovered contractility but were unaware of the patient¿s neurological status. | [5 empty] | The investigation for the reported stroke/cva issue has been completed. the impella cp device has not been returned for evaluation. as the necessary clinical information was not provided and the device was not returned for investigation, the root cause of the stroke, cva was not determined. this report is being filed as part of a retrospective review of historical records.
30495 | 1220648-2024-07273 | 18818205 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18818205&pc=OZD | 11/10/2022 | 03/01/2024 |  | 03/01/2024 |  |  | 11/10/2022 | 03/01/2024 |  |  | 03/31/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023212124 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/10/2022 |  |  |  |  | The user facility reported a 35-year-old female was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, an external bypass was performed for distal leg perfusion. | [5 empty] | The investigation of the reported issue has been completed. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
30496 | 1220648-2024-07274 | 18818210 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18818210&pc=OZD | 11/02/2022 | 03/01/2024 |  | 03/01/2024 |  |  | 11/02/2022 | 03/01/2024 |  |  | 10/31/2023 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED |  |  | US |  | 2022151384 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/02/2022 |  |  |  |  | The user facility reported a 86-year-old male patient needing the impella cp for mechanical circulatory support. the impella cp was unable to be placed due to abdominal thoracic calcification. the case was not completed. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
30497 | 1220648-2024-07275 | 18818397 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18818397&pc=OZD | 11/10/2022 | 03/01/2024 |  | 03/01/2024 |  |  | 11/10/2022 | 03/01/2024 |  |  | 07/31/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023256610 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/10/2022 |  |  |  |  | The user facility reported a 76-year-old female noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient developed limb ischemia and a fasciotomy was performed for distal leg perfusion. | [5 empty] | The investigation of the reported issue has been completed. to determine the true root cause of limb ischemia, clinical information would need to be assessed in conjunction with evidence from the returned device. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
30498 | 1220648-2024-07276 | 18818408 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18818408&pc=OZD | 11/15/2022 | 03/01/2024 |  | 03/01/2024 |  |  | 11/15/2022 | 03/01/2024 |  |  | 04/30/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023211202 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  |  | Injury | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/15/2022 |  |  |  |  | The user facility reported a 75-year-old male being placed on impella 5.5 required mechanical circulatory support. the complainant reported that resistance was met at the aorta during right axillary implant due to existing stenosis. the pump was removed due to the patient's anatomy and the impella 5.5 device was subsequently placed via direct access. there was no patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella 5.5 device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the root cause of the delivery issue was anatomy most likely due to the patient condition. this report is being filed as part of a retrospective review of historical records.
30499 | 1220648-2024-07277 | 18818516 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18818516&pc=OZD | 11/15/2022 | 03/17/2025 |  | 03/01/2024 |  |  | 11/15/2022 | 03/01/2024 |  |  | 08/31/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023262775 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/15/2022 |  |  |  |  | The user facility reported a 43-year-old male being placed on impella cp required mechanical circulatory support. after being placed on impella cp support, the patient was noted to have ¿a lot¿ of bleeding at the insertion site. the dressing was changed and insertion angle was elevated. the site stopped bleeding. the patient was transferred while on impella cp support and it had been observed that hemolysis had begun. the impella cp had been explanted and was escalated to impella 5.5 with central venous-arterial (va) extra-corporeal membrane oxygenation (ecmo). | [5 empty] | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿if the impella catheter is positioned too far into the left ventricle, the patient will not receive the benefit of impella catheter support. blood will enter the inlet area and exit the outlet area within the ventricle. obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood-colored urine.¿ impella catheter in papillary muscle: ¿if the inlet area of the catheter is lodged adjacent to the papillary muscle, the inflow may be obstructed, resulting in suction alarms. this positioning is also likely to place the outlet area too close to the aortic valve, which can cause outflow at the level of the aortic valve with blood streaming back into the ventricle, resulting in turbulent flow and hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: hemolysis: ¿patient conditions¿including catheter position, pre-existing medical conditions, and small left ventricular volumes¿may also play a role in patient susceptibility to hemolysis." "hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
30500 | 1220648-2024-07278 | 18818550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18818550&pc=OZD | 11/02/2022 | 03/01/2024 |  | 03/01/2024 |  |  | 11/02/2022 | 03/01/2024 |  |  | 07/31/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023252379 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/02/2022 |  |  |  |  | The user facility reported an impella cp implanted in a 85-year-old female patient for mechanical circulatory support. it was reported that while on support, no pulses were noted in the left foot, minor oozing from the groin site, and no signs of hematoma. manual pressure was applied, and insertion site dressings were changed. patient had significant bleeding from the right femoral artery (a-line) sheath overnight and patient received blood products. after the impella site was redressed, no bleeding or hematoma was noted from the impella site. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury" this report is being filed as part of a retrospective review of historical records.
30501 | 1220648-2024-07279 | 18818721 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18818721&pc=OZD | 11/16/2022 | 03/26/2024 |  | 03/01/2024 |  |  | 11/16/2022 | 03/01/2024 |  |  | 08/31/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023265367 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/16/2022 |  |  |  |  | The user facility reported a 58-year-old male noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that upon entering the left ventricle, the impella caused multiple rhythm issues due to a shallow ventricle. the impella insertion was aborted and the percutaneous coronary intervention (pci) was performed without support. | [5 empty] | The impella device has not been returned for evaluation. upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: potential adverse events (united states) arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported issue has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the reported issue could not be determined. correcting information reported in the initial report: there are no relevant instructions for use for the related event and product reported.
30502 | 1220648-2024-07280 | 18819018 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18819018&pc=OZD | 11/03/2022 | 09/12/2024 |  | 03/01/2024 |  |  | 11/03/2022 | 03/01/2024 |  |  | 07/31/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023254170 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/03/2022 |  |  |  |  | The user facility reported an impella cp in 82-year-old female patient for mechanical circulatory support. it was reported that while on support, the patient was experiencing distal limb ischemia. patient¿s left leg was immobilized, and foot was warmer since femorofemoral bypass, but there was still no pulse. patient remained oliguric. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07280 was provided.
30503 | 1220648-2024-07282 | 18819572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18819572&pc=OZD | 11/03/2022 | 09/12/2024 |  | 03/01/2024 |  |  | 11/03/2022 | 03/01/2024 |  |  | 08/31/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023260061 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Sepsis | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/03/2022 |  |  |  |  | The user facility reported an impella cp in 58-year-old male patient for mechanical circulatory support. it was reported that while the patient was under impella cp support, the patient experienced ventricular tachycardia (vt) when moving to the intensive care unit (icu) bed. additionally, blood work was done and suggested sepsis. no additional information was provided for the treatment of the vt and suspected sepsis. the patient continued support. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07282 was provided.
30504 | 1220648-2024-07283 | 18819671 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18819671&pc=OZD | 11/16/2022 | 03/01/2024 |  | 03/01/2024 |  |  | 11/16/2022 | 03/01/2024 |  |  | 07/31/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023257472 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 93 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/16/2022 |  |  |  |  | The user facility reported that while the impella cp support was being placed on the (b)(6) female patient, the pump was not able to pass the motor housing through the aortoiliac bifurcation. the left ventricle was wired again and the pump was inserted. the pump was able to pass pump through sheath and completely into the thoracic aorta but was hanging up on calcification. the procedure was aborted due to the patient¿s anatomy. there was no patient harm reported. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not received from the customer. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, the pump was unable to pass through aortoiliac bifurcation and vessel calcification. this report is being filed as part of a retrospective review of historical records.
30505 | 1220648-2024-07281 | 18819940 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18819940&pc=OZD | 11/16/2022 | 03/26/2024 |  | 03/01/2024 |  |  | 11/16/2022 | 03/01/2024 |  |  | 08/31/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023264044 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/16/2022 |  |  |  |  | The user facility reported a 78-year-old male in st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. after being placed on impella cp support, the patient developed a hematoma. the impella cp was repositioned, and heparin was discontinued. the patient later expired while on the impella cp support. no allegations were made towards the impella cp for cause of death. | The user facility reported a 78-year-old male in st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. after being placed on impella cp support, the patient developed a hematoma. heparin was discontinued. the patient later expired while on the impella cp support. no allegations were made towards the impella cp for cause of death. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported hematoma has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hematoma could not be determined. instructions for use for the related event are as follows: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. b5 revised to reflect that the impella was not repositioned.
30506 | 1220648-2024-07284 | 18820245 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18820245&pc=OZD | 11/03/2022 | 03/01/2024 |  | 03/01/2024 |  |  | 11/03/2022 | 03/01/2024 |  |  | 07/31/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023255788 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/03/2022 |  |  |  |  | The user facility reported an impella cp in a 77-year-old female patient for mechanical circulatory support. it was reported that while on support, the patient began to ooze from the groin site. the bleeding was resolved by re-suturing of the catheter and applying fem stop. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury" this report is being filed as part of a retrospective review of historical records.
30507 | 1220648-2024-07285 | 18820300 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18820300&pc=OZD | 11/11/2022 | 03/11/2025 |  | 03/01/2024 |  |  | 11/11/2022 | 03/01/2024 |  |  | 06/30/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023236613 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/11/2022 |  |  |  |  | The user facility reported a 70-year-old male in st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that while on impella cp support, there was pulsatile bleeding at the right femoral access site. the physician decided to wean and explant impella cp device early to resolve bleeding. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma. remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site. potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
30508 | 1220648-2024-07286 | 18820347 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18820347&pc=OZD | 11/03/2022 | 03/07/2025 |  | 03/01/2024 |  |  | 11/03/2022 | 03/01/2024 |  |  | 07/31/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023256565 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/03/2022 |  |  |  |  | The user facility reported a (b)(6) male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. unfortunately, ventricular fibrillation occurred post stent to the right coronary artery (rca) and the patient was shocked 28 times. no patient harm was reported. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | Access site bleeding: this complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. vf/vt/af/at: analysis summary: cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease ¿timing of the arrythmia event in relation to impella support (during or post-implant) ¿electrocardiogram (ekg) recordings of the arrhythmia episodes ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities ¿response to any antiarrhythmic treatments or interventions during the event ¿any documented device-related issues or complications during impella support ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand ¿review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined.
30509 | 1220648-2024-07287 | 18820808 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18820808&pc=OZD | 11/04/2022 | 09/12/2024 |  | 03/01/2024 |  |  | 11/04/2022 | 03/01/2024 |  |  | 07/31/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023255484 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Improper Flow or Infusion | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/04/2022 |  |  |  |  | The user facility reported a 54-year-old female patient with an impella cp needing mechanical circulatory support. it was reported that the pump experienced suction which caused decreased flows. the blood pressure fell drastically as both ventricles collapsed around the impella cp. volume did not help and extracorporeal membrane oxygenation (ecmo) was inserted. the intra-aortic balloon pump (iabp) was removed. the impella was used as a vent as the patient was weaned off bypass slowly with suction. | It was reported that the patient's family decided to withdraw care and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07287 was provided in sections b1, b2, b5, h1, and h6.
30510 | 1220648-2024-07288 | 18821518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18821518&pc=OZD | 11/04/2022 | 03/01/2024 |  | 03/01/2024 |  |  | 11/04/2022 | 03/01/2024 |  |  | 08/31/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023262354 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/04/2022 |  |  |  |  | The user facility reported a 86-year-old male patient needing an impella cp for mechanical circulatory support. during insertion of the impella cp, it was difficult to insert the device due to angulation of the vessel, therefore serial dilation was performed. the impella cp sheath was placed, but there was kinking observed via fluoro at the distal end of the sheath. they were able to slide the impella cp past the kink while pulling back on the sheath. the physician was unable to cross the chronic total occlusion (cto) of right coronary artery (rca). a decision was made to not to proceed with intervention of the rca at this time. the impella cp was weaned and removed. intervention could not be continued with the impella cp and it was explanted early. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
30511 | 1220648-2024-07289 | 18821905 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18821905&pc=OZD | 11/04/2022 | 03/01/2024 |  | 03/01/2024 |  |  | 11/04/2022 | 03/01/2024 |  |  | 07/31/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023259492 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 11/04/2022 |  |  |  |  | The user facility reported a 64-year-old male in acute myocardial infarction/cardiogenic shock, needing an impella cp for mechanical circulatory support . it was reported that during placement, the pump was unable to cross aortic valve (av). a trans-esoophageal (tee) revealed the av was not moving with a large clot on the sub-valvular side of the av. the impella cp placement was aborted. no patient harm was reported. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not received from the customer. however, clinical details were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, the impella was unable to pass through the aortic valve due to patient's anatomy. this report is being filed as part of a retrospective review of historical records.
30512 | 1220648-2024-07290 | 18822040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18822040&pc=OZD | 11/04/2022 | 09/12/2024 |  | 03/01/2024 |  |  | 11/04/2022 | 03/01/2024 |  |  | 06/30/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023238006 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  |  |  | Death | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 11/04/2022 |  |  |  |  | The user facility reported an 81-year-old male patient on an impella cp for mechanical circulatory support. while on support the echocardiogram indicated pericardial tamponad. a drain was placed, and 850 ml of blood drained from the patient¿s chest. the patient later expired while on the impella cp support. no allegations were made towards the impella cp for cause of death. | It was reported after impella cp support began, flows were lower than expected at 1.4l/min at p5. the patient continued to deteriorate and required increased vasopressor medication and a need for added volume. an echocardiogram was performed showing a pericardial tamponade, and a drain was placed. 850 ml of blood was drained from patients' chest. two units of blood cells were then transfused and impella cp flows improved to 2.7 l/min at p6. while preparing the patient for transfer, the patient again became hemodynamically unstable, and the flows on the impella decreased requiring increased vasopressor medication. more blood was found in the pericardial space, and the physician was able to drain the blood out of the pericardial space, which slightly improved the patient's vital signs. after a short time, the patient's blood pressure dropped and more blood was found to be collecting around the heart. flows on the impella dropped to 0.8 l/min on p2. a final attempt was made to stop the bleeding with surgiflo but was unsuccessful. the decision was made to discontinue further measures, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury. assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07290 was provided in sections b1, b2, b5, d6a, d6b, h1, and h6.
30513 | 1220648-2024-07291 | 18822199 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18822199&pc=OZD | 11/04/2022 | 09/12/2024 |  | 03/01/2024 |  |  | 11/04/2022 | 03/01/2024 |  |  | 06/30/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023247649 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Thrombosis/Thrombus | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup |  |  |  | 11/04/2022 |  |  |  |  | The user facility reported a 68-year-old male patient needing mechanical circulatory support. it was reported that while on impella cp support, the patient had runs of supraventricular tachycardia (svt) and a new left ventricle thrombus was detected via echocardiogram (echo). the patient later expired while on the impella cp support. no allegations were made towards the impella cp for cause of death. | Additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07291 was provided.
30514 | 1220648-2024-07292 | 18822278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18822278&pc=OZD | 11/05/2022 | 03/01/2024 |  | 03/01/2024 |  |  | 11/05/2022 | 03/01/2024 |  |  | 06/30/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023240887 |  | 0048-0045 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 11/05/2022 |  |  |  |  | The user facility reported that a 61-year-old female patient needing mechanical circulatory support. it was reported that post percutaneous coronary intervention, the patient started to deteriorate. the patient was found in ventricular tachycardia/ ventricular fibrillation and shocked multiple times. the patient was placed on impella for hemodynamic support. the impella was placed in the right femoral artery but post implant, the patient continued to have vt/vf. additionally, the patient was oozing from the impella site. femstop was applied which controlled the bleeding. no arrhythmia was noted overnight. patient was still vented. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
30515 | 1220648-2024-07293 | 18822322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18822322&pc=OZD | 03/30/2023 | 03/20/2024 |  | 03/01/2024 |  |  | 01/31/2024 | 03/01/2024 |  |  | 10/31/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024297631 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/31/2024 |  |  |  |  | It was reported via abiomed¿s long-term outcome and quality indicator impella registry that a 66 year old female patient with an impella cp developed an access site bleed post procedure. the patient was treated with two units of packed red blood cells (prbc). no further details were provided. | [5 empty] | The impella device has not been returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. the instructions for use states the following: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ it also states to "assess access site for bleeding and hematoma.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
30516 | 1220648-2024-07294 | 18822323 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18822323&pc=OZD | 07/16/2023 | 03/20/2024 |  | 03/01/2024 |  |  |  | 03/01/2024 |  |  | 05/31/2025 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024355567 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 45 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemolysis | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/01/2024 |  |  |  |  | It was reported via abiomed¿s long-term outcome and quality indicator impella registry that a 45 year old male patient implanted with an impella 5.5 as a bridge to transplant developed hemolysis and ventricular tachycardia (vt) during support. the hemolysis was treated with an intravenous iron infusion. the vt was treated with cardioversion. no further issues were reported. the patient was successfully transplanted. | Additional information was received via abiomed¿s long-term outcome and quality indicator impella registry indicating this patient to have endured an lv thrombus during support, with a possible relationship to the impella device. intervention was performed via bivialrudin drip. |  |  |  |  | The impella device was not returned for evaluation. a supplemental report will be submitted upon completion of the investigation. the instructions for use for hemolysis states ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected low s or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation for the reported hemolysis, thrombus, and ventricular tachycardia (vt) has been completed. the impella device was not returned for evaluation. data logs showed low flow occurred for entire case that triggered auto suction response & suction alarms. logs also showed about 2.5 hours into the case, pump was in improper position (ventricle) that triggered impella position in ventricle. the root cause of hemolysis was most likely improper positioning of the device. the root cause of the vt and thrombus could not be determined as clinical details provided were insufficient. | The investigation remains in progress. a supplemental mdr will be filed at the completion of our investigation.
30517 | 1220648-2024-07295 | 18822350 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18822350&pc=OZD | 11/05/2022 | 03/01/2024 |  | 03/01/2024 |  |  | 11/05/2022 | 03/01/2024 |  |  | 07/31/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023252314 |  | 0048-0045 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 11/05/2022 |  |  |  |  | The user facility reported a 58-year-old female patient went into cardiac arrest and was placed on an impella cp for mechanical circulatory support. after placement, the patient¿s hemoglobin dropped and the patient experienced bleeding. 2 units of blood and 1 l of crystalloids wee given to the patient. femstop was placed but the patient was still bleeding. the dressing was removed and repositioning sheath noted to not be fully advanced and there was no forward tension suture. a forward tension suture was applied and the bleeding was resolved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
30518 | 1220648-2024-07296 | 18822421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18822421&pc=OZD | 11/05/2022 | 03/01/2024 |  | 03/01/2024 |  |  | 11/05/2022 | 03/01/2024 |  |  | 06/30/2024 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023238695 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 27 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 11/05/2022 |  |  |  |  | The user facility reported a 27-year-old male patient needing mechanical circulatory support. it was reported that after being placed on impella cp support, the patient had platelets of 30,000 and had bleeding from the insertion site. the patient had successful hemostasis while the fem-stop was on. however, once removed, the site still bled. the patient was later taken to the operating room for surgical closure of the wound. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records.
30519 | 1220648-2024-07297 | 18822540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18822540&pc=OZD | 02/13/2024 | 04/17/2024 |  | 03/01/2024 |  |  | 02/13/2024 | 03/01/2024 |  |  | 09/30/2025 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024401203 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  |  | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A 72-year-old patient fell into cardiogenic shock after an elective cardiotomy for heart valve replacement. it was decided to use the impella 5.5 smartassist heart pump for hemodynamic support. the patient was examined by computed tomography and tension was applied to the impella 5.5 patient cable during positioning on the ct table. bleeding occurred at the axillary access site. the impella 5.5 heart pump was repositioned. the access site/anastomosis was surgically repaired to stop the bleeding. after the incident, the patient received blood transfusions (2 units of packed red cells and 3 units of blood plasma). the patient was then stable at a low level and was supported by the impella 5.5 heart pump for two more days. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major and the positioning issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of access site bleeding was most likely use related since traction was applied to the impella catheter leading to bleeding event during repositioning on ct table. based on clinical description, the root cause of positioning issue was most likely use related.
30520 | 1220648-2024-07298 | 18822550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18822550&pc=OZD | 02/25/2024 | 01/13/2025 |  | 03/01/2024 |  |  | 02/25/2024 | 03/01/2024 |  |  |  |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025442119 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 69 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Injury | Yes | Yes | Fluid/Blood Leak | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 02/25/2024 |  |  |  |  | The user facility reported a 69-year-old male with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that while on support, the staff noticed a tiny leakage of blood from the red impella plug. the impella was replaced, and there was no patient harm reported. | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured vascular complication with a "definite" relationship to the impella device used: ¿immediately postop in the cicu there was noted to be blood extravasating from the proximal portion of the catheter which is highly abnormal. consultation with manufacture and representative indicated pump disruption which could result in catastrophic failure. for this reason was decided to move the patient immediately back to the operating room for removal of the defective catheter and placement of a new impella 5.5¿. the patient recovered with no sequelae. |  |  |  |  | The investigation for the reported issue has not yet been completed. upon its completion, a follow up report will be submitted. | The investigation into the product damage (hole in catheter) issue has been completed since the initial report was submitted. the device was returned, visually inspected, and tested. upon investigation of the pump, blood was observed within the red handle. two puncture holes were found in the catheter. the root cause of the product damage hole in catheter was use related. | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The investigation vascular damage - peripheral vessel has not yet been completed. should new information become available a follow up report will be submitted. instructions for use: impella 5.5 with smartassist for use during section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ the following have been reported incorrectly or missing from manufacturer device report. b1 adverse event and product problem b2 added b5 event description added d6a and d6b revised from the initial submission in accordance with updated procedures. f6/f8-should have been left blank. g1 reporting contact email revised from the initial submission in accordance with updated procedures. g1 reporting contact fax number missing. g2 report source; study missing h1 type of reportable event h6 health effect - clinical code and health effect - impact code and component code h10 instructions for use.
30521 | 1220648-2024-07299 | 18822582 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18822582&pc=OZD | 02/13/2024 | 04/01/2024 |  | 03/01/2024 |  |  | 02/13/2024 | 03/01/2024 |  |  | 11/30/2025 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024428642 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/13/2024 |  |  |  |  | The user facility reported a 34-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a 5.5 pump placed for support of patient with weeks of cardiac symptoms. the patient was admitted in cgs (cardiogenic shock) and in need of cardioversion and support with ecmo (extracorporeal membrane oxygenation) and the 5.5 pump. the pump supported but with signs of hemolysis. after 8 days the pump was explanted due to the hemolysis and elevated labs. the ldh (lactate dehydrogenase) had risen. the pump was replaced, and support proceeded. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the hemolysis issue has been completed since the initial report was submitted. the product was not returned for investigation. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review.
30522 | 1220648-2024-07300 | 18822589 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18822589&pc=OZD | 02/16/2024 | 04/04/2024 |  | 03/01/2024 |  |  | 02/16/2024 | 03/01/2024 |  |  | 11/30/2025 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024433205 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/16/2024 |  |  |  |  | The user facility reported a 61-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient experienced bleeding from their access site following impella implant. revealed a right common femoral artery extravasation. blood products were transfused to treat the condition and then a surgical cutdown was performed at the site. the pump was subsequently removed and an impella 5.5 pump was placed for ongoing support. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the access site bleeding ¿ major issue has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the vascular damage was unable to be determined as no product was returned for evaluation and insufficient clinical details were provided.
30523 | 1220648-2024-07301 | 18822651 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18822651&pc=OZD | 02/18/2024 | 09/12/2024 |  | 03/01/2024 |  |  | 02/18/2024 | 03/01/2024 |  |  |  |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1263210 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic | Black Or African American | Death | No | Yes | Battery Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup | Followup |  |  | 02/18/2024 |  |  |  |  | The user facility reported a 47-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the automated impella controller (aic) had a battery failure (red alarm) at startup. troubleshooting was performed and the aic was rebooted unsuccessfully. a new aic was used. | Survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was received from the customer on 26-feb-2024. upon investigation closure, a supplemental mdr will be filed. data review: aic logs confirmed the reported failure. there were 2 instances of battery failure alarm (transport connection blown/missing ¿ event set #360). see ic4709 aic logs from complaint date in attachments section. batttest showed cell imbalance with battery 2 vcell 1 (value of 0mv) which caused the battery packs internal electronics to shut down. (see ic4709 battery analysis according to batttest in attachments). device analysis: the reported failure mode was able to be reproduced on an engineering lab bench. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure. before sending this console back to the customer, fs was instructed to replace the battery kit [0042-3016], validate charging, and perform a full functional check on the console and the optical system [0042-7316]. | The investigation into the aic - battery failure (red alarm) issue has been completed since the original report was filed. the console was returned and tested after arriving in fs. the reported failure mode was able to be reproduced on an engineering lab bench. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure. | B2: date of death is unknown. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07301 was provided.
30524 | 1220648-2024-07302 | 18822656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18822656&pc=OZD | 02/18/2024 | 04/08/2024 |  | 03/01/2024 |  |  | 02/19/2024 | 03/01/2024 |  |  | 10/08/2025 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025439309 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/19/2024 |  |  |  |  | The user facility reported a 64-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient on impella support experienced bleeding and manual pressure was being applied to the left femoral site. the transvenous pacer sheath in the left common femoral vein (cfv) had been partially dislodged so the decision was made to completely remove this sheath. on reassessment some blood pooling was noted so manual pressure held again. after extensive pressure held, new significant oozing was noted around the impella sheath with a hematoma forming subcutaneous. multiple unsuccessful attempts were made to troubleshoot to stop the bleeding. bleeding continued around the site and the leg on the impella side began to get larger. argatroban had been stopped for 2 hours at this point. impella cp explanted. one-unit prbc (packed red blood cells) being transfused during explant. hemostasis was never achieved. the physician noted there was pulsatile arterial bleeding at the site. balloon tamponade was attempted unsuccessfully, due to unable to gain wire access contralaterally. the left common femoral was then opened for exploration. at this point the patient coded, cpr administered, and rosc (return of spontaneous circulation) achieved. patient was noted to have vascular injury with dissection. manual pressure attempted, and patient coded again. this time it was unable to obtain rosc. the patient expired. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The investigation into vascular damage - peripheral vessel issue has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the vascular damage - peripheral vessel issue was not determined since product was not returned.
30525 | 1220648-2024-07304 | 18822661 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18822661&pc=OZD | 02/21/2024 | 11/13/2025 |  | 03/01/2024 |  |  | 02/21/2024 | 03/01/2024 |  |  |  |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1798913 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/21/2024 |  |  |  |  | No patient involvement: us complainant reported that the automated impella controller (aic) had display issues. words were not showing up on the display or disappearing when the selector knob would highlight it. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30526 | 1220648-2024-07305 | 18822669 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18822669&pc=OZD | 02/22/2024 | 04/08/2024 |  | 03/01/2024 |  |  | 02/22/2024 | 03/01/2024 |  |  | 09/30/2025 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025453992 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/22/2024 |  |  |  |  | The user facility reported a 77-year-old male patient of unknown race and ethnicity with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient suffered ventricular tachycardia during impella 5.5 support. the patient was defibrillated twice, and amiodarone was pushed to counteract the condition. it was noted that the impella was too deep at the time of the noted condition and positioned towards the mitral valve (mv). the pump was repositioned away from the mv and pulled back to prevent a recurrence. support continued with the implanted pump. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the ventricular tachycardia (vf/vt) issue has been completed since the initial report was submitted. the device was not returned for investigation as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined.
30527 | 1220648-2024-07306 | 18822675 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18822675&pc=OZD | 02/24/2024 | 04/23/2024 |  | 03/01/2024 |  |  | 02/24/2024 | 03/01/2024 |  |  | 08/31/2025 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP 380 SMART ASSIST SET, EU |  |  | ABIOMED, INC |  |  | US |  | 2024400663 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/24/2024 |  |  |  |  | The user facility reported a 65-year-old male patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. it was reported that patient developed hemolysis while on support and blood products were infused as treatment. the pump positioning was confirmed to be correct the physician reported it could that the patient¿s catecholamines are reduced. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the hemolysis issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product, data logs were returned for review.
30528 | 1220648-2024-07307 | 18822714 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18822714&pc=OZD | 02/24/2024 | 04/11/2024 |  | 03/01/2024 |  |  | 02/24/2024 | 03/01/2024 |  |  | 06/30/2025 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024430739 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  |  |  | Malfunction | No | Yes | Insufficient Flow or Under Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/24/2024 |  |  |  |  | The user facility reported a 52-year-old female patient of unknown race and ethnicity in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that when attempting to wean the patient from bypass, the implanted impella cp pump could not obtain reliable flow rates at p2 without suction alarms. despite undertaking proper troubleshooting steps, the low flow rates continued, and the decision was made to explant the pump. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported low pump flow has been completed. the impella device was not returned for investigation. data logs were reviewed which showed many positioning alarms occurred in aorta and left ventricle. suction alarms observed at the end of the case with low flows at p-2. no motor current spikes were found. positioning alarms were not correlated to the low flow issue. the cause of the low pump flow issue was not determined since product was not returned and inconclusive evidence based on log, clinical review. the cause of the positioning issue most likely was use related due to improper technique.
30529 | 1220648-2024-07308 | 18822998 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18822998&pc=OZD | 01/10/2024 | 09/12/2024 |  | 03/01/2024 |  |  | 02/02/2024 | 03/01/2024 |  |  | 06/30/2025 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024373986 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/02/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry indicating a 60 year old male patient presenting with cardiomyopathy for impella support had endured ventricular arrhythmia with a "possible" relationship to the impella device. intervention was required via, "calcium, amio bolus, lidocaine, magnesium, defib x 2." | It was noted the patient's family decided to withdraw care from the patient, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07308 was provided in sections b2, b5, h1, and h6. | The investigation into the ventricular arrhythmia has been completed since the initial report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the ventricular arrhythmia was not determined.
30530 | 1220648-2024-07309 | 18823037 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18823037&pc=OZD | 01/10/2024 | 03/22/2024 |  | 03/01/2024 |  |  | 02/02/2024 | 03/01/2024 |  |  | 10/31/2025 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024420774 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | No | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/02/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry indicating a 65 year old male patient presenting for high risk percutaneous coronary intervention using the impella cp for hemodynamic support had endured an access site bleed with a "possible" relationship to the impella device. blood transfusions were required. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
30531 | 1220648-2024-07310 | 18823052 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18823052&pc=OZD | 01/19/2024 | 03/22/2024 |  | 03/01/2024 |  |  | 02/02/2024 | 03/01/2024 |  |  | 07/31/2025 |  | 03/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024376652 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/02/2024 |  |  |  |  | A notification was received via abiomed¿s long-term outcome and quality indicator impella registry indicating a 78 year old male patient had endured ventricular arrhythmia with a "possible" relationship to the impella device. an amiodarone drip was started, "secondary to nsvt runs." | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the ventricular arrhythmia has been completed since the initial report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the ventricular arrhythmia was not determined.
30532 | 1220648-2024-07179 | 18806669 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18806669&pc=OZD | 09/21/2022 | 02/29/2024 |  | 02/29/2024 |  |  | 09/21/2022 | 02/29/2024 |  |  | 12/31/2023 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023163036 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/30/2024 |  |  |  |  | The user facility reported a 61-year-old male needing mechanical circulatory support. during the impella cp support, the automated impella controller (aic) displayed a red alarm, indicating 'impella position in aorta' following patient movement/nursing. it was reported that the canula made a big movement along the mitral valve with the inlet positioned very close to aortic valve. the pump was repositioned with no allegations of patient harm. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the positioning issue was patient movement based on the impella position in aorta which occurred after patient movement. instructions for use for the related event are as follows: ¿impella catheter in papillary muscle if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ this report is being filed as part of a retrospective review of historical records.
30533 | 1220648-2024-07180 | 18806728 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18806728&pc=OZD | 02/16/2024 | 04/09/2024 |  | 02/29/2024 |  |  | 02/16/2024 | 02/29/2024 |  |  | 10/31/2025 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024410610 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/16/2024 |  |  |  |  | The user facility reported that a 62-year-old male in st-elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that an implanted impella cp pump began to experience a high motor current and saturated flow during patient support. as the patient's hemodynamics were stable, the decision was made to explant the pump. there was no patient harm reported. | [5 empty] | The impella device was received from the customer and an investigation of the device is ongoing. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported saturation issue has been completed. pump was returned for evaluation. no biomaterial found in purge gap or on impeller blade. no evidence of blood ingress in purge gap. pump was run in bucket of water at p-9 for approximately 10 minutes and did have clipping of pump flow at 4.3 but saturated flow did not reproduce. pump able to purge successfully. pump teardown was performed and black blistering was found on the motor magnet. this black blistering is due to neodymium corrosion as a result of purge fluid getting under iron corrosion on the motor. data logs from the field were reviewed and it was found that a rise in motor current over a period of 12 hours before pump flows become saturated. the root cause of the saturated flow was most likely due to magnet coating separation resulting in blisters forming on the magnet. added-h6 impact code 4627.
30534 | 1220648-2024-07181 | 18806895 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18806895&pc=OZD | 02/19/2024 | 04/22/2024 |  | 02/29/2024 |  |  | 02/19/2024 | 02/29/2024 |  |  |  |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025441838 |  | 1000459 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 32 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/19/2024 |  |  |  |  | The user facility reported a 32-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant in japan had an impella 5.5 placed for support of a patient with myocarditis. the patient was also on support with extracorporeal membrane oxygenation (ecmo). the pump had a leak at the plug/shaft. the blood loss did not cause any harm or lasting injury that prompted any blood products. the team remedied with gauze only. the pump remains on for support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported device leakage has been completed. the impella device was returned for evaluation and confirmed a hole in the catheter around the 30cm mark. this hole then allowed blood to ingress through the catheter and into the red handle. according to the clinical, on the second day of impella 5.5 support blood began to leak from the catheter. no attempts were made to resolve the issue. the device remained in the patient and was wrapped in clean gauze. data log analysis was not necessary for this complaint. the cause of the hole in catheter was use related. added- d9 device return date d4-updated model number.
30535 | 1220648-2024-07182 | 18806938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18806938&pc=OZD | 09/21/2022 | 03/11/2025 |  | 02/29/2024 |  |  | 09/21/2022 | 02/29/2024 |  |  | 06/30/2024 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023238189 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/21/2022 |  |  |  |  | The user facility reported a 81-year-old female in st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported while patient was on impella cp support, the impella access site at groin was bleeding. the site was redressed and patient was transfused with 3 units of packed red blood cells. the patient continued on support. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30536 | 1220648-2024-07183 | 18807128 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18807128&pc=OZD | 02/16/2024 | 03/21/2024 |  | 02/29/2024 |  |  | 02/16/2024 | 02/29/2024 |  |  |  |  | 02/29/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1465594 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure of Device to Self-Test | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/16/2024 |  |  |  |  | The japan user facility reported that an automated impella controller (aic) in storage experienced a self-test failure and would not start up properly. the aic was rebooted to troubleshoot the issue, but the failure remained. there was no patient harm or involvement. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the reported console (aic) self-check error has been completed. the aic was returned for evaluation. logs were not available for review. during evaluation of the console, the pbm pca was physically inspected which revealed corrosion to the circuit traces that carry the communication signals between the 4-in-1 and the battery 1 circuitry on the pbm. this type of corrosion is known to be caused by leakage of the electrolytic supercapacitors. therefore, the cause of the aic issue was due to corrosion.
30537 | 1220648-2024-07184 | 18807190 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18807190&pc=OZD | 09/21/2022 | 03/11/2025 |  | 02/29/2024 |  |  | 09/21/2022 | 02/29/2024 |  |  | 06/30/2024 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235006 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/21/2022 |  |  |  |  | The user facility reported a 58-year-old male noted as high risk percutaneous cardiac insertion was implanted with an impella cp device for mechanical circulatory support. the physician made several attempts to place the device but was unable to position the device correctly. the physician decided the lesion would be better treated with a physician with more experience with chronic total occlusion and stopped the procedure. the impella device was weaned and explanted with no allegation of patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve¿ ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30538 | 1220648-2024-07185 | 18807245 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18807245&pc=OZD | 10/14/2022 | 02/29/2024 |  | 02/29/2024 |  |  | 10/14/2022 | 02/29/2024 |  |  | 05/31/2024 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023226018 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/14/2022 |  |  |  |  | The user facility reported a 70-year-old female was implanted with an impella 5.5 for mechanical circulatory support. it was reported that the impella device could not be placed during attempted delivery due to the patient's anatomy. the implant was subsequently aborted with no allegation of harm to patient. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issue was due to patient anatomy. this report is being filed as part of a retrospective review of historical records.
30539 | 1220648-2024-07186 | 18807397 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18807397&pc=OZD | 10/21/2022 | 02/29/2024 |  | 02/29/2024 |  |  | 10/21/2022 | 02/29/2024 |  |  | 05/31/2024 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023229033 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  |  |  | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/21/2022 |  |  |  |  | The user facility reported a 69-year-old female was implanted with an impella 5.5 as a bridge for left ventricular assist device work up. it was reported that the medical team attempting the impella implant had difficulty with delivery due to the patient's anatomy. the insertion of the impella 5.5 device was aborted, and was replaced with an impella cp. there was no patient harm reported. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to the patient condition. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident, stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
30540 | 1220648-2024-07187 | 18807512 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18807512&pc=OZD | 10/16/2022 | 03/10/2025 |  | 02/29/2024 |  |  | 10/16/2022 | 02/29/2024 |  |  | 06/30/2024 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023247157 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/16/2022 |  |  |  |  | The user facility reported a 76-year-old male in acute myocardial infarction cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during the post percutaneous coronary intervention (pci), the physician decided to exchange the impella cp through the left femoral artery (la), due to the impella alarmed "impella stopped motor current high" then the pump spontaneously restarted at previous p level with appropriate flow. the pump exchange was completed successfully. no patient involvement. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported pump stop issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
30541 | 1220648-2024-07188 | 18807516 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18807516&pc=OZD | 10/06/2022 | 03/11/2025 |  | 02/29/2024 |  |  | 10/06/2022 | 02/29/2024 |  |  | 05/31/2024 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023231346 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Remove | Failure to Advance | Patient-Device Incompatibility |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/06/2022 |  |  |  |  | The user facility reported a male patient was being implanted with an impella cp device for mechanical circulatory support. when the impella cp was being implanted, the physician was unable to advance the pump around the aortic arch despite multiple methods being utilized. it was decided to abort the imeplla insertion. however, when removing the pump, the wire became entangled and would not allow the impella to be withdrawn. the pump and the sheath were removed, and the physician was able to maintain wire access and replace the sheath. a coronary intervention was performed successfully without any reported patient harm. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | The investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records.
30542 | 1220648-2024-07189 | 18807699 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18807699&pc=OZD | 10/24/2022 | 05/19/2025 |  | 02/29/2024 |  |  | 10/24/2022 | 02/29/2024 |  |  | 06/30/2024 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023247036 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/24/2022 |  |  |  |  | The user facility reported an impella cp in a 85yo male patient for acute myocardial infarction and cardiogenic shock. it was reported there was bleeding around the repositioning sheath following the peel away removals. angle matching and resuturing with 2 forward tension sutures stopped bleeding from the access site. the pump was eventually explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30543 | 1220648-2024-07190 | 18807791 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18807791&pc=OZD | 09/27/2022 | 02/29/2024 |  | 02/29/2024 |  |  | 09/27/2022 | 02/29/2024 |  |  | 05/31/2024 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023227113 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/27/2022 |  |  |  |  | The user facility reported a 79-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the patient developed femoral bleeding at the access site. the bleeding was managed by adding nahco3 in purge. this issue caused a delay in systemic heparin administration. | [5 empty] | The investigation for the reported access site bleeding has been completed. no product was returned for evaluation. the cause of the access site bleeding is unknown as there is limited clinical details and no product was returned. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records.
30544 | 1220648-2024-07191 | 18807881 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18807881&pc=OZD | 09/27/2022 | 06/06/2024 |  | 02/29/2024 |  |  | 09/27/2022 | 02/29/2024 |  |  | 05/31/2024 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023222403 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/27/2022 |  |  |  |  | The user facility reported a 73-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. upon removal of the pump's sheath, the patient experienced bleeding and was treated with an 8f angioseal that was joined to the proglide; along with a pressure bandage used for closure. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding is determined to be use issue because the angioseal was added to perclose to the failure of proglide to control the bleeding and then the hemostasis is achieved by manual pressure. the instructions for use referenced in the initial report is for the impella cp with smart assist system in the following sections general patient care considerations, entering cardiac output, and potential adverse events (united states), which now includes the statement "cardiac or vascular injury (including ventricular perforation).¿ added- d6a/d6b. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
30545 | 1220648-2024-07192 | 18807918 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18807918&pc=OZD | 09/29/2022 | 02/29/2024 |  | 02/29/2024 |  |  | 09/29/2022 | 02/29/2024 |  |  | 03/31/2024 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204901 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/29/2022 |  |  |  |  | The user facility reported a 61-year-old male undergoing cardiac surgery was implanted with a impella cp device for mechanical circulatory support. the patient developed a hematoma medially on the groin site, where the nurse compressed the hematoma and held pressure above the site. partial thromboplastin time (ptt) later revealed that the patient was well above the target for anticoagulation. bicarb purged was used. the patient begun systemic heparin. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
30546 | 1220648-2024-07193 | 18807936 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18807936&pc=OZD | 10/17/2022 | 05/21/2025 |  | 02/29/2024 |  |  | 10/17/2022 | 02/29/2024 |  |  | 02/29/2024 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023201338 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper or Incorrect Procedure or Method | Activation Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/17/2022 |  |  |  |  | The user facility reported an 88-year-old male was implanted with an impella cp for mechanical circulatory support. it was reported that the case was aborted due to a damaged pump. the impella cp was started with easy guide lumen inside the pump. a new pump was used instead. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30547 | 1220648-2024-07194 | 18808231 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18808231&pc=OZD | 10/17/2022 | 02/29/2024 |  | 02/29/2024 |  |  | 10/17/2022 | 02/29/2024 |  |  | 05/31/2024 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023226470 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/17/2022 |  |  |  |  | The user facility reported that an 81-year-old male was implanted with an impella cp for mechanical circulatory support. it was reported that the patient developed an access site bleed and it was not possible to achieve sufficient control of bleeding. two proglides were used during the explant of the impella cp. however, protamine was used to successfully obtain hemostasis. | [5 empty] | The investigation for the reported access site bleeding has been completed. the impella device was not received from the customer, therefore, evaluation of the device was not possible. there was no sufficient clinical information available. therefore, the root cause of the access site bleed could not be determined. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ for access site bleeding, it ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records.
30548 | 1220648-2024-07195 | 18808310 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18808310&pc=OZD | 10/07/2022 | 09/18/2024 |  | 02/29/2024 |  |  | 10/07/2022 | 02/29/2024 |  |  | 02/29/2024 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023194694 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/07/2022 |  |  |  |  | The user facility reported a 63-year-old male patient was being implanted with an impella cp device for mechanical circulatory support. once the pump was implanted, the patient coded and cardio-pulmonary resuscitation (cpr) was performed. the patient crashed while on extra-corporeal membrane oxygenation (ecmo) and an angiogram was done. the right coronary artery was opened but the physician was unable to engage the left coronary artery. the patient was sent for a cardiac computed tomography angiography (cta) scan where an aortic dissection was noted. the patient was too sick to undergo surgery, therefore the family chose to withdraw care. | Additional information noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's (death) outcome. |  |  |  |  | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention: section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ b.2, b.5, h.1 and h.6 revised as medical safety review/death information was not reported on the initial manufacturer device report number 1220648-2024-07195. the exact date of death is unknown. d.6a and d.6b revised from the initial submission of manufacturer device report number 1220648-2024-07195 in accordance with updated procedures. f.6 and f.8 dates were reported on manufacturer device report number 1220648-2024-07195 and should not have been. g.1 reporting contact facility name and fax number were inadvertently left off the previously submitted manufacturer device report number 1220648-2024-07195 and were added. h.6 code 925 was added since the initial submission of manufacturer device report number 1220648-2024-07195 in accordance with updated procedures. h.10 additional manufacturer narrative instructions for use were revised from the initial submission of manufacturer device report number 1220648-2024-07195 in accordance with updated procedures.
30549 | 1220648-2024-07196 | 18808386 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18808386&pc=OZD | 10/04/2022 | 02/29/2024 |  | 02/29/2024 |  |  | 10/04/2022 | 02/29/2024 |  |  | 07/31/2024 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023246644 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 41 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/04/2022 |  |  |  |  | The user facility reported a 41-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was on impella cp and veno-venous extracorporeal membrane oxygenation (v-vecmo) support when the pump stopped due to high motor current. all attempts to restart the pump, including restarting the console, were unsuccessful and the impella cp device was explanted and replaced. | [5 empty] | The investigation into the pump stop has been completed. the device, data logs, nor sufficient clinical information was returned for evaluation. therefore, the root cause of the pump stop could not be determined. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
30550 | 1220648-2024-07197 | 18808454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18808454&pc=OZD | 04/21/2022 | 03/10/2025 |  | 02/29/2024 |  |  | 04/21/2022 | 02/29/2024 |  |  | 01/31/2024 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023187085 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 04/21/2022 |  |  |  |  | The user facility reported a 38- year-old male was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient was bleeding from the groin site which resolved. heparin was turned off for five hours and was restarted. three units of pack red blood cells (prbcs) were given. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleeding issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
30551 | 1220648-2024-07198 | 18808480 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18808480&pc=OZD | 10/05/2022 | 03/11/2025 |  | 02/29/2024 |  |  | 10/05/2022 | 02/29/2024 |  |  | 03/31/2024 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED |  |  | US |  | 2023214128 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/05/2022 |  |  |  |  | The user facility reported a 48-year-old male in st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient had oozing at the insertion site and low platelets. the patient's heparin was discontinued and the impella cp device was replaced with an impella 5.5 device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: "potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30552 | 1220648-2024-07199 | 18808595 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18808595&pc=OZD | 10/18/2022 | 02/29/2024 |  | 02/29/2024 |  |  | 10/18/2022 | 02/29/2024 |  |  | 04/30/2024 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023226086 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/18/2022 |  |  |  |  | The user facility reported a 77-year-old male was implanted with a impella cp. it was reported that during the procedure the impella sheath sidearm stopcock was left open after an activated clotting time (act) was obtained. this was unnoticed and the patient had a large amount of blood loss followed by hemodynamic instability, intubation, and epinephrine, levophed drip started. 3 units of red packed blood cells were given. systemic heparin had been started but was stopped due to drop in platelets and oozing at central and swan site. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records.
30553 | 1220648-2024-07200 | 18808979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18808979&pc=OZD | 10/06/2022 | 02/29/2024 |  | 02/29/2024 |  |  | 10/06/2022 | 02/29/2024 |  |  | 06/30/2024 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED |  |  | US |  | 2023234229 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/06/2022 |  |  |  |  | The user facility reported a 61-year-old female was being placed on an impella 5.5 for mechanical circulatory support. the impella 5.5 was unable to advance through the patient¿s vasculature due to the size of their axillary artery. the implant was aborted and an impella cp was placed. | [5 empty] | The investigation into the reported delivery issue has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue - anatomy was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
30554 | 1220648-2024-07201 | 18809072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18809072&pc=OZD | 10/24/2022 | 02/29/2024 |  | 02/29/2024 |  |  | 10/24/2022 | 02/29/2024 |  |  | 12/31/2022 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022021343 |  | 0048-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | Device Dislodged or Dislocated | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/24/2022 |  |  |  |  | The user facility reported that the impella cp procedure on a female patient had to be aborted due to the female patient's cardiac anatomy. additionally, the pump was noted to have dislocated during sheath change from peel-away to repositioning sheath. a new impella cp device was placed. there was no reported patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to the patient condition. this report is being filed as part of a retrospective review of historical records.
30555 | 1220648-2024-07202 | 18809228 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18809228&pc=OZD | 10/07/2022 | 02/29/2024 |  | 02/29/2024 |  |  | 10/07/2022 | 02/29/2024 |  |  | 03/31/2024 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191461 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/07/2022 |  |  |  |  | The user facility reported a male of unknown age in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced leg ischemia. the impella cp was explanted and the patient continued on extracorporeal membrane oxygenation (ecmo) support. | [5 empty] | The investigation has been completed. the root cause of the limb ischemia was unable to be determined as limited clinical details were provided and no product was returned for review. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
30556 | 1220648-2024-07203 | 18809589 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18809589&pc=OZD | 10/07/2022 | 09/19/2024 |  | 02/29/2024 |  |  | 10/07/2022 | 02/29/2024 |  |  | 02/29/2024 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195571 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Improper Flow or Infusion | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Tachycardia | Ventricular Fibrillation | [9 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/08/2022 |  |  |  |  | The procedural outcome noted that the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | The user facility reported a 68-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported the patient experienced oozing at the insertion sites and loss of pulsatility. the patient was treated with drug-eluting stents and with a sodium bicarbonate purge. the patient also experienced multiple suction events and episodes of ventricular tachycardia/ventricular fibrillation. the patient was shocked twelve times and provided treatment of amiodarone and lidocane. the patient's family withdrew care. |  |  |  |  | B2: date of patient death is unknown. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07203 was provided in sections b2, b5, h1, and h6. note: corrected information provided in h6 is an addition to previous coding. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records.
30557 | 1220648-2024-07204 | 18809599 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18809599&pc=OZD | 10/25/2022 | 04/15/2025 |  | 02/29/2024 |  |  | 10/25/2022 | 02/29/2024 |  |  | 06/30/2024 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023248301 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/25/2022 |  |  |  |  | The user facility reported a 90-year-old male needing mechanical circulatory support. it was reported that after being placed on impella cp support, a large hematoma appeared and the impella was explanted early. hematoma was resolved with manual pressure. | [5 empty] | Section d6a / d6b: added implant and explant dates. the investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
30558 | 1220648-2024-07205 | 18809762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18809762&pc=OZD | 02/07/2024 | 04/08/2024 |  | 02/29/2024 |  |  | 02/07/2024 | 02/29/2024 |  |  | 10/31/2025 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025443932 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported 64-year-old male on impella 5.5 support removed the pump himself. patient was deliriant. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ | The investigation into the positioning issues has been completed since the initial report was submitted. the device was not returned for investigation. based on the clinical data provided, the root cause of the positioning issue was use related due to improper technique as the patient pulled it out.
30559 | 1220648-2024-07206 | 18809820 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18809820&pc=OZD | 10/09/2022 | 02/29/2024 |  | 02/29/2024 |  |  | 10/09/2022 | 02/29/2024 |  |  | 06/30/2024 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023230030 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/09/2022 |  |  |  |  | The user facility reported a male patient of unknown age in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced bleeding at the impella sheath and the impella cp device was explanted. | [5 empty] | The investigation of the reported issue has been completed. the root cause of the access site bleeding issue was not determined since product was not returned. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿assess access site for bleeding and hematoma.¿ this report is being filed as part of a retrospective review of historical records.
30560 | 1220648-2024-07207 | 18810017 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18810017&pc=OZD | 02/08/2024 | 03/19/2024 |  | 02/29/2024 |  |  | 02/08/2024 | 02/29/2024 |  |  | 10/31/2025 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024421989 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Hispanic | White | Malfunction | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Embolism/Embolus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that anticoagulant was held due to bleeding from the impella access site. the clinician stated it appears repo hub is out slightly and not sutured. the physician will be notified about securing and optimizing access site. it was further reported that there were a lot of micro bubbles and turbulent flow noted on tee initially but decreased after some time. the patient was noted to be stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation for the reported embolism and access site bleeding has been completed. the impella device was not returned for evaluation. clinical details provided were sufficient to determine the root cause of the access site bleed. at that time the patient was not on anticoagulation and their act was 368. the purge solution was then switched to d5w and the pump was repositioned. the most likely cause of the major access site bleeding was anticoagulation management. however, clinical details provided were not enough to determine the root cause of embolism. the cause of the embolism was not determined because product was not returned and not enough information was given.
30561 | 1220648-2024-07208 | 18810570 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18810570&pc=OZD | 02/01/2024 | 04/23/2024 |  | 02/29/2024 |  |  | 02/01/2024 | 02/29/2024 |  |  | 10/31/2025 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP - JP |  |  | ABIOMED, INC. |  |  | US |  | 2024426983 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/01/2024 |  |  |  |  | The user facility reported a 76-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was also being supported by ecmo. during support, the patient developed an infection. the patient expired while on impella and ecmo support. the physician noted the patient expired from multi-organ failure, not impella use. the pump was a possible site of the infection, as was ecmo. | [5 empty] | The investigation for the sepsis has been completed. the impella device was not returned for evaluation. therefore, without sufficient clinical details, the root cause of the reported issues sepsis could not be determined. the instructions for use states ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿
30562 | 1220648-2024-07209 | 18810626 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18810626&pc=OZD | 10/20/2022 | 09/24/2024 |  | 02/29/2024 |  |  | 10/20/2022 | 02/29/2024 |  |  | 07/31/2024 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023254692 |  | 0550-0008 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 24 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Difficult to Insert | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/20/2022 |  |  |  |  | The user facility reported a 24-year-old male had a impella cp implanted, and the medical staff was having difficulty delivering the impella pump due to the patient's tortuous vasculature. after several failed attempts they were finally able to deliver impella successfully. there was no patient harm reported. | During support it was reported that the patient was febrile, hypotensive, tachycardia (tachy) and had increased chest tube bleeding, was given packed red blood cells. in addition, it was noted that heparin was removed from the purge due to subarachnoid hemorrhage. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. |  |  |  |  | The investigation for the reported delivery issues has been completed. the impella device was not received from the customer, therefore, evaluation of the device was not possible. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue - anatomy was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07209 was provided in sections b1, b2, b5, h1, and h6.
30563 | 1220648-2024-07210 | 18810721 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18810721&pc=PYX | 02/15/2024 | 07/12/2024 |  | 02/29/2024 |  |  | 02/16/2024 | 02/29/2024 |  |  | 01/31/2025 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024321239 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 17 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | Renal Impairment | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup |  | 02/16/2024 |  |  |  |  | The user facility reported a 17-year-old female with cardiomyopathy was implanted with an impella rp for mechanical circulatory support. while on support, the patient was experiencing hemolysis. the impella flows were reduced, volume given, pump reposition, and one unit of red blood cells was given. there was not much improvement. the pump was removed, and a large clot noted in the outlet cage of the impella rp. the patient is stable. | Additional information received on 05/20/2024 via abiomed¿s long-term outcome and quality indicator impella registry regarding the patient that endured acute renal insufficiency with a " possible " relationship to the impella device used. change in urine color (darkened, appeared blood tinged) two days post-op. suspect secondary to impella. gave additional fluid, repositioned device, and decreased p level. urine starting to clear later that afternoon, but ongoing aki and oliguria¿. the patient recovered with no sequelae. |  |  |  |  | The investigation for the reported thrombus, renal failure, and hemolysis has been completed. data logs were reviewed which showed about 7.5 hours into support, motor current spike greater than 1400 ma observed with no change in p-level and followed by drop in flows. no other issue was observed. the impella was returned for evaluation and upon visual inspection, found biomaterial in the outflow cage of rp pump. no other issues were found on the rest of the pump. pump was conducted for hemolysis test to observe if any defect was found on the pump. pump passed hemolysis test as test value was in specification. the root cause of hemolysis and thrombus was due to biomaterial ingestion likely from patient condition per return product investigation and log reviewthe root cause of renal failure was unable to be determined as limited clinical details were provided and no issue was found on the pump. the instructions for use referenced in the initial report is for the impella rp smart assist system with automated impella controller in the following sections: section: potential adverse events (united states), which includes the statement "cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. section: hemolysis. section: suction added- d6a/d6b. f6/f8 should have been left blank in the initial report. | The investigation into the hemolysis and thrombus issues has been completed since the original report was filed. the pump was returned, tested, and evaluated. visual inspection found biomaterial in the outflow cage of rp pump. the root cause of hemolysis was due to biomaterial ingestion per return product investigation and log review. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | Section b5, g2 and h6 were updated as per additional information received.
30564 | 1220648-2024-07211 | 18810770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18810770&pc=OZD | 02/20/2024 | 09/05/2024 |  | 02/29/2024 |  |  | 02/20/2024 | 02/29/2024 |  |  |  |  | 02/29/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1073779 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 13 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Computer Operating System Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/20/2024 |  |  |  |  | The user facility reported a 13-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella support when there was an automated impella controller (aic) issue. the team was alerted about ongoing suction alarms and impella position unknown alarm. staff noticed lv pressure 25/-44 and motor current flat. the aic was replaced and all alarms were resolved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The automated impella controller (aic) investigation team found that this is a pump issue, not an aic issue. a regulatory report is no longer necessary. please see manufacturer device report 1220648-2024-08693 regarding information on the pump.
30565 | 1220648-2024-07212 | 18810810 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18810810&pc=OZD | 10/31/2022 | 02/29/2024 |  | 02/29/2024 |  |  | 10/31/2022 | 02/29/2024 |  |  | 07/31/2024 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023253209 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/31/2022 |  |  |  |  | The user facility reported an impella cp in a patient for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced distal limb ischemia. the impella cp device was removed to resolve the distal limb ischemia. | [5 empty] | As the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
30566 | 1220648-2024-07213 | 18810815 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18810815&pc=OZD | 02/21/2024 | 04/08/2024 |  | 02/29/2024 |  |  | 02/21/2024 | 02/29/2024 |  |  | 10/31/2025 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024417659 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Insufficient Flow or Under Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/21/2024 |  |  |  |  | The user facility reported a 60-year-old male in acute myocardial infarction/cardiogenic shock and coronary thrombosis was implanted with an impella cp device for mechanical circulatory support. it was reported the pump demonstrated low flow after 10 minutes of support. a clot was suspected, however the impella cp was removed and assessed, and no clot was visible. no patient harm was reported | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the low pump flow issue has been completed since the original report was filed. the pump was returned, tested, and evaluated. product evaluation confirmed that there were no abnormalities with the device and the failure mode could not be reproduced. the most likely cause of the low pump flow was biomaterial.
30567 | 1220648-2024-07214 | 18810852 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18810852&pc=OZD | 10/25/2022 | 05/20/2025 |  | 02/29/2024 |  |  | 10/25/2022 | 02/29/2024 |  |  | 06/30/2024 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023235264 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/25/2022 |  |  |  |  | The user facility reported that after a 75-year-old male was placed on the impella cp for mechanical circulatory support, the peel away sheath kinked mid way through the right iliac. it was recommended to insert the dilator past the kink to help with support and the sheath came out with no issue. patient tolerated procedure well and had great results. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30568 | 1220648-2024-07215 | 18810859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18810859&pc=OZD | 02/21/2024 | 04/10/2024 |  | 02/29/2024 |  |  | 02/21/2024 | 02/29/2024 |  |  | 11/30/2025 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024433075 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/21/2024 |  |  |  |  | The user facility reported a 70-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella cp was explanted due to the patient experiencing loss of pulse in the left leg. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia ¿ reversible and the access site bleeding - major issues has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
30569 | 1220648-2024-07216 | 18810901 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18810901&pc=OZD | 10/20/2022 | 02/29/2024 |  | 02/29/2024 |  |  | 10/20/2022 | 02/29/2024 |  |  | 07/29/2024 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023256566 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/20/2022 |  |  |  |  | The user facility reported a 77-year-old female in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. after being placed on the impella cp, there were no pulses in the extremity distal to impella insertion site. patient vessels too small on contralateral limb for fem-fem bypass. the physician performed a distal perfusion run off and flow was delayed, but there was blood flow. the limb was continued to be monitored. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records.
30570 | 1220648-2024-07217 | 18810922 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18810922&pc=OZD | 02/22/2024 | 04/10/2024 |  | 02/29/2024 |  |  | 02/22/2024 | 02/29/2024 |  |  | 10/31/2025 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024421398 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Remove | Activation, Positioning or Separation Problem | Difficult to Advance |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/22/2024 |  |  |  |  | The user facility reported a 89-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that complications were encountered during attempted delivery of an impella cp pump. upon initial insertion, the sheath kinked. multiple attempts were required before the pump was able to pass with wire tension. the wire dislodged from the lv (left ventricle) prior to advancing the pump and subsequent attempts to place the pump without the wire were unsuccessful. both the introducer and pump were explanted and deemed to be too malleable/deformed for additional attempts. a new pump and sheath were subsequently used for successful implant. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues and the difficulty inserting/removing introducer issue has been completed since the original report was submitted. per the clinical information provided, the root cause of the introducer damage was most likely due to patient anatomy causing the sheath to kink. the root cause of the delivery issues was wireless insertion into the left ventricle due to guidewire becoming dislodged from the left ventricle and pump was attempted to be inserted without guidewire.
30571 | 1220648-2024-07218 | 18810958 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18810958&pc=OZD | 02/22/2024 | 04/08/2024 |  | 02/29/2024 |  |  | 02/22/2024 | 02/29/2024 |  |  | 11/30/2025 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABOIMED, INC. |  |  | US |  | 2024427569 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Insufficient Flow or Under Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/22/2024 |  |  |  |  | The user facility reported a 37-year-old female with cardiomyopathy was to be implanted with an impella 5.5 device for mechanical circulatory support. it was reported the pump was unable to get above p7. the impella was replaced with a new one. no patient harm was reported | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into low pump flow issue has been completed since the initial report was submitted. the device was not returned for investigation. per the clinical information and data logs, many suction alarms occurred. motor current spike observed around the time of the low pump flow issue observed. the cause of the low pump flow issue most likely was biomaterial ingestion.
30572 | 1220648-2024-07219 | 18810984 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18810984&pc=OZD | 02/23/2024 | 04/10/2024 |  | 02/29/2024 |  |  | 02/23/2024 | 02/29/2024 |  |  | 12/26/2025 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024434473 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/23/2024 |  |  |  |  | The user facility reported a 63-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp placed for a pci (percutaneous coronary intervention) patient. the patient's limb became ischemic with thrombus observed. the team gained contralateral access and performed aspiration and balloon angioplasty. the pump had remained on for 24 hours and was weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | The investigation into the limb ischemia - reversible issue has been completed since the original report was submitted. the device was not returned for investigation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
30573 | 1220648-2024-07220 | 18811120 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18811120&pc=OZD | 02/23/2024 | 04/11/2024 |  | 02/29/2024 |  |  | 02/23/2024 | 02/29/2024 |  |  | 11/30/2025 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024424178 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Pressure Problem | [5 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/23/2024 |  |  |  |  | The user facility reported a 57-year-old male with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported the pump had a sudden drop in purge flow and increase in purge pressure. there were no kinks in the tubing. the patient was taken for emergent extracorporeal membrane oxygenation (ecmo) with removal of 5.5 and intra-aortic balloon pump (iabp) placement. there was clot noted on inflow cage upon removal. the patient remained stable throughout. | [5 empty] | The investigation for the reported high purge pressure and thrombus has been completed. pump was returned for evaluation. dextrose buildup was present in and around purge gap. pump was primed with purge cassette that was used in the field and high purge pressure alarms were reproduced. pump teardown was performed and biomaterial was observed in purge gap under impeller blade. data logs were reviewed and a gradual increase in purge pressure over an hour and 20 minutes was observed right before high purge pressure alarms were triggered. clinical details noted that patient was possibly on sub-therapeutic levels of anti-coagulation. the root cause of the high purge pressure was due to biomaterial in the purge gap. the instructions for use states ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ added-d9 device return date. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30574 | 1220648-2024-07221 | 18811351 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18811351&pc=OZD | 10/09/2022 | 05/19/2025 |  | 02/29/2024 |  |  | 10/09/2022 | 02/29/2024 |  |  | 06/30/2024 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023247912 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Insert | Material Twisted/Bent | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/09/2022 |  |  |  |  | The user facility reported a 79 year-old female undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. during implantation, the impella cp got stuck in the descending aorta and arch. the impella cp was removed from the body and the impella cp wire was found to be kinked. the wire was replaced and the impella became stuck again in the arch. a stiff buddy wire was used and the impella cp was able to cross the arch and into the ventricle. additionally, the patient experienced a hematoma and oozing from the sheath. blood products were administered. levophed was increased initially but the bleeding resolved so it was weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30575 | 1220648-2024-07222 | 18811426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18811426&pc=OZD | 02/24/2024 | 04/10/2024 |  | 02/29/2024 |  |  | 02/24/2024 | 02/29/2024 |  |  | 11/30/2025 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC |  |  | US |  | 2024436231 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | User facility | Company representation |  |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/24/2024 |  |  |  |  | The user facility reported a 79-year-old male patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was that there were ongoing high purge pressure alarms. the cassette was swapped, and clinician had been de-airing. the physician did not want to start tissue plasminogen activator protocol, due to unrelated bleeding issues of the patient and the physician decided to explant the pump. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the high purge pressure issue has been completed since the original report was filed. the pump was returned, tested, and evaluated. biomaterial was found in the outflow around the impeller. per the product investigation and data log review, the root cause of the high purge pressure issue was biomaterial found in the outflow cage.
30576 | 1220648-2024-07224 | 18811876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18811876&pc=OZD | 02/25/2024 | 04/15/2024 |  | 02/29/2024 |  |  | 02/25/2024 | 02/29/2024 |  |  | 11/30/2025 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024423530 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Break | Activation, Positioning or Separation Problem | Difficult to Advance |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/25/2024 |  |  |  |  | The user facility reported a 60-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that impella 5.5 was unable to pass through the patient¿s calcified tortuous aorta via axillary. the impella catheter was also found to be kinked and appeared fractured upon removal. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported delivery issues has been completed. the product was returned for evaluation and a catheter kink and stretching was found by the motor. the root cause of the delivery issue is determined to be patient condition because of calcified tortuous aorta of the patient.
30577 | 1220648-2024-07223 | 18811881 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18811881&pc=OZD | 02/25/2024 | 04/11/2024 |  | 02/29/2024 |  |  | 02/25/2024 | 02/29/2024 |  |  | 11/30/2025 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024434832 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/25/2024 |  |  |  |  | The user facility reported a 74-year-old female patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was that were purge pressure high/purge system blocked alarms. alarms currently are intermittent and alternates between the two. purge cassette changed without resolving alarms. the physician did not want to start tissue plasminogen activator protocol, due to unrelated bleeding issues of the patient. no patient harm was reported. | [5 empty] | The investigation into the high purge pressure issue has been completed since the original report was submitted. the device was not returned for investigation. data log analysis confirmed a gradual rise in purge pressure and fall in purge flow over 12 hours before the alarms were triggered. the most likely cause of the high purge pressure was biomaterial. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30578 | 1220648-2024-07225 | 18812092 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18812092&pc=OZD | 02/09/2024 | 03/29/2024 |  | 02/29/2024 |  |  |  | 02/29/2024 |  |  | 11/30/2025 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024426649 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/09/2024 |  |  |  |  | The user facility reported a 67-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient had an attempted impella 5.5 implant due to decompensating. during the implant, the 5.5 implant was unable to be passed due to anatomy. decision was made to remove 5.5 and place cp for venting. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the positioning issues has been completed since the initial report was submitted. the product was not returned for investigation. as per the clinical details, the root cause of the delivery issue was most likely due to patient condition since the pump could not be implanted due to patient's anatomy.
30579 | 1220648-2024-07226 | 18812336 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18812336&pc=OZD | 10/27/2022 | 05/22/2025 |  | 02/29/2024 |  |  |  | 02/29/2024 |  |  | 07/31/2024 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023257987 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/27/2022 |  |  |  |  | The user facility reported an impella cp in a 65-year-old female patient for mechanical circulatory support. it was reported that the impella cp was turned off and explanted. there was no clot noted. the right femoral artery was surgically repaired and closed. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿cautions. physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30580 | 1220648-2024-07227 | 18812378 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18812378&pc=OZD | 10/31/2022 | 03/11/2025 |  | 02/29/2024 |  |  |  | 02/29/2024 |  |  | 04/30/2024 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023225803 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/31/2022 |  |  |  |  | The user facility reported a 63-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient coded and advanced cardiac life support protocol was initiated. it was reported after 4 minutes of cardiopulmonary resuscitation, the impella cp device had intermittent low placement signal (ps) alarms and the aortic (ao) diastolic placement signal did not correlate with arterial line. the impella cp was repositioned under echocardiography. it was reported after repositioning the impella cp device, intermittent placement signal (ps) alarms continued with no report of aortic regurgitation or aortic insufficiency. patient was monitored until impella cp device was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30581 | 1220648-2024-07228 | 18812447 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18812447&pc=OZD | 02/16/2024 | 04/10/2024 |  | 02/29/2024 |  |  |  | 02/29/2024 |  |  |  |  | 02/29/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1107437 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Inappropriate or Unexpected Reset | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/16/2024 |  |  |  |  | The user facility reported a 39-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the automated impella controller (aic) screen went completely black during patient support. the aic rebooted but the issue recurred and the decision was made to replace the console. there was no patient harm associated with the malfunction. support was maintained with the implanted pump. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into the aic - shutdown or restart issue has been completed since the initial report was submitted. the console was returned, visually inspected, and tested. the reported failure mode could not be reproduced. based on the data analysis, the root cause of the black screen was an unexpected reboot due to the calculator crashing, but the failure was not reproduced, so the cause of the calculator crash is undetermined.
30582 | 1220648-2024-07229 | 18812479 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18812479&pc=OZD | 02/18/2024 | 09/20/2024 |  | 02/29/2024 |  |  |  | 02/29/2024 |  |  | 11/30/2025 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025439291 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/18/2024 |  |  |  |  | The user facility reported a 83-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during impella cp placement, impella was not able to fully cross the aortic valve due to aortic stenosis. the case was aborted. | Survival outcome at explant noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into delivery issues has been completed since the initial report was submitted. the device was not returned for investigation. per the clinical information, the root cause of delivery issue is determined to be patient condition due to valve disease of the patient because of which impella unable to fully cross the valve. | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07229 was provided. note: corrected information provided in h6 is an addition to previous coding.
30583 | 1220648-2024-07230 | 18812517 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18812517&pc=OZD | 02/21/2024 | 04/01/2024 |  | 02/29/2024 |  |  |  | 02/29/2024 |  |  |  |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025437461 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/21/2024 |  |  |  |  | The user facility reported a 62-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a cp pump inserted via the femoral to support a patient admitted with angina and hypotension. the patient had the cp explanted after 10hours due to an access site hematoma. the hematoma was treated by blood products and removal and replacement of the pump via an upsized sheath, a 18fr. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
30584 | 1220648-2024-07231 | 18812754 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18812754&pc=OZD | 10/27/2022 | 03/11/2025 |  | 02/29/2024 |  |  |  | 02/29/2024 |  |  | 03/31/2024 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023205792 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  |  |  | Injury | Yes | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/27/2022 |  |  |  |  | The user facility reported a 36-year-old male was on impella cp support and the device was intermittently alarming. the impella cp was mispositioned and was unable to be repositioned. the decision was made to remove the impella. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30585 | 1220648-2024-07232 | 18812917 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18812917&pc=OZD | 10/27/2022 | 02/29/2024 |  | 02/29/2024 |  |  |  | 02/29/2024 |  |  | 03/31/2024 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023204929 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/27/2022 |  |  |  |  | The user facility reported a 74-year-old male in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the bedside echocardiogram indicated multiple left ventricle clots. the patient coded and one round of cpr was performed along with amino given. the clot remained in place. the patient was monitored, and the clots were no longer in the left ventricle. no additional adverse patient effects were reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ this report is being filed as part of a retrospective review of historical records.
30586 | 1220648-2024-07233 | 18812955 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18812955&pc=OZD | 10/28/2022 | 02/29/2024 |  | 02/29/2024 |  |  |  | 02/29/2024 |  |  | 07/31/2023 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022119556 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/28/2022 |  |  |  |  | The user facility reported an impella cp was implanted in a (b)(6) male patient for mechanical circulatory support. it was reported that while the patient was on the impella cp support, the patient developed a large hematoma at the right groin access site, but it was resolved with pressure. heparin was withdrawn and the patient received 1 unit of packed red blood cells (prbc). | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury" this report is being filed as part of a retrospective review of historical records.
30587 | 1220648-2024-07234 | 18812981 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18812981&pc=OZD | 02/21/2024 | 04/09/2024 |  | 02/29/2024 |  |  |  | 02/29/2024 |  |  | 11/30/2025 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024431454 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/21/2024 |  |  |  |  | The user facility reported a 77-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us complainant had a 5.5-rp support that was ongoing when on day 5 of the support there was refractory vt (ventricular tachycardia). the pump remained on for support. the team treated by medical therapy with amiodarone and the arm was moved. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined. the instructions for use states ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ added-b2 selected required intervention to prevent permanent impairment. d4-updated model number. h5-labeled for single use changed to yes.
30588 | 1220648-2024-07235 | 18812988 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18812988&pc=OZD | 02/18/2024 | 04/18/2024 |  | 02/29/2024 |  |  |  | 02/29/2024 |  |  | 09/30/2025 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2024407846 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/20/2024 |  |  |  |  | The user facility reported a 60-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the japanese complainant had a cp support that was ongoing for patient crashing with cardiomyopathy. the team observed signs of hemolysis. the ldh (lactate dehydrogenase) was rising, and poor position was noted. the "hemolytic tendency" and need for care escalation prompted the cp explant and the placement of the 5.5 pump. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ | The investigation into the hemolysis issue has been completed since the original report was submitted. the device was not returned for investigation. based on clinical description, the root cause of hemolysis was most likely due to improper positioning since pump was closer to posterior wall.
30589 | 1220648-2024-07236 | 18812993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18812993&pc=OZD | 02/21/2024 | 04/08/2024 |  | 02/29/2024 |  |  |  | 02/29/2024 |  |  | 11/30/2025 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024434180 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 22 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/21/2024 |  |  |  |  | The user facility reported a 22-year-old male with cardiomyopathy was to be implanted with an impella 5.5 device for mechanical circulatory support. it was reported that during implant the surgeon was unable to advance the device into the left ventricle due to the patient¿s anatomy. no patient harm was reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into delivery issues has been completed since the initial report was submitted. the device was not returned for investigation. per the clinical information, the root cause of the delivery issue was most likely patient condition related, as the impella was unable to advance into the left ventricle during the implant due to the patient's vessel anatomy.
30590 | 1220648-2024-07237 | 18813002 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18813002&pc=OZD | 02/22/2024 | 04/10/2024 |  | 02/29/2024 |  |  |  | 02/29/2024 |  |  | 07/31/2025 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024387978 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Ventricular Fibrillation | [9 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/22/2024 |  |  |  |  | The user facility reported a 63-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient had an attempted impella cp implant due to rapidly decompensating. during insertion the patient had access site bleeding. a covered stent and a transfusion were performed. additionally, a pleurex catheter required for right pneumothorax while attempting right axillary access. furthermore, the patient began to fully hemodynamically collapse with bradycardia and subsequent pea (pulseless electrical activity) arrest. shocks were performed for subsequent vt/vf (ventricular tachycardia/ ventricular fibrillation) while thoracic vent placed emergently. the impella implant was aborted. | [5 empty] | The investigation into the access site bleeding ¿ major, the delivery issues, the vf/vt/af/at, and the vascular damage - peripheral vessel issues has been completed since the original report was submitted. the device was not returned for investigation. per the clinical information provided, the root cause of the access site bleeding issue was use related as the sheath was inadvertently pulled back and access lost. the root cause of the delivery issue was patient condition related to small/ diseased vessels. as the necessary information was not provided, the root cause of the vt/vf was not determined. the root cause of the vascular damage - peripheral vessel issue was most likely patient condition related to occluded vessels. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿
30591 | 1220648-2024-07238 | 18813020 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18813020&pc=OZD | 02/23/2024 | 04/11/2024 |  | 02/29/2024 |  |  |  | 02/29/2024 |  |  |  |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025438788 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Activation, Positioning or Separation Problem | Overcorrection | [5 empty] | Ischemia | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/23/2024 |  |  |  |  | The user facility reported a 66-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that impella cp insertion was successful but manipulating the patient in the icu led to positioning alarms. the aic (automated impella controller) showed that the pump was in the ventricle. attempting to reposition was unsuccessful and ultimately the impella ended up under a cardiac structure and would not advance further. decision was made to go to cath lab to attempt repositioning with fluoroscopy and echo. many attempts were made to reposition the impella in the mid left ventricle (lv) cavity without success. the impella was exchanged, and a new one was inserted successfully. the physician noted the tortuous turn into the lv and that being the source of positioning difficulties. there was no patient harm. | Additional information states that pulses were absent in the right leg, monophasic in left pedal, with the right foot also cold to touch. vascular was consulted intensive team made aware. plan were made to take the patient cath lab for reperfusion sheaths bilaterally. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ | The investigation for the reported positioning issues has been completed. the impella device was not returned for investigation. clinical details stated that the repositioning guide was used and pulled back 3 cm until aortic waveform visualized. and additional 2 cm were pulled back for optimal positioning and the pump came back abruptly about 4 cm as visualized on echo. the cause of the positioning issue was use related due to patient movement per clinical review. added-b1-selected serious injury; b2-selected other serious events; b5 event narrative; h6 clinical code 1942 d4-updated model number.
30592 | 1220648-2024-07239 | 18813030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18813030&pc=OZD | 02/23/2024 | 04/11/2024 |  | 02/29/2024 |  |  |  | 02/29/2024 |  |  | 11/30/2025 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024432718 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Twisted/Bent | No Flow | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/23/2024 |  |  |  |  | The user facility reported a 32-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella support and when the impella turned on, there was no left ventricle waveform. motor current was noted to be non-pulsatile and there was constant suction. there was a kink noted near the outlet conduit. repositioning was attempted to no avail; the pump was removed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the product damage (cannula kink) issue has been completed since the original report was submitted. the device was not returned for investigation. the low flow was confirmed when pump started and remained to the rest of the case which triggered auto suction response and suction alarms. based on clinical description and data review, the root cause of low flow was most likely kinked cannula.
30593 | 1220648-2024-07240 | 18813040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18813040&pc=OZD | 02/24/2024 | 04/11/2024 |  | 02/29/2024 |  |  |  | 02/29/2024 |  |  | 10/31/2025 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024415378 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | Pressure Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/24/2024 |  |  |  |  | The user facility reported a 66-year-old male patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that after being on support for 23 days, the pump displayed high purge pressure alarms for multiple days despite administration of tissue plasminogen activator. the patient was weaned down to p4, could possibly be pump saturation. the pump was explanted and there was no patient harm reported. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the high purge pressure and the saturated flow issues has been completed since the original report was filed. the pump was returned, tested, and evaluated. once the pump was tested, a significant clot was expelled from the purge gap and the purge pressure decreased. saturated flow was not reproduced. the root cause of the high purge pressure was determined to be biomaterial.
30594 | 1220648-2024-07241 | 18813056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18813056&pc=OZD | 02/24/2024 | 04/12/2024 |  | 02/29/2024 |  |  |  | 02/29/2024 |  |  | 11/30/2025 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024427047 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Improper Flow or Infusion | Material Twisted/Bent | Pumping Problem |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/24/2024 |  |  |  |  | The user facility reported a 69-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that an implanted impella 5.5 pump began to experience high purge pressure and low purge flows during patient support. the purge cassette was replaced to troubleshoot the issue, but the pressure remained high. a tpa protocol was subsequently initiated, and plans were made to explant the pump within the next couple days. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation for the reported high purge pressure has been completed. the pump was returned for evaluation. during evaluation, a kink was observed on catheter right next to the red handle. catheter was opened and kink in purge line was observed. biomaterial was also found in purge gap, however, this was most likely due to the pump being used for a number of days after high purge pressure alarms were triggered. data logs were reviewed and it was found that a sudden increase in purge pressure occurred without any spikes in motor current. the root cause of the high purge pressure was due to a kinked purge line just distal to the red handle on the catheter side.
30595 | 1220648-2024-07242 | 18813062 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18813062&pc=OZD | 02/25/2024 | 03/08/2024 |  | 02/29/2024 |  |  |  | 02/29/2024 |  |  |  |  | 02/29/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1143008 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  |  | White | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/25/2024 |  |  |  |  | The user facility reported a 54-year-old male patient of unknown ethnicity with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that during the purge cassette change, it was noted that the purge disc insert was cracked. the patient was switched to back up automated impella controller (aic) without issue. there was no patient harm reported. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation. the logs were not relevant to the reported issue. the console was evaluated and it was confirmed that the purge disc retainer had broken off on the console. the root cause of this issue was a broken purge retainer.
30596 | 1220648-2024-07243 | 18813072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18813072&pc=OZD | 02/26/2024 | 04/15/2024 |  | 02/29/2024 |  |  |  | 02/29/2024 |  |  | 12/03/2025 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025437527 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Pressure Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/26/2024 |  |  |  |  | The user facility reported a 75-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support due to needing coronary intervention. while on support, there was high purge pressure. troubleshooting was performed unsuccessfully. the impella was replaced. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ | The investigation for the reported high purge pressure has been completed. neither the impella device nor the data logs were returned for evaluation. additionally, clinical details provided were limited to determine the root cause. therefore, the root cause of the high purge pressure could not be determined. d4-updated model number h5-correct to yes h6 medical device problem code-code 2906 selected in error. h6-investigation findings =3 211 selected in error.
30597 | 1220648-2024-07244 | 18813091 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18813091&pc=OZD | 10/28/2022 | 02/29/2024 |  | 02/29/2024 |  |  |  | 02/29/2024 |  |  | 07/31/2023 |  | 02/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022113344 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Deliver | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 10/28/2022 |  |  |  |  | The user facility reported an impella cp on a 72 year old male requiring mechanical circulatory support. it was reported that during insertion, the device was unable to be delivered due to the patient's tortuous anatomy. the impella cp was explanted. there was no reported patient harm. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to the patient condition. instructions for use for the related event are as follows: ¿potential adverse events (united states): acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records.
30598 | 1220648-2024-07127 | 18796174 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18796174&pc=OZD | 08/25/2022 | 02/28/2024 |  | 02/28/2024 |  |  | 08/25/2022 | 02/28/2024 |  |  | 05/31/2024 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023217587 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/25/2022 |  |  |  |  | The user facility reported a 84-year-old male was being placed on an impella 5.5 for mechanical circulatory support. the medical staff attempted to insert the impella 5.5 through the right subclavian artery and aborted insertion due to the patient's anatomy. after failed attempts they decided to abort the procedure. there was no patient harm reported. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device has not been returned for evaluation. the root cause of the delivery issue - anatomy was most likely due to the patient condition. this report is being filed as part of a retrospective review of historical records.
30599 | 1220648-2024-07128 | 18796228 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18796228&pc=OZD | 08/26/2022 | 02/28/2024 |  | 02/28/2024 |  |  | 08/26/2022 | 02/28/2024 |  |  | 09/30/2023 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022140748 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 08/26/2022 |  |  |  |  | The user facility reported a 61-year-old female was being placed on impella 5.5 for mechanical circulatory support. it was reported that the medical staff was unable to deliver the impella pump due to the patient's anatomy. the procedure was aborted and an impella cp was placed. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device has not been returned for evaluation. the root cause of the delivery issue - anatomy was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.. this report is being filed as part of a retrospective review of historical records.
30600 | 1220648-2024-07129 | 18796278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18796278&pc=OZD | 09/07/2022 | 02/28/2024 |  | 02/28/2024 |  |  | 09/07/2022 | 02/28/2024 |  |  | 03/31/2024 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023210195 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/07/2022 |  |  |  |  | The user facility reported a 72-year-old female in post cardiotomy cardiogenic shock¿and with low cardiac output syndrome was attempted to be implanted with an impella 5.5 device for mechanical circulatory support. all attempts were unsuccessful. the physician was unable to traverse through the patient's vasculature. the decision was made to abort the impella 5.5 implant. an impella cp pump was subsequently wired and placed without issue for patient support. there was no patient harm. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer. the root cause of the delivery issue - anatomy was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
30601 | 1220648-2024-07130 | 18796296 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18796296&pc=OZD | 09/14/2022 | 03/11/2025 |  | 02/28/2024 |  |  | 09/14/2022 | 02/28/2024 |  |  | 02/29/2024 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202622 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/14/2022 |  |  |  |  | The user facility reported a 84-year-old female noted as high risk percutaneous coronary insertion was implanted with an impella cp device for mechanical circulatory support. it was reported that while the patient was on impella cp support, the patient experienced hematomas at both groin sites. perclose was applied after the impella cp was removed. the patient received packed red blood cells (prbc) and platelets. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30602 | 1220648-2024-07131 | 18796481 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18796481&pc=OZD | 02/13/2024 | 04/01/2024 |  | 02/28/2024 |  |  | 02/13/2024 | 02/28/2024 |  |  | 11/30/2025 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024435402 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/13/2024 |  |  |  |  | The user facility reported a 84-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient developed a hematoma during impella cp support. when the bleeding worsened, the decision was made to explant the device. it was noted that the repositioning sheath was believed to be too short for the patient's large tissue track. upon removal of the pump, a medtronic sheath was inserted in the right femoral artery and the bleeding stopped. however, due to the loss of blood, the patient experienced hemorrhagic shock leading to metabolic acidosis. multiple blood products were given and six ampoules of bicarbonate were pushed to treat the metabolic acidosis. | [5 empty] | The investigation into the access site bleeding ¿ major has been completed since the initial report was submitted. the product was not returned for investigation. the root cause of the access site bleeding was determined to be patient condition as the patient was large and the repo sheath could not properly reach the vessel. the issue was resolved upon using medtronic sheath. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿
30603 | 1220648-2024-07134 | 18796700 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18796700&pc=OZD | 02/15/2024 | 03/25/2024 |  | 02/28/2024 |  |  | 02/15/2024 | 02/28/2024 |  |  | 10/31/2025 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024420474 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Activation, Positioning or Separation Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/15/2024 |  |  |  |  | The user facility reported a 70-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that upon the patient arriving in the intensive care unit, transesophageal echocardiography (tee) performed showed the impella was too deep in the ventricle. while the medical doctor was repositioning the impella at the bedside, the connection of the sterile sleeve at the blue suture hub broke which resulted in the physician accidentally pulling the pump back into the aorta. the medical doctor attempted to recross the valve using tee imaging but was unsuccessful. the decision was made to take the patient back to the operating room to replace the impella. the patient is stable. | [5 empty] | The investigation into the positioning issues has been completed since the initial report was submitted. the device was not returned for investigation. based on clinical description, the root cause of positioning issue was most likely use related. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30604 | 1220648-2024-07135 | 18796924 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18796924&pc=OZD | 02/12/2024 | 04/17/2024 |  | 02/28/2024 |  |  | 02/13/2024 | 02/28/2024 |  |  | 09/30/2025 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 - JP |  |  | ABIOMED, INC. |  |  | US |  | 2024393894 |  | 1000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/13/2024 |  |  |  |  | The user facility reported a 75-year-old male undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. the team also had patient supported by ecmo. the patient experienced pericardial effusion and bleeding during support. the physician noted the patient was prone to pericardial effusion and bleeding. due to this, the composition of the purge fluid was changed from heparin to sodium bicarbonate and the patient was kept under observation. the physician noted the sodium bicarbonate was used due to bleeding. there were no problems with the operation of the device. the patient remained on impella 5.5 support. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the patient had bleeding tendency and the patient was prone to pericardial effusion and bleeding as per the clinical description. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
30605 | 1220648-2024-07136 | 18796943 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18796943&pc=OZD | 02/16/2024 | 03/19/2024 |  | 02/28/2024 |  |  | 02/16/2024 | 02/28/2024 |  |  |  |  | 02/28/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1283308 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Damaged Prior to Use | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup |  |  |  | 02/16/2024 |  |  |  |  | The user facility reported that the automated impella controller (aic) had a damaged purge disc. the aic was removed from service. no patient harm or involvement was reported. | [5 empty] | The investigation into the aic - purge disc/flag issues has been completed since the original report was filed. the console was returned and tested after arriving in fs. upon examining aic for any visible defects, it was evident that the purge retainer was broken. the cause of the issue was a broken purge disc retainer. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
30606 | 1220648-2024-07137 | 18796981 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18796981&pc=OZD | 02/20/2024 | 04/08/2024 |  | 02/28/2024 |  |  | 02/20/2024 | 02/28/2024 |  |  | 09/30/2025 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024407026 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  |  | White | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/20/2024 |  |  |  |  | The user facility reported a 79-year-old male of unknown ethnicity in st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp exhibited a priming issue. the pump was inserted, and despite the yellow-to-yellow connection purge fluid was not seen coming out of the motor housing. the pump was replaced with a different impella cp. no patient harm was reported. | [5 empty] | The investigation into the priming issue has been completed since the original report was filed. the pump was returned, tested, and evaluated. the failure mode was not reproduced. there was no issue found during investigation. the device functioned/operated to specification. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30607 | 1220648-2024-07138 | 18797053 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18797053&pc=OZD | 02/20/2024 | 04/05/2024 |  | 02/28/2024 |  |  | 02/20/2024 | 02/28/2024 |  |  | 11/30/2025 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024432746 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | No | Other | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/20/2024 |  |  |  |  | The user facility reported a 76-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella cp insertion was aborted as the impella was not able to cross the aortic valve. the patient¿s vessel was tortuous. | [5 empty] | The investigation into the delivery issues has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the delivery issue was most likely patient condition related, the impella was unable to pass through the aortic valve due to patient¿s torturous vessels as per the clinical description provided. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30608 | 1220648-2024-07139 | 18797103 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18797103&pc=OZD | 02/05/2024 | 04/23/2024 |  | 02/28/2024 |  |  | 02/22/2024 | 02/28/2024 |  |  | 10/31/2025 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024414149 |  | 1000211 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pressure Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/22/2024 |  |  |  |  | The user facility reported a 67-year-old male in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the impella 5.5 exhibited a high purge pressure. the pump was explanted and replaced. no patient harm was reported. | [5 empty] | The investigation into the high purge pressure issue has been completed since the original report was submitted. the device was not returned for investigation. the root cause of the high purge pressure was not determined since product and data logs were not returned. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30609 | 1220648-2024-07140 | 18797179 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18797179&pc=OZD | 02/22/2024 | 04/10/2024 |  | 02/28/2024 |  |  | 02/22/2024 | 02/28/2024 |  |  | 11/30/2025 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024423573 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/22/2024 |  |  |  |  | The user facility reported a 53-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the .035 wire and pigtail were unable to cross the patient's new valve during attempted impella 5.5 implant. due to the noted resistance, the decision was made to abort the implant. there was no patient harm. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into the delivery issues has been completed since the original report was submitted. based on the clinical information provided, the root cause of the delivery issue was determined to be patient condition because the impella was unable to pass new valve and the decision was made to abort.
30610 | 1220648-2024-07141 | 18797216 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18797216&pc=OZD | 10/29/2022 | 07/09/2024 |  | 02/28/2024 |  |  | 10/29/2022 | 02/28/2024 |  |  | 07/31/2024 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023258762 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Life Threatening | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 10/29/2022 |  |  |  |  | The user facility reported a 53-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. us complainant reported the patient was on impella cp support due to going into cardiac arrest. the patient had a positive head ct scan for hemorrhagic process. it is unknown if the impella was related to the positive ct scan. | [5 empty] | The investigation for the reported stroke has been completed. the investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. the instructions for use for the impella cp with smart assist in section: potential adverse events (united states) states ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ d4-corrected model number in accordance with updated procedures, sections d6 and h10 have been revised from the initial submission. f6/f8 should have been left blank in the initial report. g1-corrected: reporting contact email. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
30611 | 1220648-2024-07142 | 18797262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18797262&pc=OZD | 02/14/2024 | 03/22/2024 |  | 02/28/2024 |  |  | 02/14/2024 | 02/28/2024 |  |  | 10/31/2025 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024421187 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/14/2024 |  |  |  |  | The user facility reported a 60-year-old male who had coded for 10-15 minutes at work, was brought to the hospital by ambulance. the patient coded for an additional 40 minutes in the ed. the patient was emergently implanted with an impella cp for circulatory support. after return of spontaneous circulation, the cp was placed. during the support of just 11 hours, there were signs of hemolysis. the hemolysis was not treated as the care was withdrawn and patient expired. | [5 empty] | The investigation into the hemolysis has been completed since the initial report was submitted. the device was not returned for investigation. based on clinical description and data analysis, the root cause of hemolysis was most likely due to pump was not in the proper position. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
30612 | 1220648-2024-07143 | 18797361 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18797361&pc=OZD | 02/16/2024 | 09/20/2024 |  | 02/28/2024 |  |  | 02/16/2024 | 02/28/2024 |  |  | 07/21/2025 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP SET, PUMP 381 (US) |  |  | ABIOMED, INC. |  |  | US |  | 2025443392 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | Adverse Event Without Identified Device or Use Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 02/16/2024 |  |  |  |  | It was noted that the patient expired after explant of the device. medical safety review of the event noted that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | The user facility reported a 79-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella was being repositioned when the impella stopped. all attempts to restart were unsuccessful. the impella was pulled back into the descending aorta and explanted. |  |  |  |  | This file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type. corrected information for the initial mfr 1220648-2024-07143 was provided in sections b1, b2, b5, d6, h1, and h6. | The investigation into the pump stop issue has been completed since the initial report was submitted. the device was not returned for investigation. per the clinical data provided, there were a few motor current spikes before the pump stop. the root cause of the pump stop issue was not determined since product was not returned and provided information was inconclusive. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ ¿impella stopped if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
30613 | 1220648-2024-07145 | 18797569 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18797569&pc=OZD | 04/07/2022 | 02/28/2024 |  | 02/28/2024 |  |  | 04/07/2022 | 02/28/2024 |  |  | 11/30/2023 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023167042 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/07/2022 |  |  |  |  | The user facility reported a 58-year-old male was implanted with an impella 5.5 device for mechanical circulatory support, and that the patient developed some oozing at the impella site. a standard bleeding protocol was followed to treat the bleeding. | [5 empty] | The investigation for the reported access site bleeding has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleeding could not be determined. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury" this report is being filed as part of a retrospective review of historical records.
30614 | 1220648-2024-07146 | 18797674 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18797674&pc=OZD | 04/08/2022 | 02/28/2024 |  | 02/28/2024 |  |  | 04/08/2022 | 02/28/2024 |  |  | 01/31/2024 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023180571 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/08/2022 |  |  |  |  | The user facility reported a 68-year-old female was being placed on impella 5.5 support for mechanical circulatory support. the medical team was unable to advance the impella 5.5 through the artery. therefore, the impella 5.5 implant was aborted and an impella cp was subsequently placed for patient support. | [5 empty] | The investigation into the reported delivery issue has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue - anatomy was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
30615 | 1220648-2024-07147 | 18797874 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18797874&pc=OZD | 09/22/2022 | 02/28/2024 |  | 02/28/2024 |  |  | 09/22/2022 | 02/28/2024 |  |  | 05/31/2024 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023217523 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Difficult to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/22/2022 |  |  |  |  | The user facility reported a female of unknown age undergoing cardiac surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the impella easily passed through the introducer and anastomosis before stopping abruptly. all further attempts with and without a wire were determined unsuccessful due to patient anatomy and insertion was aborted. the patient expired in the or due to the length of time patient was on cardiopulmonary bypass and unrecoverable metabolic burden with no allegation of patient harm due to device. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. an investigation was completed based on clinical details provided. no product defect was identified and root cause of failure to deliver was determined to be patient anatomy affecting the ability of the device to be placed. this report is being filed as part of a retrospective review of historical records.
30616 | 1220648-2024-07148 | 18798140 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18798140&pc=OZD | 09/22/2022 | 05/22/2025 |  | 02/28/2024 |  |  | 09/22/2022 | 02/28/2024 |  |  | 09/30/2023 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022139751 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Improper Flow or Infusion | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/22/2022 |  |  |  |  | The user facility reported an impella cp in 81-year-old male patient for mechanical circulatory support. it was reported that while on support, the impella was lowered to p7 to break suction. there was ongoing suction at p8 with no level changes noted. patient remained on support. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30617 | 1220648-2024-07149 | 18798264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18798264&pc=OZD | 01/27/2023 | 10/01/2025 |  | 02/28/2024 |  |  | 01/27/2023 | 02/28/2024 |  |  | 10/31/2024 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024287092 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Device Sensing Problem | Optical Problem | [5 empty] | Unspecified Vascular Problem | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 01/27/2023 |  |  |  |  | The user facility reported a 80-year-old female undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that during impella cp support, the rate of the patient's pacemaker required adjustments. however, upon attempting to adjust the rate, it was stated that the impella was blocking the bluetooth signal for the pacemaker. the pump's support level was lowered to p0 to troubleshoot and then increased again for ongoing support. there was no patient harm as a result of the noted issue. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. no product was returned for a visual inspection. data log analysis confirmed that there were no abnormalities throughout the case. there was no problem found during the case as the device operated to specification. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b2 updated to death. h1 updated to death. h6 updated to include death and placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-07149 d4 model number corrected. d6a/d6b updated with additional information. f6 and f8 should not have had entries in the initial report. g1 reporting contact email updated.
30618 | 1220648-2024-07150 | 18798317 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18798317&pc=OZD | 07/27/2023 | 08/09/2024 |  | 02/28/2024 |  |  | 07/27/2023 | 02/28/2024 |  |  | 12/31/2024 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024295685 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Activation, Positioning or Separation Problem | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 07/27/2023 |  |  |  |  | The user facility reported an unspecified patient undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported impella 5.5 positioning signal issues. | [5 empty] | The investigation of positioning issues has been completed. the root cause of the positioning issue was not determined due to a lack of clinical details and there were no abnormalities on the returned pump. d6a and d6b added. d9 date device returned to manufacturer added the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-07150: d4 model number. f6/f8-should have been left blank. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. no product was returned for investigation. the impella in ventricle alarmed because the pump was out of position and in the ventricle. the root cause of the positioning issue was not determined since product was not returned and there is a lack of clinical details. instructions for use for the related event are as follows: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿
30619 | 1220648-2024-07151 | 18798344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18798344&pc=OZD | 08/29/2023 | 09/26/2025 |  | 02/28/2024 |  |  | 12/20/2023 | 02/28/2024 |  |  | 06/30/2025 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | PUMP 550 S2 PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2024367644 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 23 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia Stroke | [12 empty] | Life Threatening | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup |  |  | 12/20/2023 |  |  |  |  | The user facility reported a 23-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the us medical center had a 5.5 pump support ongoing when the patient stood for teeth brushing and had a syncopal episode. the patient was found down and sent to ct (computed tomography) for imaging. during that time the ps was temporarily not visible. the medical team did not attribute the syncopal episode to the use of the impella. the pump remains on for support. this syncopal episode occurred on day 16 of the support. | [5 empty] | The investigation into the optical signal issue and the stroke/cva has been completed since the original report was filed. the device was returned and tested after arriving in fs. the root cause of the optical signal issue could not be determined as only the catheter was returned and no data logs were provided. as the clinical information was not provided, the root cause of the stroke/cva could not be determined. | The impella device was received from the customer on (b)(6) 2023. the root cause of the optical signal issue cannot be determined as only the catheter was returned and no data logs were provided. strokes/cvas can occur during or after impella support, and can be either ischemic or hemorrhagic. to make a determination as to the cause of the stroke, the following clinical information must be considered: -patient's medical history, including any previous strokes, cardiovascular disease, or risk factors such as hypertension, diabetes, or atrial fibrillation -timing of the stroke in relation to impella support (during or post-implant) -detailed description of the symptoms experienced by the patient -neurological examination findings and any focal deficits observed -diagnostic imaging results, such as ct scan or mri of the brain, to identify the type and location of the stroke (ischemic or hemorrhagic) -laboratory test results, including coagulation profiles and cardiac markers -any relevant procedural or device-related factors, such as duration and settings of impella support, presence of embolic protection devices, or any documented procedural complications -response to any previous anticoagulation or antiplatelet therapies -evaluation of potential alternative causes of stroke, such as arrhythmias, embolic sources, or underlying vascular pathology as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. all pertinent information available to abiomed has been submitted at this time. h6 updated. g1 reporting contact email updated. d10 concomitant medical products and therapy dates.
30620 | 1220648-2024-07152 | 18798377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18798377&pc=OZD | 02/09/2024 | 04/04/2024 |  | 02/28/2024 |  |  | 02/16/2024 | 02/28/2024 |  |  | 09/30/2025 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2024410027 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | Perforation of Vessels | Thrombosis/Thrombus | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/16/2024 |  |  |  |  | The user facility reported a 84-year-old male undergoing cardiac surgery was implanted with an impella device for mechanical circulatory support. the complainant reported that after achieving hemostasis of the left common femoral artery, the foot became cool to the touch and bluish in color. a doppler pulse was not found and the artery was found to be 100% occluded. a thrombectomy was subsequently performed and the staff was able to aspirate multiple clots. heparin was additionally added to treat the clotting. it was noted that the patient had severe peripheral and aortic disease at the time of the noted issue. while assessing the flow across the left common femoral artery, a hematoma developed in the left groin. manual pressure was applied to manage the condition and an angiogram revealed extravasation at the mid left common femoral artery. a balloon tamponade was initially applied, but a covered stent was later placed when the extravasation continued. a smaller balloon was then inflated at high pressure within the stent and the extravasation resolved. the patient¿s blood pressure began to drop and, despite hemoglobin levels being within normal limits, the decision was made to give the patient one unit of prbc (packed red blood cells). further intervention was not required. | [5 empty] | The investigation into the access site bleeding ¿ minor, the vascular damage - peripheral vessel, the limb ischemia ¿ reversible and the thrombus issues has been completed since the initial report was submitted. the device was not returned for investigation. the root cause of the access site bleeding - minor was not determined since product was not returned and there is a lack of clinical details. the root cause of the vascular damage - peripheral vessel issue was most likely patient condition related due to the patient's severe peripheral and aortic disease. as the necessary information was not provided, the root cause of the limb ischemia and the thrombus was not determined. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿be careful not to pull on the impella catheter when transferring a patient from one bed to another.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
30621 | 1220648-2024-07153 | 18798383 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18798383&pc=OZD | 02/16/2024 | 04/03/2024 |  | 02/28/2024 |  |  | 02/16/2024 | 02/28/2024 |  |  | 06/30/2025 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC, |  |  | US |  | 2024368875 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/16/2024 |  |  |  |  | The user facility reported a 68-year-old male noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was reported that during repositioning of an implanted impella cp pump, the device was pulled back into the aorta. multiple attempts to re-cross the valve was unsuccessful, and the decision was made to replace the pump. there was no patient harm reported. | [5 empty] | The investigation into the positioning issues has been completed since the initial report was submitted. the device was not returned for investigation. per the clinical data provided, the cause of the positioning issue was use related due to improper technique as the pump pulled back to aorta and unable to reinserted back. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30622 | 1220648-2024-07154 | 18798399 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18798399&pc=OZD | 04/18/2022 | 02/28/2024 |  | 02/28/2024 |  |  | 04/18/2022 | 02/28/2024 |  |  | 12/31/2023 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023176308 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 04/18/2022 |  |  |  |  | The user facility reported a 57-year-old female was on impella cp support and that the patient required defibrillation fifteen (15) times during the case. also, a left ventricle thrombus was noted in the apex during routine echocardiogram (echo). the patient was weaned and the impella cp was removed. the patient was restarted on anticoagulation therapy to address the thrombus once hemostasis of the impella site was achieved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
30623 | 1220648-2024-07155 | 18798415 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18798415&pc=OZD | 02/18/2024 | 03/28/2024 |  | 02/28/2024 |  |  | 02/18/2024 | 02/28/2024 |  |  | 04/30/2025 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024337077 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/18/2024 |  |  |  |  | The user facility reported a 69-year-old male with cardiomyopathy was implanted with an impella 5.5 as a bridge to transplant. it was reported that while the patient was on support, the systematic anticoagulant was discontinued due to the development of a hematoma. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. clinical details states that hemostasis was achieved after the heparin was stopped. therefore, the root cause of the access site bleed was most likely due to anticoagulation management. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿
30624 | 1220648-2024-07156 | 18798426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18798426&pc=OZD | 09/19/2022 | 02/28/2024 |  | 02/28/2024 |  |  | 09/19/2022 | 02/28/2024 |  |  | 09/30/2023 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022147433 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 09/19/2022 |  |  |  |  | The user facility reported a 35-year-old male needing an impella cp for mechanical circulatory support. while on the impella cp, the patient developed hemolysis and as result, the impella cp was removed. patient was then placed on extra-corporeal membrane oxygenation (ecmo) and an intra-aortic balloon pump (iabp). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left tachycardia volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
30625 | 1220648-2024-07157 | 18798432 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18798432&pc=OZD | 02/22/2024 | 04/01/2024 |  | 02/28/2024 |  |  | 02/22/2024 | 02/28/2024 |  |  |  |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025437522 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 02/22/2024 |  |  |  |  | The user facility reported a 58-year-old male in non-st elevated myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was that following impella cp implant, the peel-away sheath could not be removed naturally and had to be cut away. the repositioning sheath was subsequently placed for the planned support. there was no patient harm reported. | [5 empty] | The investigation for the reported difficulty inserting/removing introducer has been completed. the introducer was returned for evaluation. according to the clinical, before beginning impella cp support the long peel away sheath could not be removed correctly so the physician was forced to cut the sheath. product evaluation confirmed a sheath scoreline defect near the introducer hub. the cause of the mechanical introducer damage was a sheath scoreline split. added- d9 device return date, h6 impact code 4641. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
30626 | 1220648-2024-07158 | 18798566 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18798566&pc=OZD | 09/12/2022 | 03/11/2025 |  | 02/28/2024 |  |  | 09/12/2022 | 02/28/2024 |  |  | 04/30/2024 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023213608 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  |  |  | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/12/2022 |  |  |  |  | The user facility reported a 76-year-old female noted as high risk percutaneous coronary insertion had an attempted implant of the impella cp device for mechanical circulatory support. the patient had torturous vessel and the device would not advance without pressure, possibly from a kink in the sheath. the device implantation was aborted. the patient had a successful percutaneous coronary intervention¿of the left main. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30627 | 1220648-2024-07159 | 18798643 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18798643&pc=OZD | 09/12/2022 | 03/11/2025 |  | 02/28/2024 |  |  | 09/12/2022 | 02/28/2024 |  |  | 06/30/2024 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023237212 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/12/2022 |  |  |  |  | The user facility reported a 63-year-old male undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the patient experienced some oozing (bleeding) at the femoral insertion site. activated clotting time decreased and system heparin was held. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30628 | 1220648-2024-07160 | 18798664 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18798664&pc=OZD | 09/12/2022 | 03/11/2025 |  | 02/28/2024 |  |  | 09/12/2022 | 02/28/2024 |  |  | 07/31/2024 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231611 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/12/2022 |  |  |  |  | The user facility reported a 76-year-old female in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient experienced a hematoma at the insertion site. manual pressure was held. the decision was made to remove the impella and perform a surgical repair. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30629 | 1220648-2024-07161 | 18798755 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18798755&pc=OZD | 09/29/2022 | 05/28/2025 |  | 02/28/2024 |  |  | 09/29/2022 | 02/28/2024 |  |  | 03/31/2024 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023211483 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/29/2022 |  |  |  |  | The user facility reported a 54-year-old male patient was implanted with an impella cp for mechanical circulatory support. after being placed on impella cp support, the patient was observed to have bloody urine. it was unknown if there was blood in the foley or if hemolysis was occurring. the hemolysis was troubleshooted with reduction in afterload. the patient continued support until successfully weaned. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. the instructions for use states: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30630 | 1220648-2024-07162 | 18798760 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18798760&pc=OZD | 09/22/2022 | 09/25/2025 |  | 02/28/2024 |  |  | 09/22/2022 | 02/28/2024 |  |  | 05/31/2024 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023226777 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 39 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/22/2022 |  |  |  |  | The complainant reported a patient on impella cp support had on-going septic shock. patient had on-going septic shock with unknown cause. due to atrial fibrillation, 5 times cardio-version. the patient ultimately expired. the medical safety officer reviewed on (b)(6) 2024 and determined that the use of the device cannot conclusively be associated with the outcome. | The user facility reported an impella cp in 39-year-old male patient for mechanical circulatory support. it was reported while on support, the patient had on-going septic shock. |  |  |  |  | The investigation for the reported issue has been completed. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. this report is being filed as part of a retrospective review of historical records. b5 updated with additional information. d6a / d6b added implant and explant dates. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. h6 health effect - clinical code: added atrial fibrillation 1729 code. health effect - impact code: added unexpected medical intervention 4641. component code pump 925. investigation conclusions cause not established 4315. | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
30631 | 1220648-2024-07163 | 18798838 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18798838&pc=OZD | 09/23/2022 | 05/20/2025 |  | 02/28/2024 |  |  | 09/23/2022 | 02/28/2024 |  |  | 05/31/2024 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023230324 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/23/2022 |  |  |  |  | The user facility reported an impella cp was implanted in 82-year-old male patient for mechanical circulatory support. it was reported while on support, the physician wanted to reposition the device. after attempting to reposition the device, the impella was explanted as the optimal position was not achieved. the pump was replaced with impella pump 5.5. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30632 | 1220648-2024-07164 | 18799164 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18799164&pc=OZD | 09/20/2022 | 03/11/2025 |  | 02/28/2024 |  |  | 09/20/2022 | 02/28/2024 |  |  | 06/30/2024 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023237645 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  |  |  | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/20/2022 |  |  |  |  | The user facility reported a 63-year-old female in st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. after being placed on impella cp support, the patient had constant oozing at the impella insertion site. the bleeding was addressed by repositioning the sheath, placing a knee immobilizer, holding pressure at the site, and placing new sutures. the right groin was redressed, and the patient¿s pulses remained palpable. the patient was eventually weaned and the impella cp was explanted after it was discovered the impella cp was interacting with the papillary muscle. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: ¿assess access site for bleeding and hematoma.¿ ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury¿ ¿impella catheter in papillary muscle if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
30633 | 1220648-2024-07165 | 18799722 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18799722&pc=OZD | 09/17/2022 | 03/07/2025 |  | 02/28/2024 |  |  | 09/17/2022 | 02/28/2024 |  |  | 01/31/2024 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023188838 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Pain | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/17/2022 |  |  |  |  | The user facility reported an 84-year-old female patient was being implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed right leg pain throughout the procedure, therefore the impella was explanted. once explanted, the leg pain subsided immediately. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | Limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition . ¿ any concomitant medication . ¿ vascular size or variations thereof. ¿ detailed description of the symptoms experienced by the patient . ¿ physical examination findings, including pulse assessment and visual examination of the affected limb . ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography . ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders . ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension . ¿ patient's response to previous treatments or interventions for vascular conditions . with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
30634 | 1220648-2024-07166 | 18799749 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=18799749&pc=OZD | 09/18/2022 | 03/04/2024 |  | 02/28/2024 |  |  | 09/18/2022 | 02/28/2024 |  |  | 03/31/2024 |  | 02/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023205797 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 09/18/2022 |  |  |  |  | The user facility reported a 76-year-old male patient was being implanted with an impella cp device for mechanical circulatory support. upon initiation of support, the patient lost all pulsatility. there were no pulses in the feet due to the lack of pulsatility, but both feet matched in temperatures. the patient¿s extremities became cool to the touch; however pulses were unavailable. there were no orders to access or remedy. the impella was removed from the patient. | [5 empty] | The impella device was not returned for evaluation. upon completion of the investigation a supplemental report will be submitted. instructions for use for the related event are as follows: ¿potential adverse events (united states) acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and vascular injury.¿ this report is being filed as part of a retrospective review of historical records. | The investigation for the reported limb ischemia has been completed. the impella device was not returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.